Supplementary information

Supplementary Methods

Design and performance of the prototype ventilator

In this study, we created a prototype of a ventilator that can perform intermittent positive pressure ventilation and mimic the breathing cycle of the 2:2 breathing pattern during a marathon. This mode of ventilation was named the marathoners’ breathing ventilation (MBV) pattern. The inspiratory:expiratory ratio (I:E) in the MBV was set to 1:1. The prototype ventilator is a combination of digital timers, solenoid valves and existing respiratory equipment (ACOMA AR-300, ACOMA Co., Tokyo, Japan).

The ACOMA AR-300 is a volume-limited, time-cycling ventilator, in which the I:E ratio can be adjusted from 1:3 to 3:1, and the tidal volume can be set from 20 to 300 ml. The ACOMA AR-300 is a piston-driven ventilator, designed to deliver a constant inspiratory flow rate.

The I:E ratio of the ACOMA AR-300 was set to 1:1. The detector switch was attached to the piston rod of the ACOMA AR-300. The detector switch was set to emit a signal each time it detected the start of the inspiratory period of the ventilator. The signal was sent to four digital timers (H5BR-B; OMRON Co., Tokyo, Japan; Supplementary Fig. 1). Each of these timers was connected to a solenoid valve: three tri-directional valves (VDW350-6G-3-01; SMC Co., Tokyo, Japan) and one bidirectional valve (VDW31-6G-3-01; SMC Co.). These four solenoid valves were powered by 12-Volt DC supplied by the connected digital timer. All digital timers were reset as soon as the ACOMA AR-300 inspiratory start signal was received. At that point, they started to measure elapsed time. The start time and duration of the 12-Volt DC output (controlling the solenoid valves) could be set individually for each of the four digital timers.

A schematic diagram of the prototype device is shown on the left panel of Supplementary Fig. 1. The right panel of Supplementary Fig. 1 is an example of the time allocation for the four solenoid valves with the respiration rate set to 30 cycles/min. The solenoid valve operation in this example and the corresponding inspired and expired gas dynamics were as follows. One cycle was defined as the sum of eight separate actions: 1st inspiratory period, a resting period, 2nd inspiratory period, a resting period, 1st expiratory period, a resting period, 2nd expiratory period, and a final resting period. In this example, the respiration rate of the ACOMA AR 300 was set to 30 breaths/min. In response to the inspiratory start signal of the ACOMA AR 300, all digital timers were reset and began to record the time. First, the solenoid valve #1 opened for 0.3 s, and air from the ventilator flowed directly into the inspiratory circuit (1st inspiratory period: 0.3 s). When solenoid valve #1 closed, the air moved towards solenoid valve #2, but since solenoid valve #2 was closed, the air was released into the atmosphere from the exhaust pipe (resting period). After 0.5 s from the start signal, solenoid valve #2 was opened for 0.3 s, and the air from the ventilator went into the inspiratory circuit again (2nd inspiratory period: 0.3 s). When solenoid valve #2 was closed, the air was released into the atmosphere from the exhaust pipe of solenoid valve #2 (resting period). After 1.0 s from the start signal, solenoid valve #3 was opened for 0.3 s, and expired gas was released from the exhalation pipe of this valve into the atmosphere (1st expiratory period: 0.3 s). When solenoid valve #3 was closed, the expired gas moved towards solenoid valve #4, but since solenoid valve #4 was closed, the emission of the expired gas into the atmosphere was interrupted (resting period). After 1.5 s from the start signal, solenoid valve #4 was opened for 0.3 s, and expired gas was released from the exhalation pipe of solenoid valve #4 into the atmosphere (2nd expiratory period: 0.3 s). When solenoid valve #4 was closed, the emission of the expired gas into the atmosphere was interrupted (resting period). After 2.0 s from the start signal, the detector switch detected the next inspiratory cycle of the ACOMA AR 300, the digital timer was reset, and the MBV breathing cycle was repeated. In this example, it did not matter if the open duration of solenoid valve #4 was set to values above 0.5 s, because the detector switch detected the start of the next inspiratory cycle in the ACOMA AR 300. This terminated the open duration of valve #4 since the reset of the digital timers was prioritized, thus the actual open duration of solenoid valve #4 became 0.5 s.

With this prototype respiratory equipment, producing mechanical ventilation in a 2:2 breathing pattern, the respiratory rate can be adjusted in 5 cycles/min increments from 10 cycles/min to 65 cycles/min. The ventilation volume of one MBV cycle was defined as the difference between the tidal volume and respiratory rate set in the ACOMA AR 300 and the exhaust volume from the exhaust pipe of solenoid #2. Since this was a prototype, it was necessary to measure the ventilation volume of one MBV cycle using the water displacement method. In addition, the positive end-expiratory pressure (PEEP) could be set at arbitrary pressure values independently of the exhaust pipes of the solenoid valves #3 and #4.

SUPPLEMENTARY FIGURE LEGENDS

Supplementary Figure 1. The left panel is a schematic diagram of the prototype device. The right panel is an example of the time allocation for the four solenoid valves, with the respiration rate set to 30 cycles/min.

In the MBV method, the 1st inspiratory, 2nd inspiratory, 1st expiratory, and 2nd expiratory phases taken together, defined one cycle. In the IRV method, the inspiratory and subsequent expiratory phase together defined one cycle. In both MBV and IRV methods, the tidal volume was defined as the total ventilation volume during one cycle, and the unit of respiratory rate was cycles/min.