Supplementary Material

| Supplementary Table S1. Accuracy Assessment of CGM at Different Blood Glucose Levels |  |  |  |  |
| --- | --- | --- | --- | --- |
| Glucose Range | Reference Method | Paired Samples (n) | MARD (%) | Median ARD (%) |
| <70 mg/dL | ABG | 5 | 54.9 (22.0-112.5) | 31.3 (14.3-166.7) |
| 70-180 mg/dL | ABG | 636 | 12.0 (11.2-12.8) | 8.8 (8.1-9.6) |
|  | POC | 337 | 10.8 (9.7-12.0) | 8.1 (7.2-9.2) |
| >180 mg/dL | ABG | 611 | 9.6 (9.0-10.3) | 7.6 (7.2-8.3) |
|  | POC | 373 | 8.5 (7.8-9.3) | 7.3 (6.3-7.9) |

| Supplementary Table S2. CGM Accuracy Metrics in Hypoglycemic Ranges |  |  |  |
| --- | --- | --- | --- |
| Range (mg/dL) | Paired Samples (n) | MARD (%) | Median ARD (%) |
| <40 | 1 | 166.7 | 166.7 |
| 40-70 | 4 | 27.0 (18.5-33.7) | 27.7 (14.3-36.1) |
| Limited sample sizes in hypoglycemic ranges preclude robust statistical analysis and interpretation. |  |  |  |

| Supplementary Table S3. Accuracy of CGM in Non-Sepsis, Sepsis, and Septic Shock Patients |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Group | Reference | Paired Samples (n) | MARD (%) | Median ARD (%) | ISO (%) | CLSI (%) |
| Non-sepsis | ABG | 399 | 10.6 (9.7-11.4) | 8.2 (7.5-9.0) | 75.7 (70.9-79.9) | 67.9 (63.2-72.4) |
|  | POC | 333 | 10.1 (9.0-11.1) | 7.9 (7.1-8.8) | 79.9 (75.4-84.1) | 73.3 (68.5-77.8) |
| Sepsis | ABG | 449 | 10.3 (9.4-11.2) | 7.6 (6.8-8.4) | 75.5 (71.7-79.3) | 69.7 (65.3-73.9) |
|  | POC | 243 | 9.0 (8.1-9.9) | 7.5 (6.3-8.4) | 81.9 (77.0-86.4) | 75.7 (70.0-81.1) |
| Septic Shock | ABG | 404 | 12.2 (11.0-13.4) | 8.8 (8.1-10.0) | 72.5 (68.1-76.5) | 63.9 (59.4-68.3) |
|  | POC | 134 | 9.6 (8.1-11.3) | 7.1 (5.4-9.0) | 81.3 (74.6-88.1) | 78.4 (70.9-85.1) |
| Sepsis defined according to Sepsis-3 criteria; Septic shock defined as sepsis with persistent hypotension requiring vasopressors and serum lactate >2 mmol/L despite adequate volume resuscitation. ISO, ISO 15197:2013 standard; CLSI, CLSI POCT12-A3 standard. |  |  |  |  |  |  |

| Supplementary Table S4. Comparative Analysis of CGM Accuracy Intergroup Differences Based on Sepsis Stratification |  |  |
| --- | --- | --- |
| Patient Category | Reference | 95% CI of MARD Difference |
| Non-sepsis vs. Sepsis | ABG | (-0.00498, 0.01719) |
|  | POC | (-0.00759, 0.01908) |
| Non-sepsis vs. Septic Shock | ABG | (-0.02104, 0.00557) |
|  | POC | (-0.01117, 0.02710) |
| Sepsis vs. Septic Shock | ABG | (-0.02694, 0.00038) |
|  | POC | (-0.01560, 0.02458) |
| Data presented as 95% confidence intervals for MARD differences. Statistical significance (P<0.05) was determined when the 95% CI of MARD difference did not include zero. |  |  |

| Supplementary Table S5. CGM Accuracy by pH and Lactate Levels in Septic Shock Patients |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Range | Paired Samples (n) | MARD (%) | Median ARD (%) | ISO (%) | CLSI (%) |
| pH level | <7.3 | 22 | 19.3 (8.7-36.1) | 9.3 (4.8-14.9) | 72.7 (54.5-90.9) | 59.1 (36.4-77.3) |
|  | 7.3-7.5 | 337 | 12.0 (11.0-13.0) | 8.8 (8.1-10.1) | 71.5 (66.5-76.3) | 62.9 (57.6-68.2) |
|  | >7.5 | 45 | 10.3 (7.9-12.8) | 9.1 (7.2-10.8) | 80.0 (66.7–91.1) | 73.3 (60.0–84.5) |
| Lactate level (mmol/L) | ≤ 2 | 251 | 11.6 (10.4-12.9) | 8.8 (7.7-9.8) | 73.7 (67.7–78.9) | 64.9 (58.6–70.9) |
|  | 2-4 | 100 | 10.9 (9.3-12.4) | 9.1 (7.4-11.4) | 76.0 (67.0–85.0) | 67.0 (57.0–76.0) |
|  | 4-8 | 39 | 12.0 (8.9-15.2) | 9.9 (6.4-14.9) | 69.2 (53.8–82.1) | 56.4 (41.0–71.8) |
|  | > 8 | 14 | 32.9 (16.3-57.1) | 23.8 (8.1-37.1) | 42.8 (21.3–71.4) | 42.8 (21.3–71.4) |
| pH and lactate levels obtained from arterial blood gas measurements. |  |  |  |  |  |  |

| Supplementary Table S6. Comparison of CGM Accuracy (MARD) Between pH Level Groups |  |
| --- | --- |
| pH level | 95% CI of MARD Difference |
| 7.3-7.5 vs. >7.5 | (-0.01993, 0.01880) |
| Data presented as 95% confidence interval for MARD difference. Statistical significance (P<0.05) was determined when the 95% CI of MARD difference did not include zero. |  |

| Supplementary Table S7. Pairwise Comparison of CGM Accuracy (MARD) Across Lactate Level Groups |  |
| --- | --- |
| Lactate level | 95% CI of MARD Difference |
| ≤2mmol/L vs. 2-4mmol/L | (-0.02886, 0.01787) |
| ≤2mmol/L vs. 4-8mmol/L | (-0.06618, 0.01933) |
| 2-4mmol/L vs. 4-8mmol/L | (-0.06584, 0.03269) |
| Data presented as 95% confidence intervals for MARD difference. Statistical significance (P<0.05) was determined when the 95% CI of MARD difference did not include zero. |  |