Supplementary Table S1. Analyte-specific autoverification parameter ranges (October 2024 – April 2025)

| Analyte | Unit | Autoverifica-tion Range | Critical Value | History Window | Delta Check Limit | Total Tests (n) |
| --- | --- | --- | --- | --- | --- | --- |
| TP | g/L | 50–90 | — | 7 d | ±15% | 252,344 |
| ALB | g/L | 20–60 | — | 7 d | ±15% | 252,578 |
| ALT | U/L | 1–300 | — | 7 d | ±20% | 118,883 |
| AST | U/L | 1–300 | — | 7 d | ±20% | 90,193 |
| TBIL | μmol/L | 1–100 | >257 | 7 d | ±30% | 257,148 |
| DBIL | μmol/L | 0.1–20 | — | 7 d | ±30% | 157,399 |
| LDH | U/L | 100–400 | — | 7 d | ±30% | 55,365 |
| HBDH | U/L | 50–200 | — | 7 d | ±30% | 7,181 |
| GGT | U/L | 5–250 | — | 7 d | ±30% | 14,027 |
| CHE | U/L | 2,000–18,000 | — | 7 d | ±30% | 11,873 |
| ADA | U/L | 4–24 | — | 7 d | ±30% | 11,875 |
| ALP | U/L | 30–400 | — | 7 d | ±30% | 14,013 |
| TBA | μmol/L | 0.01–50 | — | 7 d | ±30% | 15,214 |
| UREA | mmol/L | 1–12 | — | 7 d | ±30% | 53,887 |
| UREA (Nephrology) | mmol/L | 1–30 | — | 7 d | ±50% | — |
| CREA | μmol/L | 20–442 | >442 | 7 d | ±30% | 107,693 |
| CREA (Nephrology) | μmol/L | 20–1,000 | >1,000 | 7 d | ±50% | — |
| UA | μmol/L | 30–800 | — | 7 d | ±30% | 49,575 |
| CYSC | mg/L | 0.5–5 | — | 7 d | ±30% | 23,117 |
| CYSC | mg/L | 5–15 (CREA>442) | — | 7 d | ±50% | — |
| GLU | mmol/L | 3–11 | Neonatal: <1.7 or >16.6Adult: <2.8 or >24.75 | 7 d | ±30% | 93,220 |
| GLU (Diabetes) | mmol/L | 3–24 | — | 7 d | ±50% | — |
| TG | mmol/L | 0.1–8 | — | 7 d | ±50% | 37,285 |
| HDL-C | mmol/L | 0.5–2.5 | — | 7 d | ±30% | 74,190 |
| LDL-C | mmol/L | 1–6 | — | 7 d | ±30% | 74,238 |
| TC | mmol/L | 1–8 | — | 7 d | ±30% | 74,234 |
| APOA1 | g/L | 0.3–2.6 | — | 7 d | ±30% | 3,798 |
| APOB | g/L | 0.2–2.4 | — | 7 d | ±30% | 3,798 |
| LPA | mg/L | 0–1,000 | — | 7 d | ±15% | 3,799 |
| Potassium | mmol/L | 3.2–5.5 | <3.0 or >6.5 | 7 d | ±10% | 35,815 |
| NA | mmol/L | 134–150 | <120 or >160 | 7 d | ±5% | 24,916 |
| CL | mmol/L | 95–114 | — | 7 d | ±5% | 12,866 |
| CO₂ | mmol/L | 18–33 | — | 7 d | ±20% | 11,059 |
| CA | mmol/L | 1.87–2.77 | <1.75 or >3.5 | 7 d | ±15% | 24,798 |
| P | mmol/L | 0.77–2 | — | 7 d | ±20% | 3,886 |
| MG | mmol/L | 0.68–1.2 | — | 7 d | ±20% | 3,354 |
| CK | U/L | 20–500 | — | 7 d | ±30% | 4,299 |
| CK-MB | U/L | 0–25 | — | — | — | 4,175 |
| HCY | μmol/L | 0–25 | — | 7 d | ±20% | 3,747 |
| RF | IU/mL | 0–20 | — | 7 d | ±30% | 1,574 |
| ASO | IU/mL | 0–200 | — | 7 d | ±30% | 1,348 |

Note: “—” indicates not applicable. Pass rates for UREA (Nephrology), CREA (Nephrology), CYSC (CREA>442), GLU (Diabetes), and other population-specific configurations are not reported separately. Abbreviations: ADA, Adenosine deaminase; ALB, Albumin; ALP, Alkaline phosphatase; ALT, Alanine aminotransferase; APOA1, Apolipoprotein A1; APOB, Apolipoprotein B; ASO, Antistreptolysin O; AST, Aspartate aminotransferase; CA, Calcium; CHE, Cholinesterase; CK, Creatine kinase; CK-MB, Creatine kinase-MB isoenzyme; CL, Chloride; CO₂, Carbon dioxide; CREA, Creatinine; CYSC, Cystatin C; DBIL, Direct bilirubin; GGT, Gamma-glutamyl transferase; GLU, Glucose; HBDH, Hydroxybutyrate dehydrogenase; HCY, Homocysteine; HDL-C, High-density lipoprotein cholesterol; LDH, Lactate dehydrogenase; LDL-C, Low-density lipoprotein cholesterol; LPA, Lipoprotein (a); MG, Magnesium; NA, Sodium; P, Phosphorus; RF, Rheumatoid factor; TBA, Total bile acids; TBIL, Total bilirubin; TC, Total cholesterol; TG, Triglycerides; TP, Total protein; UA, Uric acid; UREA, Urea.

Supplementary Table S2. Classification of intercepted samples during autoverification system validation

| Interception Reason | Intercepted Samples | % of Intercepted | % of Total |
| --- | --- | --- | --- |
| Approved after direct manual review | 939 | 46.49% | 1.77% |
| Exceeded delta check limit | 491 | 24.31% | 0.92% |
| Exceeded autoverification range (clinically consistent) | 424 | 20.99% | 0.80% |
| Required auto-dilution (exceeded linear range) | 82 | 4.06% | 0.15% |
| Critical value alert triggered | 74 | 3.66% | 0.14% |
| Required instrument change and retesting | 10 | 0.50% | 0.02% |
| Total intercepted | 2,020 | 100% | 3.80% |