## 1. General Info
Study_ID | Full_Title | First_Author | Year | Paper_Type | N_Participants | Age | Sex | BMI | Population | Motor_Task | Clothing | Mass_kg | Height_cm | N_Participants_Notes | Clothing_Notes
Salehi_2026 | Evaluation of a markerless motion capture to measure 3D joint kinematics during occupational lifting tasks using mobile devices | Salehi M | 2026 | Validation | 21 | 33.8 ± 4.3 | 11M/10F | NR | Healthy adults | Occupational lifting tasks (3 heights x 2 asymmetry angles x 2 weights) | Not specified | 74.1 ± 13.0 | 171.6 ± 9.1
Verheul_2026 | Markerless motion capture for running: validity and reliability of whole-body, joint, and muscle kinematics | Verheul J | 2026 | Validation | 30 | 34 ± 10 | 19M/11F | NR | Healthy experienced runners (training >=2x/week for >=1 year) | Running (Treadmill, 5 speeds: 8, 10, 12, 14, 16 km/hr) | Retroreflective markers attach | 71.6 ± 11.4 |  |  | Retroreflective markers attached (71 markers full body). No specific clothing mentioned beyond standard lab attire
Natraj_2025 | 3D pose estimation for scalable remote gait kinematics assessment | Natraj S | 2025 | Validation | 225 | 46.06 ± 17.23 (SCI) / 30 ± 5 (Healthy controls) | 171M/54F (SCI) + 6M/5F (Healthy) = 177M/59F total | NR | Pathological (Spinal Cord Injury - tetraplegic/paraplegic) / Healthy controls | Marche | Not specified (ecological vide |  | 1.75 ± 0.10 | 225 SCI + 11 healthy controls = 236 total | Not specified (ecological video recordings from clinical gait analysis)
Zhu_2023 | Monocular 3D Human Pose Markerless Systems for Gait Assessment | Zhu X | 2023 | Methodological | 78 | Not specified (MoVi dataset used) | Not specified (MoVi dataset) |  | Healthy (MoVi dataset participants) | Marche | Not specified |  |  | 78 (from MoVi dataset)
Zago_2020 | 3D Tracking of Human Motion Using Visual Skeletonization and Stereoscopic Vision | Zago M | 2020 | Methodological | 2 | 24 (both participants) | 2M/0F | ~20.4 | Sain | Walking (overground, straight and diagonal directions) | Minimal/Tight (Spandex) |  |  |  | Minimal, close-fitting clothes
Young_2023 | Internet-of-Things-Enabled Markerless Running Gait Assessment from a Single Smartphone Camera | Young F | 2023 | Validation | 31 | 34.5 ± 9.7 | 20M/11F | NR | Healthy experienced runners | Running (Treadmill, 8-14 km/h) | Standardized neutral cushionin | 76.2 ± 4.1 |  |  | Standardized neutral cushioning running shoe provided (Saucony Guide Runner)
Verhoeven_2025 | From marker to markerless: Validating DeepLabCut for 2D sagittal plane gait analysis in adults and newly walking toddlers | Verhoeven M | 2025 | Validation | 31 | 23.6 ± 2.2 years (Adults) / 13.7 ± 2.2 months (Toddlers) | 15M/16F (Adults: 8M/8F; Toddlers: 7M/8F) | NR | Healthy young adults + Typically developing toddlers (newly walking) | Marche (Overground + Treadmill) | Not specified (barefoot walkin | 175.0 ± 9.5 |  | 31 total (16 adults + 15 toddlers) | Not specified (barefoot walking mentioned for adults in overground trials)
Turner_2024 | Validation of OpenCap: A low-cost markerless motion capture system for lower-extremity kinematics during return-to-sport tasks | Turner J.A | 2024 | Validation | 12 | 18.2 ± 3.3 | 6M/6F | NR | Post-ACLR athletes (12-18 months post-surgery) | Jump-landing rebound, Single-leg forward hop, Single-leg lateral-vertical hop | Minimal/Tight (Spandex) | 75.9 |  |  | Black spandex shirts and shorts
Uhlrich_2023 | OpenCap: Human movement dynamics from smartphone videos | Uhlrich S.D | 2023 | Validation | 10 | 27.7 ± 3.8 (lab study, n=10) | 4M/6F (lab); field study sex distribution not specified | 22.8 ± 2.7 | Sain | Walking, Squatting, Sit-to-stand, Drop jump (lab); Walking (field) | Streetwear (Loose) |  |  | 10 (lab validation) + 100 (field study) = 110 total | Mixed (lab: standard lab attire; field: everyday clothing)
Torvinen_2024 | Evaluation of 3D Markerless Motion Capture System Accuracy during Skate Skiing on a Treadmill | Torvinen P | 2024 | Validation | 10 | 22.7 ± 7.5 | 5M/5F |  | Healthy cross-country skiers | Skate Skiing (Treadmill) | Not specified
Takeda_2020 | Artificial Intelligence-Assisted motion capture for medical applications: a comparative study between markerless and passive marker | Takeda I | 2020 | Validation | 1 | Not specified | 1M/0F | NR | Sain | Walking (Treadmill, normal gait + simulated lower limb dysfunction with AFO and crutch) | Minimal/Tight (Spandex)
Stenum_2021 | Two-dimensional video-based analysis of human gait using pose estimation | Stenum J | 2021 | Validation | 31 (analyzed) / 32 (recruited) | Not specified (public dataset) | 22M/10F |  | Sain | Walking (overground, left and right sagittal views) | Not specified
Pagnon_2021 | Pose2Sim: An End-to-End Workflow for 3D Markerless Sports Kinematics—Part 1: Robustness | Pagnon D | 2021 | Methodological | 1 | Not specified | 1M/0F |  | Sain | Walking, Running, Cycling | Minimal/Tight (Spandex) |  |  |  | Minimal (visible in figures - shorts)
Pagnon_2022 | Pose2Sim: An End-to-End Workflow for 3D Markerless Sports Kinematics—Part 2: Accuracy | Pagnon D | 2022 | Validation | 1 | Not specified | 1M/0F | NR | Sain | Walking, Running, Cycling | Minimal/Tight (Spandex) |  |  |  | Minimal (shorts/underwear, visible in Fig 1 and 2)
Needham_2021 | Development, evaluation and application of a novel markerless motion analysis system to understand push-start technique in elite skeleton athletes | Needham L | 2021 | Validation | 12 | Not specified | 7M/5F |  | Sain (Elite Athletes) | Skeleton push-start (Sprinting) | Minimal/Tight (Spandex)
Molteni_2025 | Comparing the Accuracy of Markerless Motion Analysis and Optoelectronic System for Measuring Gait Kinematics of Lower Limb | Molteni L.E | 2025 | Validation | 20 | 11.1 ± 3.4 | 12M/8F | 18.3 ± 3.3 | Pathologique (CP, Para) / Sain | Marche | Not specified
Martis_2024 | A Step Forward Understanding Directional Limitations in Markerless Smartphone-Based Gait Analysis: A Pilot Study | Martiš P | 2024 | Validation | 10 | 29.7 ± 8.6 | 6M/4F | 23.5 ± 2.0 | Sain | Marche (TUG test) | Not specified
Sabo_2022 | Concurrent validity of Zeno instrumented walkway and video-based gait features in adults with Parkinson’s disease | Sabo A | 2022 | Validation | 25 | 65.1 ± 7.3 | 16M/9F |  | Pathologique (Parkinson) | Marche | Streetwear (Loose) |  |  |  | Streetwear - 'clothes they wore to the clinic'
Panconi_2025 | DeepLabCut custom-trained model and the refinement function for gait analysis | Panconi G | 2025 | Validation | 40 | ~23.1 | 21M/19F | NR | Sain | Marche | Streetwear (Loose) | 74.1 ± 10.3
Ota_2021 | Verification of validity of gait analysis systems during treadmill walking and running using human pose tracking algorithm | Ota M | 2021 | Validation | 24 | 26.1 ± 4.6 | 17M/7F | 21.7 | Sain | Marche et Course (Treadmill) | Minimal/Tight (Spandex) |  |  |  | Minimal/Tight - 'appropriately sized light-colored garments that were not loose'
Nguyen_2023 | Deep learning-based estimation of whole-body kinematics from multi-view images | Nguyen K.X | 2023 | Methodological | 7 | Not specified | 7M/0F |  | Sain (Roofers) | Shingle installation (Roofing) | Not specified
Needham_2022 | The development and evaluation of a fully automated markerless motion capture workflow | Needham L | 2022 | Validation | 15 | Not specified | 7M/8F |  | Sain | Running, Walking, Jumping | Not specified
Moro_2020 | Markerless gait analysis in stroke survivors based on computer vision and deep learning: a pilot study | Moro M | 2020 | Feasibility | 10 | 62.75 ± 12.29 | 3M/7F |  | Pathologique (Stroke) | Marche | Not specified
Moro_2022 | Markerless vs. Marker-Based Gait Analysis: A Proof of Concept Study | Moro M | 2022 | Validation | 16 | 27.0 ± 2.0 | 10M/6F |  | Sain | Marche | Not specified
Mercadal_2024 | Exercise quantification from single camera view markerless 3D pose estimation | Mercadal-Baudart C | 2024 | Validation | 12 | 20-30 | 8M/4F |  | Sain | Squats, Jumps, Deadlifts | Mixed
Menychtas_2023 | Gait analysis comparison between manual marking, 2D pose estimation algorithms, and 3D marker-based system | Menychtas D | 2023 | Validation | 17 | 69 ± 5 | Not specified |  | Sain | Marche (Treadmill) | Not specified
Mehdizadeh_2021 | Concurrent validity of human pose tracking in video for measuring gait parameters in older adults: a preliminary analysis with multiple trackers, viewing angles, and walking directions | Mehdizadeh S | 2021 | Validation | 11 | 85.2 ± 5.6 | 2M/9F |  | Sain (Older Adults) | Marche | Not specified
Lim_2022 | Markerless Motion Capture System Based on Webcams Using OpenPose | Lim W | 2022 | Feasibility | 1 | 29 | 1M/0F | NR | Sain | Squat | Not specified
Embry_2022 | Video-Based Pose Estimation for Gait Analysis in Stroke Survivors during Clinical Assessments: A Proof-of-Concept Study | Embry K | 2022 | Validation | 8 | 57.9 ± 7.8 | 5M/3F |  | Pathologique (Stroke) | Marche | Mixed |  |  |  | Mixed/Varied - participants could wear variety of clothing and leg braces
Lima_2024 | Validity and reliability of trunk and lower-limb kinematics during squatting, hopping, jumping and side-stepping using OpenCap markerless motion capture application | Lima Y.L | 2024 | Validation | 19 | 26.3 ± 4.5 | 10M/9F | 23.3 ± 2.6 | Sain | Squatting, hopping, jumping, side-stepping | Minimal/Tight (Spandex) |  |  |  | Own shorts, sports bra (for women), athletic shoes
Li_2024 | GaitFormer: Leveraging dual-stream spatial-temporal Vision Transformer via a single low-cost RGB camera for clinical gait analysis | Li J | 2024 | Methodological | 80 | 24 ± 1 | Not specified | NR | Sain (Simulating pathology) | Marche | Not specified |  | 1.75 ± 0.05
Ino_2023 | Validity of AI-Based Gait Analysis for Simultaneous Measurement of Bilateral Lower Limb Kinematics Using a Single Video Camera | Ino T | 2023 | Validation | 21 | 20.7 ± 1.0 | 10M/11F | 21.6 ± 2.6 | Sain | Marche | Not specified
Lambricht_2025a | Impact of Running Clothes on Accuracy of Smartphone-Based 2D Joint Kinematic Assessment During Treadmill Running Using OpenPifPaf | Lambricht N | 2025 | Validation | 30 | 23.1 ± 4.0 | 15M/15F | NR | Sain | Running (Treadmill) | Minimal/Tight (Spandex) | 66.6 ± 9.1 |  |  | Minimal (Lab) vs Streetwear (Running clothes)
Huang_2024 | Comparison of kinematics and joint moments calculations for lower limbs during gait using markerless and marker-based motion capture | Huang T | 2024 | Validation | 16 | Males: 30.3±9.3, Females: 23.6±1.7 | 11M/5F | NR | Sain | Marche | Not specified | 76.6
Lambricht_2025b | Quantifying performance and joint kinematics in functional tasks crucial for anterior cruciate ligament rehabilitation using smartphone video and pose detection | Lambricht N | 2025 | Validation | 12 | 24 ± 3.4 | 6M/6F | NR | Sain | SEBT, Single/Vertical Hop | Minimal/Tight (Spandex) | 63.7 ± 8.4 |  |  | Minimal - 'only a dark boxer short', barefoot
Lagomarsino_2025 | Video-based markerless assessment of bilateral upper limb motor activity following cervical spinal cord injury | Lagomarsino B | 2025 | Validation | 24 | 34.7 ± 13.5 | 16M/8F |  | Pathologique (cSCI) / Sain | Arm Stabilization Test | Not specified
Kumar_2022 | Prediction of lower limb kinematics from vision-based system using deep learning approaches | Kumar K.S | 2022 | Methodological | 33 | Not specified | 25M/8F |  | Sain | Squat, Marche (Treadmill) | Not specified
Kakavand_2025 | Comparison of kinematics and kinetics between OpenCap and a marker-based motion capture system in cycling | Kakavand R | 2025 | Validation | 10 | 29.5 ± 3.3 | 5M/5F | NR | Sain | Cycling | Minimal/Tight (Spandex) | 70.6 ± 11.8 | 176.0 ± 0.08 |  | Tight, minimal clothing + cycling cleats (Santic S3-KMS20025)
John_2024 | Accuracy of Video-Based Gait Analysis Using Pose Estimation During Treadmill Walking Versus Overground Walking in Persons After Stroke | John K | 2024 | Validation | 30 | 63 ± 9 | 21M/9F | NR | Pathologique (Stroke) | Marche (Overground + Treadmill) | Not specified | 85.1 ± 19.2 | 173.0 ± 0.11
Ino_2024 | Validity and Reliability of OpenPose-Based Motion Analysis in Measuring Knee Valgus during Drop Vertical Jump Test | Ino T | 2024 | Validation | 21 | 20.7 ± 1.0 | 10M/11F | 21.6 ± 2.6 | Sain | Drop Vertical Jump (DVJ) | Not specified
Horsak_2025 | Validity and reliability of monocular 3D markerless gait analysis in simulated pathological gait: A comparative study with OpenCap | Horsak B | 2025 | Validation | 19 | 30 ± 9 | 9M/10F | 23.2 ± 3.3 | Sain (Simulating pathology) | Marche | Minimal/Tight (Spandex) |  |  |  | Minimal / Streetwear
Horsak_2024 | Inter-trial variability is higher in 3D markerless compared to marker-based motion capture: Implications for data post-processing and analysis | Horsak B | 2024 | Reliability | 21 | 30.2 ± 8.5 | 9M/12F (of 21 recruited) | 23.2 ± 3.4 | Sain (Simulating pathology) | Marche | Minimal/Tight (Spandex) |  |  | 21 recruited (9M/12F), 2 excluded, 18-19 analyzed depending on condition
Horsak_2023 | Concurrent validity of smartphone-based markerless motion capturing to quantify lower-limb joint kinematics in healthy and pathological gait | Horsak B | 2023 | Validation | 21 | 30.2 ± 8.5 | 9M/12F (of 21 recruited) | 23.2 ± 3.4 | Sain (Simulating pathology) | Marche | Minimal/Tight (Spandex)
Harrison_2025 | Evaluation of the Ergomechanic Markerless Motion Capture System for Lower Body Kinematics During Standing, Squatting, and Walking | Harrison S.M | 2025 | Validation | 18 | 23.3 ± 2.6 | 13M/5F | NR | Sain | Marche, Squat, Standing | Not specified | 67.6 ± 10.5 | 171.9 ± 7.9
Emmerson_2025 | Evaluation of a markerless motion capture system for measuring mechanical work during tennis strokes | Emmerson J | 2025 | Validation | 16 | Not specified | 6M/10F | NR | Sain (Tennis players) | Tennis strokes | Not specified
Dunn_2023 | Application of video frame interpolation to markerless, single-camera gait analysis | Dunn M | 2023 | Methodological | 32 | Not specified | 22M/10F (of 32 recruited; sex breakdown of 28 analyzed not specified) |  | Sain | Marche | Not specified |  |  | 32 recruited, 28 analyzed (4 excluded for labelling inconsistencies)
Drazan_2021 | Moving outside the lab: Markerless motion capture accurately quantifies sagittal plane kinematics during the vertical jump | Drazan J.F | 2021 | Validation | 15 | Not specified | Not specified |  | Sain | Saut vertical | Not specified
DHaene_2024 | Validation of a 3D Markerless Motion Capture Tool Using Multiple Pose and Depth Estimations for Quantitative Gait Analysis | D'Haene M | 2024 | Validation | 15 | 27.8 ± 4.5 | 7M/8F | NR | Sain | Marche (Treadmill) | Minimal/Tight (Spandex) | 70 ± 12 | 175.0 ± 10
deBorba_2025 | Effect of Walking Speed on the Reliability of a Smartphone-Based Markerless Gait Analysis System | de Borba E.F | 2025 | Validation | 15 | 32.3 ± 8.3 | 8M/7F | NR | Sain | Marche | Minimal/Tight (Spandex) | 71.1 ± 21.6
DAntonio_2020 | A markerless system for gait analysis based on OpenPose library | D'Antonio E | 2020 | Feasibility | 1 | 28 | 1M/0F |  | Sain | Marche (Treadmill) | Not specified
DAntonio_2021 | Validation of a 3D Markerless System for Gait Analysis based on OpenPose and Two RGB Webcams | D'Antonio E | 2021 | Validation | 6 | 27.6 ± 3.4 | 5M/1F |  | Sain | Marche et Course (Treadmill) | Not specified
DAntonio_2020b | Characterization of a low-cost markerless system for 3D gait analysis | D'Antonio E | 2020 | Feasibility | 1 | 28 | 1M/0F |  | Sain | Marche (Treadmill) | Not specified
CarrascoPlaza_2022 | Evaluation of Human Pose Estimation in 3D with Monocular Camera for Clinical Application | Carrasco-Plaza J | 2022 | Methodological | 11 | Not specified | 6M/5F |  | Sain | Activities of daily living (Human3.6M dataset) | Not specified
Boldo_2024 | On the reliability of single-camera markerless systems for overground gait monitoring | Boldo M | 2024 | Methodological | 8 | 28.0 ± 3.7 | 6M/2F | NR | Sain | Marche | Not specified | 68.1 ± 14.9
Bittner_2023 | Towards Single Camera Human 3D-Kinematics | Bittner M | 2023 | Methodological | 82 | Not specified | Mixed |  | Sain | Activities of daily living (BML-Movi dataset) | Not specified
Bertozzi_2025 | Concurrent validity of IMU and phone-based markerless systems for lower-limb kinematics during cognitively-challenging landing tasks | Bertozzi F | 2025 | Validation | 30 | 23.9 ± 2.8 | 17M/13F | NR | Sain (Competitive athletes) | Jump-land-jump task | Not specified | 66.1 ± 9.4
Barzyk_2024a | AI-smartphone markerless motion capturing of hip, knee, and ankle joint kinematics during countermovement jumps | Barzyk P | 2024 | Validation | 11 | 28.4 ± 9.2 | 10M/1F | NR | Sain | Countermovement jumps (CMJ) | Minimal/Tight (Spandex) | 69.8 ± 12.0
Barzyk_2024b | Steps to Facilitate the Use of Clinical Gait Analysis in Stroke Patients: The Validation of a Single 2D RGB Smartphone Video-Based System for Gait Analysis | Barzyk P | 2024 | Validation | 8 | 59.4 ± 13.7 | 6M/2F | NR | Pathologique (Stroke) | Marche | Not specified | 80.6 ± 18.2
Bak_2025 | Validity of Markerless Motion Capture System and Its Correlation with Physical Characteristics for Hip Range of Motion Measurement: A Pilot Study | Bak S.Y | 2025 | Validation | 11 | 23.82 ± 3.25 | 4M/7F | NR | Sain | Hip ROM exercises | Minimal/Tight (Spandex) | 64.9 ± 13.9
Bae_2024 | Concurrent validity and test reliability of the deep learning markerless motion capture system during the overhead squat | Bae K | 2024 | Validation | 31 | 32.20 ± 5.90 | 22M/9F | NR | Sain | Overhead squat | Minimal/Tight (Spandex) | 66.8 ± 12.8
Azhand_2021 | Algorithm based on one monocular video delivers highly valid and reliable gait parameters | Azhand A | 2021 | Validation | 43 | 73.9 ± 6.0 | 18M/15F | NR | Sain (Older Adults) | Marche | Not specified | 73.8 ± 15.6
Andreoni_2024 | Comparison of the Accuracy of Markerless Motion Analysis and Optoelectronic System for Measuring Lower Limb Gait Kinematics | Andreoni G | 2024 | Validation | 4 | 28.4 ± 3.7 | 2M/2F | NR | Sain | Marche | Not specified | 60.5 ± 6.9
Anderson_2025 | Validation of markerless video-based gait analysis using pose estimation in toddlers with and without neurodevelopmental disorders | Anderson J.T | 2025 | Validation | 112 | 2.5 ± 0.6 | 74M/38F |  | Pathologique (NDDs) / Sain (Toddlers) | Marche | Not specified
Aleksic_2024 | Validation of Automated Countermovement Vertical Jump Analysis: Markerless Pose Estimation vs. 3D Marker-Based Motion Capture System | Aleksic J | 2024 | Validation | 12 | 25.6 ± 3.4 | 7M/5F | NR | Sain | Countermovement Jump (CMJ) | Minimal/Tight (Spandex) | 76.0 ± 16.0
Ahn_2024 | Simultaneous Validity and Intra-Test Reliability of Joint Angle Measurement through Novel Multi-RGB Sensor-Based Three-Joint-Continuous-Motion Analysis: A Pilot Study | Ahn J | 2024 | Validation | 5 | 25.4 ± 2.0 | 5M/0F | NR | Sain | Squat, Gait (Treadmill) | Minimal/Tight (Spandex) | 68.9 ± 6.8 | 174.4 ± 5.0
Wang_2024 | Markerless gait analysis through a single camera and computer vision | Wang H | 2024 | Validation | 14 | 27.0 ± 4.7 | 7M/7F | NR | Sain | Marche | Not specified | 71.9 ± 14.4
Verheul_2024 | Jumping towards field-based ground reaction force estimation and assessment with OpenCap | Verheul J | 2024 | Validation | 15 | 22.4 ± 3.6 | 9M/6F | NR | Sain (Recreational athletes) | Vertical Jumps (CMJ, Squat, Drop) | Not specified | 77.9 ± 12.6
Wang_2025 | Evaluation of a smartphone-based markerless system to measure lower-limb kinematics in patients with knee osteoarthritis | Wang J | 2025 | Validation | 83 | OA: 64.5±6.5 / HC: 55.2±3.3 | 67M/16F | NR | Pathologique (Knee OA) / Sain | Marche | Minimal/Tight (Spandex) | 62.7
Bousigues_2025 | The effects of markerless inconsistencies are at least as large as the effects of the marker-based soft tissue artefact | Bousigues S | 2025 | Methodological | 16 | 25 ± 3 | 9M/7F | NR | Sain | Marche, Sit-to-Stand, CMJ | Minimal/Tight (Spandex) | 68.4 ± 13.5
Ueno_2024 | Calibrationless monocular vision musculoskeletal simulation during gait | Ueno R | 2024 | Validation | 10 | 27.7 ± 3.8 | 4M/6F | NR | Sain | Marche | Not specified | 69.2 ± 11.6
Yagi_2020 | Gait Measurement at Home Using A Single RGB Camera | Yagi K | 2020 | Feasibility | 29 | 21-54 | Not specified |  | Sain | Marche | Not specified
Tahara_2025 | Predicting Walkway Spatiotemporal Parameters Using a Markerless, Pixel-Based Machine Learning Approach | Tahara A.K | 2025 | Validation | 17 | 18-40 | 9M/8F |  | Sain | Marche | Not specified
Francia_2026 | Dual-Camera System for AI Markerless 3D Lower Limb Motion Analysis | Francia C | 2026 | Validation | 10 | 25 ± 2 | 5M/5F | NR | Sain | Timed-Up and Go test (TUG) | Not specified | 63.3 ± 6.8
Koleini_2025 | BioPose: Biomechanically-accurate 3D Pose Estimation from Monocular Videos | Koleini F | 2025 | Methodological | 82 | Not specified | Mixed |  | Sain | Activities of daily living (BML-Movi / OpenCap datasets) | Not specified
Yamamoto_2021 | Accuracy of Temporo-Spatial and Lower Limb Joint Kinematics Parameters Using OpenPose for Various Gait Patterns With Orthosis | Yamamoto M | 2021 | Validation | 11 | 23.2 ± 1.4 | 11M/0F | NR | Sain (Walking with AFO) | Marche (Treadmill) | Not specified | 61.2 ± 9.7
Hulleck_2024 | BlazePose-Seq2Seq: Leveraging Regular RGB Cameras for Robust Gait Assessment | Hulleck A.A | 2024 | Methodological | 10 | Not specified | 10M/0F | NR | Sain | Marche | Not specified | 69.5 ± 7.4
Templin_2024 | Evaluation of drop vertical jump kinematics and kinetics using 3D markerless motion capture in a large cohort | Templin T | 2024 | Validation | 127 | 15.6 ± 1.4 | 0M/127F | NR | Sain (Adolescent female athletes) | Drop Vertical Jump (DVJ) | Not specified | 64.4 ± 12.3
Wang_2023 | A webcam-based machine learning approach for three-dimensional range of motion evaluation | Wang X.M | 2023 | Validation | 25 | Not specified | 12M/13F |  | Sain | ROM exercises (Neck, spine, limbs) | Motion capture suit
Wade_2023 | Examination of 2D frontal and sagittal markerless motion capture: Implications for markerless applications | Wade L | 2023 | Validation | 15 | 26 ± 5 | 7M/8F | NR | Sain | Marche | Not specified | 73 ± 14
Kim_2025 | Assessment of temporospatial and kinematic gait parameters using human pose estimation in patients with Parkinson's disease: A comparison between near-frontal and lateral views | Kim J | 2025 | Validation | 24 | 66.9 ± 7.9 | 12M/12F |  | Pathologique (Parkinson) | Marche (Treadmill + Overground) | Not specified
Ho_2024 | Pathological Gait Analysis With an Open-Source Cloud-Enabled Platform Empowered by Semi-Supervised Learning-PathoOpenGait | Ho M.Y | 2024 | Validation | 23 | 65 (approx) | Not specified |  | Pathologique (Parkinson, MSA) / Sain | Marche (8-meter walk) | Not specified
He_2024 | A novel multi-level 3D pose estimation framework for gait detection of Parkinson's disease using monocular video | He R | 2024 | Validation | 59 | 62.7 ± 10.6 (PD) / 55.4 ± 5.8 (HC) | Not specified | NR | Pathologique (Parkinson) / Sain | Marche | Not specified | 61.5 ± 6.5
Balci_2026 | Reliability assessment of markerless technologies in biomechanical motion analysis. a performance comparison | Balci I.C | 2026 | Methodological | 1 | 20 | 0M/1F | NR | Sain | Isolated joint movements (upper & lower limb) | Minimal/Tight (Spandex) | 53
Hii_2023 | Automated Gait Analysis Based on a Marker-Free Pose Estimation Model | Hii C.S.T | 2023 | Validation | 31 | Not specified | 22M/10F |  | Sain | Marche | Not specified
Chen_2025 | Gait anomaly detection based on video-derived 3D pose estimation | Chen L | 2025 | Validation | 10 (GPOSE) + 230 (TUG dataset) | Not specified (Elderly focus) | Not specified |  | Sain (Simulating abnormal gait) | Marche (TUG test) | Not specified
Fukushima_2024 | The potential of human pose estimation for motion capture in sports: a validation study | Fukushima T | 2024 | Validation | 5 | 30.2 ± 6.6 | 5M/0F | NR | Sain | Athletic and sports movements | Not specified | 74.2 ± 9.1
Ge_2025 | Reliability and validity of OpenPose for measuring HKA angle in dynamic walking videos in patients with knee osteoarthritis | Ge F | 2025 | Validation | 50 | 64.6 ± 5.2 | 18M/32F |  | Pathologique (Knee OA) | Marche | Not specified
Asaeda_2024 | Reliability and validity of knee valgus angle calculation at single-leg drop landing by posture estimation using machine learning | Asaeda M | 2024 | Validation | 15 | 19.9 ± 1.0 | 15M/0F |  | Sain | Single-leg drop landing | Minimal/Tight (Spandex)
Liang_2022 | The reliability and validity of gait analysis system using 3D markerless pose estimation algorithms | Liang S | 2022 | Validation | 30 | 56.6 ± 2.53 (Elderly) / 27.27 ± 4.31 (Young) | Not specified |  | Sain | Marche (Treadmill) | Minimal/Tight (Spandex) |  |  |  | Minimal/Tight
Svetek_2025 | Validation of OpenCap on lower extremity kinematics during functional tasks | Svetek A | 2025 | Validation | 20 | 21.3 ± 1.5 | 2M/18F | NR | Sain (Ice hockey players) | Marche, Course, Squat, Sauts (CMJ, Drop landing) | Minimal/Tight (Spandex) | 71.08 ± 7.42 | 171.0 ± 8.0
Hu_2025 | Dual-Camera Markerless Motion Capture System for Precise Lower-Limb Kinematic Analysis in Osteoarthritis | Hu B | 2025 | Validation | 152 | 62.3 ± 11.5 | 57M/95F | NR | Pathologique (Knee OA / Hip OA) | Marche | Not specified | 63.0 ± 10.4 | 158.4 ± 8.9
Haberkamp_2022 | Validity of an artificial intelligence, human pose estimation model for measuring single-leg squat kinematics | Haberkamp L.D | 2022 | Validation | 22 | 16.5 ± 1.2 | 8M/14F |  | Sain | Single-leg squat | Not specified
Vafadar_2022 | Assessment of a novel deep learning-based marker-less motion capture system for gait study | Vafadar S | 2022 | Validation | 41 | 8-44 | 23M/18F | 21.0 (Training) / 21.9 (Test) | Pathologique / Sain | Marche | Not specified
Washabaugh_2022 | Comparing the accuracy of open-source pose estimation methods for measuring gait kinematics | Washabaugh E.P | 2022 | Validation | 31 | Not specified | 22M/10F |  | Sain | Marche | Not specified
Yeom_2026 | Validation of OpenCap, a low-cost markerless motion capture system, during the overhead squat | Yeom S | 2026 | Validation | 22 | 25.5 (SD:2.7) | 22M | NR | Sain | Overhead Squat | Minimal/Tight (Spandex) | 75.7 ± 9.0
Gao_2025 | Identifying gait deviation in children with cerebral palsy using BlazePose | Gao J | 2025 | Validation | 10 | 15 ± 3 | 7M/3F |  | Pathologique (CP) | Marche (Overground) | Not specified
Kumtekar_2024 | Mobile application for upper extremity range of motion measurement using MediaPipe and AWS cloud | Kumtekar R | 2024 | Validation | 15 | 28.3 (SD:5.1) | 9M/6F |  | Sain | ROM Membre Supérieur | Not specified
Nishizawa_2024 | OpenPose-based gait analysis for hemiplegia rehabilitation | Nishizawa K | 2024 | Validation | 20 | Healthy: 23.3 ± 0.94; L-Hemi: 58.0 ± 13.5; R-Hemi: 55.4 ± 9.5 | Not specified |  | Pathologique (Hémiplégie) / Sain | Marche (Overground) | Minimal/Tight (Spandex)
Todesca_2025 | YOLO11x-pose for cyclist posture estimation | Todesca M | 2025 | Validation | 2 | 24 and 31 | 2M/0F | NR | Sain (Cyclistes) | Cyclisme (Stationnaire) | Minimal/Tight (Spandex) | 78.5
Lin_2024 | Estimating 3D kinematics from video with biomechanical model and synthetic data (CVPRW) | Lin D | 2024 | Methodological | 56 ODAH (synthetic) + 10 OpenCap + 90 BMLmovi (eval datasets) |  |  |  | N/A (Benchmark) | Marche, Saut | Not specified
Zhang_2024 | Multi-view markerless kinematics with GNN and OpenSim (IEEE PRML) | Zhang H | 2024 | Methodological | 0 |  |  |  | N/A (Benchmark) | Marche | Not specified
Hii_2023b | Frontal plane gait analysis using OpenPose, YOLOv7, and MediaPipe | Hii C.S.T | 2023 | Validation | 31 | 24.3 (SD:3.1) | 22M/10F |  | Sain | Marche (Overground) | Minimal/Tight (Spandex)
Pappas_2024 | MediaPipe + ZoeDepth for Parkinson's gait assessment in frontal plane | Pappas I | 2024 | Validation | 62 (28 GPJATK + 17 healthy + 17 PD) | Validation: not specified; HC: 64 ± 9; PD: 62 ± 8 | 38M/24F |  | Pathologique (Parkinson) / Sain | Marche (Overground) | Not specified
Irfan_2025 | Reliable gait measurements using smartphone vision sensor with BlazePose and Bezier curves | Irfan M | 2025 | Validation | 31 healthy (GPJATK Kwolek 2019 dataset) | Patients: 30-65; healthy 20-60 | 21M/10F | NR | Sain | Marche (Overground, left-right and right-left) | Own clothing (no markers) | NR | NR | Validation cohort: 31 healthy from GPJATK (Kwolek 2019) dataset, 21M/10F. KIMS patients used only qualitatively for real-environment testing (not in RMSE table)
Liang_2023 | 3D gait analysis for the elderly mobility based on multiple RGB cameras | Liang Y | 2023 | Validation | 19 | Elderly: 67.2; Young: 24.1 | 8M/8F (per group) |  | Sain (Older Adults) / Sain | Marche (Overground) | Not specified
Cerciola_2024 | Comparative analysis of markerless motion capture systems for measuring human kinematics | Cerciola L | 2024 | Validation | 10 | 25 (SD:4) | 10M |  | Sain | Marche et Course (Treadmill) | Minimal/Tight (Spandex)
Zhang_2025 | Quantitative TUG test with markerless MoCap and ML for balance and fall risk | Zhang Y | 2025 | Validation | 70 | Able-bodied: 45.8 ± 20.0; Post-stroke: 65.6 ± 5.6; Knee OA: 62.3 ± 9.2 | 22M/48F | NR | Sain (Older Adults) | Marche (TUG test) | Not specified | 60.6 ± 12.5
Kidzinski_2020 | Deep neural networks enable quantitative movement analysis using single-camera videos | Kidziński Ł | 2020 | Validation | 1026 unique patients (1792 videos) | Cerebral palsy patients (pediatric) | NR | NR | Pathologique (Cerebral palsy) | Marche (Overground) | Not specified | NR | NR | Train: 1792 videos / 1026 patients; CP clinical gait analysis dataset (Gillette Children's Specialty Healthcare)
Hwang_2026 | Sit-to-Stand Power From 2D Pose Estimation as an Indicator of Muscle Strength in Older Adults | Hwang W | 2026 | Validation | 129 (validation subgroup n=20) | 71.8 ± NR (community-dwelling older adults ≥65 yr) | NR | NR | Sain (elderly community-dwelling) | Sit-to-stand (30-s STS) | Validation: barefoot, swim cap, shorts | NR | NR | 129 full cohort, 20 validation subgroup with concurrent MoCap+force plate
Shah_2025 | Biomechanics of Knee Kinematics During Sit-Ups Using MoCap Technology | Shah E.-E.-Z | 2025 | Validation | 15 | Adults (varied) | NR | NR (heights 1.5-1.8m, mass 50-90kg) | Sain (South Asian) | Sit-ups | Reflective markers required (Qualisys) | 50-90 | 150-180 | South Asian cohort, varied heights and body masses
Rode_2025 | Assessment of monocular human pose estimation models for clinical movement analysis | Rode D | 2025 | Validation | 25 | 26 [20-33] | 11M/14F | NR (Height: M 1818[1730-1920] / F 1649[1580-1740] mm; Weight: M 79[65-94] / F 60[47-80] kg) | Sain | Multiple exercises: squat, hip abduction, bird dog, single leg shoulder bridge, shoulder rotation + elbow flexion, knee extension, quadriceps stretch | Markers on skin/tight clothing (Plug-in Gait) | M 79 / F 60 (mean) | M 181.8 / F 164.9 | Physio2.2M dataset, 2.2 million frames, 3 sets × 5 reps per exercise
Irfan_2025b | Kinematic Gait Analysis Using Markerless System to Determine Joint Angles | Irfan M | 2025 | Validation | 10 patients (6M/4F) + 20 healthy from GPJATK dataset | Patients: 30-65; healthy 20-60 | Patients 6M/4F; Healthy 8F/12M | NR | Sain + Pathologique (elderly with stick) | Marche (Overground, left-right and right-left) | Own clothing (no markers) | NR | NR | 10 patients KIMS Hyderabad (4 with walking stick) + 20 healthy subjects from public dataset [Kwolek et al. 2019]
Talaa_2025 | Markerless Knee Angle Measurement: Advancements in Gait Analysis Using Computer Vision | Talaa S | 2025 | Validation | 11 (subset of GPJATK n=32) | Young adults | 5F/6M | NR | Sain | Marche (straight line 6.5 m, 28 walking cycles) | Not specified | NR | NR | Subset of GPJATK dataset (Kwolek 2019)
Ang_2024 | Markerless Human Motion Analysis for Telerehabilitation: A Case Study on Squat | Ang Y.H | 2024 | Validation | 1 (case study) | NR | NR | NR | Sain | Deep Squat (3 speeds: fast, normal, slow) | Not specified | NR | NR | Case study on squatting for telerehabilitation context. Springer chapter 2024 (ASTI book series 'Advancing Smart Cities')
Hou_2026 | A Marker-Free Motion Capture System Built on Unsynchronized Cameras | Hou H | 2026 | Methodological | Multi-subject dataset (n exact NR ; 8:1:1 train/val/test split) | NR | NR |  | Sain | Walking + Running (5 speeds) + In-place (squat, jump, stretch) | Tight (marker-based reference) |  |  | Springer LNAI ICIRA 2025 (publication 2026). Peking University. Custom dataset captured indoors
Athama_2025 | Wireless Single-Camera Markerless Motion Capture System for Healthcare Applications | Athama A | 2025 | Methodological | MoVi dataset (90 subjects, public) | NR (MoVi dataset varied) | NR |  | Sain (MoVi public dataset) | Various (MoVi: walking, multiple actions) | Not specified |  |  | IEEE HPCC 2025. Brunel University London + Univ of Essex. Focus = wireless transmission impact on markerless MoCap. Uses MoVi public dataset for training/validation
Tsintzira_2025 | Low-Cost Markerless Gait Analysis Using a Minimal Human Model | Tsintzira K | 2025 | Feasibility | 1 | 30 | 0M/1F |  | Sain | Marche (Treadmill walking) | Not specified | 58 | 164 | IEEE BIBE 2025. National Technical University of Athens. Proof of concept on 1 subject. Comparison vs Fukuchi 2018 public dataset

## 2. Video Input
Study_ID | N_Cameras | Positioning | Camera_Type | Resolution | Frequency_Hz | Calibration | Synchronization | Environment
Salehi_2026 | 3 | ±45° front-diagonal (+ front) | Smartphone | 720p | 60 | Checkerboard | Software (Cross-correl) | Laboratory
Verheul_2026 | 3 | Around treadmill | Tablet (iPad) | 1080p | 240 | Checkerboard | Hardware | Laboratory
Natraj_2025 | 1 | Not specified (1.5m height) | Webcam | 1080p | 50 | None (Monocular) |  | Clinic
Zhu_2023 | 1 | Sagittal (Perpendicular) | Industrial Camera | 800x600 | 30 | None (Monocular) |  | Laboratory
Zago_2020 | 2 | Frontal + Sagittal | Webcam | 1080p | 30 | Checkerboard | Software (Cross-correl) | Laboratory
Young_2023 | 1 | Sagittal (Side) | Smartphone | 1080p | 240 | None (Monocular) |  | Laboratory
Verhoeven_2025 | 1 | Sagittal | Industrial Camera | 720p | 100 | None (Monocular) | Hardware | Laboratory
Turner_2024 | 2 | ±45° front-diagonal | Smartphone | 720p | 60 | Checkerboard | Software (Cross-correl) | Laboratory
Uhlrich_2023 | 2 | ±45° front-diagonal | Smartphone | 720p | 60 | Checkerboard | Software (Cross-correl) | Laboratory
Torvinen_2024 | 2 | Sagittal + Diagonal (Right side) | Industrial Camera | 720p | 100 | Manual | Manual (Event) | Laboratory
Takeda_2020 | 1 | Sagittal | Hybrid/DSLR | Not specified | 60 | None (Monocular) |  | Laboratory
Stenum_2021 | 2 | Bilateral Sagittal | Industrial Camera | Not specified | 30 | None (Monocular) |  | Laboratory
Pagnon_2021 | 8 | 360° Multi-view (Virtual) | Industrial Camera | Not specified | 30 | Checkerboard | Hardware | Laboratory
Pagnon_2022 | 8 | 360° Multi-view | Industrial Camera | Not specified | 30 | Checkerboard | Hardware | Laboratory
Needham_2021 | 9 | Around push track | Industrial Camera | 1080p | 200 | Binary dot matrix + bundle adjustment | Hardware | Laboratory
Molteni_2025 | 2 | Frontal + Sagittal | Webcam | 480p | 25 | None (Monocular) | Software | Laboratory
Martis_2024 | 2 | Frontal/Diagonal | Smartphone | 720p | 60 | Checkerboard | Software | Laboratory
Sabo_2022 | 1 | Frontal (Towards/Away) | Webcam | 480p | 30 | None (Monocular) | Software | Clinic
Panconi_2025 | 1 | Sagittal | Industrial Camera | 480p | 25 | None (Monocular) | Hardware | Laboratory
Ota_2021 | 2 | Posterior + Sagittal (Right side) | Digital Video Camera | Not specified | 60 | None (Monocular) | Not specified | Laboratory
Nguyen_2023 | 3 | Multi-view | Industrial Camera | 1280×720 | 100 | Checkerboard | Hardware | Laboratory
Needham_2022 | 9 | Around capture volume | Industrial Camera | 1080p | 200 | Checkerboard | Hardware | Laboratory
Moro_2020 | 2 | Sagittal + Frontal | Industrial Camera | 480p | 25 | None (Monocular) | Hardware | Laboratory
Moro_2022 | 3 | Multi-view | Industrial Camera | 1292x964 | 30 | Checkerboard | Hardware | Laboratory
Mercadal_2024 | 1 | Sagittal | Smartphone | 1080p | 60 | None (Monocular) |  | Laboratory
Menychtas_2023 | 1 | Sagittal | Industrial Camera | Not specified | 100 | None (Monocular) | Hardware | Laboratory
Mehdizadeh_2021 | 2 | Front + Back | Smartphone | 1080p | 30 | None (Monocular) |  | Clinic
Lim_2022 | 4 | Multi-view | Webcam | Not specified | 25 | Checkerboard | Software (auto-alignment to Vicon) | Laboratory
Embry_2022 | 1 | Sagittal (Below waist) | Digital RGB camera (tripod) | 720p (1280x720) | 30 | None (Monocular) |  | Clinic
Lima_2024 | 2 | ±45° front-diagonal | Smartphone | 1080p | 60 | Checkerboard | Software (Cross-correl) | Laboratory
Li_2024 | 1 | Sagittal | Smartphone | 1080p | 60 | None (Monocular) |  | Laboratory
Ino_2023 | 1 | Sagittal | Industrial Camera | 720p | 120 | T-wand (Vicon) |  | Laboratory
Lambricht_2025a | 1 | Sagittal | Smartphone | 720p | 60 | None (Monocular) |  | Laboratory
Huang_2024 | 8 | Multi-view | Industrial Camera (Z-CAM E2) | Not specified | 120 | Wand | Hardware | Laboratory
Lambricht_2025b | 1 | Sagittal | Smartphone (iPhone 10) | 720p | 60 | None (Monocular) |  | Laboratory
Lagomarsino_2025 | 1 | Frontal | Industrial Camera (BTS VIXTA) | 480p | 25 | None (Monocular) |  | Clinic
Kumar_2022 | 1 | Multi-view | Action Cam | 1080p | 60 | Checkerboard | Hardware | Laboratory
Kakavand_2025 | 4 | ±45° front-diagonal | Smartphone | Not specified | 60 | Checkerboard | Software | Laboratory
John_2024 | 2 | Frontal + Sagittal (Separate) | Tablet (Samsung Galaxy A7) | 1920x1080 | 30 | None (Monocular) | Software (Cross-correl) | Clinic
Ino_2024 | 1 | Frontal | Industrial Camera (Vicon Bonita Video 720C) | 720p | 120 | Wand | Hardware | Laboratory
Horsak_2025 | 1 | Random perspective | Smartphone | Not specified | Not specified | OpenCap calibration (checkerboard) | Software (Cross-correl) | Laboratory
Horsak_2024 | 2 | ±35° front-diagonal | Smartphone | 720x1280 | 60 | Checkerboard | Software | Laboratory
Horsak_2023 | 2 | ±35° front-diagonal | Smartphone | 720x1280 | 60 | Checkerboard | Software | Laboratory
Harrison_2025 | 4 | Around capture volume | Webcam | Not specified | 60 | Checkerboard | Hardware | Laboratory
Emmerson_2025 | 8 | Around capture volume | Industrial Camera | 1080p | 200 | Binary dot matrix (bundle adjustment) | Hardware | Laboratory
Dunn_2023 | 1 | Sagittal (from multi-camera dataset) | Industrial Camera | 540p | 25 | None (Monocular) | Hardware (Vicon MX Giganet) | Laboratory
Drazan_2021 | 2 | Frontal + Sagittal (Separate) | Industrial Camera | 800×600 | 30 | None (Monocular) |  | Laboratory
DHaene_2024 | 3 | Frontal + Sagittal + Back | RGB-D Camera (StereoLabs ZED 2) | 1080p | 30 | ZED360 (StereoLabs) | Hardware | Laboratory
deBorba_2025 | 2 | ±35° front-diagonal | Smartphone | 1080p | 120 | Checkerboard | Software (Cross-correl) | Laboratory
DAntonio_2020 | 2 | Sagittal | Webcam | Not specified | 60 | Checkerboard | Software (Cross-correl) | Laboratory
DAntonio_2021 | 2 | Multi-view | Webcam | 1280×720 | 60 | Checkerboard | Software (Cross-correl) | Laboratory
DAntonio_2020b | 2 | Multi-view | Webcam | Not specified | 60 | Checkerboard | Software (Cross-correl) | Laboratory
CarrascoPlaza_2022 | 1 | Monocular (from Multi-view Dataset) | Dataset (Human3.6M) | 1080p | 50 | Checkerboard |  | Laboratory
Boldo_2024 | 1 | Sagittal | Industrial Camera (RealSense RGB-D) | 480p | 60 | None (Monocular) | Hardware (LED trigger) | Laboratory
Bittner_2023 | 1 | Frontal + Sagittal (Separate) | Industrial Camera (Point Grey) | Not specified | Not specified | None (Monocular) |  | Laboratory
Bertozzi_2025 | 2 | ±30° front-diagonal | Smartphone | 1080p | 60 | Checkerboard | Software (Cross-correl) | Laboratory
Barzyk_2024a | 1 | Frontal | Smartphone (iPhone X) | 4K | 60 | None (Monocular) |  | Laboratory
Barzyk_2024b | 1 | Sagittal | Smartphone (Google Pixel 6a) | 1080p | 60 | None (Monocular) |  | Laboratory
Bak_2025 | 5 | Frontal + Bilateral Sagittal | RGB Camera (4D EYE) | 1080p | 12 | Proprietary (4D EYE) | Software (4D EYE) | Laboratory
Bae_2024 | 18 | ±45° front-diagonal | Machine-vision / Industrial (18x FLIR) | 1080p | 60 | Checkerboard | Software | Laboratory
Azhand_2021 | 1 | Frontal | Tablet (iPad) | 1080p | 30 | None (Monocular) |  | Clinic
Andreoni_2024 | 2 | Frontal + Sagittal (Separate) | Industrial Camera | 480p | 25 | Checkerboard | Software (Cross-correl) | Laboratory
Anderson_2025 | 1 | Frontal | Webcam (monocular RGB, unspecified) | Not specified | 30 | None (Monocular) |  | Laboratory
Aleksic_2024 | 2 | Bilateral Sagittal | Smartphone | 1080p | 240 | Manual | Software (Cross-correl) | Laboratory
Ahn_2024 | 5 | Frontal + Bilateral Sagittal | RGB Camera (4DEYE) | 1080p | 12 | Checkerboard | Software | Laboratory
Wang_2024 | 1 | Multiple angles tested (0°, 45°, 90°, 135°, 180°) | Action Cam (GoPro Hero 8 Black) | 1080p | 60 | None (Monocular) | Hardware (Remote controller) | Laboratory
Verheul_2024 | 3 | Around capture volume | Tablet (iPad Pro 11-inch) | Not specified | 240 | Checkerboard | Manual (LED light bar trigger) | Laboratory
Wang_2025 | 2 | ±35° front-diagonal | Smartphone | Not specified | 120 | Checkerboard | Software | Laboratory
Bousigues_2025 | 10 | Around capture volume | Industrial Camera | 1920x1088 | 60 | QTM calibration (Qualisys) | Hardware (QTM) | Laboratory
Ueno_2024 | 1 | Multiple angles tested (0°, ±45°, ±70°) | Smartphone (iPhone 12 Pro) | Not specified | 60 | None (Monocular) |  | Laboratory
Yagi_2020 | 1 | Frontal + Sagittal (Separate) | RGB Camera (unspecified) | 2704x1500 | 60 | Homography (Floor points) | Software | Laboratory
Tahara_2025 | 1 | Not specified | Action Cam (GoPro Hero 10) | 1920x1080 | 60 | None (Monocular) |  | Laboratory
Francia_2026 | 2 | Frontal + Sagittal (Orthogonal) | Webcam (Insta360 Link) | Not specified | Not specified | Not specified | Hardware | Laboratory
Koleini_2025 | Varies (benchmark datasets) | Multi-view (benchmark datasets) | Industrial Camera | Varies (benchmark datasets) | 50 | Known (benchmark datasets) | Known (benchmark datasets) | Laboratory
Yamamoto_2021 | 1 | Sagittal | Industrial Camera (Basler acA720-290gc) | 720x520 | 100 | None (Monocular) |  | Laboratory
Hulleck_2024 | 1 | Sagittal | Action Cam (GoPro) | Not specified | 120 | None (Monocular) | Software (Cross-correl) | Laboratory
Templin_2024 | 12 | Around capture volume | Industrial Camera (Qualisys Miqus) | Not specified | 120 | Checkerboard/Wand (Qualisys) | Hardware | Laboratory
Wang_2023 | 1 | Multiple angles tested (Frontal, Sagittal, Coronal) | Webcam (Logitech C922) | 1080p | 15 | None (Monocular) |  | Laboratory
Wade_2023 | 2 | Frontal + Sagittal (Separate) | Industrial Camera (JAI sp5000c) | 1080p | 200 | Binary dot matrix | Hardware | Laboratory
Kim_2025 | 2 | Frontal + Sagittal (Separate) | Smartphone (Samsung Galaxy A Quantum) | Not specified | 30 | None (Monocular) | Software (Color-change sync) | Laboratory
Ho_2024 | 2 | Frontal (Binocular) | Industrial Camera (ZED Stereo) | 1080p | 30 | None (Binocular depth) | Software | Clinic
He_2024 | 1 | Frontal | Industrial Camera | 720p | 30 | None (Monocular) | Software (Cross-correl) | Laboratory
Balci_2026 | 5 | Around capture volume | RGB-D Camera (StereoLabs ZED 2i) | 720p | 60 | ZED360 | Hardware | Laboratory
Hii_2023 | 1 | Sagittal | Industrial Camera (Basler Pilot piA1900-32gc) | 960x540 | 25 | None (Monocular) | Hardware (Dataset pre-synchronized) | Laboratory
Chen_2025 | 2 (TUG dataset, primary) | Multi-view (2 cameras) | RGB-D Camera (Kinect) | 460x480 (TUG dataset) | 24 | Multi-view matching optimization | Software | Laboratory / Clinical
Fukushima_2024 | 12 | 360° Multi-view | Industrial Camera | 1080p | 100 | Calibration cage | Hardware | Laboratory
Ge_2025 | 1 | Frontal | Smartphone (Xiaomi Redmi) | 600x1000 | 30 | None (Monocular) |  | Clinic
Asaeda_2024 | 1 | Frontal | Industrial Camera (Vicon Vue) | Not specified | 100 | None (Monocular) | Hardware (Vicon MX Giganet) | Laboratory
Liang_2022 | 1 | Sagittal | Industrial Camera (Vicon Vue) | 1080p | 60 | None (Monocular) | Hardware | Laboratory
Svetek_2025 | 2 | ±45° front-diagonal | Tablet (iPad) | Not specified | 60 | Checkerboard | Software (Cross-correl) | Laboratory
Hu_2025 | 2 | Frontal + Sagittal (Orthogonal) | Industrial Camera (Blackfly S) | Not specified | 150 | Checkerboard | Hardware (Vicon lock sync device) | Laboratory
Haberkamp_2022 | 2 | Frontal + Sagittal (Separate) | Smartphone | 1080p | 30 | None (Monocular) | Software (Manual sync) | Laboratory
Vafadar_2022 | 4 | Multi-view | Industrial Camera | 1080p | 100 | Checkerboard | Hardware | Laboratory
Washabaugh_2022 | 1 | Sagittal | Industrial Camera (Basler Pilot piA1900-32gc) | 960x540 | 25 | None (Monocular) |  | Laboratory
Yeom_2026 | 2 | Front 45° | Smartphone | 1920x1080 | 60 | Checkerboard (210x175 mm) | OpenCap native | Laboratory
Gao_2025 | 1 | Sagittal (side view) | RGB Camera | 320x240 | 30 | None (Monocular) |  | Laboratory
Kumtekar_2024 | 1 | Frontal | Smartphone | 1280x720 | 30 | None (Monocular) |  | Laboratory
Nishizawa_2024 | 1 | Sagittal (bilateral) | Webcam (Sanwa CMS-V52S) | 1920x1080 | 25 | None (Monocular) |  | Clinic
Todesca_2025 | 7 | Around cycling setup | Industrial Camera (IDS U3-3160CP-M-GL) | 1984x1264 | 60 | Checkerboard | Hardware (OptiTrack sync) | Laboratory
Lin_2024 | 1 (SCAI-Gait) + Multi (benchmarks) | Monocular (SCAI-Gait) + Multi-view (benchmarks) | Industrial Camera (SCAI-Gait) / Industrial Camera (benchmarks) | 1920x1080 and 720x576 (SCAI-Gait) | 50 and 25 (SCAI-Gait) | Known (benchmark datasets) | Known (benchmark datasets) | Clinic (SCAI-Gait) + Laboratory (benchmarks)
Zhang_2024 | Multi | Multi-view | Industrial Camera | Varies | Varies | Uncalibrated (Learnable triangulation) | Sync provided | Laboratory
Hii_2023b | 1 | Frontal | Industrial Camera (Basler Pilot piA1900-32gc) | 960x540 | 25 |  |  | Laboratory
Pappas_2024 | 1 | Frontal | Industrial Camera (AViPAS AV-1360) | 1920x1080 (clinical) / 960x540 (GPJATK) | 30 (clinical) / 25 (GPJATK) | None (Monocular) |  | Clinic + Laboratory (GPJATK)
Irfan_2025 | 1 | Smartphone single camera | Smartphone (Redmi 2201117TI, Snapdragon 8 Gen 2) | 1080×2400 | 14 | None | N/A | Clinical (KIMS Hyderabad hospital) + dataset
Liang_2023 | 4 | Surrounding 4m walkway | Industrial Camera (Noraxon NiNOX 125) | 1088x704 | 60 | Chessboard (MATLAB stereo calibration) | Noraxon hardware sync | Laboratory
Cerciola_2024 | 2 | Surrounding treadmill | Webcam | 1920x1080 | 60 | Stereo calibration (Zhang's method) |  | Laboratory
Zhang_2025 | 8 | Surrounding walkway | Industrial Camera (RRIS40 system) |  | 120 | Multi-view weighted triangulation (custom) | Hardware (RRIS40 native) | Laboratory + Clinic (Rehabilitation center)
Kidzinski_2020 | 1 | Sagittal | Commodity RGB camera | NR | NR | None (single uncalibrated camera) | N/A | Laboratory (gait lab)
Hwang_2026 | 1 | Sagittal | iPad Pro 11 RGB | 800×600 px | 30 | Single camera, perpendicular sagittal, 3 m distance, 0.8 m height | N/A | Laboratory
Shah_2025 | 1 | Lab setup | RGB camera | NR | 100 | Spatial accuracy ±0.6 mm (Qualisys) | Within Qualisys system | Laboratory
Rode_2025 | 4 | 27 OMC + 4 RGB perspectives | Machine-vision / RGB (4 cameras) | NR (4 RGB perspectives) | NR | OMC calibrated; RGB perspectives noted | Yes (concurrent OMC + RGB) | Swiss Center for Movement Analysis, Zurich
Irfan_2025b | 1 | Frontal (laptop on table, 4 m from subject — visual: looks frontal/sagittal of full walking bout) | Built-in laptop camera (Windows OS, 64-bit browser) | NR | NR | None | N/A | Clinical (KIMS Hyderabad hospital) + dataset
Talaa_2025 | 2 | Left (C1) + Right (C3) sagittal | Standard RGB | 960×540 px | 25 | Yes (calibrated + synchronized via Vicon MXGiganet) | Yes (MXGiganet) | Laboratory
Ang_2024 | 1 | Sagittal | Smartphone | NR | NR | None (Monocular) | N/A | Home
Hou_2026 | 4 | Multi-view (4 cameras, unsynchronized) | Industrial Camera (image-based, unsynchronized) | NR | 30 Hz (c1,c2) / 10 Hz (c3,c4) — heterogeneous | Zhang's method (intrinsic + extrinsic) | Software (proposed temporal augmentation + multi-view post-process) | Laboratory (indoor)
Athama_2025 | 1 | Monocular | Single camera (generic RGB) | NR | NR | None (Monocular) | N/A | MoVi dataset (lab + IMU + video)
Tsintzira_2025 | 1 | Sagittal | Kinect/Depth (ZED-2 stereo) | NR | NR | ZED-2 factory (stereo) | N/A | Laboratory (treadmill)

## 3. CV Processing
Study_ID | Pose_2D_Algo | Backbone_Architecture | Model_Training | Training_Dataset | Method_3D | Pipeline_Name | Occlusion_Handling | Filtering | Marker_Augmentation | Real_Time | Filtering_Notes
Salehi_2026 | OpenPose / HRNet (via OpenCap) | LSTM (Augmenter) | Fine-tuned | Custom (lifting MoCap) | Triangulation (Multi-view) | OpenCap (Modified) |  | Butterworth | Yes | No (Post-process)
Verheul_2026 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Natraj_2025 | Multi-estimator benchmark (Detectron2, AlphaPose, YOLOv3+HRNet, RTMPose, MediaPipe) | Detectron2 (Mask R-CNN) | Pre-trained | Mixed | Lifting (2D->3D NN) | VideoPose3D | Linear completion | Not specified | No | No (Post-process)
Zhu_2023 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | Lifting (2D->3D NN) | Custom (BlazePose + KF) | Kalman Filter | Other | No | Not specified | Custom (FDF)
Zago_2020 | OpenPose (COCO-18) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom |  | Butterworth | No | No (Post-process)
Young_2023 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | None (2D Pure) | Custom (IoT App) |  | Moving Average | No | No (Post-process) | Moving Average
Verhoeven_2025 | DeepLabCut (Custom) | ResNet-50 | Fine-tuned | Custom | None (2D Pure) | DeepLabCut | Cubic Hermite Interpolation | Butterworth | No | No (Post-process)
Turner_2024 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Uhlrich_2023 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap | Built-in | Butterworth | Yes | No (Post-process)
Torvinen_2024 | DeepLabCut (Custom) | ResNet-50 | Fine-tuned | Custom | Triangulation (Multi-view) | Custom (DLC + DLT) |  | Butterworth | No | No (Post-process)
Takeda_2020 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Moving Average | No | No (Post-process) | Moving Average
Stenum_2021 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom | Gap filling (trajectory interpolation) | Butterworth | No | No (Post-process) | Low-pass filter (not explicitly named as Butterworth)
Pagnon_2021 | OpenPose (BODY_25B) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Pose2Sim |  | Butterworth | No | No (Post-process)
Pagnon_2022 | OpenPose (BODY_25B) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Pose2Sim |  | Butterworth | No | No (Post-process)
Needham_2021 | Proprietary / Unknown | CNN (DensePose) | Pre-trained | Mixed | Triangulation (Multi-view) | Custom |  | Kalman | No | No (Post-process) | Kalman Filter
Molteni_2025 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom | Linear Interpolation (short gaps) | Other | No | No (Post-process) | Linear Interpolation
Martis_2024 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Sabo_2022 | OpenPose (BODY_25) / AlphaPose / Detectron | Mixed | Pre-trained | Mixed | None (2D Pure) | Custom | Confidence-based interpolation | Butterworth | No | No (Post-process)
Panconi_2025 | DeepLabCut (Custom) | ResNet-101 | Fine-tuned | Custom | None (2D Pure) | DeepLabCut | Linear interpolation (low-confidence points) | Butterworth | No | No (Post-process)
Ota_2021 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process)
Nguyen_2023 | N/A (End-to-End) | ResNet-50 | Fine-tuned | Mixed | Direct Angle Regression (End-to-End) | Volumetric Net |  | Not specified | No | No (Post-process)
Needham_2022 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom | Robust RANSAC | Kalman | No | No (Post-process) | Bi-directional Kalman
Moro_2020 | DeepLabCut (Custom) | ResNet-50 | Fine-tuned | Custom | None (2D Pure) | Custom |  | Not specified | No | No (Post-process)
Moro_2022 | Pose ResNet-152 | Pose ResNet-152 | Fine-tuned | Mixed | Triangulation (Multi-view) | Custom (AdaFuse) | Epipolar Geometry | Butterworth | No | No (Post-process)
Mercadal_2024 | Detectron2 (Mask R-CNN) | Strided Transformer | Pre-trained | COCO (2D) + Custom VICON (3D) | Lifting (2D->3D NN) | Custom |  | Not specified | No | No (Post-process)
Menychtas_2023 | OpenPose (BODY_25) / MediaPipe | VGG-19 / MobileNetV2 | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process)
Mehdizadeh_2021 | OpenPose (BODY_25) / AlphaPose / Detectron | Mixed | Pre-trained | Mixed | None (2D Pure) | Custom | Confidence-based linear interpolation + manual label swap correction | Butterworth | No | No (Post-process)
Lim_2022 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom (OpenCV) |  | Butterworth | No | No (Post-process)
Embry_2022 | DeepLabCut (Custom) | ResNet-50 | Fine-tuned | Custom | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process) | High-pass Butterworth 8th-order 0.25 Hz + Median (5 frames) + Gaussian
Lima_2024 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap | Built-in | Butterworth | Yes | No (Post-process)
Li_2024 | Stacked Hourglass (CNN) | Vision Transformer (ViT) | Fine-tuned | Custom | Direct Angle Regression (End-to-End) | GaitFormer | Attention Mechanism | Butterworth | No | No (Post-process)
Ino_2023 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process)
Lambricht_2025a | OpenPifPaf | OpenPifPaf | Pre-trained | Mixed | None (2D Pure) | Custom | Cubic spline interpolation (gap filling) | Butterworth | No | No (Post-process)
Huang_2024 | Pixmotion (TsingVA) | N/A (TsingVA / SMPL) | Pre-trained | Not specified | Model Fitting (SMPL/HMR) | Custom |  | Butterworth | No | No (Post-process)
Lambricht_2025b | OpenPifPaf | OpenPifPaf | Pre-trained | Mixed | None (2D Pure) | Custom | Cubic spline interpolation (gaps < 0.25s) | Butterworth | No | No (Post-process)
Lagomarsino_2025 | MediaPipe (BlazePose 33) / AlphaPose | Mixed | Pre-trained | Mixed | Lifting (2D->3D NN) | MediaPipe Pose 3D | Confidence-based (threshold 0.7) + pchip interpolation (gaps < 2s) | Butterworth | No | No (Post-process)
Kumar_2022 | N/A (End-to-End) | ResNet-18 + LSTM/GRU | Fine-tuned | Custom | Direct Angle Regression (End-to-End) | DeepConvGRU |  | Not specified | No | No (Post-process)
Kakavand_2025 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap | LSTM | Butterworth | Yes | No (Post-process)
John_2024 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Other | No | No (Post-process) | Not clearly specified for OpenPose data
Ino_2024 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process)
Horsak_2025 | CameraHMR (intégré - SMPL) | CameraHMR (SMPL) | Pre-trained | Not specified | Model Fitting (SMPL/HMR) | CameraHMR |  | Butterworth | No | No (Post-process)
Horsak_2024 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Horsak_2023 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Harrison_2025 | HALPE26 (MMPose) via Ergomechanic | HRNet (MMPose HALPE26) | Pre-trained | Mixed | Triangulation (Multi-view) | Ergomechanic |  | Butterworth | No | No (Post-process)
Emmerson_2025 | OpenPose (BODY_25) / HRNet | VGG-19 / HRNet | Pre-trained | COCO + COCO-Wholebody | Triangulation (Multi-view) | Custom |  | Butterworth | No | No (Post-process) | Bi-directional Kalman filter (3D keypoints) + 4th-order Butterworth 6 Hz (joint angles)
Dunn_2023 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | Video Frame Interpolation
Drazan_2021 | DeepLabCut (Custom) | ResNet-50 | Fine-tuned | Custom | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process)
DHaene_2024 | YOLOv8 | YOLO | Pre-trained | Mixed | Depth Sensor (RGB-D) | Custom |  | Butterworth | No | No (Post-process)
deBorba_2025 | HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap | Built-in | Butterworth | Yes | No (Post-process)
DAntonio_2020 | OpenPose (COCO-18) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom |  | Not specified | No | No (Post-process)
DAntonio_2021 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom |  | Not specified | No | No (Post-process)
DAntonio_2020b | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom |  | Not specified | No | No (Post-process)
CarrascoPlaza_2022 | N/A (End-to-End) | Martinez 3D Pose Estimator | Fine-tuned | Human3.6M | Lifting (2D->3D NN) | Custom |  | Not specified | No | No (Post-process)
Boldo_2024 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process) | 3rd-order Savitzky-Golay + 4th-order Butterworth 6 Hz
Bittner_2023 | N/A (End-to-End) | ResNeXt-50 + LSTM/Transformer | Fine-tuned | BML-Movi | Direct Angle Regression (End-to-End) | D3KE |  | Not specified | No | No (Post-process)
Bertozzi_2025 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap | LSTM | Butterworth | Yes | No (Post-process)
Barzyk_2024a | Sbsq-pose (Custom Deep Learning) | Sbsq-pose | Pre-trained | Mixed | Lifting (2D->3D NN) | Custom |  | Not specified | No | No (Post-process)
Barzyk_2024b | SMARTGAIT (Custom Deep Learning) | Sbsq-pose | Pre-trained | Mixed | Lifting (2D->3D NN) | Custom |  | Butterworth | No | No (Post-process)
Bak_2025 | OpenPose (via 4D EYE system) | VGG-19 (OpenPose) | Pre-trained | Not specified | Triangulation (Multi-view) | Custom |  | Butterworth | No | No (Post-process)
Bae_2024 | YoloX + HRNet/DARK + TCFormer | HRNet/DARK + TCFormer | Pre-trained | COCO 2017 + COCO-Wholebody | Triangulation (Multi-view) | Ergo (Custom) |  | Butterworth | No | No (Post-process)
Azhand_2021 | CNN (Custom - algorithm name not specified) | CNN + TCN | Fine-tuned | Mixed | Lifting (2D->3D NN) | Custom |  | Not specified | No | No (Post-process)
Andreoni_2024 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom | Interpolation (short gaps) | Butterworth | No | No (Post-process) | Butterworth (may be for reference optoelectronic data only)
Anderson_2025 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Not specified | No | No (Post-process)
Aleksic_2024 | MMPose | RTMPose (MMPose) | Pre-trained | Mixed | None (2D Pure) | Custom | MOSSE tracker + deepocsort | Moving Average | No | No (Post-process) | Exponential Moving Average
Ahn_2024 | OpenPose (via 4DEYE system) | VGG-19 (OpenPose) | Pre-trained | Not specified | Triangulation (Multi-view) | Custom |  | Butterworth | No | No (Post-process)
Wang_2024 | OpenPose (BODY_25) | VGG-19 | Pre-trained | AMASS | Model Fitting (SMPL/HMR) | HuMoR |  | Butterworth | No | No (Post-process)
Verheul_2024 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Wang_2025 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap | LSTM | Butterworth | Yes | No (Post-process)
Bousigues_2025 | OpenPose (BODY_25B) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) + Marker Augmentation | Custom (Pose2Sim v0.3 + Uhlrich 2023 LSTM augmenter) |  | Not specified | Yes | No (Post-process)
Ueno_2024 | MMPose | CNN (HuMoR) | Pre-trained | AMASS | Model Fitting (SMPL/HMR) | HuMoR + CusToM |  | Not specified | No | No (Post-process)
Yagi_2020 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Not specified | No | No (Post-process)
Tahara_2025 | MediaPipe (BlazePose 33) | BlazePose (MobileNetV2) | Fine-tuned | Not specified | None (2D Pure) | Custom |  | Not specified | No | No (Post-process)
Francia_2026 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | Dual-camera EKF fusion | Custom (EKF Fusion) | Extended Kalman Filter | Butterworth | No | No (Post-process)
Koleini_2025 | N/A (End-to-End) | ViT-H/16 + MQ-HMR | Fine-tuned | Mixed | Direct Angle Regression (End-to-End) | BioPose |  | Not specified | No | No (Post-process)
Yamamoto_2021 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Not specified | No | No (Post-process)
Hulleck_2024 | MediaPipe (BlazePose 33) | BlazePose + Seq2Seq (LSTM/GRU) | Fine-tuned | Custom | Monocular Lifting 2D→3D (Seq2Seq) | Custom |  | Not specified | No | No (Post-process)
Templin_2024 | Custom CNN (SwRI, trained on synthetic SMPL-X data) | CNN (SwRI) | Fine-tuned | Mixed (Synthetic + MoCap) | Triangulation (Multi-view) | ENABLE (SwRI) |  | Butterworth | No | No (Post-process)
Wang_2023 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | Monocular Lifting 2D→3D (BlazePose 3D) | Custom | Visibility threshold (0.5) | Not specified | No | No (Post-process)
Wade_2023 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom | Kalman Filter | Kalman | No | No (Post-process) | Bi-directional Kalman filter
Kim_2025 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | None (2D Pure) | Custom | Acceleration-threshold denoising + cubic interpolation | Butterworth | No | No (Post-process)
Ho_2024 | Detectron2 + OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | Stereo Depth (ZED Stereo Camera) | PathoOpenGait (Custom) | Confidence threshold (0.5) | Moving Average | No | No (Post-process) | Moving average (window 5)
He_2024 | HRNet | HRNet + GLA-GCN | Fine-tuned | Mixed | Lifting (2D->3D NN) | Custom |  | Not specified | No | No (Post-process)
Balci_2026 | MediaPipe (BlazePose 33) / YOLOv8 | BlazePose / YOLO | Pre-trained | Mixed | StereoLabs ZED depth (3D) + MediaPipe monocular 3D | Custom |  | Not specified | No | No (Post-process)
Hii_2023 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | None (2D Pure) | TG-Gait (Custom) | Cubic spline interpolation (gap-filling) | Butterworth | No | No (Post-process) | 10th-order Butterworth low-pass (normalized cutoff 0.1752)
Chen_2025 | OpenPose (BODY_25) | VGG-19 + ShuffleNet | Fine-tuned | Mixed | Triangulation (Multi-view) | Custom |  | Not specified | No | No (Post-process)
Fukushima_2024 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | Triangulation (Multi-view) | Custom |  | Butterworth | No | No (Post-process)
Ge_2025 | OpenPose (BODY_25) | VGG-19 | Pre-trained | COCO | None (2D Pure) | Custom |  | Not specified | No | No (Post-process)
Asaeda_2024 | MediaPipe (BlazePose 33) | BlazePose | Pre-trained | Not specified | None (2D Pure) | Custom |  | Low-pass (unspecified) | No | No (Post-process) | Low-pass filter 2 Hz (for MediaPipe data)
Liang_2022 | OpenPose (BODY_25) | VGG-19 + 3DPoseNet | Pre-trained | Mixed | Lifting (2D->3D NN) | Custom | Linear interpolation (gap-filling) | Gaussian | No | No (Post-process) | Linear interpolation + 1D Gaussian smoothing
Svetek_2025 | OpenPose / HRNet (via OpenCap) | HRNet | Pre-trained | COCO | Triangulation (Multi-view) | OpenCap |  | Butterworth | Yes | No (Post-process)
Hu_2025 | ViTPose (ViT-base) | Vision Transformer (ViT) | Fine-tuned | Custom (OA patient Vicon projections) | Triangulation (Multi-view) | Custom |  | Not specified | No | No (Post-process)
Haberkamp_2022 | OpenPose (BODY_25) | VGG-19 (OpenPose BODY_25) | Pre-trained | COCO | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process) | Butterworth (4th order, 6 Hz)
Vafadar_2022 | Custom CNN (ENSAM marker-less system) | Custom CNN | Fine-tuned | Mixed (ENSAM + Human3.6M) | Triangulation (Multi-view) | Custom |  | Butterworth | No | No (Post-process)
Washabaugh_2022 | OpenPose / MoveNet / DeepLabCut | Mixed | Pre-trained | Mixed | None (2D Pure) | Custom |  | Butterworth | No | No (Post-process)
Yeom_2026 | HRNet (via OpenCap) | HRNet | Pre-trained + fine-tuned | COCO + custom | Triangulation (multi-view) | OpenCap |  | Butterworth | Yes (OpenSim markers) | No (Post-process) | Butterworth low-pass
Gao_2025 | BlazePose | MobileNetV2 + custom | Pre-trained | BlazePose dataset | None (2D only) | Custom (BlazePose + gait features) |  | Median | No | No (Post-process) | Median filter
Kumtekar_2024 | MediaPipe Pose | BlazePose (MediaPipe) | Pre-trained | MediaPipe dataset | None (2D only) | Custom (mobile app + AWS) |  | Moving Average | No | No (Post-process) | Moving average
Nishizawa_2024 | OpenPose | VGG-19 / CMU | Pre-trained | COCO + MPII | None (2D only) | Custom (OpenPose + gait analysis) |  | Butterworth | No | No (Post-process) | Butterworth low-pass
Todesca_2025 | YOLO11x-Pose | CSPDarknet (YOLO11x) | Pre-trained + fine-tuned | COCO + custom cycling | Triangulation (multi-view) | Custom (YOLO11x + triangulation) | Confidence-based weighting + displacement propagation | Other | No | No (Post-process) | Alpha-beta-gamma filter
Lin_2024 | Custom (biomech-aware) | Custom CNN + biomechanical model | Trained on synthetic data | Synthetic (OpenSim-generated) | Direct 3D end-to-end (biomechanics-aware) | Custom | Biomechanical constraints |  | No | No (Post-process) | Not applicable (end-to-end biomechanical model with built-in temporal constraints)
Zhang_2024 | ResNet152 (Simple Baseline) | HRNet | Fine-tuned | COCO WholeBody | GNN-based 3D reconstruction + OpenSim | Custom (GNN + OpenSim IK) | GNN occlusion handling | Not specified | Yes (OpenSim markers) | No (Post-process)
Hii_2023b | OpenPose + YOLOv7-Pose + MediaPipe | Various (VGG/YOLO/BlazePose) | Pre-trained | COCO | None (2D frontal only) | Custom (comparison study) | Cubic spline interpolation (gap-filling) | Moving Average | No | Yes (MediaPipe) | Moving average
Pappas_2024 | MediaPipe Pose | BlazePose (MediaPipe) | Pre-trained | MediaPipe dataset | ZoeDepth (monocular depth) | Custom (MediaPipe + ZoeDepth) |  | Butterworth | No | No (Post-process) | Butterworth low-pass (1 Hz) + Savitzky-Golay
Irfan_2025 | BlazePose | MobileNetV1 | Pretrained (MediaPipe BlazePose) | COCO superset (33 landmarks) | Not 3D (2D keypoints only) | Custom in-house | NR | NR | No | Yes (smartphone-capable) | Joint angle via law of cosines from 2D keypoints
Liang_2023 | HRNet | HRNet-W48 | Pre-trained | COCO WholeBody | Triangulation (multi-view) | Custom (HRNet + OpenSim) |  | Median | No | No (Post-process) | Median filter (9 frames)
Cerciola_2024 | AP/BZ/TC/RTMC/RTMH/MP + BA | Various (AlphaPose/BlazePose/TC-Former/RTMPose/MediaPipe) | Pre-trained | COCO/HALPE/BlazePose | Triangulation (DLT, multi-view) | Custom (comparison, 7 methods) |  | Butterworth | No | No (Post-process) | Butterworth 10 Hz
Zhang_2025 | Custom CNN (RRIS40 dataset) | Custom (data-driven + anatomical priors) | Pre-trained (RMSE40) | RMSE40 | Triangulation + model fitting | Custom (trained on RRIS40 dataset) | Theia3D native | Other | Yes (Theia3D virtual markers) | No (Post-process) | Theia3D native
Kidzinski_2020 | OpenPose | CMU OpenPose CNN | Pretrained (OpenPose) + CNN/RF/RR regressors trained on CP gait dataset | 1792 videos CP clinical gait + ground truth from optical MoCap | Not applicable (2D → clinical metrics regression) | mobile-gaitlab (github stanfordnmbl) | NR | NR | No | No | CNN time-series regression direct from 2D keypoint trajectories
Hwang_2026 | ViFive proprietary (custom) | Stacked Hourglass (modified) + random forest classifier | Trained on diverse age/body type/environment dataset | Proprietary dataset (ViFive Inc.) | Pinhole-camera perspective correction (2D only) | ViFive Inc | NR | 4th-order Butterworth low-pass 4 Hz on keypoints | No (direct 2D keypoints) | Designed for handheld/tablet device | Pinhole perspective correction using anthropometric ratios (Winter)
Shah_2025 | MediaPipe (via OpenPose mention) | MediaPipe Pose CNN | Pretrained MediaPipe | COCO/MediaPipe pretrained | OpenCap (multi-view triangulation) + MediaPipe 2D | MediaPipe + OpenCap comparison | Reports occlusion sensitivity | NR | OpenCap uses virtual markers | Yes (MediaPipe) | Comparative: Qualisys (gold) vs OpenCap vs MediaPipe
Rode_2025 | 11 open-source pose estimators benchmarked (BlazePose, RTMPose, OpenPose, etc. — direct + lifting) | Multiple (stacked hourglass, HRNet, transformers — varies by estimator) | All pretrained on COCO/Human3.6M etc | Open-source pretrained models | Direct 3D and 2D→3D lifting methods (both evaluated) | Physio2.2M benchmark | NR | NR | No | Direct 25-200 FPS; Lifting 117-9341 FPS | Per-pose-estimator analysis
Irfan_2025b | BlazePose | MobileNetV1 | Pretrained (MediaPipe BlazePose) | COCO superset (33 landmarks) | Not 3D (2D keypoints only) | Custom in-house | NR | NR | No | Yes (smartphone-capable) | Joint angle via law of cosines from 2D keypoints
Talaa_2025 | MediaPipe Pose | MediaPipe (CNN-based, bottom-up, 33 landmarks) | Pretrained MediaPipe | MediaPipe default | Not 3D (2D per-camera angles) | MediaPipe + OpenCV + Python custom | NR | NR | No | Yes | Angle via arctan2 of 2D landmarks (hip-knee-ankle triplet)
Ang_2024 | OpenPose (BODY_25) | OpenPose CMU CNN | Pretrained | COCO | None (2D Pure) | Custom | NR | Normalization | No | No (Post-process) | Joint angles computed via vector geometry from 2D keypoints (BODY_25 model)
Hou_2026 | Proprietary / Unknown | Transformer-based multi-view temporal fusion | Trained on custom dataset (8:1:1 split) | Custom (4 cameras, walking/running/in-place) | Model Fitting (SMPL/HMR) | Custom (Virtual Marker Extraction + Temporal Augmentation + Transformer Post-process) |  |  | Yes (virtual markers extracted from SMPL surface, ¦M¦=56 landmarks) | No (Post-process) | Novel multi-view temporal post-processing + temporal augmentation training to handle unsynchronized cameras
Athama_2025 | PoseFormerV2 | PoseformerV2 (transformer-based) + Deep-JSCC encoder | Pretrained PoseformerV2 + custom Deep-JSCC trained on MoVi | MoVi (Large Multipurpose Motion and Video) | Lifting (2D→3D NN) | Deep-JSCC wireless framework + SMPL rendering |  |  | Virtual (SMPL+AMASS for rendering only) | Yes (Near real-time) | Deep Joint Source-Channel Coding for wireless transmission. SMPL+AMASS animation for privacy-preserving rendering
Tsintzira_2025 | ZED-2 SDK Body Tracking | ZED-2 proprietary body tracking | Pretrained (ZED SDK) | ZED proprietary | Triangulation (Multi-view) [via stereo ZED] | Custom 2D human model (14 DoF, sagittal) |  |  | Anatomical landmarks as keypoints | Not specified | 2D human model 14 DoF: hip(X,Y,θ) + each leg(hip,knee,ankle) + each arm(shoulder,elbow) + head(rel.torso)

## 4. Biomechanics
Study_ID | Bio_Software | Kinematics_Method | Skeletal_Model | Scaling | Kinetics_Forces
Salehi_2026 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Beaucage-Gauvreau | Scaling | No
Verheul_2026 | OpenCap (OpenSim) | Inverse Kinematics (IK) | LaiUhlrich2022 | Scaling | No
Natraj_2025 | Custom Code (Python) | Direct Prediction (Deep Learning) | N/A (3D keypoints) |  | No
Zhu_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Zago_2020 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Young_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Verhoeven_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Turner_2024 | OpenCap (OpenSim) | Inverse Kinematics (IK) | LaiUhlrich2022 | Scaling | No
Uhlrich_2023 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Lai 2017 / Uhlrich | Scaling | Yes
Torvinen_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Takeda_2020 | Kinovea / Custom Code | Direct (Geometric Vectors) |  |  | No
Stenum_2021 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Pagnon_2021 | OpenSim | Inverse Kinematics (IK) | Rajagopal | Scaling | No
Pagnon_2022 | OpenSim | Inverse Kinematics (IK) | Rajagopal | Scaling | No
Needham_2021 |  | N/A (Center of Mass) |  |  | No
Molteni_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Martis_2024 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Rajagopal | Scaling | No
Sabo_2022 | Custom Code (Python/Matlab) |  |  |  | No
Panconi_2025 | Custom Code (Python/Matlab) |  |  |  | No
Ota_2021 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Nguyen_2023 | OpenSim | Direct Prediction (Deep Learning) |  |  | No
Needham_2022 | OpenSim | Inverse Kinematics (IK) | Not specified | Scaling | No
Moro_2020 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Moro_2022 | OpenSim | Inverse Kinematics (IK) | Rajagopal | Scaling | No
Mercadal_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Menychtas_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Mehdizadeh_2021 | Custom Code (Python/Matlab) |  |  |  | No
Lim_2022 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Embry_2022 | Custom Code (Python/Matlab) | Direct Prediction (Deep Learning) |  |  | No
Lima_2024 | OpenCap (OpenSim) | Inverse Kinematics (IK) | LaiUhlrich2022 | Scaling | No
Li_2024 | Custom Code (Python/Deep Learning) | Direct Prediction (Deep Learning) |  |  | No
Ino_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Lambricht_2025a | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Huang_2024 | Visual3D (Pixmotion/TsingVA) | Inverse Kinematics (IK) | SMPL-based (6 DOF kinematic chain) | Scaling | Yes
Lambricht_2025b | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Lagomarsino_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Kumar_2022 | Custom Code (Python/Deep Learning) | Direct Prediction (Deep Learning) |  |  | No
Kakavand_2025 | OpenSim | Inverse Kinematics (IK) | Lai 2017 / Uhlrich | Scaling | Yes
John_2024 | Custom Code (Python/Matlab) | Direct Prediction (Deep Learning) |  |  | No
Ino_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Horsak_2025 | OpenSim | Inverse Kinematics (IK) | Lai 2017 / Uhlrich | Scaling | No
Horsak_2024 | OpenSim | Inverse Kinematics (IK) | Lai 2017 / Uhlrich | Scaling | No
Horsak_2023 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Lai 2017 / Rajagopal | Scaling | No
Harrison_2025 | OpenSim | Inverse Kinematics (IK) | Rajagopal | Scaling | No
Emmerson_2025 | OpenSim | Inverse Kinematics (IK) | Rajagopal | Scaling | No
Dunn_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Drazan_2021 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
DHaene_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
deBorba_2025 | OpenSim | Inverse Kinematics (IK) | Lai 2017 / Uhlrich | Scaling | No
DAntonio_2020 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
DAntonio_2021 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
DAntonio_2020b | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
CarrascoPlaza_2022 | Custom Code (Python/TensorFlow) | Direct Prediction (Deep Learning) |  |  | No
Boldo_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Bittner_2023 | OpenSim | Direct Prediction (Deep Learning) |  |  | No
Bertozzi_2025 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Lai 2017 / Uhlrich | Scaling | No
Barzyk_2024a | Custom Code (Python/Matlab) | Direct Prediction (Deep Learning) |  |  | No
Barzyk_2024b | Custom Code (Python/Matlab) | Direct Prediction (Deep Learning) |  |  | No
Bak_2025 | Custom (OpenPose MLS) | Direct (Geometric Vectors) |  |  | No
Bae_2024 | Ergo (Custom DL) | Direct (Geometric Vectors) |  |  | No
Azhand_2021 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Andreoni_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Anderson_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Aleksic_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Ahn_2024 | Custom (OpenPose-based) | Direct (Geometric Vectors) |  |  | No
Wang_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Verheul_2024 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Lai 2017 / Rajagopal | Scaling | No
Wang_2025 | OpenCap (OpenSim) | Inverse Kinematics (IK) | Lai 2017 / Rajagopal | Scaling | No
Bousigues_2025 | Custom Code (Python/Matlab) | Inverse Kinematics (IK) | Duprey 2010 | Scaling | No
Ueno_2024 | CusToM (Matlab Toolbox) | Inverse Kinematics (IK) | Full-body (42 DOF, 82 muscles) | Scaling | Yes
Yagi_2020 | Custom Code (Python/Matlab) |  |  |  | No
Tahara_2025 | Custom Code (Python/vaila Toolbox) | Direct Prediction (Deep Learning) |  |  | No
Francia_2026 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Koleini_2025 | Custom Code (Python/OpenSim-based) | Direct Prediction (Deep Learning) |  |  | No
Yamamoto_2021 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Hulleck_2024 | Custom Code (Python/Keras) | Direct Prediction (Deep Learning) |  |  | No
Templin_2024 | OpenSim | Inverse Kinematics (IK) | Rajagopal | Scaling | Yes
Wang_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Wade_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Kim_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Ho_2024 | Custom Code (Python/PyTorch) | Direct Prediction (Deep Learning) |  |  | No
He_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Balci_2026 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Hii_2023 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Chen_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Fukushima_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Ge_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Asaeda_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Liang_2022 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Svetek_2025 | OpenCap (OpenSim) | Inverse Kinematics (IK) | LaiUhlrich2022 | Scaling | No
Hu_2025 | OpenSim | Inverse Kinematics (IK) | Generic OpenSim (CGM lower-limb) | Median-value scaling (per trial) | No
Haberkamp_2022 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Vafadar_2022 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Washabaugh_2022 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) |  |  | No
Yeom_2026 | OpenCap (OpenSim) | Inverse Kinematics | LaiUhlrich2022 | Scaling | No
Gao_2025 | None (2D angles) | Direct 2D angle computation | None (2D keypoints) |  | No
Kumtekar_2024 | None (2D angles) | Direct 2D angle computation | None (2D keypoints) |  | No
Nishizawa_2024 | None (2D angles) | Direct 2D angle computation | None (2D keypoints) |  | No
Todesca_2025 | Custom | Direct 3D angle computation | Custom (cyclist model) | Manual scaling | No
Lin_2024 | Custom Code (Python/OpenSim-based) | Direct Prediction (Deep Learning) | Custom (biomechanical) | Synthetic scaling | No
Zhang_2024 | OpenSim | Inverse Kinematics | Rajagopal full-body | Marker-based scaling | No
Hii_2023b | None (2D angles) | Direct 2D angle computation (frontal) | None (2D keypoints) |  | No
Pappas_2024 |  | Direct 2D + depth angle computation |  |  | No
Irfan_2025 | Custom (CSV export) | Law of cosines from 2D coordinates | BlazePose 33 keypoints | N/A | No
Liang_2023 | OpenSim | Inverse Kinematics | Gait2354 | Manual (anthropometric) | Yes
Cerciola_2024 | Custom (MATLAB) | Direct 2D angle computation (sagittal) | Segment-based (body segments) |  | No
Zhang_2025 | OpenCap (OpenSim) | Inverse Kinematics | Full-body (custom 40 landmarks) | Custom auto-scaling | Yes
Kidzinski_2020 | None (DL regression directly to clinical metrics) | ML regression from 2D pose trajectories | N/A | N/A | No
Hwang_2026 | Custom MATLAB (R2023b) | Direct angle from 2D keypoints (sagittal plane); CoM weighted average (Winter) | Anthropometric weights (Winter) | Pixel-to-distance via D_peak = 0.961·h (anthropometric scaling) | Yes — STS power computed from CoM acceleration × mass (P = m(a-g)·v)
Shah_2025 | Mokka (open-source) | C3D files, marker triplets (thigh + shank) | Free body segments (thigh, shank as rigid links) | Anthropometric (height, mass) | Newton-Euler equations mentioned (∑F=ma, ∑M=Iα) but no force plate
Rode_2025 | None (direct keypoint comparison) | Joint angles computed from keypoints (knee, elbow flexion) | Plug-in Gait (Conventional Gait Model) — 46 reflective markers | NR (markers anatomical) | No
Irfan_2025b | Custom (CSV export) | Law of cosines from 2D coordinates | BlazePose 33 keypoints | N/A | No
Talaa_2025 | Custom Python (NumPy) | Arctan2 from 2D landmarks | MediaPipe 33 landmarks | N/A | No
Ang_2024 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) | BODY_25 keypoints | N/A | No
Hou_2026 | OpenSim (used for downstream skeleton reconstruction) | Virtual marker positions → OpenSim downstream (potential angles via IK, not directly validated) | SMPL mesh + 56 virtual whole-body landmarks (OpenSim compatible) |  | No (future direction mentioned)
Athama_2025 | Custom Code (Python/Matlab) | Direct (Geometric Vectors) | 17 joints (PoseformerV2 output) + SMPL for rendering |  | No
Tsintzira_2025 | Custom Code (Python/Matlab) + OpenSim Gait2392 (anthropometric parameters) | Inverse Kinematics (IK) — closed-form analytical | 2D 14-DoF model (sagittal plane), Gait2392 anthropometrics | Personalized via subject mass/inertia (Gait2392 estimates) | Yes — Inverse Dynamics (ID) + GRF estimation via Lagrangian dynamics, NO force plate needed

## 5. Global & Spatio_Validation
Study_ID | Gold_Standard | Validates_Kinematics | Validates_Spatiotemp | Validates_Kinetics | Reported_Metrics | Global_Kinematic_Error (Deg) | Global_Kinematic_Correl | Spatio_Speed_Error | Spatio_StepLength_Error | Spatio_StepTime_Error | Authors_Conclusion | Validation_Output_Type | Gold_Standard_Type | Gold_Standard_Details | Extraction_Notes
Salehi_2026 | OptiTrack | Yes | No | No | RMSE only | 9.5 |  |  |  |  | Custom lifting model reduced kinematic errors compared to the original OpenCap. | Angle | MoCap | Optitrack Prime 13 (8 cameras, 100 Hz) | Global_Kinematic_Error (Deg): 9.5 (RMSE, proposed model)
Verheul_2026 | Vicon | Yes | No | No | SPM + RMSE | 3.6-8.6 |  |  |  |  | Caution recommended due to systematic differences with marker-based systems. | Angle | MoCap | Vicon Vantage (11 cameras, 200 Hz) | Global_Kinematic_Error (Deg): 3.6-8.6 (muRMSD range across joints, 5 running speeds)
Natraj_2025 | MoCap (Other) | Yes | No | No | Other | 90.7 mm (PA-MPJPE) |  |  |  |  | VideoPose3D is the best-performing model for pathological gait in SCI. | Position 3D (mm) | MoCap | MoCap at SPZ (20 markers) | Global_Kinematic_Error (Deg): 90.7 mm (PA-MPJPE) CONFIRMED.
Zhu_2023 | Qualisys | Yes | Yes | No | Other | 25.40% |  | Pearson r=0.79 |  |  | Post-processing algorithm improved the BlazePose performance by 10.7%. | Angle + Spatiotemporel | MoCap | Qualisys (MoVi) | Global_Kinematic_Error (Deg): 25.40% (PE overall, best model KF2+FDF). The 17% refers to knee PE (17.33%), not an overall lower bound. ¦ Spatio_Speed_Error: Pearson r=0.79 (gait speed, monocular BlazePose vs Qualisys)
Zago_2020 | BTS | No | Yes | No | RMSE only |  |  |  | 1.23-7.42 cm | 0.02-0.08 s | Feasibility of tracking spatiotemporal parameters with OpenPose confirmed. | Spatiotemporel | MoCap | BTS Smart
Young_2023 | Vicon | Yes | Yes | No | Other | 1.4 (L) / 12.5 (R) | 0.96-0.98 |  |  | 0.01-0.03s | Good to excellent agreement in running gait outcomes. | Angle + Spatiotemporel | MoCap | Vicon Vertex (14 cameras, 200 Hz) | Global_Kinematic_Error (Deg): 1.4 (L) / 12.5 (R) — Knee Flexion mean error
Verhoeven_2025 | Vicon | Yes | No | No | RMSE + ICC | 0.78-6.74 | 0.60-1.00 |  |  |  | DeepLabCut proved a valid tool; higher validity found for adults. | Angle | MoCap | Vicon Vero v2.2 (10 infrared cameras, 100 Hz) | Global_Kinematic_Error (Deg): 0.78-6.74 (75% Model)
Turner_2024 | Vicon | Yes | No | No | CMC + RMSE | MAE 3.85° / RMSE 4.44° | Grand mean CMC 0.76; sagittal CMC range higher, frontal/transverse 0.51-0.60 |  |  |  | OpenCap errors are comparable to other markerless systems, but caution is needed for frontal/transverse planes. | Angle | MoCap | Vicon (10 cameras, 250 Hz) | Global_Kinematic_Error (Deg): MAE 3.85° / RMSE 4.44° (body text p.4); ⚠️ abstract p.1 says RMSE 4.34° — internal paper inconsistency, body value retained
Uhlrich_2023 | Motion Analysis | Yes | No | Yes | MAE + Pearson | 4.5 (MAE) |  |  |  |  | THE founding paper of OpenCap. Accurately predicts dynamic measures (forces/moments) from video. | Kinetics + Angle (global) | MoCap | Motion Analysis Corp (8 cameras, 100 Hz) + Bertec Force Plates (2000 Hz) | Global_Kinematic_Error (Deg): 4.5 (MAE). CONFIRMED from Table 1.
Torvinen_2024 | Vicon | Yes | No | No | Other | -2.47 to 3.69 |  |  |  |  | Custom DeepLabCut method is accurate for 3D kinematic analysis of skate skiing on a treadmill. | Angle (global) | MoCap | Vicon | Global_Kinematic_Error (Deg): -2.47 to 3.69 (Mean Diff)
Takeda_2020 | Motion Analysis | Yes | No | No | Other |  | 0.51-0.97 ICC |  |  |  | OpenPose is an adequate substitute for conventional PMMC for sagittal plane kinematics. | Cinematique (non specifie) | Autre |  | Global_Kinematic_Correl: 0.51-0.97 ICC (MAC3D vs OpenPose: Hip 0.97, Knee 0.92, Ankle 0.51)
Stenum_2021 | MoCap (Other) | Yes - validates sagittal plane hip, knee, and ankle angles | Yes | No | RMSE + ICC |  |  | MAE<0.04 m/s; ICC=0.964-0.986 | MAE=0.049 m; ICC=0.927-0.973 | MAE=0.02 s; ICC=0.995-0.998 | OpenPose can estimate 2D spatiotemporal parameters, but perspective errors occur with a single camera. | Spatiotemporel | MoCap | HM-GPJATK public dataset (motion capture) | Spatio_Speed_Error: MAE<0.04 m/s; ICC=0.964-0.986 (OpenPose vs Optitrack/Vicon) ¦ Spatio_StepLength_Error: MAE=0.049 m; ICC=0.927-0.973 (OpenPose vs Optitrack/Vicon) ¦ Spatio_StepTime_Error: MAE=0.02 s; ICC=0.995-0.998 (OpenPose vs Optitrack/Vicon)
Pagnon_2021 | None (Reliability) | Yes | No | No | Other | 1.5 |  |  |  |  | Introduces Pose2Sim workflow. Demonstrates high robustness to image degradation. | Angle (global) | MoCap | Virtual cameras generated from 68-camera 3D mesh capture (Panoptic Studio). NOT a traditional optoelectronic MoCap system | Global_Kinematic_Error (Deg): 1.5 (MAE between degraded and reference Pose2Sim conditions - ROBUSTNESS metric, NOT accuracy vs gold standard). Walking accuracy vs normative: Hip MAE=9.0, Knee MAE=5.7, Ankle MAE=5.4 from Fig 8.
Pagnon_2022 | Qualisys | Yes | No | No | CMC + RMSE | Walking RMSE: Hip 5.6, Knee 5.1, Ankle 4.0; Mean Errors: Walk 3.0, Run 4.1, Cycle 4.0 | > 0.90 |  |  |  | Pose2Sim is highly accurate but reveals a systematic 15-degree offset for the hip during running due to OpenPose model biases. | Angle (global) | MoCap | Qualisys (20 optoelectronic cameras) | Global_Kinematic_Error (Deg): Walking RMSE: Hip 5.6, Knee 5.1, Ankle 4.0 (Table 4); Mean Errors: Walk 3.0, Run 4.1, Cycle 4.0
Needham_2021 | Qualisys | No | Yes | No | Other |  |  | 0.015 m/s | 0.001 m | 0.008 - 0.022s | Excellent agreement for spatio-temporal variables using 9 cameras for elite skeleton push-start. | Spatiotemporel | MoCap | Qualisys Oqus (15 cameras, 200 Hz)
Molteni_2025 | BTS | Yes | Yes | No | MAE + ICC |  |  | 0.16 m/s | Stride length MAE=0.18-0.19 m | 1.63 step/min | Openpose is accurate for sagittal plane gait kinematics and spatiotemporal parameters. | Angle + Spatiotemporel | MoCap | BTS SMART DX (8 high-resolution infrared cameras, 100 Hz) | Spatio_StepLength_Error: Stride length MAE=0.18-0.19 m (OpenPose vs BTS Smart)
Martis_2024 | Vicon | Yes | Yes | No | RMSE only | 4.5 |  | Mean diff ~0.00-0.01 m/s | Mean diff ~0.00 m | n/a (see Extraction_Notes) | Good accuracy in spatial-temporal parameters but challenges in kinematic parameters, especially when walking away from cameras. | Angle + Spatiotemporel | MoCap | Vicon (17 cameras, 150 Hz) | Global_Kinematic_Error (Deg): 4.5 (Grand Mean WTC) ¦ Spatio_Speed_Error: Mean diff ~0.00-0.01 m/s (r=0.99-1.0, OpenCap vs Vicon, walking toward camera best) ¦ Spatio_StepLength_Error: Mean diff ~0.00 m (r=0.98-0.99, OpenCap vs Vicon) ¦ Spatio_StepTime_Error: confirmed_absent (step time not explicitly reported as separate error metric)
Sabo_2022 | Zeno walkway (PKMAS) | No | Yes | No | Other |  |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | Spearman ρ 0.55-0.94 | 2D pose-estimation libraries can be used for longitudinal analysis of spatiotemporal gait features, outperforming the 3D ROMP method. | Spatiotemporel | Walkway/Force |  | Spatio_Speed_Error: confirmed_absent (pixel-based distance; only Spearman correlations reported, no absolute error in m/s) ¦ Spatio_StepLength_Error: confirmed_absent (pixel-based distance; Spearman rho only) ¦ Spatio_StepTime_Error: Spearman ρ 0.55-0.94 (cadence; best detectors 0.81-0.93)
Panconi_2025 | Force Plates | No | Yes | No | MAE + Pearson |  |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | Pearson r=0.90 | Custom trained DeepLabCut outperformed OpenPose pre-trained for spatiotemporal parameters. | Spatiotemporel | Walkway/Force |  | Spatio_Speed_Error: confirmed_absent (force platforms used; no spatial calibration for speed measurement) ¦ Spatio_StepLength_Error: confirmed_absent (force platforms; no spatial calibration) ¦ Spatio_StepTime_Error: Pearson r=0.90 (DLCCT380R, step time vs force platforms)
Ota_2021 | Vicon | Yes | No | No | ICC |  | n/a (see Extraction_Notes) |  |  |  | OpenPose is valid for measuring sagittal plane lower limb ROM but lacks full 3D capabilities. | Cinematique (non specifie) | MoCap | Vicon Nexus (8 cameras, 60 Hz) | Global_Kinematic_Correl: Reported
Nguyen_2023 | Vicon | Yes | No | No | MAE | 7.19 MPJAE on Roofing dataset |  |  |  |  | End-to-end volumetric pose direct angle estimation works well on roofing and H3.6M datasets. | Angle | MoCap | Vicon Vantage V16 (14 cameras, 100 Hz) | Global_Kinematic_Error (Deg): 7.19 MPJAE on Roofing dataset. CONFIRMED. Also 8.41 on Human3.6M.
Needham_2022 | Qualisys | Yes | No | No | Other | [5 empty] | Fully automated OpenPose to OpenSim workflow with 9 cameras at 200 Hz shows high agreement with marker-based. | Angle | MoCap | Qualisys Oqus (15 cameras, 200 Hz)
Moro_2020 | BTS | Yes | Yes | No | Other |  |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | Pilot study on stroke survivors showing that DeepLabCut can estimate 2D elevation angles, though 25 fps causes ankle errors. | Spatiotemporel | MoCap | BTS SMART DX (8 cameras, 250 Hz) | Spatio_Speed_Error: confirmed_absent (qualitative comparison only; no numerical error values) ¦ Spatio_StepLength_Error: confirmed_absent (qualitative comparison only) ¦ Spatio_StepTime_Error: confirmed_absent (qualitative comparison only)
Moro_2022 | OptiTrack | Yes | Yes | No | Other | 23.65 mm (MPJPE) |  | Mean diff ~0.04 m/s | Mean diff ~0.05 m | Mean diff ~0.02 s | Validates a multi-camera pipeline using epipolar geometry (AdaFuse) to refine 2D joints before 3D reconstruction. | Position 3D (mm) + Spatiotemporel | MoCap | OptiTrack Flex 13 (8 cameras, 100 Hz) | Spatio_Speed_Error: Mean diff ~0.04 m/s (computable from Table data; AdaFuse multi-camera vs Optitrack) ¦ Spatio_StepLength_Error: Mean diff ~0.05 m (stride length, computable from Table data) ¦ Spatio_StepTime_Error: Mean diff ~0.02 s (stride time, computable from Table data)
Mercadal_2024 | Vicon | Yes | No | No | RMSE only | < 10 |  |  |  |  | Uses Strided Transformer to lift 2D to 3D from a single iPhone. Very good accuracy for exercises (squat, CMJ). | Angle + Spatiotemporel | MoCap | VICON (Strided Transformer model)
Menychtas_2023 | Vicon | Yes | No | No | SPM + RMSE | [5 empty] | Demonstrates through SPM analysis that MediaPipe has significant errors and that OpenPose struggles with stance phase due to crossing legs. | Angle | MoCap | Vicon (10 infrared cameras, 100 Hz)
Mehdizadeh_2021 | Xsens | No | Yes | No | ICC |  |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | Step time Pearson R=0.71; Cadence R=0.97-1.00 | AlphaPose and Detectron had higher temporal agreement than OpenPose. | Spatiotemporel | MoCap | Xsens IMU (7 sensors) - NOT optical MoCap | Spatio_Speed_Error: confirmed_absent (pixel-normalized spatial; no calibrated speed in m/s) ¦ Spatio_StepLength_Error: confirmed_absent (pixel-normalized spatial; no calibrated step length) ¦ Spatio_StepTime_Error: Step time Pearson R=0.71; Cadence R=0.97-1.00 (AlphaPose/Detectron vs Xsens IMU)
Lim_2022 | Vicon | Yes | No | No | Other | [5 empty] | Very small sample (N=1) for squat assessment. Showed similar trajectory profiles but specific biases in ROM. | Angle | MoCap | Vicon (13 cameras: 6 Vantage5, 1 MX-T160, 5 MX-T40S, 1 MX-T40, 100 Hz)
Embry_2022 | GAITRite (instrumented walkway) | No | Yes | No | Other |  |  | Accuracy=-0.02 +/- 0.11 m/s |  | n/a (see Extraction_Notes) | Excellent validation of purely spatiotemporal parameters on stroke survivors using a custom DeepLabCut model. | Spatiotemporel | Walkway/Force | GAITRite Gold instrumented walkway (CIR Systems) - NOT optical MoCap | Spatio_Speed_Error: Accuracy=-0.02 +/- 0.11 m/s (r=0.92, DeepLabCut vs GAITRite) ¦ Spatio_StepTime_Error: confirmed_absent (the -0.25 value is cadence error in steps/min, NOT step time; no explicit step time error reported)
Lima_2024 | Vicon | Yes | No | No | CMC + RMSE | 2.0 - 12.8 (RMSE) | 0.05 - 0.99 (CMC) |  |  |  | OpenCap holds promise as a scalable tool, though task and plane-specific errors should be considered. | Angle | MoCap | Vicon (11 cameras, 200 Hz) + OpenSim
Li_2024 | NOKOV multi-camera | Yes | No | No | MAE + Pearson | 2.76 - 4.16 (MAE) | 0.90 - 0.98 (r) |  |  |  | GaitFormer achieves high accuracy in kinematic estimation using a single RGB camera, comparable to marker-based systems. | Angle | MoCap | NOKOV multi-camera optical tracking system (NOT Vicon)
Ino_2023 | Vicon | Yes | No | No | CMC + MAE | 2.3 - 4.1 (MAE) | 0.89 - 0.99 |  |  |  | Validates OpenPose for 2D sagittal gait. Shows that the leg opposite to the camera has higher errors due to occlusion. | Angle | MoCap | Vicon Nexus 2.10 (14 cameras: 8 MX T10-S + 6 Vero v2.2, 120 Hz)
Lambricht_2025a | Vicon | Yes | No | No | SPM + RMSE | 2.9-5.6 |  |  |  |  | Shows running clothes have minimal impact on 2D kinematics accuracy extracted by OpenPifPaf vs tight clothing. | Angle | MoCap | Vicon V5 Motion Systems (8 cameras, 100 Hz) | Global_Kinematic_Error (Deg): 2.9-5.6 (RMSE range across joints)
Huang_2024 | Nokov Mars 4H (13 cameras) | Yes | No | Yes | CMC + RMSE | [5 empty] | Validates moments (kinetics) and reports RMSE in Nm, showing high similarity between markerless and marker-based systems. | Kinetics | Autre | Nokov Mars 4H (13 cameras, 120 Hz) - NOT Vicon. MMC is TsingVA/Pixmotion (8 cameras, 120 Hz)
Lambricht_2025b | Vicon | Yes | No | No | RMSE + ICC |  | ICC > 0.97 for performance, peak knee flexion, and knee ROM across all tasks |  |  |  | OpenPifPaf successfully assesses knee and hip flexion in ACL-focused dynamic jumping tasks from smartphone video. | Angle | MoCap | Vicon V5 (8 cameras, 100 Hz) + Gaitway 3D | Global_Kinematic_Correl: Table 1: ICC > 0.97 for performance, peak knee flexion, and knee ROM across all tasks. CONFIRMED.
Lagomarsino_2025 | BTS | No | No | No | Other | [5 empty] | 2D methods lack sensitivity for upper limb motor assessment in cSCI; 3D MediaPipe improves compensation detection. | Non specifie | MoCap | BTS Smart
Kumar_2022 | OptiTrack | Yes | No | No | MAE + Pearson | [5 empty] | Predicts joint angular velocities using a CNN + LSTM/GRU model. Highly accurate for sagittal knee and hip movements. | Cinematique (non specifie) | MoCap | Optitrack
Kakavand_2025 | Vicon | Yes | No | Yes | RMSE + Pearson | 3.2-12.4 |  |  |  |  | Validates OpenCap kinetics and kinematics during cycling. Reveals RMSE up to 10 degrees due to markerless tracking limitations on the bike. | Angle + Kinetics | MoCap | Vicon Motion Systems Ltd. (10 cameras, 250 Hz) + Sensix pedal force sensors | Global_Kinematic_Error (Deg): 3.2-12.4 (RMSE range across joints during cycling)
John_2024 | Vicon | No | Yes | No | ICC |  |  |  | Pearson r=0.97-0.98; ICC(1,1)=0.964-0.981 | Pearson r=0.98-0.99; ICC(1,1)=0.977-0.994 | Validates purely spatiotemporal parameters. Shows that frontal video overground introduces more errors than sagittal or treadmill. | Spatiotemporel | MoCap | Vicon (10 cameras, Denver CO) - stroke patients, OpenPose | Spatio_StepLength_Error: Pearson r=0.97-0.98; ICC(1,1)=0.964-0.981 (treadmill, all steps; OpenPose vs Vicon; no MAE/RMSE) ¦ Spatio_StepTime_Error: Pearson r=0.98-0.99; ICC(1,1)=0.977-0.994 (treadmill, all steps; OpenPose vs Vicon; no MAE/RMSE)
Ino_2024 | Vicon | Yes | No | No | MAE + ICC | 2.4 (MAE) | 0.86-1.00 |  |  |  | OpenPose is highly accurate (MAE 2.4°) for estimating knee valgus during dynamic drop jump. | Angle | MoCap | Vicon Nexus 2.10 (14 cameras, 120 Hz) - DVJ knee valgus only
Horsak_2025 | Vicon | Yes | No | No | RMSE + Pearson | 5.5 (RMSD) |  |  |  |  | CameraHMR (Monocular) achieves similar validity to OpenCap (Stereo), but struggles with ankle kinematics during non-crouch gaits. | Angle | MoCap | Vicon Nexus 2.14 (16 cameras, 120 Hz) - from Horsak 2023/2024a datasets
Horsak_2024 | Vicon | Yes | No | No | RMSE | [5 empty] | Evaluates inter-trial variability, not accuracy. Found that markerless introduces 14% more variability (noise) than markers. | Cinematique (non specifie) | MoCap | Vicon
Horsak_2023 | Vicon | Yes | No | No | RMSE + Pearson | 5.8 |  |  |  |  | OpenCap errors are between 2 and 11 degrees. Pathological simulation creates significantly higher errors than normal gait. | Angle | MoCap | Vicon Nexus 2.14 (16 cameras, 120 Hz)
Harrison_2025 | Vicon | Yes | No | No | RMSE + Pearson | 7.5 |  |  |  |  | Validates Ergomechanic system using HALPE26 pose model. Good accuracy for knee but poor for hip due to ambiguous pelvis tracking. | Angle | MoCap | Vicon Vero 2.2 (12 cameras, 120 Hz) | Global_Kinematic_Error (Deg): 7.5 (mean RMS difference, walking, base model)
Emmerson_2025 | Qualisys | Yes | No | Yes | RMSE | NR (audit 2026-06) |  |  |  |  | HRNet produces more stable tracking than OpenPose, reducing high frequency errors in internal mechanical work estimations. | Kinetics | MoCap | Qualisys Arqus 12-camera + Kistler 9287 force plates (1000 Hz)
Dunn_2023 | Vicon | Yes | Yes | No | RMSE + FLOA + Pearson | 10.20-20.90 |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | 0.02 - 0.03s | VFI (Video Frame Interpolation) artificially increases FPS to 100Hz, improving spatiotemporal parameter precision, but angular RMSE remains very high. | Angle + Spatiotemporel | MoCap | Vicon MX-T40 (10 cameras, 100 Hz) | Spatio_Speed_Error: confirmed_absent (no spatial calibration; single camera without depth reference) ¦ Spatio_StepLength_Error: confirmed_absent (no spatial calibration)
Drazan_2021 | Qualisys | Yes | No | No | RMSE + CMC | < 3.22 (RMSE) | > 0.85 (CMC) |  |  |  | A custom-trained DeepLabCut model offers excellent agreement for sagittal 2D kinematics during jumping. | Angle | MoCap | Qualisys (15 cameras, MoVe dataset)
DHaene_2024 | OptiTrack | Yes | No | No | RMSE + LCC | < 5.4 | 0.91 - 0.99 (LCC) |  |  |  | YOLOv8 paired with ZED2 depth cameras has a constant bias (offset) against MOCAP, but is highly correlated. | Angle | MoCap | OptiTrack Motive 2.2.0 (9 infrared cameras, 100 Hz)
deBorba_2025 | BTS | Yes | Yes | No | SPM + RMSE + ICC | 1.0 - 1.77 (RMSE) |  |  |  |  | Validates OpenCap across multiple walking speeds. Shows very low global RMSE but higher discrete ROM errors. | Angle | MoCap | BTS SMART DX 100 (8 cameras, 100 Hz)
DAntonio_2020 | Xsens | Yes | No | No | Mean Diff (Deg) | 1.1 - 9.9 |  |  |  |  | Feasibility study against IMUs. Large errors found at the maxima and minima (peaks) of the joint angles. | Angle | MoCap | Xsens MTw IMUs (7 sensors, 60 Hz) - NOT optical MoCap
DAntonio_2021 | Xsens | Yes | No | No | ICC | 1.6 - 14.0 |  |  |  |  | Extension of the 2020 studies. Validates OpenPose + 2 webcams against IMUs. Optimal conditions yield 1.6° error, but can reach 14°. | Angle | MoCap | Xsens MTw IMUs - NOT optical MoCap (same group as DAntonio_2020) | Global_Kinematic_Error (Deg): 1.6 - 14.0 (Mean Diff)
DAntonio_2020b | IMU | Yes | No | No | Other | 1.1 - 9.9 |  |  |  |  | Another conference paper evaluating OpenPose with 2 webcams vs IMUs on 1 subject. High errors at peak angles. | Angle | MoCap | IMU (Inertial Sensors) | Global_Kinematic_Error (Deg): 1.1 - 9.9 (Mean Diff)
CarrascoPlaza_2022 | MoCap (Other) | Yes | No | No | Other | > 5.0 |  |  |  |  | Theoretical evaluation of a 2D to 3D neural network. Conclut que les erreurs dépassent souvent le seuil clinique de 5°. | Angle (global) | MoCap | Human3.6M database (4 high-res + 10 IR optoelectronic cameras, 50 Hz) | Global_Kinematic_Error (Deg): > 5.0 (Mean Joint Angle Error)
Boldo_2024 | Vicon | Yes | No | No | Pearson | 0.2 - 11.8 |  |  |  |  | Provides guidelines for clinical 2D overground gait analysis. Ankle angles show largest bias. | Angle (global) | MoCap | Vicon Nexus v2.14 (MX 13 cameras, 8 IR, 120 Hz) | Global_Kinematic_Error (Deg): 0.2 - 11.8 (Mean Bias)
Bittner_2023 | Qualisys | Yes | No | No | MAE + Pearson | 3.54 - 5.44 | 0.66 - 0.77 |  |  |  | Introduces D3KE (Direct 3D Kinematic Estimation) which reduces MAE by 35% compared to multi-step OpenPose+OpenSim pipelines. | Angle | MoCap | Qualisys OMC from BML-MoVi database (processed through AMASS + OpenSim)
Bertozzi_2025 | BTS | Yes | No | No | SPM + RMSE | 5.0 - 15.6 |  |  |  |  | Validates OpenCap on highly dynamic jump landings. OpenCap shows large RMSE for the hip (11-15°), struggling with high speeds at 60 fps. | Angle | MoCap | SMART-DX EVO (BTS, 8 cameras, 100 Hz). Also Xsens Awinda (7 sensors, 60 Hz) and OpenCap v0.2
Barzyk_2024a | Vicon | Yes | No | No | SPM + MAE | 2.6 - 4.5 (MAE) | 0.94 - 0.96 (r) |  |  |  | Validation of the Sbsq-pose commercial app. Shows acceptable knee/ankle tracking from frontal view, but hip errors exceed 10° (RMSE). | Angle | MoCap | Vicon (12 T40S cameras, 100 Hz)
Barzyk_2024b | Vicon | Yes | Yes | No | SPM + RMSE | 4.8 - 6.8 | 0.94 - 0.98 | 0.01 m/s (MAE) | 0.03 m (3 cm) MAE | 0.05 s (MAE) | Excellent validation of Sbsq-pose algorithm for post-stroke gait analysis using a single smartphone. | Angle + Spatiotemporel | MoCap | Vicon (12 T40S cameras, 100 Hz)
Bak_2025 | Vicon | Yes | No | No | RMSE + ICC | 1.6-10.2 | 0.87-1.00 |  |  |  | Demonstrates that BMI and thigh circumference strongly correlate with increased markerless tracking errors. | Angle | MoCap | Vicon MX-T10 (6 cameras, 100 Hz). 4DEYE (5 RGB cameras, OpenPose) | Global_Kinematic_Error (Deg): 1.6-10.2 (individual RMSE range from Table 1, subjects 3 to 2). Subgroup means: 4.1-7.6 (Table 3). The current 2.4-7.4 may be imprecise. ¦ Global_Kinematic_Correl: 0.87-1.00 (ICC range from Table 1: 0.870 to 0.996). The 0.74 does not appear in Table 1.
Bae_2024 | Qualisys (MIQUS M3, 8 cameras, 100 Hz) | Yes | No | No | RMSE + ICC | 2.33-6.25 | > 0.94 (r) |  |  |  | Ergo system (using MediaPipe and OpenSim) showed excellent agreement for time-series and peak joint angles during overhead squat. | Angle | MoCap | Qualisys MIQUS M3 (8 cameras, 100 Hz) - NOT Vicon. Ergo system (18 RGB cameras, 30 Hz) | Global_Kinematic_Error (Deg): 2.33-6.25 (RMSE range, hip int/ext rot min to shoulder flex/ext max)
Azhand_2021 | None (Reliability) | No | Yes | No | ICC + Pearson |  |  | 0.04 m/s | Hand: 0.15 cm / Stand: 1.20 cm | Hand: 0.0018 s / Stand: 0.0031 s | Excellent validation for spatiotemporal parameters on older adults using a single frontal iPad video. | Spatiotemporel | Walkway/Force | GAITRite pressure-sensitive walkway (5.2 m). NOT optical MoCap. Spatiotemporal only | Spatio_StepLength_Error: Hand: 0.15 cm / Stand: 1.20 cm (Mean diff) ¦ Spatio_StepTime_Error: Hand: 0.0018 s / Stand: 0.0031 s (Mean diff)
Andreoni_2024 | BTS | Yes | Yes | No | Other | 2.7 - 9.9 |  | < 0.11 m/s | < 0.14 cm |  | Validates OpenPose against BTS. The absolute error on joints frequently exceeds the 5-degree clinical limit. | Angle + Spatiotemporel | MoCap | SMART DX (BTS SpA, Milan, 8 cameras, 100 Hz). 22 markers Davis protocol | Global_Kinematic_Error (Deg): 2.7 - 9.9 (Absolute Error)
Anderson_2025 | None (Reliability) | No | Yes | No | Other |  |  | 0.04 m/s | 0.03 m | 0.01 s | Excellent pediatric application. Proves that OpenPose can extract accurate spatiotemporal parameters from a frontal view. | Spatiotemporel | Walkway/Force | ProtoKinetics Zeno Walkway (16 ft, 120 Hz pressure mat). NOT optical MoCap
Aleksic_2024 | Qualisys | Yes | Yes | No | RMSE + Pearson r | [5 empty] | Validates jump height + CoM + lower-limb joint angles (Sagittal RMSE 5.4-8.8°, Pearson r 0.98-0.997). Markerless OpenPose-based system shows excellent validity vs Vicon. | CoM / Performance | MoCap | Qualisys Miqus M3 (5 cameras, 300 Hz). MMPose/RTMPose markerless
Ahn_2024 | Vicon | Yes | No | No | ICC only | 1.8 - 3.8 (MAE) | ICC(3,k) Hip 0.925-0.998, Knee 0.842-0.989, Trunk Rot 0.795-0.962 |  |  |  | OpenPose + custom OpenCV analysis (no OpenSim). ICC-based validation only. Measures Hip + Knee + Trunk Rotation (NOT ankle). | Angle | MoCap | Vicon MXT (8 IR cameras, Nexus 2.x, Shogun 1.x, 250 fps). 4DEYE (5 RGB cameras, 30 fps, SYM Healthcare)
Wang_2024 | Motion Analysis Cortex (14 IR cams, 60 Hz) | Yes | Yes | No | Other | 5.79 - 10.23: Hip 10.23, Knee 6.90, Ankle 5.79 from Fig 4 | 0.80 - 0.95 (Rxy) | 0.10 m/s | 4.0 cm | 0.001 s | Tests different camera angles and distances with HuMoR. 45°/135° angles introduce massive errors compared to strictly sagittal or frontal. | Angle + Spatiotemporel | MoCap | Cortex Ver 7.02.1815, Motion Analysis Corp (14 IR cameras, 60 Hz) - NOT Vicon | Global_Kinematic_Error (Deg): 5.79 - 10.23 (RMSE Joint Sag): Hip 10.23, Knee 6.90, Ankle 5.79 from Fig 4
Verheul_2024 | Vicon | No | No | Yes | Other | [5 empty] | Uses OpenCap kinematics to predict Ground Reaction Forces (GRF). OpenCap smoothing underestimates impact peaks during drop jumps. | Kinetics | Walkway/Force | Vicon Nexus + Kistler force platform. OpenCap GRF estimation study - NOT joint angle validation
Wang_2025 | Vicon | Yes | No | No | RMSE + ICC | 4.7-8.2 | 0.50-0.75 |  |  |  | OpenCap has higher RMSE in patients with Knee OA than healthy controls. Walking away from camera also increases error. | Angle (global) | MoCap | Vicon (10 cameras, Oxford Metrics, 150 Hz). OpenCap markerless. Knee OA + healthy controls | Global_Kinematic_Error (Deg): 4.7-8.2 (Healthy WTC 4.7, Healthy WAC 6.7-6.9, OA WTC 6.0-6.2, OA WAC 8.0-8.2) ¦ Global_Kinematic_Correl: 0.50-0.75 CONFIRMED from Table 3. OA WAC=0.50, Controls WTC=0.75.
Bousigues_2025 | Qualisys | No | No | No | Other | [5 empty] | Proves that 2D-3D triangulation (markerless) generates skeletal inconsistencies (joint dislocations, length variations) equivalent or larger than skin soft-tissue artifacts. | Non specifie | MoCap | Qualisys (10+10 cameras, 300 Hz). OpenPose + Pose2Sim. Skeletal inconsistencies study, NOT joint angle RMSE/MAE
Ueno_2024 | Motion Analysis | Yes | No | Yes | MAE only | 6.6 - 8.5 (MAE) |  |  |  |  | Validates HuMoR (calibration-less single camera) on OpenCap dataset. MAE of 8.5° overall, drops to 6.6° if ankle/hip rotation excluded. | Kinetics + Angle (global) | MoCap | Motion Analysis Corp (8 cameras, 100 Hz). HuMoR single camera. OpenCap dataset
Yagi_2020 | OptiTrack | No | Yes | No | Other |  |  | 0.04 m/s | Stride length 3.20-5.59 cm; Step length 5.12-7.60 cm |  | Proof-of-concept for home monitoring. Uses a homography matrix on the floor plane to estimate spatiotemporal parameters. | Spatiotemporel | MoCap | OptiTrack (16 cameras). Spatiotemporal ONLY. Single RGB camera + OpenPose
Tahara_2025 | GAITRite | No | Yes | No | RMSE + MAE |  |  | XGBoost: MAE 0.115 m/s, RMSE 0.21 m/s, R²=0.834 | Step length: MAE=0.028 m, RMSE=0.038 m, R2=0.709 (XGBoost); Stride length: MAE=0.118 m, RMSE=0.198 m, R2=0.864 (GBR) | XGBoost stride: MAE 0.077 s, RMSE 0.161 s, R²=0.848; GBR stride: MAE 0.095 s | Uses Machine Learning on MediaPipe 2D coordinates to predict spatiotemporal parameters previously acquired by a GAITRite mat. | Spatiotemporel | Walkway/Force | GAITRite (CIR Systems). Spatiotemporal ONLY. MediaPipe + ML. No joint angles | Spatio_Speed_Error: XGBoost: MAE 0.115 m/s, RMSE 0.21 m/s, R²=0.834 (best model)
Francia_2026 | Xsens | Yes | No | No | RMSE + ICC | 22.62-22.63 | 0.67-0.71 (ICC) |  |  |  | Uses an Extended Kalman Filter to merge two 2D MediaPipe outputs. Validates Knee Flexion against IMUs. | Angle | MoCap | XSens IMU (Movella, Enschede). NOT optical MoCap. MediaPipe (2 Insta360 cameras). Knee FE only
Koleini_2025 | Qualisys | Yes | No | No | MAE only | 3.14 - 4.84 |  |  |  |  | BioPose introduces a state-of-the-art End-to-End network leveraging Neural Inverse Kinematics to predict accurate joint angles from a single image. | Angle (global) | MoCap | BML-MoVi dataset (Qualisys OMC, processed through AMASS + OpenSim). Also evaluated on OpenCap dataset (Qualisys + OpenSim). BioPose framework: MQ-HMR + NeuriK + OpenSim
Yamamoto_2021 | OptiTrack | Yes | Yes | No | MAE + ICC | n/a (see Extraction_Notes) | 0.60 - 0.98 (ICC) | MAE=0.01 +/- 0.02 m/s | Step length MAE=0.04 +/- 0.03 m (ICC=0.89-0.94); Stride length MAE=0.03 +/- 0.03 m (ICC=0.98) | Stride time MAE=0.01 +/- 0.01 s | Single-camera OpenPose evaluation. High reliability for hip/knee but fails to accurately capture the ankle (ICC < 0.60) during gait with an orthosis. | Angle + Spatiotemporel | MoCap | OptiTrack (7 infrared cameras, NaturalPoint, Corvallis, OR, 100 Hz). OpenPose. Gait with orthosis | Global_Kinematic_Error (Deg): Reported ¦ Spatio_Speed_Error: MAE=0.01 +/- 0.02 m/s (ICC=0.98, OpenPose vs OptiTrack) ¦ Spatio_StepTime_Error: Stride time MAE=0.01 +/- 0.01 s (ICC=0.98, OpenPose vs OptiTrack)
Hulleck_2024 | Vicon | Yes | No | No | MAE only | 5.3 - 10.8 |  |  |  |  | Incorporating a Seq2Seq deep learning model to BlazePose data significantly improves joint kinematic accuracy from a single camera. | Angle | MoCap | Vicon (10 cameras, 120 Hz). BlazePose + Seq2Seq deep learning
Templin_2024 | Qualisys | Yes | No | Yes | RMSE + Pearson | 2.52-9.61 | n/a (see Extraction_Notes) |  |  |  | Demonstrates that a CNN trained heavily on synthetic (SMPL-X) data yields highly accurate 3D kinematics and kinetics in OpenSim. | Angle + Kinetics | MoCap | Qualisys Miqus (12 cameras, 120 Hz for markerless; 80 cameras, 240 Hz for marker-based). Emory University SPARC. ENABLE markerless system (SwRI). 127 adolescent female athletes. DVJ task | Global_Kinematic_Error (Deg): 2.52-9.61 (RMSE range across all DoFs) ¦ Global_Kinematic_Correl: Reported
Wang_2023 | OptiTrack | Yes | No | No | ICC |  | 0.50 - 0.99 (ICC) |  |  |  | Demonstrates that BlazePose from a single webcam is reliable for 3D ROM estimation, though accuracy drops at the apex of movements. | Cinematique (non specifie) | MoCap | Optitrack
Wade_2023 | Qualisys | Yes | No | No | Other | [5 empty] | Demonstrates that the error of 2D joint angles extracted from the occluded side of the body is double that of the visible side. | Angle | MoCap | Qualisys Oqus (15 cameras, 200 Hz). OpenPose (2D). Two machine vision cameras (JAI sp5000c, 1920x1080). Sagittal and frontal plane. 15 healthy participants
Kim_2025 | OptiTrack | No | Yes | No | MAE + Pearson |  |  | 0.07 m/s | 0.09 m | 0.01 s | Evaluates spatiotemporal parameters in Parkinson's patients. Shows that lateral view is essential for spatial parameters like step length. | Spatiotemporel | MoCap | OptiTrack prime 41/17W (8 cameras, 240 Hz). Motive 2.0.2 software. MediaPipe Pose (smartphone). 24 PD patients. Sagittal plane joint angles from lateral view
Ho_2024 | Stereo Camera | No | Yes | No | MAE + ICC |  |  | 4.9% (velocity) |  |  | Focuses entirely on turn time and spatiotemporal parameter estimation in pathological gait using a stereo camera. | Spatiotemporel | Manuel | ZED camera + OpenPose. Gold standard comparison: independent clinical gait lab measurements. Spatiotemporal only
He_2024 | Vicon | Yes | Yes | No | ICC |  |  | n/a (see Extraction_Notes) | r=0.88-0.92; mean diff d=53.7-63.2 mm; ICC(2,k)=0.87-0.89 | r=0.95-0.96; d=0.02 s; ICC(2,k)=0.93-0.97 | Focuses on spatiotemporal parameters and classification of Parkinson's disease. Validates accuracy of the classification model, but no joint kinematics (angles) are reported. | Spatiotemporel | MoCap | Vicon (12 MX cameras, 100 Hz) + AMTI force plates (1000 Hz). Multi-level 3D pose for PD gait | Spatio_Speed_Error: confirmed_absent (gait speed not directly reported as validated parameter) ¦ Spatio_StepLength_Error: r=0.88-0.92; mean diff d=53.7-63.2 mm (large error); ICC(2,k)=0.87-0.89 (Transformer+GCN vs Vicon) ¦ Spatio_StepTime_Error: r=0.95-0.96; d=0.02 s; ICC(2,k)=0.93-0.97 (single step time; Transformer+GCN vs Vicon)
Balci_2026 | OptiTrack | Yes | No | No | ICC |  | 0.61 - 0.85 |  |  |  | Compares multiple markerless pipelines. MediaPipe 2D outperforms YOLOv8 and MediaPipe 3D for basic joint angle measurements. | Cinematique (non specifie) | MoCap | OptiTrack (16 cameras, 240 FPS, 49 markers). Reliability study of 6 markerless methods
Hii_2023 | Vicon | No | Yes | No | Other |  |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | MAE=0.03-0.04 s; ICC(2,1)=0.832-0.846 | Presents a smoothing algorithm for MediaPipe but lacks a gold standard optical tracking system for angle validation. | Spatiotemporel | MoCap | Vicon | Spatio_Speed_Error: confirmed_absent (temporal parameters only; no spatial calibration for speed) ¦ Spatio_StepLength_Error: confirmed_absent (temporal parameters only; no step length validation) ¦ Spatio_StepTime_Error: MAE=0.03-0.04 s; ICC(2,1)=0.832-0.846 (MediaPipe vs Vicon)
Chen_2025 | Vicon | Yes | Yes | No | Other | 3.63 |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | Proposes an abnormal gait recognition network based on 3D pose estimation. Claims a joint error within 3.63°. | Spatiotemporel + Angle (global) | MoCap | Multi-view RGB cameras + OpenPose + modified 3DPS. Vicon 10 cameras for GPOSE public dataset validation. TUG clinical data has no MoCap gold standard for joint angles | Spatio_Speed_Error: confirmed_absent (classification/anomaly detection focus; gait parameters used as features, no individual spatio validation error metrics) ¦ Spatio_StepLength_Error: confirmed_absent (classification focus; no individual step length error reported) ¦ Spatio_StepTime_Error: confirmed_absent (classification focus; no individual step time error reported)
Fukushima_2024 | Vicon | Yes | No | No | MAE only | 9.0 - 9.7 (Mean) |  |  |  |  | Validates OpenPose on highly dynamic sports movements. MAE is roughly 9 degrees, with the elbow being the main source of error. | Angle | MoCap | Vicon Nexus 2.8.2 (10 IR cameras, 100 Hz). OpenPose (12 RGB cameras, 100 fps). Sports/athletic movements
Ge_2025 | Radiography (X-ray) | Yes | No | No | ICC |  | 0.89 |  |  |  | Validates OpenPose against clinical weight-bearing X-rays for measuring the frontal Hip-Knee-Ankle (HKA) angle. Found very good consistency. | Angle | Autre | X-ray radiography (NOT optical MoCap). HKA angle in KOA patients. OpenPose via smartphone (30 fps) | Global_Kinematic_Correl: 0.89 (ICC vs X-Ray)
Asaeda_2024 | Vicon | Yes | No | No | ICC + Error | 18.8 - 19.6 (Error) | n/a (see Extraction_Notes) |  |  |  | Validates MediaPipe for dynamic knee valgus. MediaPipe significantly overestimates the valgus angle by ~19 degrees. | Angle | MoCap | Vicon MX (100 Hz) + Vicon Vue (100 Hz). MediaPipe Pose. Knee valgus at single-leg drop landing. FRONTAL plane only | Global_Kinematic_Correl: Poor (ICC)
Liang_2022 | Vicon | Yes | No | No | ICC |  | 0.538-0.765 |  |  |  | Validates 3DPoseNet (lifting network) applied to OpenPose 2D data for estimating sagittal plane joint kinematics and sample entropy. | Cinematique (non specifie) | MoCap | Vicon MX (6 cameras, 60 Hz) + 1 Vue video camera (60 Hz). OpenPose + 3DPoseNet. 30 participants (15 elderly + 15 young). Treadmill walking | Global_Kinematic_Correl: 0.538-0.765 (ICC concurrent validity Table 3)
Svetek_2025 | Vicon | Yes | No | No | RMSE + Pearson | 3.16-16.41 |  |  |  |  | Highlights that safety bars on treadmills occlude the smartphone cameras and artificially increase the RMSE of hip/knee angles. | Angle | MoCap | Vicon (10 cameras, 240 Hz). OpenCap (2 iPads, 60 Hz). 20 collegiate ice hockey players. Walking, running, squat, CMJ, drop landing | Global_Kinematic_Error (Deg): 3.16-16.41 (RMSE full range across joints/tasks)
Hu_2025 | Vicon | Yes | No | No | RMSE + ICC | 3.45 - 4.71 | 0.24-0.90 (ICC) |  |  |  | Validates a customized dual-camera Vision Transformer model specifically on osteoarthritic patients. Demonstrates excellent accuracy. | Angle | MoCap | Vicon Vero2.2 (10 cameras, 150 Hz). Dual-camera markerless (Blackfly S USB3, FLIR, triangulation). 152 OA patients (67 hip, 85 knee). OpenSim CGM model
Haberkamp_2022 | Motion Analysis | Yes | No | No | SPM + Pearson |  | n/a (see Extraction_Notes) |  |  |  | Validates OpenPose during a single-leg squat. Finds significant waveform differences at the beginning and end of the squat using SPM. | Cinematique (non specifie) | Autre | Motion Analysis Corp (12 cameras, 120 Hz). OpenPose 2D (2 iPod Touch, 30 Hz). 22 adolescent runners (14F, 8M). Single-leg squat | Global_Kinematic_Correl: Reported
Vafadar_2022 | Vicon | Yes | Yes | No | Other | Joint center position: 13.1 mm |  | MD=0.00 m/s | MD=-0.02 cm; Stride length MD=-0.06 cm (SD=0.81 cm) | MD=0 ms; Stance time MD=1 ms (SD=7 ms) | Evaluates a custom 4-camera deep learning system trained on EOS bi-planar X-ray data. Shows high accuracy for joint centers. | Spatiotemporel + Angle (global) | MoCap | Vicon (7 cameras, 100 Hz). Custom 4-camera DL system (ENSAM dataset). 41 subjects (24 asymptomatic + 15 XLH children + 2 adults). Walking | Global_Kinematic_Error (Deg): Joint center position: 13.1 mm (mean diff, SD 10.2 mm) — N=41 study ¦ Spatio_Speed_Error: MD=0.00 m/s (SD=0.00, MaxD=0.02; custom 4-camera DL vs Vicon) ¦ Spatio_StepLength_Error: MD=-0.02 cm (SD=0.85 cm; step length); Stride length MD=-0.06 cm (SD=0.81 cm) ¦ Spatio_StepTime_Error: MD=0 ms (SD=6 ms, MaxD=20 ms; step time); Stance time MD=1 ms (SD=7 ms)
Washabaugh_2022 | Vicon | Yes | Yes | No | SPM + MAE | 3.7 - 5.1 (MAE) |  | MAE~0.04-0.07 m/s | MAE~0.04 m | MAE~0.02-0.03 s | Compares OpenPose, MoveNet, and DeepLabCut. OpenPose was the most accurate for sagittal hip and knee kinematics. | Angle + Spatiotemporel | MoCap | Vicon MX-T40 (10 cameras, 100 Hz). HM-GPJATK public dataset. 32 able-bodied participants. OpenPose, MoveNet (Lightning/Thunder), DeepLabCut compared | Spatio_Speed_Error: MAE~0.04-0.07 m/s (OpenPose/MoveNet Thunder best; vs Vicon) ¦ Spatio_StepLength_Error: MAE~0.04 m (OpenPose; vs Vicon) ¦ Spatio_StepTime_Error: MAE~0.02-0.03 s (OpenPose/MoveNet Thunder; vs Vicon)
Yeom_2026 | OptiTrack (7 cams, 200 Hz) | Yes | No | No | RMSE, CMC, AUC | Lower body grand mean RMSE 5.45°; Upper body 36.7° | Lower body grand mean CMC 0.93; Upper body grand mean CMC 0.48 |  |  |  | OpenCap accurate for lower body kinematics during overhead squat; upper body accuracy questionable. Classification AUC comparable to marker-based. | Angle | MoCap | Optitrack | Global_Kinematic_Error (Deg): Lower body grand mean RMSE 5.45° (Sag 3.81-8.16; Front 2.24-3.26; Trans 5.02-5.36); Upper body 36.7° ¦ Global_Kinematic_Correl: Lower body grand mean CMC 0.93 (Sag 0.97-0.99, Front 0.76-0.97, Trans 0.66-0.83); Upper body grand mean CMC 0.48
Gao_2025 | Vicon | Yes | Yes | No | MAE, Accuracy, Sensitivity, Specificity | n/a (see Extraction_Notes) |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | BlazePose can identify gait deviations in CP children; sensitivity >80% for most gait features. | Spatiotemporel | MoCap | Vicon | Global_Kinematic_Error (Deg): NR (deviation detection focus) ¦ Spatio_Speed_Error: confirmed_absent (deviation identification focus; sensitivity/specificity reported, no individual spatio error metrics) ¦ Spatio_StepLength_Error: confirmed_absent (deviation identification focus) ¦ Spatio_StepTime_Error: confirmed_absent (deviation identification focus)
Kumtekar_2024 | Qualisys 10-camera | Yes | No | No | Peak-angle differences (deg) | Shoulder peak diff ~8-10°; Elbow ~1-5° | ICC > 0.90 |  |  |  | Mobile app with MediaPipe provides clinically acceptable ROM measurements for upper extremity. | Angle (global) | MoCap
Nishizawa_2024 | OptiTrack | Yes | Yes | No | RMSE, correlation | Hip: ~5°; Knee: ~7°; Ankle: ~6° | r > 0.85 | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | OpenPose can distinguish hemiplegic gait patterns; useful for rehab monitoring. | Angle + Spatiotemporel | MoCap | Optitrack | Global_Kinematic_Correl: r > 0.85 (hip, knee) ¦ Spatio_Speed_Error: confirmed_absent (not validated; focus is on joint angles and classification) ¦ Spatio_StepLength_Error: confirmed_absent (step length not validated as individual error metric; paper focuses on knee joint angle and hemiplegia classification) ¦ Spatio_StepTime_Error: confirmed_absent (not validated; focus is on joint angles and classification)
Todesca_2025 | OptiTrack | Yes | No | No | RMSE, MAE | Upper body 3.0-6.9°; Knees 8.7-11.4° (MAJAE) | r > 0.90 |  |  |  | YOLO11x-pose suitable for cycling posture estimation with acceptable accuracy. | Angle | MoCap | Optitrack
Lin_2024 | Marker-based MoCap + OpenSim IK | Yes | No | No | MPJPE, joint angle error | Mean MAE 8.06° across 3 datasets |  |  |  |  | Biomechanical model + synthetic data improves 3D kinematics estimation from video. | Angle + Position 3D (mm) | Dataset |  | Global_Kinematic_Error (Deg): Mean MAE 8.06° across 3 datasets (OpenCap 9.65 / BMLmovi 9.73 / ODAH 4.81)
Zhang_2024 | None (Benchmark) | Yes | No | No | MPJPE, joint angle error | Joint angle error: ~4-6° (GNN) |  |  |  |  | GNN-based multi-view 3D reconstruction with OpenSim IK shows promise for markerless kinematics. | Angle + Position 3D (mm) | Dataset
Hii_2023b | Vicon | No | n/a (see Extraction_Notes) | No | MAE, RMSE, ICC | Knee frontal: OpenPose 5.2°, YOLOv7 4.8°, MediaPipe 4.1° | ICC > 0.80 |  |  |  | MediaPipe best for frontal plane knee angle; all 3 methods clinically acceptable for screening. | Angle | MoCap | Vicon | Validates_Spatiotemp: Yes (step time only)
Pappas_2024 | Vicon MX-T40 (validation dataset GPJATK) | Yes | Yes | No | Correlation, AUC, sensitivity, specificity | n/a (see Extraction_Notes) | r > 0.70 | ICC(2,1) 0.82-0.94 | ICC(2,1) 0.60-0.85 | ICC(2,1)=0.99; Pearson r=0.99 | MediaPipe + ZoeDepth can distinguish PD patients from healthy controls using frontal-view gait features; AUC > 0.85. | Spatiotemporel | Aucun |  | Global_Kinematic_Error (Deg): NR (gait feature focus) ¦ Global_Kinematic_Correl: r > 0.70 (gait features vs UPDRS) ¦ Spatio_Speed_Error: ICC(2,1) 0.82-0.94 (Table 1) ¦ Spatio_StepLength_Error: ICC(2,1) 0.60-0.85 (Table 1) ¦ Spatio_StepTime_Error: ICC(2,1)=0.99; Pearson r=0.99 (both PCM and ZoeDepth; vs MoCap)
Irfan_2025 | MoCap (GPJATK dataset — Kwolek et al. 2019) | Yes (hip, knee, ankle angles) | No | No | RMSE (deg) | Average RMSE 7.8±3.9° across all joints | NR | NR | NR | NR | Markerless BlazePose system viable for clinical gait angle assessment; left hip most accurate, right ankle least accurate. | Hip / knee / ankle joint angles (sagittal) | Multi-view MoCap dataset | GPJATK dataset: 4 calibrated synchronized cameras 960×540@25fps + 10 MoCap MX-T40 cameras 2352×1728@100Hz | Global_Kinematic_Error (Deg): Average RMSE 7.8±3.9° across all joints (Table III, 31 subjects)
Liang_2023 | Vicon | Yes | Yes | No | RMSE, Pearson correlation | Hip F/E: ~4°; Knee F/E: ~5°; Ankle: ~5° | r = 0.97 (hip F/E); r = 0.92 (knee F/E); r = 0.85 (ankle) | n/a (see Extraction_Notes) | Mean diff ~1 cm | n/a (see Extraction_Notes) | RGB cameras with triangulation + OpenSim can distinguish elderly gait patterns; correlation strong for hip/knee. | Angle + Spatiotemporel | MoCap | Vicon | Spatio_Speed_Error: confirmed_absent (no explicit speed error metric; paper reports mean values with t-test showing no significant difference) ¦ Spatio_StepLength_Error: Mean diff ~1 cm (elderly: system 55cm vs Vicon 56cm; t-test p>0.05, no significant difference; no MAE/RMSE reported) ¦ Spatio_StepTime_Error: confirmed_absent (step time not reported as separate error metric; stride length validated but not temporal parameters individually)
Cerciola_2024 | Vicon | Yes | No | No | RMSE, ROM error, ANOVA | Hip: IMU 2.8-3.4°, MP 3.3-4.8°, RTMH 4.3-6.8°; Knee: IMU 3.3-3.7°, RTMH 4.0-4.8°, MP 4.8-5.6° |  |  |  |  | IMU best overall; among vision methods MP and RTMH (RTMPose) best for hip/knee. Ankle unreliable with video. Running degrades all vision methods. | Angle | MoCap | Vicon
Zhang_2025 | Vicon / OptiTrack / Qualisys (multi-site) | Yes | Yes | No | RMSE, MAE, R², BBS prediction | Hip F/E: 9.14° ; Knee F/E: 7.72° ; Ankle D/P: 3.03° |  | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) | Theia3D MoCap + ML predicts BBS scores (R²=0.82) from TUG kinematics; identifies fall risk in older adults. | Angle + Spatiotemporel | MoCap | Vicon (200Hz) + OptiTrack (200Hz) + Qualisys Miqus M3 (200Hz) | Spatio_Speed_Error: NR (gait speed extracted as ML feature for BBS prediction but not validated individually against gold standard; RMSE for kinematic angles only: hip 9.14, knee 7.72, ankle 3.05 degrees) ¦ Spatio_StepLength_Error: NR (step length extracted as ML feature for BBS prediction but not validated individually) ¦ Spatio_StepTime_Error: NR (temporal parameters extracted as ML features but not individually validated)
Kidzinski_2020 | Optical motion capture (clinical gait analysis) | Yes (knee flexion at max ext) | n/a (see Extraction_Notes) | No | Correlation (r), MAE | 4.8° MAE knee flexion at max ext | r=0.83 knee flexion; r=0.75 GDI | r=0.73 walking speed (MAE 0.13 m/s) | NR | r=0.79 cadence | Video-based DL achieves correlations approaching theoretical upper bounds; viable for clinical/home use | Continuous gait metrics + clinical scores (GDI, GMFCS) | Optical MoCap (marker-based, full body) | Clinical 3D MoCap, Gillette Children's; reflective markers; gait cycle detection semi-manual; IK | Validates_Spatiotemp: Yes (walking speed, cadence) ¦ Spatio_StepTime_Error: r=0.79 cadence (MAE 0.08 strides/s)
Hwang_2026 | Optical motion capture + force plate (Motion Analysis Corp. + Kistler 9260AA6) | Yes (knee flexion, trunk flexion angles) | n/a (see Extraction_Notes) | Yes (peak STS power vs force plate) | RMSE, ICC, Pearson r | ICC 0.83-0.89 joint angles | ICC peak STS power = 0.94 | NR | NR | NR | Peak STS power from 2D pose correlates strongly with muscle strength (knee extensor r=0.64, handgrip r=0.64, ASMI r=0.70); supports AWGS framework for sarcopenia | Joint angles + kinetics (peak power) | Optical MoCap full body + force plate | Motion Analysis Corp. 3D OMC, Helen Hayes 31-marker set, Orthotrak 6.6 sync, 120 Hz; Kistler force plate single under both feet | Validates_Spatiotemp: Yes (peak STS power timing)
Shah_2025 | Qualisys (marker-based MoCap) | Yes (knee joint angles) | No | No (theoretical equations only) | RMSE, percentage error | Qualisys < 5% error; MediaPipe higher variability at peak flexion | n/a (see Extraction_Notes) | N/A | N/A | N/A | Qualisys most accurate; markerless (MediaPipe/OpenCap) viable but slightly less precise at peak flexion. Anthropometry affects knee biomechanics significantly. | Knee joint angles (sagittal) | Qualisys marker-based MoCap | Qualisys QTM, reflective markers, 100 Hz, 0.6 mm calibration accuracy | Global_Kinematic_Error (Deg): Qualisys < 5% error (gold std); MediaPipe higher variability at peak flexion ¦ Global_Kinematic_Correl: NR (qualitative comparison)
Rode_2025 | Optical motion capture (passive marker, 27 IR cameras, Plug-in Gait) | Yes (knee flexion, elbow flexion) | No | No | MPJPE (mm), MAE joint angles (deg), Inference speed (FPS) | Knee MAE 9.3-21.9 (2D) / 14.1-25.8 (3D); Elbow MAE 21.5-28.9 (2D) / 16.3-26.0 (3D) | NR | N/A | N/A | N/A | Some 2D estimators achieve accuracy comparable to visual assessment; large variability across estimators and image dimensions. Useful for clinical applications under constraints. | MPJPE (mm) + joint angles (knee, elbow flexion) | Marker-based OMC (Plug-in Gait, 46 markers) | 27 IR cameras, Swiss Center for Movement Analysis, Plug-in Gait/Conventional Gait Model, full body, 2 expert taping | Global_Kinematic_Error (Deg): Knee MAE 9.3-21.9 (2D) / 14.1-25.8 (3D); Elbow MAE 21.5-28.9 (2D) / 16.3-26.0 (3D) — range across 11 estimators
Irfan_2025b | MoCap (GPJATK dataset — Kwolek et al. 2019) | Yes (hip, knee, ankle angles) | No | No | MAE (deg) | Hip R 3.6°, Hip L 7.8°, Knee R 9.1°, Knee L 3.8°, Ankle R 9.6°, Ankle L 7.8° | NR | NR | NR | NR | Markerless BlazePose system viable for clinical gait angle assessment; left hip most accurate, right ankle least accurate. | Hip / knee / ankle joint angles (sagittal) | Multi-view MoCap dataset | GPJATK dataset: 4 calibrated synchronized cameras 960×540@25fps + 10 MoCap MX-T40 cameras 2352×1728@100Hz
Talaa_2025 | Vicon MX-T40 (10 cameras) | Yes (knee angles) | No | No | MAE (deg) | Knee MAE 5.67° (avg). Per condition: C1-LeftKnee 6.61°, C1-RightKnee 5.17°, C3-LeftKnee 6.61°, C3-RightKnee 3.88° | NR | NR | NR | NR | MediaPipe-based markerless method achieves MAE 5.67° on knee angles vs Vicon; viable for biomechanics/sports medicine/physiotherapy. | Knee angles only (sagittal) | Vicon MoCap | Vicon MX-T40, 10 cameras, 2352×1728 px @ 100 Hz; MXGiganet sync with video
Ang_2024 | Vicon Optical-passive MoCap | Yes (joint angles + angular velocities) | No | No | Correlation coefficient + R² (linear regression) | n/a (see Extraction_Notes) | 0.96-0.99; 0.84-0.94 |  |  |  | OpenPose-based markerless provides comparable joint angles to Vicon at 3 squat speeds. Suitable for home-based telerehabilitation. | Angle + Angular Velocity | Optical MoCap (Vicon) | Vicon Optical-passive MoCap system (specifics NR) | Global_Kinematic_Error (Deg): NR (correlation-based validation, no degree errors reported) ¦ Global_Kinematic_Correl: 0.96-0.99 (joint angles); 0.84-0.94 (angular velocities)
Hou_2026 | Vicon Valkyrie V16 (12 infrared cameras) | Validates virtual marker positions (downstream angles via OpenSim possible but not directly reported in deg) | No | No | Marker position error (cm) | n/a (see Extraction_Notes) |  | Tested at 5 walking + 5 running speeds |  |  | Multi-view unsynchronized markerless system achieves 3.05 cm overall marker localization error. Innovative: temporal augmentation + virtual markers + transformer fusion. Future work to integrate OpenSim downstream for clinical use. | 3D virtual marker positions (cm) | Optical MoCap (Vicon Valkyrie V16) | Vicon Valkyrie V16, 12 infrared cameras | Global_Kinematic_Error (Deg): NR (positions only) ¦ Spatio_Speed_Error: Tested at 5 walking + 5 running speeds (not validated as parameter)
Athama_2025 | MoVi dataset ground truth (MoCap + IMU) | Yes (joint angles, 4 joints: L/R Elbows + L/R Knees) | No | No | MPJAE + DTW (Dynamic Time Warping) | n/a (see Extraction_Notes) |  |  |  |  | Wireless single-camera markerless framework achieves viable joint angle estimation across SNR conditions. Privacy-preserving rendering via SMPL+AMASS. | Angle (MPJAE) + temporal consistency (DTW) | MoCap (via MoVi public dataset) | MoVi public dataset (90 subjects, MoCap + IMU + video). No direct MoCap deployment | Global_Kinematic_Error (Deg): MPJAE varies with SNR (Figs 2-3, joint-specific values at different channel conditions)
Tsintzira_2025 | Fukuchi 2018 public dataset (treadmill kinematics + kinetics averages) | Yes (joint trajectories vs database averages) | No | Yes (Joint torques + GRF vs database) | RMSE (% Body Weight for GRF) + qualitative trajectory comparison | n/a (see Extraction_Notes) |  | NR |  |  | Low-cost stereo camera + minimal 2D model achieves IK + ID + GRF estimation without force plate. Results match Fukuchi public dataset averages. Future: multi-subject validation, direct MoCap comparison. | Angle + Kinetics (torques + GRF) | Public dataset averages (Fukuchi) | Fukuchi et al. 2018 — public dataset of overground and treadmill walking kinematics and kinetics in healthy individuals (PeerJ) | Global_Kinematic_Error (Deg): Qualitative trajectory match (no per-joint RMSE in degrees)

## 6. Joint_Validation
Study_ID | Error_Metric | Correl_Metric | Pelvis_Err_Sag | Pelvis_Err_Front | Pelvis_Err_Trans | Hip_Err_Sag | Hip_Err_Front | Hip_Err_Trans | Knee_Err_Sag | Knee_Err_Front | Knee_Err_Trans | Ankle_Err_Sag | Ankle_Err_Front | Ankle_Err_Trans | Pelvis_Corr_Sag | Pelvis_Corr_Front | Pelvis_Corr_Trans | Hip_Corr_Sag | Hip_Corr_Front | Hip_Corr_Trans | Knee_Corr_Sag | Knee_Corr_Front | Knee_Corr_Trans | Ankle_Corr_Sag | Ankle_Corr_Front | Ankle_Corr_Trans | Extraction_Notes
Salehi_2026 | RMSE |  | ~2.0 | ~2.0 | ~2.0 | ~3.0 | ~2.0 | ~7.0 | ~5.0 |  |  | ~4.0
Verheul_2026 | RMSE |  |  |  |  | 6.7-8.2 | 3.6-5.0 |  | 5.4-5.9 |  |  | 5.1-7.2 | 8.0-8.6
Natraj_2025 | PA-MPJPE | [16 empty] | Hip Pearson r=0.768-0.816 (R/L) |  |  | Knee Pearson r=0.840-0.884 (R/L)
Zhu_2023 | PE (%) |  |  |  |  | 25.57 |  |  | 17.33
Zago_2020
Young_2023 | Mean Err | ICC | [6 empty] | 1.4 (L) / 12.5 (R) | [11 empty] | 0.96-0.98
Verhoeven_2025 | RMSE | ICC |  |  |  | Adults 75%: 2.40-2.93 ; Toddlers 75%: 7.09-10.57 |  |  | Adults 75%: 0.78-0.97 ; Toddlers 75%: 5.71-21.35 |  |  | Adults 75%: 2.22-6.74 ; Toddlers 75%: 9.07-17.37 | [5 empty] | 0.94-0.96 (Adults 75%) |  |  | 1.00 (Adults 75%) |  |  | 0.86-0.96 (Adults 75%)
Turner_2024 | RMSE (Deg) | CMC |  |  |  | 4.2-5.7 | 2.4-5.2 | 3.73-5.16 | 5.0-6.0 |  |  | 5.9-6.9 | 4.8-5.4 |  |  |  |  | 0.94-0.98 | 0.65-0.78 | 0.51-0.60 | 0.94-0.98 |  |  | 0.84-0.93 | 0.47-0.54
Uhlrich_2023 | MAE (Deg)
Torvinen_2024 | Mean Diff (Deg) |  |  |  |  | 4.21 (G1) / 3.56 (G3) |  |  | 3.27 (G1) / 3.01 (G3) |  |  | 4.96 (G1) / 4.69 (G3) | [5 empty] | 0.94 (G1) / 0.97 (G3) |  |  | 0.96 |  |  | 0.82-0.84 (Pearson) / 0.77-0.80 (ICC)
Takeda_2020 |  | ICC | [15 empty] | 0.97 |  |  | 0.92 |  |  | 0.51
Stenum_2021 | [5 empty] | 4.0 |  |  | 5.6 |  |  | 7.4
Pagnon_2021
Pagnon_2022 |  | CMC |  |  |  | 5.6 |  |  | 5.1 |  |  | 4.0 | [5 empty] | 0.96 (Walk) / 0.65 |  |  | > 0.90 |  |  | 0.90 (Walk) / 0.75 (Cycle) |  |  | Hip_Err_Sag: 5.6 (Walking RMSE) ¦ Knee_Err_Sag: 5.1 (Walking RMSE) ¦ Ankle_Err_Sag: 4.0 (Walking RMSE) ¦ Hip_Corr_Sag: 0.96 (Walk) / 0.65 (Run, hip 15° offset) — CMC ¦ Ankle_Corr_Sag: 0.90 (Walk) / 0.75 (Cycle) — CMC
Needham_2021
Molteni_2025 | MAE (Deg) | ICC |  |  |  | 5.41 (L) / 3.19 (R) |  |  | Right 5.81 / Left 4.98 |  |  | Right 3.61 / Left 4.97
Martis_2024 | RMSE (Deg) |  | 6.6 (WTC) / 6.9 (WAC) | 2.8 (WTC) / 5.3 (WAC) [obliquity] | 2.9 (WTC) / 3.8 (WAC) [rotation] | 8.9 (WTC) / 8.0 (WAC) | 3.5 (WTC) / 6.1 (WAC) [adduction] | 4.0 (WTC) / 4.7 (WAC) [rotation] | 5.6 (WTC) / 5.8 (WAC) |  |  | 4.7 (WTC) / 11.9 (WAC) | [14 empty] | Ankle_Err_Sag: 4.7 (WTC) / 11.9 (WAC) — large direction dependence
Sabo_2022
Panconi_2025
Ota_2021 |  | ICC | [15 empty] | 0.016-0.864 |  |  | 0.234-0.934 |  |  | 0.107-0.677 |  |  | Knee_Corr_Sag: 0.234-0.934 (ICC across walking/running) ¦ Ankle_Corr_Sag: 0.107-0.677 (ICC range across DF/PF/ROM)
Nguyen_2023 | MPJAE (degrees) |  |  |  |  | 6.90 (L) / 7.18 (R) |  |  | 3.66 (R) / 4.61 (L) | [17 empty] | Hip_Err_Sag: 6.90 (L) / 7.18 (R) — MPJAE Ours, Roofing dataset ¦ Knee_Err_Sag: 3.66 (R) / 4.61 (L) — MPJAE Ours
Needham_2022 | Mean Diff (Deg) |  |  |  |  | 0.1-2.8 |  |  | 2.3-3.9 (bias) |  |  | 0.7-1.9 | [14 empty] | Hip_Err_Sag: 0.1-2.8 (Flex/Ext) ¦ Ankle_Err_Sag: 0.7-1.9 (Plantar/Dorsi)
Moro_2020
Moro_2022
Mercadal_2024 | RMSE (Deg) |  |  |  |  | < 10 |  |  | < 10 |  |  | < 10
Menychtas_2023 | [5 empty] | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) | [14 empty] | Hip_Err_Sag: Reported ¦ Knee_Err_Sag: Reported ¦ Ankle_Err_Sag: Reported
Mehdizadeh_2021
Lim_2022 | Mean Bias (Deg) |  |  |  |  | NR (audit 2026-06) |  |  | NR (audit 2026-06)
Embry_2022
Lima_2024 | RMSE (Deg) | ICC | 5.1-9.4 |  |  | 5.7-12.2 |  |  | 4.2-7.1 |  |  | 7.9-15.7 |  |  | 0.41-0.98 |  |  | 0.80-0.99 |  |  | 0.67-0.99 |  |  | 0.86-0.98 |  |  | Pelvis_Err_Sag: 5.1-9.4 (RMSE Trunk FE across 6 tasks, Fig 1) ¦ Hip_Err_Sag: 5.7-12.2 (RMSE Hip FE across 6 tasks, Fig 1)
Li_2024 | RMSE (Deg) | n/a (see Extraction_Notes) |  |  |  | 1.425-1.504 |  |  | 1.917-1.927 |  |  | 2.565-3.117 | [5 empty] | 0.98 |  |  | 0.98 |  |  | 0.90 |  |  | Correl_Metric: Not reported (paper uses % frames <5° error, NOT Pearson)
Ino_2023 | MAE (Deg) | CMC |  |  |  | 2.5-3.5 |  |  | 2.3-3.1 (MAE R/L) |  |  | 3.1-4.1 | [5 empty] | 0.945-0.978 (CMC R/L) |  |  | 0.988-0.994 |  |  | 0.890-0.936 (CMC R/L)
Lambricht_2025a | RMSE (Deg) |  |  |  |  | 5.6 |  |  | 2.9 |  |  | 3.5
Huang_2024 |  | CMC |  |  |  | 5.3 (RMSD) | 6.1 (RMSD) | 8.5 (RMSD) | 6.8 (RMSD) | 6.4 (RMSD) | 9.5 (RMSD) | 6.5 (RMSD) | 10.5 (RMSD) | 11.1 (RMSD) |  |  |  | >0.75 |  |  | >0.75 | 0.58 |  | >0.75 | 0.45 |  | Hip_Corr_Sag: >0.75 (CMC sagittal) ¦ Knee_Corr_Sag: >0.75 (CMC sagittal) ¦ Knee_Corr_Front: 0.58 (CMC ab/ad moment) ¦ Ankle_Corr_Sag: >0.75 (CMC sagittal) ¦ Ankle_Corr_Front: 0.45 (CMC ab/ad moment)
Lambricht_2025b | RMSE (Deg) | ICC |  |  |  | 3.1 |  |  | 2.8 |  |  | 3.9 | [8 empty] | 0.936-0.981
Lagomarsino_2025
Kumar_2022 | MAE (rad/s), Pearson r | Pearson |  |  |  | 15.13 |  |  | 18.09 |  |  | 5.83 | [5 empty] | 0.94 |  |  | 0.96 |  |  | 0.92 |  |  | Hip_Err_Sag: 15.13 (MAE angular velocity, GRU) ¦ Knee_Err_Sag: 18.09 (MAE angular velocity, GRU) ¦ Ankle_Err_Sag: 5.83 (MAE angular velocity, GRU)
Kakavand_2025 | RMSE (Deg) | Pearson |  |  |  | 7.9 | 3.2 | 3.0 | 10.2 |  |  | 12.4
John_2024 | [17 empty] | Cross-correlation r (treadmill): Hip=0.86-0.96 |  |  | Knee=0.94-0.99 |  |  | 0.51-0.77
Ino_2024 | MAE (Deg) | ICC | [7 empty] | 2.4 | 2.4 | [10 empty] | 0.83 | n/a (see Extraction_Notes) |  |  |  | Knee_Corr_Front: 0.83 (CMC OpenPose, range 0.53-0.99) ¦ Knee_Corr_Trans: Reported
Horsak_2025 | RMSE (Deg) |  | 5.2 |  | 3.1 | 6.7 | 4.5 | 5.6 | 4.3 |  |  | 8.4
Horsak_2024
Horsak_2023 | RMSE (Deg) |  | 4.2-6.0 |  |  | 5.4-7.6 | 10.2 | 7.0 | 5.7-8.5 |  |  | 4.6-7.9 | [14 empty] | Pelvis_Err_Sag: 4.2-6.0 (RMSE range across 4 gait conditions) ¦ Hip_Err_Sag: 5.4-7.6 (RMSE range across 4 gait conditions) ¦ Ankle_Err_Sag: 4.6-7.9 (RMSE range across 4 gait conditions)
Harrison_2025 | RMSE (Deg) |  |  |  |  | 11.8 | 5.0 | 6.0 | 6.5 |  |  | 8.9
Emmerson_2025
Dunn_2023 | RMSE (Deg) | [7 empty] | 10.20-11.85 |  |  | 16.52-20.90
Drazan_2021 | RMSE (Deg) | CMC |  |  |  | < 3.22 |  |  | < 3.22 |  |  | < 3.22 | [5 empty] | 0.85 |  |  | 0.96 |  |  | 0.97
DHaene_2024 | RMSE (Deg) | LCC |  |  |  | 4.14-5.40 |  |  | 4.10-4.85 | [8 empty] | 0.91 |  |  | 0.96
deBorba_2025 | RMSE (Deg) |  |  |  |  | 1.0-1.52 |  |  | 1.0-1.52 |  |  | 1.52-1.77
DAntonio_2020 | Mean Diff (Deg) |  |  |  |  | 1.1-6.1 |  |  | Knee FE sagittal: 3.8-9.1
DAntonio_2021 | Mean Diff (Deg) |  |  |  |  | 2.1-6.1 |  |  | 6.4-9.9
DAntonio_2020b | Mean Diff (Deg) |  |  |  |  | 1.1-6.1 |  |  | 3.8-9.1
CarrascoPlaza_2022 | [5 empty] | 15.7 (R) / 12.7 (L) |  |  | n/a (artefact) |  |  | 9.5 (R) / 5.6 (L) | [14 empty] | Knee_Err_Sag: 0.0004 (R) / 0.00004 (L) — unusually low, paper-specific calc
Boldo_2024 | Mean Bias (Deg)
Bittner_2023 | MAE (Deg) | Pearson |  |  |  | 2.54 |  |  | 3.96 | [8 empty] | 0.77 |  |  | 0.84
Bertozzi_2025 | RMSE (Deg) | CMC |  |  |  | 21.6-23.7 | 5.2-5.4 | 8.4-9.5 | 6.9-7.8 |  |  | 9.9-11.1 | [5 empty] | 0.85-0.87 | 0.57-0.71 | 0.45 - 0.55 | 0.96-0.98 |  |  | 0.93-0.94 |  |  | Hip_Err_Front: 5.2-5.4 (OpenCap Hip Add/Abd) ¦ Hip_Err_Trans: 8.4-9.5 (OpenCap Hip Int/Ext Rot RMSE) ¦ Knee_Err_Sag: 6.9-7.8 (OpenCap Knee Flex/Ext RMSE Tables 5-6) ¦ Ankle_Err_Sag: 9.9-11.1 (OpenCap Ankle Flex/Ext RMSE Tables 5-6) ¦ Hip_Corr_Sag: 0.85-0.87 (CMC OpenCap) ¦ Hip_Corr_Front: 0.57-0.71 (CMC OpenCap) ¦ Knee_Corr_Sag: 0.96-0.98 (CMC OpenCap) ¦ Ankle_Corr_Sag: 0.93-0.94 (CMC OpenCap)
Barzyk_2024a | MAE (Deg) | Pearson |  |  |  | 2.6 |  |  | 2.1 |  |  | 4.5 | [5 empty] | 0.96 |  |  | 0.99 |  |  | 0.87
Barzyk_2024b | RMSE (Deg) | Pearson |  |  |  | 3.5 | 4.2 |  | 3.7 | 3.9 |  | 4.6 | [5 empty] | 0.93 |  |  | 0.94 |  |  | 0.79
Bak_2025 | RMSE (Deg) | ICC |  |  |  | 1.634-10.181 | 2.3-5.1 | [10 empty] | 0.870-0.996 | 1) for hip flexion ROM: 0.870-0.996 across subjects | [7 empty] | Hip_Corr_Sag: 0.870-0.996 (ICC range, pooled) ¦ Hip_Corr_Front: 1) for hip flexion ROM: 0.870-0.996 across subjects. The sheet range 0.86-0.89 is too narrow - it misses the high-end ICCs (up to 0.996). Table 1 also includes Cronbach alpha and other reliability metrics. The Hip Front ICC=0.74-0.98 is NOT in the paper - no frontal plane ICC is reported.
Bae_2024 | RMSE (Deg) | R² + ICC(2,1) | NR (audit 2026-06) |  |  | 4.7 | 2.5 | 2.33 | 3.25 |  |  | NR (audit 2026-06) | [5 empty] | ICC(2,1)=0.99 | [8 empty] | Hip_Corr_Sag: ICC(2,1)=0.99 (Table 4)
Azhand_2021
Andreoni_2024 | Mean Bias (Deg) |  |  |  |  | 4.71-6.98 | 4.7-14.07 |  | 3.75-3.98 |  |  | 1.92-2.15 | [14 empty] | Hip_Err_Front: 4.7-14.07 (poor ICC 0.108-0.293)
Anderson_2025
Aleksic_2024 | [5 empty] | 8.0 (L) / 8.8 (R) |  |  | 6.9 (L) / 7.163 (R) |  |  | 5.4 (L) / 6.49 (R) | [5 empty] | 0.997 (L) / 0.996 (R) |  |  | 0.994 (L+R) |  |  | 0.984 (L) / 0.981 (R)
Ahn_2024 | ICC (3,k) | ICC |  |  |  | NR (audit 2026-06) | NR (audit 2026-06) |  | 2.2 |  |  | NR (audit 2026-06) | [5 empty] | 0.925-0.998 (ICC) | n/a (see Extraction_Notes) |  | 0.842-0.989 (ICC) |  |  | n/a (trunk, not ankle) |  |  | Hip_Corr_Front: The paper does NOT measure hip frontal plane angles. Only hip flexion (sagittal), knee flexion (sagittal), and trunk rotation (transverse) are measured. The 0.94 is not traceable to any metric. ¦ Ankle_Corr_Sag: The paper does NOT measure ankle angles. The third measurement is TRUNK ROTATION (transverse plane). Table 2 ICC(3,k) for Trunk Rotation: Left=0.909-0.982, Right=0.832-0.992 across subjects. The 0.85 could come from the lower end of trunk rotation ICC (Subject 2 Right=0.832). Sheet incorrectly labels this as 'Ankle'.
Wang_2024 | RMSE | Pearson |  |  |  | 10.23 |  |  | 6.90 |  |  | 5.79 | [5 empty] | 0.90 |  |  | 0.95 |  |  | 0.80
Verheul_2024
Wang_2025 | RMSE (Deg) | ICC |  |  |  | 5.1-9.3 |  |  | 4.0-8.6 |  |  | 4.6-10.4
Bousigues_2025
Ueno_2024 | MAE (Deg)
Yagi_2020
Tahara_2025
Francia_2026 | RMSE (Deg) | ICC | [6 empty] | Right KFE RMSE=22.62(+/-3.45), Left KFE RMSE=22.63(+/-3.18) | [11 empty] | Table 1 ICC: Right=0.71(+/-0.07), Left=0.67(+/-0.08) | [5 empty] | Knee_Err_Sag: Table 1: Right KFE RMSE=22.62(+/-3.45), Left KFE RMSE=22.63(+/-3.18). Values CONFIRMED. However, gold standard is XSens IMU (NOT optical MoCap). RMSE is very high (~22 deg) indicating poor accuracy. ¦ Knee_Corr_Sag: Table 1 ICC: Right=0.71(+/-0.07), Left=0.67(+/-0.08). Values CONFIRMED. Also Spearman rho: Right=0.69(+/-0.08), Left=0.68(+/-0.09). Bias: Right=7.5(+/-6.70), Left=8.90(+/-5.37).
Koleini_2025 | MAE (Deg)
Yamamoto_2021 | MAE (Deg) | ICC |  |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) | [5 empty] | 0.35-0.85 |  |  | 0.41-0.87 |  |  | 0.01-0.37 |  |  | Hip_Err_Sag: Reported ¦ Knee_Err_Sag: Reported ¦ Ankle_Err_Sag: Reported ¦ Hip_Corr_Sag: 0.35-0.85 (ICC self-selected) ¦ Knee_Corr_Sag: 0.41-0.87 (ICC self-selected) ¦ Ankle_Corr_Sag: 0.01-0.37 (ICC self-selected, poor)
Hulleck_2024 | MAE (Deg) | [7 empty] | 8.04-12.74 |  |  | 5.33-8.06
Templin_2024 | RMSE (Deg) | Pearson |  |  |  | 6.93 | 2.52 (Ad/Ab RMSE) | 5.82 | 5.55 | 4.75 (Ad/Ab RMSE) | 5.86 | 5.77 | 9.61 |  |  |  |  | 0.98 | 0.69 (Pearson r) | 0.67 (Pearson r) | 0.99 | 0.45 (Pearson r) | 0.59 (Pearson r) | 0.97 | n/a (see Extraction_Notes) |  | Hip_Err_Trans: 5.82 (Int/Ext Rot RMSE) ¦ Knee_Err_Trans: 5.86 (Int/Ext Rot RMSE) ¦ Ankle_Err_Front: 9.61 (Subtalar RMSE) ¦ Ankle_Corr_Front: Reported
Wang_2023 | SEM (deg) / ICC | n/a (see Extraction_Notes) |  |  |  | 4.27-5.67 (SEM) | [11 empty] | 0.64-0.84 (ICC) |  |  | n/a (see Extraction_Notes) | [5 empty] | Correl_Metric: ICC CONFIRMED but these are test-retest and inter-rater reliability ICCs, not concurrent validity ICCs ¦ Knee_Corr_Sag: NO KNEE ROM TESTED. Knee flexion was not isolated as a separate ROM movement. Only 'Hip Flexion (Knee Flexed)' was tested, which is a hip movement.
Wade_2023 | Mean Bias (Deg) |  |  |  |  | -3.6 (camera) / -4.6 (occluded) |  |  | 1.5 (camera) / 1.6 (occluded) |  |  | -8.4 (camera) / -9.3 (occluded)
Kim_2025 | [5 empty] | 1.41-2.76 |  |  | 3.11-4.31 |  |  | 1.81-3.47 | [5 empty] | 0.82-0.99 |  |  | 0.94-0.98 |  |  | 0.27-0.90
Ho_2024
He_2024
Balci_2026 |  | ICC |  |  |  | MediaPipe 2D Bias 35.32°; multi-pipelines compared | MediaPipe 2D Bias 2.13° |  | MediaPipe 2D Bias 32.19° | [17 empty] | Hip_Err_Sag: MediaPipe 2D Bias 35.32° (RHFE Table 3); multi-pipelines compared ¦ Hip_Err_Front: MediaPipe 2D Bias 2.13° (RHAA Table 3) ¦ Knee_Err_Sag: MediaPipe 2D Bias 32.19° (RKFE Table 4)
Hii_2023
Chen_2025 | Mean Diff (Deg) | [7 empty] | 3.77 |  |  | 4.39
Fukushima_2024 | MAE (Deg) |  |  |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) | [17 empty] | Hip_Err_Sag: Reported ¦ Knee_Err_Sag: Reported
Ge_2025 |  | ICC | [15 empty] | n/a (see Extraction_Notes) |  |  |  | 0.897 |  |  |  |  | Hip_Corr_Sag: WRONG COLUMN. The 0.89 is HKA angle ICC(2,1)=0.897, which is a FRONTAL plane alignment measure (hip-knee-ankle angle), not hip sagittal. Gold standard is X-ray radiography, NOT optical MoCap. Should be in a frontal plane column or reclassified entirely. ¦ Knee_Corr_Front: 0.897 (HKA angle)
Asaeda_2024 | Mean Bias (Deg) | ICC | [7 empty] | 18.83 - 19.68 | [11 empty] | ICC inter-rater <0.4 / Pearson r=0.554-0.757 |  |  |  |  | Knee_Corr_Front: ICC inter-rater <0.4 / Pearson r=0.554-0.757 (concurrent validity)
Liang_2022 |  | ICC
Svetek_2025 | RMSE (Deg) | Bland-Altman only |  |  |  | 7.12-16.41 | 3.16-5.83 |  | 4.56-10.91 | [17 empty] | Correl_Metric: NO Pearson r reported in paper. Only Bland-Altman plots (Figs 1-5) with mean difference and limits of agreement. No correlation coefficient in any table. Correl_Metric should be blank or removed.
Hu_2025 | RMSE (Deg) | ICC |  |  |  | 4.68 - 4.71 |  |  | 3.56 - 3.71 |  |  | 3.45 - 3.70 | [5 empty] | 0.90-0.91 |  |  | 0.94-0.95 |  |  | 0.84-0.86 |  |  | Hip_Corr_Sag: 0.90-0.91 (ICC WTC/WAC) ¦ Knee_Corr_Sag: 0.94-0.95 (ICC WTC/WAC) ¦ Ankle_Corr_Sag: 0.84-0.86 (ICC WTC/WAC)
Haberkamp_2022 |  | Pearson |  | -4.9 |  | 3.7 |  |  | 1.2 (AMD) |  |  | -14.5 | [5 empty] | 0.78 |  |  | 0.94
Vafadar_2022 | [5 empty] | 4.5 (±8.2) |  |  | -3.0 (±2.7) |  |  | 1.1 | [14 empty] | Ankle_Err_Sag: 1.1 (Table 3 ankle flex/ext)
Washabaugh_2022 | MAE (Deg) |  |  |  |  | 3.7 |  |  | 5.1
Yeom_2026 | RMSE (°) | CMC |  |  |  | 3.81-8.16 (Sag); grand mean lower body 5.45 | 2.24-3.26 | 5.02-5.36 | 3.81-8.16 |  |  | 4.82-7.35 | [5 empty] | 0.97-0.99 | 0.76-0.97 | 0.66-0.83 | 0.97-0.99 |  |  | 0.48-0.73 |  | 0.66-0.83 | Hip_Corr_Sag: 0.97-0.99 (lower body Sag CMC) ¦ Hip_Corr_Front: 0.76-0.97 (lower body Front CMC) ¦ Hip_Corr_Trans: 0.66-0.83 (lower body Trans CMC) ¦ Knee_Corr_Sag: 0.97-0.99 (lower body Sag CMC) ¦ Ankle_Corr_Trans: 0.66-0.83 (lower body Trans range)
Gao_2025 | RMSE (°) | CMC |  |  |  | 9.90 |  |  | 11.54 |  |  | 15.74 | [5 empty] | 0.94 |  |  | 0.93 |  |  | 0.78
Kumtekar_2024 | Peak angle difference (deg) | ICC
Nishizawa_2024 | RMSE (°) | n/a (see Extraction_Notes) |  |  |  | 5.9 (RMSE) / 5.0 (MAE) |  |  | 7.0 (RMSE) |  |  | 13.4 (RMSE) / 10.7 (MAE) | [5 empty] | 0.91 (R²) |  |  | >0.85 |  |  | 0.03 (R²) |  |  | Correl_Metric: R² (coefficient of determination) ¦ Ankle_Corr_Sag: 0.03 (R²) — poor
Todesca_2025 | n/a (see Extraction_Notes) | n/a (see Extraction_Notes) |  |  |  | 6.2-6.9 |  |  | 8.7-11.4 |  |  | n/a (see Extraction_Notes) | [5 empty] | >0.90 |  |  | >0.90 | [5 empty] | Error_Metric: MAJAE (Mean Absolute Joint Angle Error) in degrees ¦ Correl_Metric: PCC (Pearson Correlation Coefficient) - same metric, different label. Confirmed. ¦ Hip_Err_Sag: 6.2-6.9 (MAJAE Table II) ¦ Knee_Err_Sag: 8.7-11.4 (MAJAE Table II) ¦ Ankle_Err_Sag: Ankle is NOT reported in this cycling study. Paper only reports shoulder, elbow, hip, knee angles.
Lin_2024 | n/a (see Extraction_Notes) | NR |  |  |  | NR (audit 2026-06) |  |  | ~5-8 |  |  | ~5-8 | [14 empty] | Error_Metric: MAE (deg) - Mean Absolute Error of joint angles, but reported only as global average across all joints
Zhang_2024 | n/a (see Extraction_Notes) | NR |  |  |  | Mean MPJAE 8.60° (range 7.43-13.35° across OpenCap+BMLmovi datasets |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) | [14 empty] | Hip_Err_Sag: Mean MPJAE 8.60° (range 7.43-13.35° across OpenCap+BMLmovi datasets — no per-joint breakdown) ¦ Knee_Err_Sag: NR per-joint (only global MPJAE 8.60° reported) ¦ Ankle_Err_Sag: NR per-joint (only global MPJAE 8.60° reported)
Hii_2023b | MAE (°) | ICC | [7 empty] | NR (audit 2026-06) | [11 empty] | >0.80
Pappas_2024 | AUC, sensitivity | Pearson r
Irfan_2025 | RMSE (deg) | NR |  |  |  | 4.3-9.6 (L 4.3-7.7, R 4.8-9.6) |  |  | 5.8-9.1 (L 5.8-7.9, R 8.1-9.1) |  |  | 1.0-14.7 | [14 empty] | Ankle_Err_Sag: 1.0-14.7 (high variance)
Liang_2023 | n/a (see Extraction_Notes) | Pearson r |  |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) |  |  | n/a (see Extraction_Notes) | [5 empty] | 0.91 | 0.92 |  | 0.94 |  |  | 0.89 |  |  | Error_Metric: No RMSE reported for joint angle waveforms. Paper reports ROM comparison and Pearson r only. ¦ Hip_Err_Sag: NOT REPORTED (no joint angle RMSE) ¦ Knee_Err_Sag: NOT REPORTED (no joint angle RMSE) ¦ Ankle_Err_Sag: NOT REPORTED (no joint angle RMSE)
Cerciola_2024 | e_ROM (deg) | NR |  |  |  | 3.5-5.1 |  |  | 4.8-5.6 |  |  | IMU ankle RMSE IS reported: 2.1-2.8 degrees across tasks | [14 empty] | Error_Metric: Table II reports e_ROM (absolute ROM error in degrees), NOT RMSE of waveforms. RMSE is discussed in text but Table II is e_ROM. These are different metrics. ¦ Hip_Err_Sag: 3.5-5.1 (MediaPipe; multiple pipelines W3.5/W5.0/R7.0) ¦ Knee_Err_Sag: 4.8-5.6 (MediaPipe; multiple pipelines) ¦ Ankle_Err_Sag: IMU ankle RMSE IS reported: 2.1-2.8 degrees across tasks. Vision systems cannot measure ankle (explicitly stated in Results). No e_ROM for ankle in Table II.
Zhang_2025 | RMSE (°) | n/a (see Extraction_Notes) |  |  |  | 9.14 (F/E RMSE) |  |  | 7.72 |  |  | 3.03 (D/P RMSE) | [14 empty] | Correl_Metric: ICC is reported for markerless vs marker-based joint kinematics
Kidzinski_2020 | MAE (deg) | Pearson r | [6 empty] | 4.8° | [11 empty] | 0.83
Hwang_2026 | ICC, RMSE | ICC | [6 empty] | 5.9 (±3.7) | [11 empty] | 0.83-0.84
Shah_2025 | RMSE, % | [7 empty] | n/a (see Extraction_Notes) | [17 empty] | Knee_Err_Sag: NR (qualitative: higher variability at peak flexion for MediaPipe)
Rode_2025 | MAE (deg) |  |  |  |  | NR |  |  | 9.3-21.9 (2D) / 14.1-25.8 (3D) | [17 empty] | Knee_Err_Sag: 9.3-21.9 (2D) / 14.1-25.8 (3D) — range across 11 estimators
Irfan_2025b | MAE (deg) |  |  |  |  | 3.6 (L) / 7.8 (R) |  |  | 9.1 (L) / 3.8 (R) |  |  | 9.6 (L) / 7.8 (R)
Talaa_2025 | MAE (deg) | [7 empty] | 5.67 | [17 empty] | Knee_Err_Sag: 5.67 (average MAE)
Ang_2024 | n/a (see Extraction_Notes) | Pearson r | [15 empty] | 0.966-0.982 |  |  | 0.956-0.984 |  |  | 0.924-0.976 |  |  | Error_Metric: Correlation coefficient (no degree error reported) ¦ Hip_Corr_Sag: 0.966-0.982 (Pearson r, 3 speeds) ¦ Knee_Corr_Sag: 0.956-0.984 (Pearson r, 3 speeds) ¦ Ankle_Corr_Sag: 0.924-0.976 (Pearson r, 3 speeds)
Hou_2026 | Marker position error (cm) | NR | [24 empty] | Error_Metric: Marker position error (cm) — NOT joint angle error
Athama_2025 | MPJAE (deg) | DTW | [6 empty] | 2.20-15.29 | [17 empty] | Knee_Err_Sag: 2.20-15.29 (MPJAE knee R/L, transmitted across SNR 0-15 dB)
Tsintzira_2025 | Qualitative trajectory match + GRF RMSE | n/a (see Extraction_Notes) | [24 empty] | Correl_Metric: NR (visual comparison Fig 3)

## 7. Architecture Synthesis
Study_ID | AUTO_SUMMARY (Generated) | Final_Classification | Comments
Salehi_2026 |  | 2. Multi-view Triangulation | Retrained OpenCap marker augmenter specifically for lifting tasks, reducing upper-body and trunk errors significantly.
Verheul_2026 |  | 2. Multi-view Triangulation | Highlights OpenCap's potential for running kinematics but warns about systematic differences vs marker-based systems.
Natraj_2025 |  | 3. Monocular Lifting (2D→3D) | Large scale study for remote gait assessment in SCI patients. Uses VideoPose3D for 2D to 3D lifting.
Zhu_2023 |  | 3. Monocular Lifting (2D→3D) | Presents a post-processing algorithm (Kalman Filter + Custom Filter) to improve BlazePose 3D output.
Zago_2020 |  | 2. Multi-view Triangulation | Early study applying 3D triangulation to OpenPose for gait spatiotemporal parameters.
Young_2023 |  | 1. 2D Monocular | Proposes an IoT architecture to capture 2D running gait from a smartphone and stream it to a cloud server.
Verhoeven_2025 |  | 1. 2D Monocular | Custom-trained DeepLabCut for 2D sagittal gait analysis in both adults and toddlers.
Turner_2024 |  | 2. Multi-view Triangulation | Validates OpenCap for return-to-sport dynamic tasks. Notes lower agreement in frontal/transverse planes.
Uhlrich_2023 |  | 2. Multi-view Triangulation | The seminal OpenCap paper. Highlights the use of two smartphones, HRNet, and OpenSim for robust 3D kinematics/kinetics.
Torvinen_2024 |  | 2. Multi-view Triangulation | Validates a custom 3D system using DeepLabCut and DLT for skate skiing.
Takeda_2020 |  | 1. 2D Monocular | One of the earliest feasibility studies using OpenPose to track 2D sagittal walking kinematics.
Stenum_2021 |  | 1. 2D Monocular | Demonstrates 2D OpenPose tracking for gait but highlights out-of-plane perspective errors.
Pagnon_2021 |  | 2. Multi-view Triangulation | Part 1 of Pose2Sim. Focuses on the robustness of the workflow across varying camera setups and image quality.
Pagnon_2022 |  | 2. Multi-view Triangulation | Part 2 of Pose2Sim. Validates the accuracy and identifies systematic biases (e.g., hip offset) originating from the OpenPose model itself.
Needham_2021 |  | 2. Multi-view Triangulation | Uses 9 cameras and DensePose to track the center of mass during elite skeleton push-start.
Molteni_2025 |  | 1. 2D Monocular | Simple 2D OpenPose validation for spatiotemporal parameters and basic kinematics.
Martis_2024 |  | 2. Multi-view Triangulation | Highlights OpenCap limitations when walking direction changes (e.g., walking away from cameras during TUG).
Sabo_2022 |  | 1. 2D Monocular | Validates 2D pose estimation specifically for longitudinal spatiotemporal parameters in Parkinson's disease.
Panconi_2025 |  | 1. 2D Monocular | Demonstrates that a custom-trained DeepLabCut model outperforms pre-trained OpenPose for gait parameters.
Ota_2021 |  | 1. 2D Monocular | Extracts 2D sagittal plane angles from OpenPose but notes tracking issues due to overlapping limbs in 2D.
Nguyen_2023 |  | 5. End-to-End Direct | Uses a Volumetric neural network to directly predict 3D joint centers from multi-view images in occupational settings.
Needham_2022 |  | 2. Multi-view Triangulation | A fully automated pipeline from OpenPose to OpenSim IK using a robust RANSAC triangulation method.
Moro_2020 |  | 1. 2D Monocular | Pilot study evaluating DeepLabCut for 2D gait tracking in stroke survivors.
Moro_2022 |  | 2. Multi-view Triangulation | Uses AdaFuse (epipolar geometry constraints) to refine 2D joints before 3D reconstruction.
Mercadal_2024 |  | 3. Monocular Lifting (2D→3D) | Validates a Strided Transformer network lifting 2D to 3D from a single smartphone for resistance exercises.
Menychtas_2023 |  | 1. 2D Monocular | Applies SPM to compare 2D trajectories. Shows crossing legs during stance phase severely affects OpenPose.
Mehdizadeh_2021 |  | 1. 2D Monocular | Compares AlphaPose, Detectron, and OpenPose for purely temporal gait parameters in older adults.
Lim_2022 |  | 2. Multi-view Triangulation | Very basic multi-webcam OpenPose triangulation setup evaluated on a single subject doing squats.
Embry_2022 |  | 1. 2D Monocular | Validates DeepLabCut strictly for spatiotemporal parameters in stroke survivors in a clinical setting.
Lima_2024 |  | 2. Multi-view Triangulation | Comprehensive OpenCap validation on athletic tasks. Highlights that frontal/transverse plane kinematics are less reliable.
Li_2024 |  | 3. Monocular Lifting (2D->3D) | Introduces GaitFormer to directly predict 3D angles from single-camera video via spatial-temporal transformers.
Ino_2023 |  | 1. 2D Monocular | Confirms that the leg opposite to the 2D sagittal camera suffers from significantly higher tracking errors.
Lambricht_2025a |  | 1. 2D Monocular | Demonstrates that the type of clothing (tight vs loose) has minimal impact on 2D OpenPifPaf tracking accuracy.
Huang_2024 |  | 4. Body Mesh Recovery (SMPL) | Validates kinematics and kinetics (moments) using a SMPL-based markerless method.
Lambricht_2025b |  | 1. 2D Monocular | OpenPifPaf validation for ACL-R rehabilitation tasks (hops) from a sagittal smartphone view.
Lagomarsino_2025 |  | 3. Monocular Lifting (2D→3D) | Shows that 3D lifting (MediaPipe Pose 3D) outperforms 2D for detecting compensations in upper limb movements in cSCI.
Kumar_2022 |  | 5. End-to-End Direct | Uses a CNN-RNN (LSTM/GRU) architecture to predict continuous joint angular velocities from video.
Kakavand_2025 |  | 2. Multi-view Triangulation | Applies OpenCap to cycling. Identifies large errors because the generic IK model doesn't handle the bike constraints well.
John_2024 |  | 1. 2D Monocular | Spatiotemporal validation in stroke patients. Frontal overground views yield the highest errors.
Ino_2024 |  | 1. 2D Monocular | Validates OpenPose specifically for 2D frontal knee valgus during dynamic drop jumps.
Horsak_2025 |  | 4. Body Mesh Recovery (SMPL) | Compares monocular CameraHMR with multi-cam OpenCap. Shows HMR struggles with ankle kinematics.
Horsak_2024 |  | 2. Multi-view Triangulation | Shows that markerless (OpenCap) inherently adds ~14% more inter-trial variability (noise) compared to marker-based.
Horsak_2023 |  | 2. Multi-view Triangulation | Validates OpenCap on healthy vs simulated pathological gait. Pathological gait yields larger errors.
Harrison_2025 |  | 2. Multi-view Triangulation | Validates the Ergomechanic system. Pelvis tracking is ambiguous leading to high hip errors.
Emmerson_2025 |  | 2. Multi-view Triangulation | Validates internal mechanical work estimation for tennis using HRNet. Better stability than OpenPose.
Dunn_2023 |  | 1. 2D Monocular | Tests Video Frame Interpolation (VFI) to boost framerate. Improves timing accuracy but not spatial joint angles.
Drazan_2021 |  | 1. 2D Monocular | Excellent DeepLabCut custom-trained validation for sagittal jumping kinematics.
DHaene_2024 |  | 2. Multi-view Triangulation | Uses YOLOv8 and ZED2 depth cameras. Highly correlated with Optitrack but features a constant bias.
deBorba_2025 |  | 2. Multi-view Triangulation | Validates OpenCap across multiple walking speeds. High RMSE for specific discrete ROM metrics despite good waveform correlation.
DAntonio_2020 |  | 2. Multi-view Triangulation | Feasibility study vs IMUs. Large errors found at the maxima and minima of the joint angles.
DAntonio_2021 |  | 2. Multi-view Triangulation | Extension of the 2020 study. 2 webcams vs IMUs. Errors vary heavily depending on the condition.
DAntonio_2020b |  | 2. Multi-view Triangulation | Another conference paper evaluating OpenPose with 2 webcams vs IMUs.
CarrascoPlaza_2022 |  | 3. Monocular Lifting (2D→3D) | Theoretical evaluation of a 2D to 3D neural network. Concludes clinical thresholds (5°) are often exceeded.
Boldo_2024 |  | 1. 2D Monocular | Provides guidelines for clinical 2D overground gait analysis. Highlights ankle angle bias.
Bittner_2023 |  | 6. End-to-End Direct (Video->Angles) | Introduces D3KE (Direct 3D Kinematic Estimation) which significantly outperforms standard OpenPose+OpenSim pipelines.
Bertozzi_2025 |  | 2. Multi-view Triangulation | Shows OpenCap struggles to track hip accurately during highly dynamic jump landings at 60 fps.
Barzyk_2024a |  | 3. Monocular Lifting (2D→3D) | Validation of the Sbsq-pose app. Acceptable knee/ankle tracking from frontal view, high hip errors.
Barzyk_2024b |  | 3. Monocular Lifting (2D→3D) | Excellent validation of Sbsq-pose for post-stroke gait analysis using a single smartphone.
Bak_2025 |  | 2. Multi-view Triangulation | Demonstrates that BMI and thigh circumference strongly correlate with increased markerless tracking errors.
Bae_2024 |  | 2. Multi-view Triangulation | Ergo system (using MediaPipe) showed excellent agreement for time-series and peak joint angles during overhead squat.
Azhand_2021 |  | 3. Monocular Lifting (2D→3D) | Excellent validation for spatiotemporal parameters on older adults using a CNN+TCN lifting network from an iPad.
Andreoni_2024 |  | 1. 2D Monocular | Validates OpenPose against BTS. The absolute error on joints frequently exceeds the 5-degree clinical limit.
Anderson_2025 |  | 1. 2D Monocular | Excellent pediatric application. Proves that OpenPose can extract accurate spatiotemporal parameters from a frontal view.
Aleksic_2024 |  | 1. 2D Monocular | Validates jump height and Center of Mass tracking with MMPose. No joint angle validation.
Ahn_2024 |  | 2. Multi-view Triangulation | OpenPose BODY_25 + custom OpenCV analysis (NOT MediaPipe + OpenSim). ICC(3,k)-based validation only. Measures Hip + Knee + Trunk Rotation (NOT ankle).
Wang_2024 |  | 3. Monocular Lifting (2D→3D) | Tests different camera angles and distances with HuMoR. 45°/135° angles introduce massive errors compared to purely sagittal/frontal.
Verheul_2024 |  | 2. Multi-view Triangulation | Uses OpenCap kinematics to predict Ground Reaction Forces (GRF). Over-smooths impact peaks.
Wang_2025 |  | 2. Multi-view Triangulation | OpenCap has higher RMSE in patients with Knee OA than healthy controls.
Bousigues_2025 |  | 2. Multi-view Triangulation | Proves that 2D-3D triangulation generates skeletal inconsistencies equivalent to skin soft-tissue artifacts.
Ueno_2024 |  | 3. Monocular Lifting (2D→3D) | Validates HuMoR on OpenCap dataset. MAE of 8.5° overall.
Yagi_2020 |  | 1. 2D Monocular | Proof-of-concept for home monitoring using a homography matrix to estimate spatiotemporal parameters.
Tahara_2025 |  | 1. 2D Monocular | Uses ML on MediaPipe 2D coordinates to predict spatiotemporal parameters (GAITRite).
Francia_2026 |  | 1. 2D Monocular | Uses an EKF to merge two 2D MediaPipe outputs to track Knee Flexion.
Koleini_2025 |  | 4. Body Mesh Recovery (SMPL) | BioPose: HMR (SMPL-H vertices) + Neural Inverse Kinematics post-processing. Mesh is the core intermediate representation. [Reclassified from End-to-End 2026-04-15]
Yamamoto_2021 |  | 1. 2D Monocular | Uses OpenPose without biomechanical constraints on subjects wearing AFOs.
Hulleck_2024 |  | 3. Monocular Lifting (2D→3D) | Takes 2D BlazePose output sequences and translates them to 3D joint angles using a Seq2Seq NLP model.
Templin_2024 |  | 2. Multi-view Triangulation | Trains keypoint detector on synthetic SMPL-X dataset before triangulation and IK.
Wang_2023 |  | 3. Monocular Lifting (2D→3D) | Calculates 3D ROM angles using MediaPipe 3D joint outputs from a single camera.
Wade_2023 |  | 1. 2D Monocular | Highlights the limits of 2D OpenPose tracking when evaluating the contralateral (occluded) leg.
Kim_2025 |  | 1. 2D Monocular | Evaluates spatiotemporal parameters in Parkinson's disease. No angular error metrics.
Ho_2024 |  | 2. Multi-view Triangulation | Employs a binocular ZED stereo camera to get depth directly without a 3D lifting network.
He_2024 |  | 3. Monocular Lifting (2D→3D) | Trains a Random Forest classifier on 2D->3D lifted gait parameters to predict Parkinson's disease.
Balci_2026 |  | 1. 2D Monocular | Highlights the fact that MediaPipe 2D tracking provides better ICCs than MediaPipe 3D lifting or YOLOv8 out-of-the-box.
Hii_2023 |  | 1. 2D Monocular | Strictly a signal processing paper focused on extracting spatio-temporal parameters using MediaPipe. No Vicon comparison.
Chen_2025 |  | 2. Multi-view Triangulation | Improves OpenPose multi-view tracking using ShuffleNet for cross-view matching.
Fukushima_2024 |  | 2. Multi-view Triangulation | Reconstructs 3D kinematics directly via OpenPose triangulation. Does not use a biomechanical model (IK).
Ge_2025 |  | 1. 2D Monocular | Evaluates coronal alignment (valgus/varus) from frontal video using geometric vectors.
Asaeda_2024 |  | 1. 2D Monocular | Calculates dynamic knee valgus purely from 2D coordinates in the frontal plane.
Liang_2022 |  | 3. Monocular Lifting (2D→3D) | Trains 3DPoseNet to lift 2D OpenPose keypoints into 3D prior to angle calculation.
Svetek_2025 |  | 2. Multi-view Triangulation | Shows OpenCap struggles when physical obstacles (treadmill bars) partially occlude the participant.
Hu_2025 |  | 2. Multi-view Triangulation | ViT 2D keypoints → triangulation 3D → OpenSim Inverse Kinematics → joint angles. n=152 OA patients (67 hip + 85 knee).
Haberkamp_2022 |  | 1. 2D Monocular | Assesses kinematics independently in the 2D frontal and 2D sagittal planes.
Vafadar_2022 |  | 2. Multi-view Triangulation | Custom system trained specifically with bi-planar X-ray (EOS) to correct soft tissue artifacts.
Washabaugh_2022 |  | 1. 2D Monocular | Benchmark study on 2D sagittal pose estimation algorithms. Demonstrates SPM waveform analysis on time-series errors.
Yeom_2026 |  | 2. Multi-view Triangulation | Clinically relevant; first OpenCap validation for overhead squat movement quality screening
Gao_2025 |  | 1. 2D Monocular | Clinical screening tool for CP; no joint angle validation, focuses on gait deviation detection
Kumtekar_2024 |  | 1. 2D Monocular | Telehealth-oriented; upper extremity focus; limited to 2D angles
Nishizawa_2024 |  | 1. 2D Monocular | Clinical rehab application; moderate accuracy; useful for asymmetry detection
Todesca_2025 |  | 2. Multi-view Triangulation | Sport-specific (cycling); good accuracy; near real-time capable
Lin_2024 |  | 3. Monocular Lifting (2D→3D) | MICRO-oriented: methodological contribution; synthetic data training approach
Zhang_2024 |  | 2. Multi-view Triangulation | MICRO-oriented: GNN method for multi-view 3D; limited validation
Hii_2023b |  | 1. 2D Monocular | Frontal plane only; comparative study; MediaPipe recommended; IEEE NBEC 2023
Pappas_2024 |  | 3. Monocular Lifting (2D→3D) | Clinical screening (Parkinson's); frontal view; novel depth estimation approach
Irfan_2025 | Smartphone (Redmi 2201117TI, Snapdragon 8 Gen 2) + BlazePose 2D → hip/knee/ankle angles via law of cosines, validated on 31-subject GPJATK MoCap dataset | 2D Monocular | Smartphone vision sensor (NOT laptop). Mean RMSE 7.8±3.9° across joints (Table III, 31 subjects). Hardware-deployment focus.
Liang_2023 |  | 2. Multi-view Triangulation | Elderly-focused; good correlations for hip/knee; step/stride length validation; IEEE MVIP 2023
Cerciola_2024 |  | 1. 2D Monocular | Most comprehensive vision method comparison; IMU still superior; IEEE Sensors Journal 2024
Zhang_2025 |  | 2. Multi-view Triangulation | Clinical application; combines MoCap with ML for functional assessment; IEEE JBHI 2025
Kidzinski_2020 | Single sagittal video → OpenPose 2D → CNN regression → clinical gait metrics (speed, cadence, knee flexion at max ext, GDI). CP cohort n=1026 | 2D Monocular + Temporal DL | Landmark Nature Comm paper. End-to-end ML regression from 2D pose trajectories to clinical gait metrics (GDI), bypassing 3D reconstruction. Source code public.
Hwang_2026 | iPad Pro RGB single camera sagittal + custom 2D pose (ViFive, stacked hourglass) → CoM trajectory + sit-to-stand power; validated vs optical MoCap + force plate on 20 elderly subjects | 2D Monocular + Kinetics | Strong elderly clinical population (n=129, mean age 71.8). Combines kinematics + computed kinetics (STS power). Published JCSM. Force plate validation rare in MMC literature.
Shah_2025 | Comparison of Qualisys (gold), OpenCap (multi-view), and MediaPipe (monocular) for knee kinematics during sit-ups on 15 South Asian adults | 2D Monocular (MediaPipe) + Multi-view (OpenCap) comparative | Compares 3 systems on same subjects. Knee flexion during sit-ups — rare task. Anthropometric diversity (South Asian). MediaPipe peak flexion limitation reported.
Rode_2025 | Benchmark of 11 open-source monocular pose estimators (direct + 2D→3D lifting) on new Physio2.2M dataset (25 healthy adults, 7 physical exercises) with 27-camera OMC ground truth | 2D Monocular + Monocular Lifting 2D→3D (benchmark) | Massive benchmark paper. Physio2.2M dataset released. Wide MAE range across estimators (some viable, others not). Useful reference for selecting monocular models.
Irfan_2025b | Laptop built-in camera + BlazePose 2D → hip/knee/ankle angles via law of cosines, validated on GPJATK dataset + tested on 10 elderly patients | 2D Monocular | Smartphone-class hardware (laptop webcam). MAE 3.6-9.6° on lower limbs. Targets in-clinic daily-use deployment.
Talaa_2025 | MediaPipe 2D on 2 calibrated cameras → knee angle via arctan2 from hip-knee-ankle 2D landmarks; validated on 11 walks from GPJATK dataset vs Vicon MX-T40 | 2D Monocular | Single-knee target only. Per-camera angle (not triangulated). MAE 5.67° avg.
Ang_2024 | Single video camera + OpenPose BODY_25 2D → joint angles via vector geometry. Case study (n=1) of squat at 3 speeds; validation against Vicon by correlation coefficient and linear regression. Telerehabilitation context | 1. 2D Monocular | Springer chapter 2024. Case study only (n=1). Strong correlations (>0.96 joint angles, >0.84 angular velocities) but no absolute error in degrees reported. Validates_Kinematics=Yes by correlation criterion.
Hou_2026 | Multi-view (4 unsynchronized cameras) + SMPL Human Mesh Recovery + virtual markers (56 whole-body landmarks from SMPL surface) + Transformer-based multi-view temporal fusion. Validation: marker position error 3.05 cm overall (Walk 2.62, Run 3.77, In-place 3.44) vs Vicon Valkyrie V16 (12 cams). OpenSim downstream for kinematic/kinetic analyses (not directly validated) | 4. Body Mesh Recovery (SMPL) - multi-view | INNOVATIVE PAPER. Multi-camera UNSYNCHRONIZED + SMPL Mesh Recovery + Transformer temporal fusion. ⚠️ Validation reports POSITION error in cm, not joint angles in degrees. Authors mention OpenSim downstream for clinical analyses. Final classification could be 2. Multi-view Triangulation + SMPL Mesh Recovery overlay, OR 4. Body Mesh Recovery extended to multi-view. Worth highlighting in MICRO/MACRO bridge as novel architecture.
Athama_2025 | Single camera + PoseformerV2 (2D→3D lifting) + Deep-JSCC wireless transmission + SMPL/AMASS animated rendering. Validation: MPJAE on 4 joints (Elbows + Knees) under varying SNR using MoVi dataset. Pipeline focuses on wireless healthcare applications with privacy | 4. Monocular Lifting (2D→3D) | IEEE HPCC 2025. Novel use case: wireless transmission for telehealth + privacy-preserving rendering. Joint angle validation limited to 4 joints (Elbows + Knees) — no hip/ankle. Gold standard = MoVi public dataset (no direct MoCap deployment by authors).
Tsintzira_2025 | ZED-2 stereo + ZED Body Tracking + 2D 14-DoF human model (sagittal) → analytical closed-form IK + Lagrangian-based ID + GRF estimation WITHOUT force plate. Validation: 1 subject treadmill walking, comparison vs Fukuchi 2018 public dataset (kinematics + kinetics + GRF). GRF RMSE 5% BW (horiz), 11.3% BW (vert) | 1. 2D Analysis (Monocular) + IK/MSK + Physics-based ID + GRF | IEEE BIBE 2025. Proof of concept (n=1). UNIQUE: estimates GRF WITHOUT force plate via Lagrangian dynamics. Pipeline le plus complet du corpus (kinematics + kinetics + GRF). Gold standard = comparison vs public dataset averages, not direct MoCap. Validation Kinetics=Yes (GRF vs database).

## Config_Lists
Paper_Type | Clothing | Camera_Type | Environment | Calibration | Algo_Pose_2D | Training_Dataset | Method_3D | Real_Time | Bio_Software | Kinematics_Method | Yes_No | Reported_Metrics | Final_Classification
Validation | Minimal/Tight (Spandex) | Smartphone | Laboratory | Checkerboard | OpenPose (BODY_25) | COCO | None (2D Pure) | Yes (Real-Time) | OpenSim | Inverse Kinematics (IK) | Yes | RMSE + ICC | 1. 2D Analysis (Monocular)
Methodological | Streetwear (Loose) | Webcam | Clinic | Wand | OpenPose (COCO-18) | MPII | Triangulation (Multi-view) | No (Post-process) | Visual3D | Direct (Geometric Vectors) | No | MAE + Pearson | 2. 3D Geometric (Triangulation)
Clinical Applied | Mixed | Industrial Camera | Outdoors (In the wild) | Human (Static Pose) | DeepLabCut (Custom) | Human3.6M | Lifting (2D->3D NN) | Not specified | Pyomeca | Direct Prediction (Deep Learning) | Not specified | RMSE only | 3. Model-Based Fusion (Triangulation + IK)
Review | Not specified | Action Cam (GoPro) | Home | Automatic (Self-calib) | MediaPipe (BlazePose 33) | Custom | Model Fitting (SMPL/HMR) |  | BTK | Physics-Informed DL |  | CMC + RMSE | 4. Monocular Lifting (2D->3D)
Feasibility |  | Hybrid/DSLR |  | None (Monocular) | HRNet (COCO-17) | Mixed | Volumetric (Visual Hull) |  | Custom Code (Python/Matlab) |  |  | SPM + RMSE | 5. Monocular Shape/SMPL
|  | Kinect/Depth |  |  | AlphaPose (COCO-17) | Not specified | Direct Angle Regression (End-to-End) |  | AnyBody |  |  | Other | 6. End-to-End Direct (Video->Angles)
|  | Drone |  |  | MoveNet (Thunder) |  | Physics-Informed Optimization |  | Integrated in Pipeline (Black Box)
|  | Tablet (iPad) |  |  | YOLO-Pose
OpenPifPaf (COCO-17)
Proprietary / Unknown
Manual
N/A (End-to-End)