Appendix A: Detailed Validation Study Parameters

This appendix reproduces the full phase-by-phase specifications summarized in Section 11. Exact sample-size and analytic specifications should be finalized in protocol-specific work; the tables below retain only the phase logic, primary endpoints, and approximate planning ranges needed to distinguish technical validation from educational validation.

A.1 Phase 1: Bench Validation of Fidelity and Interoperability

| Item | Recommendation |
| --- | --- |
| Purpose | Show that playback is technically faithful enough for the intended teaching task, using pre-specified equivalence testing rather than descriptive comparison alone |
| Stimulus set and scale | 20–30 curated cases spanning normal sinus rhythm, atrial fibrillation, bundle-branch block, STEMI-pattern findings, brady-/tachycardia, and artifact-challenge cases, repeated across 2–3 named monitor families if interoperability is claimed; PTB-XL/PTBDB-class 12-lead data should be used if 12-lead claims are made, not MIT-BIH alone |
| Primary endpoint | RMS voltage error against source, evaluated against a pre-specified equivalence margin rather than judged as descriptively “low.” The margin itself should be defined prospectively from clinical ECG standards, measurement uncertainty, and preservation of curriculum-relevant morphology, with input from cardiology and biomedical-engineering expertise; the 20.8–28.4 µV reconstruction error reported in a comparable published system [10] is contextual evidence that such a margin is achievable, not itself a validated or educationally justified threshold; that figure derives from a ten-patient proof-of-concept validation in which reconstruction error was greatest for bundle-branch-block recordings and was reduced two- to three-fold when a simultaneous-sampling converter was used for re-digitisation [10] |
| Secondary endpoints | Timing and ST-segment deviation error, rhythm-classification agreement, named-monitor recognition success |
| Statistical approach | Equivalence-testing logic (e.g., repeated-measures Bland-Altman or hierarchical agreement models); exact margins and models to be finalized in protocol-specific work |

A.2 Phase 2: Randomized Crossover Educational Trial

| Item | Recommendation |
| --- | --- |
| Purpose | Test whether the module improves performance relative to comparator exposure, using each learner as their own control |
| Design and participants | Randomized crossover trial with three arms: standardized playback hardware, a matched screen-based tool using the same waveform library, and static tracings or existing simulator display in senior medical students, residents, or a mixed novice/intermediate cohort, with period/sequence effects modeled explicitly. The hardware-vs-screen comparison isolates the physical-integration variable identified in Section 3.7 from the standardization variable that both hardware and screen arms share |
| Planning scale | 40–60 learners, a planning range only; not based on a prior effect-size estimate specific to this intervention, since none yet exists |
| Single primary endpoint (by design) | ECG interpretation accuracy on a standardized test bank scored against repository ground truth, kept as the sole primary endpoint to avoid multiplicity problems |
| Secondary/exploratory endpoints | Time to diagnosis, confidence, cognitive workload, state anxiety, retention at 2–6 weeks; hypothesis-generating only, not sufficient alone to claim benefit if the primary endpoint is null |

A.3 Phase 3: Multisite Pragmatic Implementation Study

| Item | Recommendation |
| --- | --- |
| Purpose | Test cross-site reproducibility and implementation feasibility using a shared waveform library and faculty package |
| Design and scale | 4–8 simulation centers, approximately 20–30 learners per site as a naive planning figure; a stepped-wedge or staged rollout should be considered if implementation timing is itself of interest. Clustering by site will inflate the true required sample size beyond this naive figure; cluster-aware power planning is required before execution but is not detailed here |
| Primary endpoint | Between-site variability in performance on a common ECG case set, operationalized as the intraclass correlation coefficient (or an equivalent variance component) from a multilevel model with learner nested within site, combined with successful delivery of a predefined proportion (e.g., ≥90%) of standardized cases without technical deviation. A variance measure alone is difficult to interpret in isolation, so the two are reported jointly rather than the variance component serving as a sole primary endpoint. |
| Secondary endpoints | Implementation fidelity, faculty usability (Section 9), setup time, technical failure rate, learner workload, scenario-version concordance across sites |