Table 1. Physiological variables and polygraph-derived measurements

| Parameter | Physiological definition | Polygraph-derived measurement | Unit |
| --- | --- | --- | --- |
| Asynchrony Index (AI) | Proportion of breaths affected by patient–ventilator asynchrony, including ineffective efforts, auto-triggering, double triggering, premature cycling, and late cycling | AI = (number of asynchronous events) ÷ (number of ventilator cycles + ineffective efforts) × 100, identified by comparison of Pes deflections with airway pressure (Paw) and flow signals | % |
| Esophageal pressure–time product (PTPes) | Integrated inspiratory Pes over time, reflecting overall patient inspiratory effort per minute | Sum of inspiratory Pes–time areas below baseline, normalized per minute | cmH₂O·s/min |
| Inspiratory esophageal pressure swing (ΔPes) | Difference between end-expiratory and peak inspiratory Pes, surrogate of inspiratory effort | Breath-by-breath difference between Pes at end-expiration and inspiratory nadir | cmH₂O |
| Inspiratory trigger delay | Time between onset of patient inspiratory effort and ventilator pressurization | Interval between onset of negative Pes deflection and rise in Paw or inspiratory flow | ms |
| Expiratory trigger delay | Time between termination of patient inspiratory effort and termination of ventilator assistance | Interval between return of Pes toward baseline (end of inspiration, see legend) and opening of the expiratory valve / cessation of inspiratory flow | ms |
| Synchronization time | Temporal overlap between patient inspiratory effort and ventilator assistance | Overlap between negative Pes deflection and positive Paw or inspiratory flow, expressed as a percentage of the patient’s (Pes) inspiratory time | % of insp. time |
| Peak inspiratory flow | Maximum inspiratory airflow during inspiration | Maximum value of the inspiratory flow signal | mL/s |
| Peak inspiratory pressure (PIP) | Maximum airway pressure during inspiration | Peak Paw measured during the inspiratory phase | cmH₂O |
| Positive end-expiratory pressure (PEEP) | Airway pressure at end-expiration | Mean Paw over the final 200 ms of expiration | cmH₂O |
| Rise time | Time required to reach target inspiratory pressure after triggering | Interval between onset of Paw rise and achievement of inspiratory pressure plateau | ms |
| Inspiratory time, patient (Pes) | Duration of patient inspiratory (neural) effort | Interval from onset of negative Pes deflection (Pes initial decay) to the end of inspiration (point at which 25% of the time from the maximal Pes deflection to its return to baseline has elapsed) | ms |
| Expiratory time, patient (Pes) | Duration of patient expiratory (neural) phase | Interval from the end of inspiration (as defined above) to the onset of the next inspiratory Pes deflection | ms |
| Mechanical inspiratory time | Duration of ventilator-delivered inspiratory flow | Interval from onset of inspiratory flow (>0) to return to zero flow | ms |
| Mechanical expiratory time | Duration of ventilator expiratory phase | Interval from end of inspiratory flow to onset of the next inspiratory flow | ms |
| Respiratory rate, patient (neural) | Number of spontaneous (neural) inspiratory efforts per minute | Count of Pes inspiratory deflections per minute | breaths/min |
| Respiratory rate, machine | Number of ventilator-delivered breaths per minute | Count of inspiratory flow cycles per minute | breaths/min |
| Patient duty cycle (Ti/Ttot, neural) | Fraction of the breath cycle occupied by neural inspiration | Ratio of Pes inspiratory time to total Pes cycle time (Ti/Ttot) | ratio |
| Machine duty cycle (Ti/Ttot, mechanical) | Fraction of the breath cycle occupied by ventilator insufflation | Ratio of mechanical inspiratory time to total mechanical cycle time (Ti/Ttot) | ratio |
| Inspiratory time in excess (premature cycling-off) | Extent to which neural (Pes) inspiration outlasts ventilator insufflation | (Pes inspiratory time − mechanical inspiratory time) ÷ Pes inspiratory time × 100; a positive value indicates that neural inspiration outlasts mechanical insufflation (premature cycling-off) | % of Pes Ti |

Legend.

Pes, esophageal pressure; Paw, airway pressure; AI, asynchrony index; PTP, pressure–time product; PTPes, inspiratory esophageal pressure–time product per minute; ΔPes, inspiratory esophageal pressure swing; PIP, peak inspiratory pressure; PEEP, positive end-expiratory pressure; Ti/Ttot, inspiratory-to-total-time ratio (duty cycle); RR, respiratory rate.

Patient (neural) inspiratory and expiratory times, duty cycle and respiratory rate are derived from the Pes signal, whereas the corresponding mechanical variables are derived from the ventilator flow signal. Synchronization time and the inspiratory time in excess are expressed as percentages of the patient’s (Pes) inspiratory time. Asynchrony events were identified by visual inspection of synchronized Pes, flow, and Paw tracings. Time-based variables are expressed in milliseconds, pressures in cmH₂O, flow in mL/s, and rate in breaths/min.