Supplementary table 2: Study protocol temporal bone MRI protocols

|  | 0.55-Tesla (0.55T) |  | 1.5-Tesla (1.5T) |  |
| --- | --- | --- | --- | --- |
|  | CISS | 2D T2-FSE | CISS | 2D T2-FSE |
| Repetition time (TR) | 7.02 ms | 2000 ms | 5.66 ms a | 3800 ms b |
| Echo time (TE) | 3.06 ms | 99 ms | 2.52 ms | 88 ms |
| Number of excitations | 1 | 4 | 1 | 4 |
| Refocusing flip angle (FA) | 80° | 180° | 70° a | 150° b |
| Pixel band width | 425 Hz/Px | 252 Hz/Px | 504 Hz/Px | 438 Hz/Px |
| Turbo factor | NA | 17 | NA | 17 |
| Pixel spacing | 0.74 mm | 0.81 mm (zero-filling interpolation 0.4 mm) | 0.6 mm | 0.62 mm (zero-filling interpolation 0.3 mm) |
| Slice thickness | 0.7 mm | 1.5 mm | 0.6 mm | 1.2 mm |
| Matrix size | 256 x 256 | 272X272 | 320X320 | 272X272 |
| Displayed field of view | 190 x 190 mm | 220 mm x 220 mm | 200 mm x 181 mm | 170 mm x 170 mm |
| Acquisition time | 4.21 min | 3.34 min | 4.46 min | 6.16 min |
| Acceleration | GRAPPA factor 2 | GRAPPA 2/Deep Resolve Boost with high denoising | None | None |
| Phase encoding direction | R>L | R>L | R>L | R>L |
| RF pulse type | Fast | Normal | Normal | Normal |

CISS: Constructive interference in the steady state

a In order to stay within Specific Absorption Ratio (SAR) limits for the HiRes 90K™ CI at 1.5T the FA was reduced to 33-39° in 3 patients and in one patient the FA was reduced to 59° with the TR increased to 7.88ms (increasing acquisition time to 7.14 min)

bIn order to stay within Specific Absorption Ratio (SAR) limits for the HiRes 90K™ CI at 1.5T the TR was increased to 4450-5620 which increased the acquisition time to 7.13 min-9.07 min.