| Supplementary Table 1. Comparison of Cerebrovascular Adverse Events Reported in the MAUDE Database Before and After the VARIPULSE™ Safety Communication (Recall Period). |  |  |
| --- | --- | --- |
| Outcome | Pre-recall (n= 58 reports) | Post-recall (n= 329 reports) |
| Stroke/CVA | 22 (37.93%) | 107(32.52%) |
| TIA | 3 (5.17%) | 15 (4.56%) |
| Ischemic stroke | 2 (3.45%) | 6 (1.8%) |
| Combined cerebrovascular events* | 27 (46.55%) | 124 (37.69%) |
| *Combined cerebrovascular events include stroke, transient ischemic attack (TIA), and ischemic stroke. Abbreviations: CVA, cerebrovascular accident; TIA, transient ischemic attack. |  |  |

| Supplementary Table 2. Device Problems Excluded from The Analysis. |  |  |
| --- | --- | --- |
| Device-related problems | Number | Percentage (%) |
| Adverse Event Without Identified Device or Use Problem | 6 | 85.71% |
| Blank | 1 | 14.28 % |
| Total | 7 | 100% |

| Supplementary Table 3. Patient Problems Excluded from The Analysis. |  |  |
| --- | --- | --- |
| Patient-related problems | Number | Percentage (%) |
| No Clinical Signs, Symptoms or Conditions | 81 | 93.11% |
| Insufficient Information | 3 | 3.45% |
| Appropriate Term / Code Not Available | 1 | 1.15% |
| Total | 85 | 100 |

| Supplementary Table 4. Overview of Reported Death Cases in the FDA MAUDE Database Associated with Pulsed Field Ablation Systems (VARIPULSE™ Platform). |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| MDR Report ID | Patient Demographics | Event Description | Reported Cause of Death | Device Evaluation | Manufacturer Conclusion |
| 24090030 | 75-year-old female | Death reported following ablation procedure; limited clinical details available | Not specified | Device not returned; no evaluation performed | Causality cannot be determined due to insufficient information |
| 23968953 | 83-year-old female | Patient developed hemorrhagic stroke post-procedure and did not regain consciousness | Hemorrhagic stroke | No device available for analysis; procedural factors reviewed | Stroke considered multifactorial; no confirmed device malfunction |
| 23479787 | 76-year-old male | Patient developed intracranial hemorrhage weeks after procedure following fall from bicycle | Traumatic intracranial bleed (post-fall) | No device evaluation; additional follow-up obtained | Death not related to device or procedure; attributed to comorbidities and trauma |
| 25015553 | Not specified | Death reported following an idiopathic ventricular tachycardia ablation. After a successfully completed concomitant atrial fibrillation and left atrial appendage occlusion procedure, the patient experienced low blood pressure and asystole during removal of sheaths and preparation for wake-up/transport, and did not survive. No evidence of steam pop. | Possible drug reaction (per physician opinion) | Device not returned; no evaluation performed | Causality cannot be determined due to insufficient information. |
| 21020743 | 77-year-old Female | Death reported the day after a paroxysmal atrial fibrillation ablation. Procedure was successful, but the patient had pre-existing heart failure (EF 20-40%) and potassium issues. Post-procedure, potassium increased to >7 mmol/L, was managed temporarily, but the patient later went into cardiac arrest with pulseless electrical activity and could not be resuscitated. No autopsy. | Heart failure due to high potassium (per physician opinion) | Device not returned; no evaluation performed. | Causality cannot be determined due to insufficient information. The physician did not believe death was related to the pulsed field energy or catheter, but to the patient's poor health. |