Supplementary Table 1: Summary protocol comparing standard Incubation times as described in the product insert to the shortened protocol used in this study.

| Step | Description | Components | Plate addition | Standard Incubation Times | Shortened Steps |
| --- | --- | --- | --- | --- | --- |
| S-PLEX Wash | 3X PBS-T wash |  |  |  |  |
| Plate Coating | Biotin Capture antibody with S-PLEX Coating Reagent C1 | Capture Antibody | 50 μL/well | 1h at RT with shaking or overnight at 4C without shaking | 1h at RT with shaking or overnight at 4C without shaking |
|  |  | Coating Reagent C1 |  |  |  |
|  |  | Diluent 100 |  |  |  |
| S-PLEX Wash | 3X PBS –T wash |  |  |  |  |
| Sample incubation | Cal Curve and sample addition | Assay Diluent | 25 μL/well Assay Diluent + 25 μL/well Sample | 1.5h at RT with shaking | 1 hr at 27C with shaking |
|  |  | Samples |  |  |  |
|  |  | Calibrator |  |  |  |
| S-PLEX Wash | 3X PBS –T wash |  |  |  |  |
| Antibody Binding | Antibody with TURBO-TAG label | Antibody Diluent | 50 μL/well | 1h at RT with shaking | 40 min at 27C with shaking |
|  |  | TURBO-TAG Conjugated Antibody |  |  |  |
| S-PLEX Wash | 3X PBS –T wash |  |  |  |  |
| Enhancement Step | Ligation | Enhance E1 | 50 μL/well | 30 min at RT with shaking | 10 min at 27C with shaking |
|  |  | Enhance E2 |  |  |  |
|  |  | Enhance E3 |  |  |  |
|  |  | Molecular Grade Water |  |  |  |
| S-PLEX Wash | 3X PBS –T wash |  |  |  |  |
| Detection Step | Amplification/Detection | Detect D1 | 50 μL/well | 1h at 27C with shaking | 40 min at 27C with shaking |
|  |  | Detect D2 |  |  |  |
|  |  | Molecular Grade Water |  |  |  |
| S-PLEX Wash | 3X PBS –T wash |  |  |  |  |
| Read |  | Read Buffer | 150 μL/well | None | None |
| S-PLEX Wash | Gentle wash with lower dispense rate with tip away from spot |  |  |  |  |
| Total Time |  |  |  | 5 hours | 3.5 hours |

Supplementary Table 2: Sample description of all contrived samples for the multi-site, multi-lot precision study. Spike levels were prepared by blending purified analytes at high (01), mid (02-03) and low (04) levels along the quantitative range and adding to pooled samples before aliquoting. Normal or diseased samples were not spiked unless noted. All samples were commercially-obtained then modified as needed.

| Sample ID | Sample Description | Matrix | Disease State or Spike Level | Sample Type |
| --- | --- | --- | --- | --- |
| Sample 01 | Hemolyzed Plasma | Plasma | Hemolyzed | Challenge |
| Sample 02 | BioIVT Plasma AD Pool | Plasma | AD | Clinical |
| Sample 03 | Cord Serum 3 | Serum | Normal | Clinical/Challenge |
| Sample 04 | Lupus Serum Pool 1 | Serum | Lupus | Challenge |
| Sample 05 | Cord Serum 2 | Serum | Normal | Clinical/Challenge |
| Sample 06 | Normal Serum 2 | Serum | Normal | Clinical |
| Sample 07 | Normal Male Plasma 1 | Plasma | Normal | Clinical |
| Sample 08 | BioIVT Serum AD Pool 1 | Serum | AD | Clinical |
| Sample 09 | Cord Serum 6 | Serum | Spike 03 | Clinical |
| Sample 10 | Plasma Spike 04 | Plasma | Spike 04 | Clinical |
| Sample 11 | Serum Spike 04 | Serum | Spike 04 | Clinical |
| Sample 12 | CBH Serum AD Pool | Serum | AD | Clinical |
| Sample 13 | Lupus Serum Pool 2 | Serum | Lupus | Challenge |
| Sample 14 | Normal Serum 1 | Serum | Normal | Clinical/Challenge |
| Sample 15 | Serum Spike 01 | Serum | Spike 01 | Clinical/Challenge |
| Sample 16 | Hemolyzed Serum | Serum | Hemolyzed | Challenge |
| Sample 17 | Plasma Spike 01 | Plasma | Spike 01 | Clinical/Challenge |
| Sample 18 | Cord Plasma 1 | Plasma | Normal | Clinical |
| Sample 19 | HAMA+ | Plasma | HAMA+ | Challenge |
| Sample 20 | Plasma Spike 02 | Plasma | Spike 02 | Clinical |
| Sample 21 | Cord Plasma 2 | Plasma | Spike 03 | Clinical |
| Sample 22 | Serum Spike 03 | Serum | Spike 03 | Clinical |
| Sample 23 | Cord Serum 7 | Serum | Spike 04 | Clinical |
| Sample 24 | MS Plasma Pool | Plasma | MS | Clinical |
| Sample 25 | Lupus Plasma Pool | Plasma | Lupus | Challenge |
| Sample 26 | BioIVT Serum AD Pool 2 | Serum | AD | Clinical |
| Sample 27 | Serum Spike 02 | Serum | Spike 02 | Clinical |
| Sample 28 | HAMA+ | Serum | HAMA+ | Challenge |
| Sample 29 | Cord Serum 5 | Serum | Spike 02 | Clinical/Challenge |
| Sample 30 | Normal Plasma 2 | Plasma | Normal | Clinical |
| Sample 31 | Normal Plasma 3 | Plasma | Normal | Clinical |
| Sample 32 | Plasma Spike 03 | Plasma | Spike 03 | Clinical |
| Sample 33 | Cord Serum 1 | Serum | Normal | Clinical/Challenge |
| Sample 34 | Cord Serum 4 | Serum | Spike 01 | Clinical/Challenge |
| Sample 35 | MS Serum Pool | Serum | MS | Clinical |

Supplementary Table 3: Full AONVA results detailing the contribution of lot, site, run and residual variance to total reproducibility for every sample/analyte combination. *Sample excluded from overall precision analysis due to mean in-well concentration falling outside reportable range.

|  |  |  |  |  |  |  | Residual (Error) |  | Between-Lot |  | Between-Site |  | Within-Laboratory |  | Total Reproducibility |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analyte | Sample | Mean (pg/mL) | N |  | Balanced | Outliers | %CV | SD | %CV | SD | %CV | SD | %CV | SD | %CV | SD |
| GFAP | QC 01 | 3,609.84 | 360 |  | balanced | 0 | 5.9 | 213.38 | 3.6 | 128.40 | 8.9 | 320.57 | 9.6 | 347.19 | 14.8 | 534.15 |
| GFAP | QC 02 | 711.75 | 360 |  | balanced | 0 | 7.0 | 50.06 | 3.1 | 21.81 | 7.9 | 56.18 | 10.8 | 77.19 | 15.5 | 109.98 |
| GFAP | QC 03 | 145.54 | 360 |  | balanced | 0 | 6.3 | 9.17 | 2.9 | 4.18 | 0.0 | 0.00 | 12.0 | 17.49 | 13.9 | 20.19 |
| GFAP | QC 04 | 32.64 | 360 |  | balanced | 0 | 9.0 | 2.94 | 1.2 | 0.40 | 0.0 | 0.00 | 15.5 | 5.06 | 18.0 | 5.87 |
| GFAP | QC 05* | 6.24 | 360 |  | balanced | 2 | 17.8 | 1.11 | 1.9 | 0.12 | 0.0 | 0.00 | 18.7 | 1.17 | 25.9 | 1.62 |
| GFAP | Sample 01 | 17.44 | 360 |  | balanced | 0 | 8.6 | 1.49 | 1.8 | 0.31 | 7.3 | 1.27 | 11.9 | 2.08 | 16.5 | 2.88 |
| GFAP | Sample 02 | 48.47 | 360 |  | balanced | 0 | 8.3 | 4.03 | 0.2 | 0.10 | 8.2 | 3.97 | 10.2 | 4.92 | 15.5 | 7.50 |
| GFAP | Sample 03 | 177.88 | 360 |  | balanced | 0 | 7.0 | 12.38 | 4.2 | 7.39 | 8.0 | 14.22 | 7.9 | 14.11 | 13.9 | 24.68 |
| GFAP | Sample 04 | 45.09 | 360 |  | balanced | 0 | 8.3 | 3.75 | 0.6 | 0.28 | 6.2 | 2.79 | 9.1 | 4.11 | 13.8 | 6.23 |
| GFAP | Sample 05 | 175.98 | 360 |  | balanced | 0 | 7.3 | 12.80 | 4.6 | 8.16 | 8.5 | 14.97 | 9.2 | 16.11 | 15.2 | 26.72 |
| GFAP | Sample 06 | 59.99 | 360 |  | balanced | 0 | 14.4 | 8.61 | 10.7 | 6.42 | 16.4 | 9.86 | 18.4 | 11.04 | 30.5 | 18.29 |
| GFAP | Sample 07 | 21.08 | 360 |  | balanced | 0 | 8.6 | 1.81 | 5.1 | 1.07 | 5.3 | 1.12 | 13.1 | 2.76 | 17.3 | 3.65 |
| GFAP | Sample 08 | 102.99 | 360 |  | balanced | 0 | 7.6 | 7.82 | 4.0 | 4.10 | 11.1 | 11.40 | 9.6 | 9.87 | 17.0 | 17.47 |
| GFAP | Sample 09 | 273.29 | 360 |  | balanced | 0 | 6.4 | 17.45 | 2.3 | 6.20 | 9.6 | 26.18 | 9.1 | 24.86 | 14.8 | 40.58 |
| GFAP | Sample 10 | 94.48 | 360 |  | balanced | 0 | 13.7 | 12.99 | 9.8 | 9.29 | 14.8 | 13.98 | 17.4 | 16.47 | 28.4 | 26.86 |
| GFAP | Sample 11 | 46.11 | 360 |  | balanced | 0 | 8.9 | 4.09 | 1.7 | 0.78 | 11.0 | 5.09 | 9.8 | 4.50 | 17.3 | 7.97 |
| GFAP | Sample 12 | 218.57 | 360 |  | balanced | 0 | 6.8 | 14.86 | 4.4 | 9.70 | 7.5 | 16.43 | 9.8 | 21.48 | 14.8 | 32.34 |
| GFAP | Sample 13 | 24.84 | 360 |  | balanced | 0 | 7.8 | 1.94 | 3.8 | 0.94 | 6.9 | 1.71 | 11.6 | 2.89 | 16.1 | 3.99 |
| GFAP | Sample 14 | 20.79 | 360 |  | balanced | 0 | 9.4 | 1.95 | 4.2 | 0.88 | 8.1 | 1.69 | 13.2 | 2.75 | 18.6 | 3.87 |
| GFAP | Sample 15 | 3,565.03 | 360 |  | balanced | 0 | 7.4 | 264.96 | 4.9 | 174.00 | 11.9 | 425.88 | 8.7 | 311.26 | 17.3 | 615.42 |
| GFAP | Sample 16 | 21.98 | 358 |  | Unbalanced | 0 | 8.5 | 1.87 | 3.3 | 0.73 | 1.7 | 0.38 | 12.7 | 2.78 | 15.7 | 3.45 |
| GFAP | Sample 17 | 2,948.10 | 360 |  | balanced | 0 | 9.3 | 275.23 | 4.6 | 135.92 | 12.4 | 364.51 | 9.2 | 269.95 | 18.6 | 547.69 |
| GFAP | Sample 18 | 47.46 | 338 |  | Unbalanced | 0 | 4.8 | 2.28 | 3.3 | 1.59 | 2.9 | 1.36 | 14.4 | 6.85 | 15.8 | 7.51 |
| GFAP | Sample 19 | 35.22 | 360 |  | balanced | 0 | 8.1 | 2.87 | 4.3 | 1.51 | 5.0 | 1.77 | 10.4 | 3.67 | 14.8 | 5.21 |
| GFAP | Sample 20 | 657.03 | 360 |  | balanced | 0 | 8.3 | 54.31 | 5.1 | 33.83 | 10.0 | 65.88 | 11.1 | 73.25 | 17.9 | 117.48 |
| GFAP | Sample 21 | 155.01 | 358 |  | Unbalanced | 0 | 6.0 | 9.35 | 3.7 | 5.76 | 5.0 | 7.75 | 12.4 | 19.23 | 15.1 | 23.47 |
| GFAP | Sample 22 | 122.28 | 360 |  | balanced | 0 | 7.5 | 9.20 | 4.6 | 5.58 | 13.1 | 16.03 | 10.8 | 13.18 | 19.1 | 23.37 |
| GFAP | Sample 23 | 188.99 | 360 |  | balanced | 0 | 6.6 | 12.42 | 3.8 | 7.10 | 8.1 | 15.23 | 10.3 | 19.54 | 15.1 | 28.61 |
| GFAP | Sample 24 | 64.25 | 360 |  | balanced | 0 | 8.9 | 5.69 | 5.8 | 3.71 | 8.8 | 5.66 | 10.4 | 6.71 | 17.3 | 11.10 |
| GFAP | Sample 25 | 69.03 | 352 |  | Unbalanced | 0 | 10.7 | 7.36 | 6.7 | 4.61 | 10.5 | 7.26 | 10.8 | 7.43 | 19.6 | 13.54 |
| GFAP | Sample 26 | 198.76 | 360 |  | balanced | 0 | 6.6 | 13.14 | 2.5 | 4.97 | 7.6 | 15.11 | 11.0 | 21.93 | 15.2 | 30.11 |
| GFAP | Sample 27 | 736.95 | 360 |  | balanced | 0 | 8.6 | 63.32 | 3.2 | 23.43 | 11.4 | 84.38 | 10.9 | 80.56 | 18.3 | 134.79 |
| GFAP | Sample 28 | 26.86 | 360 |  | balanced | 0 | 8.2 | 2.19 | 4.0 | 1.07 | 6.2 | 1.67 | 15.2 | 4.08 | 18.7 | 5.03 |
| GFAP | Sample 29 | 732.66 | 360 |  | balanced | 0 | 7.5 | 54.97 | 5.5 | 40.43 | 9.5 | 69.74 | 10.7 | 78.54 | 17.1 | 125.25 |
| GFAP | Sample 30 | 83.03 | 352 |  | Unbalanced | 6 | 28.4 | 23.54 | 16.9 | 14.06 | 15.7 | 13.07 | 20.0 | 16.59 | 41.7 | 34.61 |
| GFAP | Sample 31 | 41.96 | 360 |  | balanced | 0 | 7.4 | 3.12 | 6.1 | 2.57 | 8.6 | 3.62 | 12.1 | 5.08 | 17.7 | 7.43 |
| GFAP | Sample 32 | 156.29 | 360 |  | balanced | 0 | 7.7 | 11.96 | 1.9 | 2.90 | 10.2 | 15.88 | 10.5 | 16.45 | 16.6 | 25.96 |
| GFAP | Sample 33 | 183.53 | 344 |  | Unbalanced | 0 | 7.3 | 13.45 | 4.9 | 8.96 | 8.8 | 16.11 | 11.6 | 21.37 | 17.0 | 31.26 |
| GFAP | Sample 34 | 3,241.04 | 360 |  | balanced | 0 | 7.0 | 226.71 | 3.8 | 121.71 | 10.8 | 349.33 | 12.5 | 406.55 | 18.3 | 594.58 |
| GFAP | Sample 35 | 66.71 | 360 |  | balanced | 0 | 8.6 | 5.72 | 6.3 | 4.23 | 7.8 | 5.21 | 14.0 | 9.34 | 19.3 | 12.85 |
| NF-L | QC 01 | 13,266.06 | 344 |  | Unbalanced | 0 | 6.5 | 860.26 | 2.8 | 368.69 | 7.4 | 982.59 | 7.2 | 955.09 | 12.5 | 1,659.42 |
| NF-L | QC 02 | 2,230.55 | 344 |  | Unbalanced | 0 | 9.1 | 202.68 | 2.5 | 54.70 | 5.4 | 120.16 | 9.5 | 210.82 | 14.4 | 320.87 |
| NF-L | QC 03 | 463.54 | 344 |  | Unbalanced | 0 | 7.3 | 34.01 | 1.6 | 7.44 | 1.4 | 6.56 | 9.7 | 45.10 | 12.4 | 57.35 |
| NF-L | QC 04 | 94.35 | 344 |  | Unbalanced | 0 | 11.3 | 10.68 | 3.3 | 3.14 | 6.0 | 5.65 | 13.3 | 12.54 | 18.7 | 17.69 |
| NF-L | QC 05* | 21.55 | 344 |  | Unbalanced | 2 | 25.5 | 5.49 | 0.0 | 0.00 | 21.0 | 4.52 | 30.1 | 6.50 | 44.7 | 9.63 |
| NF-L | Sample 01 | 64.99 | 344 |  | Unbalanced | 0 | 11.1 | 7.23 | 0.0 | 0.00 | 8.4 | 5.47 | 16.1 | 10.44 | 21.3 | 13.83 |
| NF-L | Sample 02 | 113.56 | 344 |  | Unbalanced | 0 | 9.4 | 10.70 | 4.4 | 4.98 | 4.1 | 4.66 | 11.1 | 12.65 | 15.8 | 17.91 |
| NF-L | Sample 03 | 119.17 | 344 |  | Unbalanced | 0 | 7.5 | 8.96 | 3.9 | 4.66 | 4.4 | 5.25 | 11.2 | 13.29 | 14.7 | 17.50 |
| NF-L | Sample 04 | 101.34 | 344 |  | Unbalanced | 0 | 10.6 | 10.71 | 3.4 | 3.44 | 3.9 | 3.92 | 11.1 | 11.24 | 16.2 | 16.38 |
| NF-L | Sample 05 | 111.91 | 344 |  | Unbalanced | 0 | 8.5 | 9.47 | 2.4 | 2.70 | 1.0 | 1.14 | 12.4 | 13.89 | 15.3 | 17.07 |
| NF-L | Sample 06 | 183.42 | 344 |  | Unbalanced | 0 | 12.8 | 23.49 | 17.2 | 31.52 | 13.8 | 25.34 | 19.9 | 36.42 | 32.3 | 59.28 |
| NF-L | Sample 07 | 67.83 | 344 |  | Unbalanced | 0 | 11.3 | 7.68 | 5.2 | 3.53 | 3.1 | 2.09 | 15.2 | 10.32 | 19.9 | 13.50 |
| NF-L | Sample 08 | 341.38 | 344 |  | Unbalanced | 2 | 9.4 | 31.98 | 2.5 | 8.61 | 4.8 | 16.28 | 10.5 | 35.99 | 15.1 | 51.55 |
| NF-L | Sample 09 | 1,881.55 | 344 |  | Unbalanced | 0 | 8.6 | 162.42 | 1.3 | 25.26 | 10.7 | 201.13 | 8.7 | 162.90 | 16.3 | 306.61 |
| NF-L | Sample 10 | 425.58 | 344 |  | Unbalanced | 0 | 9.8 | 41.64 | 2.1 | 8.91 | 10.6 | 45.26 | 9.3 | 39.72 | 17.3 | 73.75 |
| NF-L | Sample 11 | 480.22 | 344 |  | Unbalanced | 0 | 9.6 | 45.93 | 2.4 | 11.59 | 11.7 | 56.40 | 8.1 | 38.99 | 17.4 | 83.34 |
| NF-L | Sample 12 | 1,085.41 | 344 |  | Unbalanced | 0 | 6.6 | 71.17 | 4.7 | 50.81 | 2.8 | 30.78 | 9.9 | 107.16 | 13.1 | 141.70 |
| NF-L | Sample 13 | 65.46 | 344 |  | Unbalanced | 0 | 13.0 | 8.52 | 3.7 | 2.39 | 4.2 | 2.73 | 13.1 | 8.55 | 19.3 | 12.60 |
| NF-L | Sample 14* | 30.48 | 344 |  | Unbalanced | 0 | 19.7 | 6.01 | 2.7 | 0.82 | 3.6 | 1.08 | 21.1 | 6.44 | 29.2 | 8.91 |
| NF-L | Sample 15* | 55,612.82 | 344 |  | Unbalanced | 0 | 8.9 | 4,927.08 | 8.6 | 4,770.00 | 2.8 | 1,575.74 | 22.4 | 12,456.66 | 25.7 | 14,306.65 |
| NF-L | Sample 16 | 303.95 | 342 |  | Unbalanced | 0 | 8.3 | 25.14 | 5.5 | 16.80 | 7.4 | 22.63 | 11.7 | 35.43 | 17.0 | 51.78 |
| NF-L | Sample 17* | 47,633.78 | 344 |  | Unbalanced | 4 | 13.0 | 6,180.90 | 11.1 | 5,268.23 | 6.6 | 3,135.93 | 28.2 | 13,419.59 | 33.6 | 15,996.16 |
| NF-L | Sample 18 | 83.50 | 324 |  | Unbalanced | 0 | 9.1 | 7.63 | 0.0 | 0.00 | 6.8 | 5.71 | 19.1 | 15.93 | 22.2 | 18.56 |
| NF-L | Sample 19 | 42.67 | 344 |  | Unbalanced | 0 | 12.4 | 5.31 | 4.0 | 1.69 | 3.9 | 1.65 | 15.4 | 6.55 | 20.5 | 8.76 |
| NF-L | Sample 20 | 10,514.95 | 344 |  | Unbalanced | 0 | 10.0 | 1,052.59 | 3.5 | 373.06 | 9.6 | 1,013.21 | 10.1 | 1,061.35 | 17.5 | 1,843.96 |
| NF-L | Sample 21 | 1,655.65 | 342 |  | Unbalanced | 0 | 7.4 | 122.63 | 0.5 | 8.31 | 1.4 | 22.53 | 11.0 | 182.71 | 13.4 | 221.35 |
| NF-L | Sample 22 | 1,825.34 | 344 |  | Unbalanced | 0 | 8.8 | 160.99 | 1.7 | 30.50 | 15.9 | 289.43 | 8.6 | 156.80 | 20.1 | 367.71 |
| NF-L | Sample 23 | 493.80 | 344 |  | Unbalanced | 0 | 8.2 | 40.27 | 0.0 | 0.00 | 11.2 | 55.24 | 9.9 | 49.10 | 17.0 | 84.17 |
| NF-L | Sample 24 | 53.83 | 344 |  | Unbalanced | 0 | 12.2 | 6.58 | 3.1 | 1.66 | 11.5 | 6.19 | 15.3 | 8.22 | 22.9 | 12.32 |
| NF-L | Sample 25 | 74.45 | 336 |  | Unbalanced | 0 | 10.5 | 7.81 | 0.0 | 0.00 | 15.5 | 11.51 | 22.4 | 16.65 | 29.1 | 21.70 |
| NF-L | Sample 26 | 1,956.46 | 344 |  | Unbalanced | 0 | 6.1 | 119.18 | 5.9 | 115.57 | 6.4 | 124.57 | 9.8 | 192.51 | 14.5 | 283.09 |
| NF-L | Sample 27 | 13,059.37 | 344 |  | Unbalanced | 0 | 8.9 | 1,159.53 | 3.9 | 514.80 | 10.6 | 1,387.33 | 11.1 | 1,448.62 | 18.2 | 2,373.33 |
| NF-L | Sample 28 | 48.08 | 344 |  | Unbalanced | 0 | 13.6 | 6.54 | 2.7 | 1.31 | 9.4 | 4.54 | 23.4 | 11.24 | 28.8 | 13.84 |
| NF-L | Sample 29 | 10,910.49 | 344 |  | Unbalanced | 0 | 8.4 | 917.71 | 2.0 | 220.86 | 12.8 | 1,394.73 | 8.7 | 951.06 | 17.7 | 1,934.10 |
| NF-L | Sample 30 | 75.16 | 338 |  | Unbalanced | 0 | 14.2 | 10.68 | 3.7 | 2.80 | 17.2 | 12.96 | 19.2 | 14.39 | 29.7 | 22.29 |
| NF-L | Sample 31 | 59.67 | 344 |  | Unbalanced | 0 | 11.6 | 6.94 | 2.8 | 1.68 | 13.2 | 7.88 | 16.1 | 9.58 | 24.0 | 14.31 |
| NF-L | Sample 32 | 1,808.85 | 344 |  | Unbalanced | 0 | 9.2 | 166.28 | 3.3 | 60.34 | 12.2 | 220.10 | 8.7 | 157.69 | 17.9 | 323.42 |
| NF-L | Sample 33 | 121.34 | 328 |  | Unbalanced | 0 | 11.5 | 14.00 | 1.7 | 2.01 | 5.5 | 6.67 | 13.6 | 16.47 | 18.7 | 22.71 |
| NF-L | Sample 34* | 50,372.88 | 344 |  | Unbalanced | 2 | 14.1 | 7,089.77 | 10.3 | 5,174.51 | 4.3 | 2,170.04 | 37.2 | 18,741.97 | 41.3 | 20,808.91 |
| NF-L | Sample 35 | 53.55 | 344 |  | Unbalanced | 0 | 18.1 | 9.69 | 5.6 | 3.01 | 4.4 | 2.37 | 14.3 | 7.65 | 24.1 | 12.92 |
| Tau | QC 01 | 2,256.48 | 360 |  | balanced | 0 | 6.4 | 143.44 | 2.5 | 57.29 | 8.7 | 197.07 | 7.9 | 177.26 | 13.6 | 306.79 |
| Tau | QC 02 | 453.66 | 360 |  | balanced | 0 | 6.5 | 29.39 | 4.4 | 19.78 | 4.6 | 20.86 | 10.5 | 47.82 | 13.9 | 63.07 |
| Tau | QC 03 | 86.55 | 360 |  | balanced | 0 | 6.9 | 5.98 | 4.1 | 3.56 | 6.1 | 5.31 | 9.5 | 8.23 | 13.9 | 12.02 |
| Tau | QC 04 | 18.49 | 360 |  | balanced | 0 | 6.1 | 1.13 | 5.2 | 0.96 | 0.0 | 0.00 | 13.5 | 2.50 | 15.7 | 2.90 |
| Tau | QC 05 | 3.64 | 360 |  | balanced | 0 | 8.7 | 0.32 | 4.2 | 0.15 | 2.7 | 0.10 | 18.9 | 0.69 | 21.4 | 0.78 |
| Tau | Sample 01 | 23.65 | 360 |  | balanced | 0 | 6.9 | 1.63 | 2.6 | 0.62 | 0.0 | 0.00 | 11.4 | 2.69 | 13.6 | 3.21 |
| Tau | Sample 02 | 24.94 | 360 |  | balanced | 0 | 6.3 | 1.58 | 3.1 | 0.78 | 0.0 | 0.00 | 12.0 | 2.98 | 13.9 | 3.46 |
| Tau | Sample 03* | 0.63 | 360 |  | balanced | 4 | 33.8 | 0.21 | 19.0 | 0.12 | 10.4 | 0.07 | 26.9 | 0.17 | 48.3 | 0.31 |
| Tau | Sample 04 | 7.85 | 360 |  | balanced | 0 | 7.9 | 0.62 | 5.0 | 0.39 | 6.8 | 0.53 | 17.2 | 1.35 | 20.7 | 1.63 |
| Tau | Sample 05* | 0.59 | 360 |  | balanced | 2 | 20.6 | 0.12 | 18.6 | 0.11 | 9.0 | 0.05 | 20.1 | 0.12 | 35.4 | 0.21 |
| Tau | Sample 06 | 6.80 | 360 |  | balanced | 0 | 11.5 | 0.78 | 7.1 | 0.48 | 9.0 | 0.62 | 13.4 | 0.91 | 21.0 | 1.43 |
| Tau | Sample 07 | 24.13 | 360 |  | balanced | 2 | 9.6 | 2.31 | 4.9 | 1.19 | 3.8 | 0.92 | 12.4 | 2.98 | 16.8 | 4.06 |
| Tau | Sample 08 | 20.09 | 360 |  | balanced | 0 | 9.2 | 1.84 | 1.8 | 0.36 | 2.3 | 0.46 | 11.1 | 2.22 | 14.7 | 2.95 |
| Tau | Sample 09 | 590.76 | 360 |  | balanced | 0 | 6.5 | 38.41 | 1.5 | 8.89 | 9.3 | 54.93 | 9.7 | 57.08 | 15.0 | 88.49 |
| Tau | Sample 10 | 111.60 | 360 |  | balanced | 0 | 7.7 | 8.55 | 2.4 | 2.63 | 5.7 | 6.33 | 9.4 | 10.50 | 13.6 | 15.18 |
| Tau | Sample 11 | 120.79 | 360 |  | balanced | 0 | 7.3 | 8.77 | 2.5 | 2.97 | 8.6 | 10.41 | 9.5 | 11.43 | 14.9 | 18.02 |
| Tau | Sample 12 | 128.80 | 360 |  | balanced | 0 | 6.7 | 8.58 | 1.3 | 1.73 | 3.2 | 4.08 | 10.3 | 13.26 | 12.7 | 16.40 |
| Tau | Sample 13 | 12.00 | 360 |  | balanced | 0 | 10.1 | 1.21 | 5.8 | 0.69 | 8.6 | 1.03 | 15.5 | 1.86 | 21.2 | 2.55 |
| Tau | Sample 14* | 2.59 | 360 |  | balanced | 2 | 13.1 | 0.34 | 1.0 | 0.03 | 3.4 | 0.09 | 14.7 | 0.38 | 20.0 | 0.52 |
| Tau | Sample 15* | 5,823.22 | 360 |  | balanced | 0 | 5.5 | 319.97 | 5.1 | 297.55 | 3.5 | 205.84 | 13.2 | 766.02 | 15.6 | 905.58 |
| Tau | Sample 16 | 20.88 | 358 |  | Unbalanced | 0 | 7.9 | 1.66 | 0.0 | 0.00 | 5.0 | 1.05 | 13.5 | 2.82 | 16.5 | 3.44 |
| Tau | Sample 17* | 6,045.10 | 360 |  | balanced | 0 | 6.9 | 419.18 | 8.1 | 488.85 | 7.3 | 441.85 | 18.0 | 1,089.97 | 22.2 | 1,340.88 |
| Tau | Sample 18 | 109.67 | 338 |  | Unbalanced | 0 | 4.9 | 5.32 | 0.0 | 0.00 | 7.7 | 8.42 | 15.6 | 17.13 | 18.1 | 19.81 |
| Tau | Sample 19 | 16.76 | 360 |  | balanced | 0 | 6.7 | 1.12 | 4.8 | 0.80 | 0.8 | 0.13 | 11.8 | 1.98 | 14.4 | 2.42 |
| Tau | Sample 20 | 2,429.20 | 360 |  | balanced | 0 | 6.7 | 162.59 | 6.1 | 148.58 | 9.0 | 218.92 | 14.0 | 340.33 | 19.0 | 460.72 |
| Tau | Sample 21 | 676.10 | 358 |  | Unbalanced | 0 | 5.2 | 34.95 | 2.9 | 19.44 | 3.6 | 24.65 | 12.7 | 86.17 | 14.5 | 98.15 |
| Tau | Sample 22 | 627.28 | 360 |  | balanced | 2 | 7.0 | 43.74 | 3.2 | 19.88 | 4.0 | 25.20 | 10.4 | 65.54 | 13.6 | 85.09 |
| Tau | Sample 23 | 100.37 | 360 |  | balanced | 0 | 6.3 | 6.32 | 0.5 | 0.54 | 2.4 | 2.44 | 11.6 | 11.62 | 13.4 | 13.46 |
| Tau | Sample 24 | 22.63 | 360 |  | balanced | 0 | 7.4 | 1.68 | 5.3 | 1.20 | 3.9 | 0.89 | 16.8 | 3.81 | 19.5 | 4.42 |
| Tau | Sample 25 | 21.46 | 352 |  | Unbalanced | 0 | 9.7 | 2.08 | 3.6 | 0.77 | 6.9 | 1.47 | 16.7 | 3.59 | 20.8 | 4.47 |
| Tau | Sample 26 | 26.04 | 360 |  | balanced | 0 | 6.4 | 1.68 | 5.5 | 1.44 | 3.3 | 0.85 | 12.4 | 3.22 | 15.3 | 4.00 |
| Tau | Sample 27 | 2,168.33 | 360 |  | balanced | 0 | 7.4 | 160.42 | 5.1 | 110.99 | 13.2 | 286.38 | 12.8 | 277.00 | 20.5 | 443.61 |
| Tau | Sample 28* | 0.78 | 360 |  | balanced | 0 | 18.8 | 0.15 | 8.7 | 0.07 | 13.7 | 0.11 | 36.3 | 0.28 | 44.0 | 0.34 |
| Tau | Sample 29 | 1,934.30 | 360 |  | balanced | 0 | 7.3 | 141.90 | 2.4 | 46.84 | 9.4 | 181.41 | 11.0 | 212.76 | 16.4 | 317.03 |
| Tau | Sample 30 | 6.81 | 352 |  | Unbalanced | 0 | 7.3 | 0.49 | 13.6 | 0.93 | 3.6 | 0.24 | 16.8 | 1.14 | 23.1 | 1.57 |
| Tau | Sample 31 | 20.16 | 360 |  | balanced | 0 | 7.7 | 1.55 | 4.0 | 0.80 | 2.4 | 0.47 | 12.5 | 2.52 | 15.4 | 3.10 |
| Tau | Sample 32 | 690.04 | 360 |  | balanced | 0 | 7.7 | 52.80 | 5.4 | 37.11 | 11.3 | 78.31 | 11.1 | 76.47 | 18.4 | 127.06 |
| Tau | Sample 33* | 0.62 | 344 |  | Unbalanced | 4 | 37.7 | 0.24 | 16.1 | 0.10 | 14.6 | 0.09 | 18.6 | 0.12 | 47.3 | 0.30 |
| Tau | Sample 34* | 5,401.19 | 360 |  | balanced | 0 | 5.7 | 308.43 | 8.4 | 454.22 | 12.1 | 650.97 | 16.6 | 894.67 | 22.9 | 1,235.17 |
| Tau | Sample 35 | 5.17 | 360 |  | balanced | 0 | 9.3 | 0.48 | 7.6 | 0.39 | 4.1 | 0.21 | 18.8 | 0.97 | 22.7 | 1.17 |

Supplementary Table 4: Substances tested for interference with assay measurements, and the final in-sample concentration tested.

| Interferent | Test Concentration (mg/ml) |
| --- | --- |
| Ampicillin | 5 |
| Benzodiazapines | 0.0125 |
| Bilirubin (conjugated) | 1 |
| Bilirubin (unconjugated) | 0.25 |
| Erythropoietin | 0.005 |
| Gentamicin | 0.03 |
| Levetiracetam | 0.18 |
| Melatonin | 0.595 |
| Morphine | 0.05 |
| Phenobarbital | 0.05 |
| Phenytoin | 0.06 |
| Piperacillin | 2.5 |
| Tazobactam | 2.5 |
| Topiramate | 0.06 |
| Vancomycin | 5 |