Supplementary Information

Recombinant Allergen-Based Quantum Dot Lateral Flow Immunoassays for
Component-Specific IgE Detection in House Dust Mite Allergy

Applied Biochemistry and Biotechnology

Yun-Zhi Lia, Zi-Xin Wanga, Xiao-Ya Wanga, Yan-Yan Sunc*, Xin-lei Chub*, Chun-Ming Donga*

a College of Marine and Environmental Sciences, Tianjin University of Science and Technology, Tianjin 300457, China

b Department of Epidemiology and Biostatistics, Tianjin Medical University Cancer Institute and Hospital, Tianjin Medical University, Tianjin 300060, China

c Jinan Deheng Medical Technology Co., Ltd., Jinan 250031, Shandong Province, China

*Corresponding authors: Chun-Ming Dong (mingchundongjy@tust.edu.cn), Xin-lei Chu (chuxinlei2014@163.com), and Yan-Yan Sun (sunyanyanggg@163.com)

Supplementary Material 1 This file contains Supplementary Tables S1-S10 and Supplementary Figs. S1-S7 supporting the bioinformatics analyses, QD-LFIA accelerated storage assessment, MBP-control testing, and comparison with ImmunoCAP

Table S1 ProtParam-predicted physicochemical properties of Der p 1, Der p 2, Der p 10, and Der p 23

| Parameter | Der p 1 | Der p 2 | Der p 10 | Der p 23 |
| --- | --- | --- | --- | --- |
| Number of amino acids | 300 | 130 | 284 | 70 |
| Molecular weight (Da) | 34,033.92 | 14,252.52 | 32,900.74 | 8,112.66 |
| Theoretical pI | 5.40 | 6.50 | 4.78 | 4.32 |
| Charged residues (negative/positive) | 36/27 | 17/16 | 74/50 | 16/6 |
| Instability index | 36.94 | 35.71 | 46.78 | 61.56 |
| Aliphatic index | 74.87 | 89.92 | 79.15 | 18.14 |
| GRAVY | −0.453 | −0.205 | −1.060 | −1.344 |

Note. GRAVY, grand average of hydropathicity. Values were calculated using the ExPASy ProtParam tool.

Table S2 PSIPRED-predicted secondary-structure composition of Der p 1, Der p 2, Der p 10, and Der p 23

| Allergen | α-Helix | β-Strand | Random coil |
| --- | --- | --- | --- |
| Der p 1 | 98 (32.67%) | 43 (14.33%) | 159 (53.00%) |
| Der p 2 | 0 (0.00%) | 69 (53.08%) | 61 (46.92%) |
| Der p 10 | 281 (98.94%) | 0 (0.00%) | 3 (1.06%) |
| Der p 23 | 2 (2.86%) | 12 (17.14%) | 56 (80.00%) |

Note. Values are shown as the number of residues (percentage of the amino acid sequence).

Table S3 Linear B-cell epitope regions predicted using the IEDB Analysis Resource

| Allergen | Predicted region(s) |
| --- | --- |
| Der p 1 | 6–25: FEEYKKAFNK SYATFEDEEA 35–36: KY 38–102: QSNGGAINHL SDLSLDEFKN RFLMSAEAFE HLKTQFDLNA ETNACSINGN APAEIDLRQM RTVTP 104: R 128–135: YRNQSLDL 147–148: HG 161–176: IQHNGVVQES YYRYVA 178–179: EQ 181–194: CRRPNAQRFG ISNY 227: A 229–245: RHYDGRTIIQ RDNGYQP 256–263: SNAQGVDY 272–276: TNWGD 278: G |
| Der p 2 | 5–17: DVKDCANHEI KKV 22–25: CHGS 64–76: VDVPGIDPNA CHY 78–90: KCPLVKGQQY DIK 93–101: WNVPKIAPK |
| Der p 10 | 5–171: KNKMQAMKLE KDNAIDRAEI AEQKARDANL RAEKSEEEVR ALQKKIQQIE NELDQVQEQL SAANTKLEEK EKALQTAEGD VAALNRRIQL IEEDLERSEE RLKIATAKLE EASQSADESE RMRKMLEHRS ITDEERMEGL ENQLKEARMM AEDADRKYDE VARKLAM 174–279: ADLERAEERA ETGESKIVEL EEELRVGNNL KSLEVSEEKA QQREEAHEQQ IRIMTTKLKE AEARAEFAER SVQLQKEVGR LEDELVHEKE KYKSISDELD QTFA |
| Der p 23 | 5–30: NDDDPTTTVH PTTTEQPDDK FECPSR 34–39: FADPKD 50–67: EAVHKDCPGN TRWNEDEE |

Note. Residue positions are one-based. Spaces within amino acid sequences are included only to improve readability.

Table S4 Linear B-cell epitope regions predicted using BepiPred 2.0

| Allergen | Predicted region(s) |
| --- | --- |
| Der p 1 | 7–20: EEYKKAFNKS YATF 33–53: SVKYVQSNGG AINHLSDLSL D 55: F 57–151: NRFLMSAEAF EHLKTQFDLN AETNACSING NAPAEIDLRQ MRTVTPIRMQ GGCGSAWAFS GVAATESAYL AYRNQSLDLA EQELVDCASQ HGCHG 178–192: EQSCRRPNAQ RFGIS 207–217: EALAQTHSAI A 236–244: IIQRDNGYQ 253–281: VGYSNAQGVD YWIVRNSWDT NWGDNGYGY 283: A 285: N 295–296: PY |
| Der p 2 | 5–14: DVKDCANHEI 24–31: GSEPCIIH 35: P 39–64: EAVFEANQNT KTAKIEIKAS IDGLEV 74–76: CHY 78–79: KC 81–93: LVKGQQYDIK YTW 98–127: IAPKSENVVV TVKVMGDDGV LACAIATHAK |
| Der p 10 | 5–42: KNKMQAMKLE KDNAIDRAEI AEQKARDANL RAEKSEEE 48–91: KKIQQIENEL DQVQEQLSAA NTKLEEKEKA LQTAEGDVAA LNRR 93–144: QLIEEDLERS EERLKIATAK LEEASQSADE SERMRKMLEH RSITDEERME GL 146–280: NQLKEARMMA EDADRKYDEV ARKLAMVEAD LERAEERAET GESKIVELEE ELRVVGNNLK SLEVSEEKAQ QREEAHEQQI RIMTTKLKEA EARAEFAERS VQKLQKEVGR LEDELVHEKE KYKSISDELD QTFAE |
| Der p 23 | 6: D 62–66: WNEDE |

Note. Residue positions are one-based. Spaces within amino acid sequences are included only to improve readability.

Table S5 Potential linear B-cell epitope regions identified from DNAStar Protean profiles

| Allergen | Predicted region(s) |
| --- | --- |
| Der p 1 | 14: N 27–28: RK 55–56: FK 131: Q 153–154: TI 157: G 170: S 179–187: QSCRRPNAQ 201: N 203–205: NKI 232–233: DG 239–244: RDNGYQ 268–271: NSWD 274: W |
| Der p 2 | 32–34: RGK 45–48: NQNT 85: Q 98–102: IAPKS |
| Der p 10 | 14–17: EKDN 69–71: TKL 117–120: SQSA 157–160: DADR 203–204: NL |
| Der p 23 | 5–10: NDDDPT 17–24: TTEQPDDK 30: R 36–41: DPKDPH 59–67: NTRWNEDEE |

Note. Residue positions are one-based. Spaces within amino acid sequences are included only to improve readability.

Table S6 Sixteen-residue linear B-cell epitope windows predicted using ABCpred

| Allergen | Predicted region(s) |
| --- | --- |
| Der p 1 | 4–19; 15–30; 24–39; 34–49; 46–61; 62–77 69–84; 87–102; 96–111; 107–122; 116–131; 127–142 138–153; 148–163; 155–170; 165–180; 178–193; 189–204 204–219; 219–234; 230–245; 236–251; 242–257; 258–273 266–281; 273–288; 284–299 |
| Der p 2 | 5–20; 11–26; 21–36; 31–46; 38–53; 56–71 65–80; 73–88; 82–97; 88–103; 94–109; 110–125 |
| Der p 10 | 1–16; 12–27; 21–36; 32–47; 38–53; 48–63 64–79; 75–90; 82–97; 92–107; 99–114; 106–121 117–132; 123–138; 132–147; 142–157; 149–164; 155–170 162–177; 169–184; 179–194; 191–206; 215–230; 224–239 232–247; 243–258; 254–269; 262–277; 268–283 |
| Der p 23 | 6–21; 12–27; 22–37; 32–47; 42–57; 48–63 55–70 |

Note. Each predicted window contains 16 residues; residue positions are one-based.

Table S7 Results of the 10-day accelerated storage comparison for the Der p 1 QD-LFIA

| Sample | ImmunoCAP-assigned concentration (IU/mL) | QD-LFIA result after 10 days at 4 °C (IU/mL) | QD-LFIA result after 10 days at 37 °C (IU/mL) | Relative deviation (%) |
| --- | --- | --- | --- | --- |
| 1 | 94.92 | 98.72 | 94.27 | −4.50% |
| 2 | 61.97 | 62.59 | 58.92 | −5.87% |
| 3 | 29.52 | 30.11 | 31.32 | 4.01% |
| 4 | 79.87 | 83.06 | 76.22 | −8.24% |
| 5 | 7.05 | 6.77 | 6.89 | 1.86% |
| 6 | 3.83 | 3.91 | 4.00 | 2.28% |
| 7 | 0.98 | 0.96 | 0.93 | −3.48% |
| 8 | 0.47 | 0.46 | 0.48 | 4.25% |

Note. Relative deviation (%) = [(C37 °C/C4 °C) − 1] × 100. Deviations were calculated using the unrounded mean assay values.

Table S8 Results of the 10-day accelerated storage comparison for the Der p 2 QD-LFIA

| Sample | ImmunoCAP-assigned concentration (IU/mL) | QD-LFIA result after 10 days at 4 °C (IU/mL) | QD-LFIA result after 10 days at 37 °C (IU/mL) | Relative deviation (%) |
| --- | --- | --- | --- | --- |
| 1 | 85.48 | 88.04 | 82.46 | −6.34% |
| 2 | 61.20 | 61.20 | 59.82 | −2.25% |
| 3 | 30.31 | 29.40 | 29.86 | 1.57% |
| 4 | 115.32 | 118.78 | 111.42 | −6.20% |
| 5 | 5.12 | 5.02 | 5.14 | 2.39% |
| 6 | 3.38 | 3.21 | 3.29 | 2.60% |
| 7 | 1.45 | 1.45 | 1.37 | −5.50% |
| 8 | 0.48 | 0.47 | 0.43 | −7.64% |

Note. Relative deviation (%) = [(C37 °C/C4 °C) − 1] × 100. Deviations were calculated using the unrounded mean assay values.

Table S9 Results of the 10-day accelerated storage comparison for the Der p 10 QD-LFIA

| Sample | ImmunoCAP-assigned concentration (IU/mL) | QD-LFIA result after 10 days at 4 °C (IU/mL) | QD-LFIA result after 10 days at 37 °C (IU/mL) | Relative deviation (%) |
| --- | --- | --- | --- | --- |
| 1 | 92.61 | 95.39 | 92.22 | −3.32% |
| 2 | 66.25 | 64.93 | 62.22 | −4.16% |
| 3 | 39.56 | 39.56 | 36.05 | −8.88% |
| 4 | 37.36 | 36.24 | 33.00 | −8.94% |
| 5 | 7.08 | 7.08 | 7.11 | 0.37% |
| 6 | 3.21 | 3.21 | 3.18 | −0.94% |
| 7 | 1.96 | 1.86 | 1.91 | 2.34% |
| 8 | 0.38 | 0.37 | 0.34 | −7.49% |

Note. Relative deviation (%) = [(C37 °C/C4 °C) − 1] × 100. Deviations were calculated using the unrounded mean assay values.

Table S10 Results of the 10-day accelerated storage comparison for the Der p 23 QD-LFIA

| Sample | ImmunoCAP-assigned concentration (IU/mL) | QD-LFIA result after 10 days at 4 °C (IU/mL) | QD-LFIA result after 10 days at 37 °C (IU/mL) | Relative deviation (%) |
| --- | --- | --- | --- | --- |
| 1 | 86.27 | 81.96 | 78.63 | −4.06% |
| 2 | 59.45 | 60.04 | 61.29 | 2.08% |
| 3 | 25.84 | 25.06 | 25.70 | 2.55% |
| 4 | 114.04 | 111.76 | 112.49 | 0.65% |
| 5 | 5.37 | 5.16 | 4.87 | −5.45% |
| 6 | 3.01 | 2.95 | 2.83 | −4.20% |
| 7 | 1.61 | 1.61 | 1.60 | −0.41% |
| 8 | 0.49 | 0.51 | 0.47 | −7.94% |

Note. Relative deviation (%) = [(C37 °C/C4 °C) − 1] × 100. Deviations were calculated using the unrounded mean assay values.

Fig. S1 AlphaFold 3-predicted tertiary structure of Der p 1 (pTM = 0.94). The predicted structure and predicted aligned error plot are shown

Fig. S2 AlphaFold 3-predicted tertiary structure of Der p 2 (pTM = 0.88). The predicted structure and predicted aligned error plot are shown

Fig. S3 AlphaFold 3-predicted tertiary structure of Der p 10 (pTM = 0.38). The predicted structure and predicted aligned error plot are shown; the low pTM value indicates limited confidence in the global model

Fig. S4 AlphaFold 3-predicted tertiary structure of Der p 23 (pTM = 0.60). The predicted structure and predicted aligned error plot are shown; structural interpretation should account for the modest model confidence

Fig. S5 DNAStar Protean profiles used to identify candidate linear B-cell epitope regions in (A) Der p 1, (B) Der p 2, (C) Der p 10, and (D) Der p 23. The displayed tracks include hydrophilicity, flexibility, surface probability, and antigenicity profiles

Fig. S6 Distribution of apparent IgE concentrations obtained with the MBP-control QD-LFIA in 112 residual serum samples

Fig. S7 Distribution of component-specific IgE concentrations measured using ImmunoCAP (A, C, E, G) and QD-LFIA (B, D, F, H): Der p 1 (A, B; n = 84), Der p 2 (C, D; n = 81), Der p 10 (E, F; n = 82), and Der p 23 (G, H; n = 91). The cutoff for a positive result was 0.35 IU/mL