A LABEL-FREE MULTIPLEX SALIVARY BIOSENSING PLATFORM FOR NON-INVASIVE INFLAMMATORY PROFILING

Bianca Elizabeth David¹, Kai-Chun Lin¹, Shreya Parulekar¹, Akash Kumar¹, Sumana Karmakar¹, Sasya Madhurantakam¹, Vikram Narayanan Dhamu², Sriram Muthukumar², Shalini Prasad¹*

¹Department of Bioengineering, The University of Texas at Dallas, Richardson, TX, United States

²EnLiSense LLC, Allen, TX, United States

Supplementary Figure S1. Calibrated dose-response (CDR) curves for CRP, calprotectin, TNF-α, lactoferrin, IL-6, IL-12, IL-22, and IL-23 obtained using the multiplex READ platform. These calibration models were used for quantitative biomarker determination and subsequent spike-and-recovery validation in saliva.

Supplementary Table S1. Spike-and-recovery performance of the multiplex READ platform in human saliva.

| CRP |  |  |  |
| --- | --- | --- | --- |
| Spiked Concentration (µg/mL) | Recovered Concentration (µg/mL) | % Recovery | % Accuracy |
| 4.6875 | 4.05 | 86.38 | 115.76 |
| 9.375 | 8.42 | 89.82 | 111.34 |
| 18.75 | 19.05 | 101.61 | 98.41 |
| 37.5 | 37.65 | 100.39 | 99.61 |
| Calprotectin |  |  |  |
| Spiked Concentration (µg/mL) | Recovered Concentration (µg/mL) | % Recovery | % Accuracy |
| 1 | 0.91 | 91.44 | 109.37 |
| 3 | 2.73 | 90.99 | 109.90 |
| 9 | 8.43 | 93.62 | 106.82 |
| 18 | 15.62 | 86.80 | 115.20 |
| 81 | 74.49 | 91.96 | 108.75 |
| TNF-alpha |  |  |  |
| Spiked Concentration (pg/mL) | Recovered Concentration (pg/mL) | % Recovery | % Accuracy |
| 5 | 5.43 | 108.59 | 92.09 |
| 15 | 15.31 | 102.06 | 97.99 |
| 30 | 27.81 | 92.72 | 107.86 |
| 60 | 51.74 | 86.24 | 115.96 |
| 90 | 81.82 | 90.91 | 110.00 |
| Lactoferrin |  |  |  |
| Spiked Concentration (µg/mL) | Recovered Concentration (µg/mL) | % Recovery | % Accuracy |
| 0.9375 | 0.81 | 86.51 | 115.60 |
| 1.875 | 1.66 | 88.61 | 112.86 |
| 3.75 | 3.13 | 83.36 | 119.96 |
| 7.5 | 6.45 | 85.99 | 116.30 |
| IL-6 |  |  |  |
| Spiked Concentration (pg/mL) | Recovered Concentration (pg/mL) | % Recovery | % Accuracy |
| 5 | 5.14 | 102.77 | 97.30 |
| 15 | 13.30 | 88.66 | 112.79 |
| 30 | 26.87 | 89.57 | 111.64 |
| 60 | 52.75 | 87.92 | 113.74 |
| 90 | 102.51 | 113.90 | 87.80 |
| IL-12 |  |  |  |
| Spiked Concentration (pg/mL) | Recovered Concentration (pg/mL) | % Recovery | % Accuracy |
| 0.5 | 0.51 | 101.81 | 98.22 |
| 1.5 | 1.61 | 107.60 | 92.94 |
| 2.5 | 2.30 | 92.10 | 108.58 |
| 3.5 | 3.72 | 106.35 | 94.03 |
| IL-22 |  |  |  |
| Spiked Concentration (pg/mL) | Recovered Concentration (pg/mL) | % Recovery | % Accuracy |
| 5 | 5.41 | 108.15 | 92.46 |
| 15 | 16.07 | 107.14 | 93.33 |
| 30 | 28.61 | 95.36 | 104.86 |
| 60 | 55.67 | 92.78 | 107.78 |
| IL-23 |  |  |  |
| Spiked Concentration (pg/mL) | Recovered Concentration (pg/mL) | % Recovery | % Accuracy |
| 5 | 4.79 | 95.71 | 104.48 |
| 15 | 14.14 | 94.28 | 106.07 |
| 30 | 26.31 | 87.70 | 114.03 |
| 60 | 61.59 | 102.65 | 97.42 |

Supplementary Table S2. Intra-assay and inter-assay coefficient of variation (CV) values for the multiplex READ platform.

| CRP |  |  |
| --- | --- | --- |
| Concentration (µg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0.0 | 12.6 | 14.2 |
| 3.1 | 19.1 | 16.6 |
| 6.3 | 19.0 | 15.0 |
| 12.5 | 10.4 | 15.6 |
| 25.0 | 10.0 | 17.0 |
| 50.0 | 10.7 | 15.6 |
| Calprotectin |  |  |
| Concentration (µg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0.5 | 3.0 | 7.0 |
| 1.5 | 14.0 | 13.0 |
| 4.5 | 13.0 | 16.0 |
| 13.5 | 12.0 | 13.0 |
| 40.5 | 10.0 | 10.0 |
| 121.5 | 9.0 | 8.0 |
| TNF-alpha |  |  |
| Concentration (pg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0 | 13.7 | 17.2 |
| 1 | 16.9 | 19.8 |
| 10 | 17.6 | 18.8 |
| 20 | 17.8 | 18.8 |
| 40 | 17.2 | 19.0 |
| 80 | 17.0 | 19.1 |
| 120 | 17.0 | 19.3 |
| Lactoferrin |  |  |
| Concentration (µg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0 | 13.9 | 16.2 |
| 0.625 | 11.9 | 15.6 |
| 1.25 | 11.6 | 15.3 |
| 2.5 | 11.6 | 15.1 |
| 5 | 11.6 | 14.9 |
| 10 | 11.4 | 14.7 |
| IL-6 |  |  |
| Concentration (pg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0 | 19.1 | 17.4 |
| 1 | 18.7 | 15.7 |
| 10 | 18.3 | 16.2 |
| 20 | 18.3 | 17.7 |
| 40 | 17.9 | 19.0 |
| 80 | 18.6 | 18.2 |
| 120 | 19.1 | 18.9 |
| IL-12 |  |  |
| Concentration (pg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0.0 | 10.0 | 17.0 |
| 0.2 | 12.1 | 19.7 |
| 1.2 | 12.6 | 19.7 |
| 2.1 | 12.4 | 17.6 |
| 3.1 | 13.1 | 18.4 |
| 4.0 | 14.3 | 18.4 |
| IL-22 |  |  |
| Concentration (pg/mL) | Intra-Assay CV (%) | Inter-Assay CV (%) |
| 0 | 19.8 | 18.4 |
| 1 | 15.6 | 19.8 |
| 10 | 19.1 | 18.5 |
| 20 | 17.2 | 19.6 |
| 40 | 15.4 | 18.3 |
| 80 | 12.2 | 18 |
| IL-23 |  |  |
| Concentration (pg/mL) | Inter-Assay CV (%) | Intra-Assay CV (%) |
| 0 | 18.1 | 14.2 |
| 1 | 16.0 | 15.4 |
| 10 | 13.7 | 15.7 |
| 20 | 16.7 | 18.8 |
| 40 | 17.8 | 15.3 |
| 80 | 7.0 | 19.3 |

Supplementary Table S3. Characteristics of the Preliminary Clinical Cohort

| Characteristic | Non-IBD Subjects | IBD Subjects |
| --- | --- | --- |
| Participants (N) | 9 | 2 |
| Longitudinal Saliva Samples/Participant | ~28 | ~28 |
| Approximate Total Samples | ~252 | ~56 |
| Disease Type | No IBD | Ulcerative colitis |
| Clinical Disease Severity | Not collected | Not collected |
| Female, N (%) | 5 (55.6%) | 2 (100%) |
| Male, N (%) | 4 (44.4%) | 0 |
| Median Age, Years (Range) | 28 (18-51) | 38.5 (19-58) |