Supplementary table:

Operational conditions and standard (Applichem®) preparations (as per company guidelines) employed in Atomic Absorption Spectrophotometer for determination of heavy metals AOAC (1990)

| Heavy metals | Wavelength (nm) | Slit Width (nm) | Lamp Current (mA) | Burner Head | Flame | Burner Height (mm) | Oxidant gas pressure (kpa)/ Flow rate (l/min) | Fuel gas pressure (kpa)/ Flow rate (l/min) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Cadmium | 228.8 | 1.3 | 7.5 | Standard type | Air-C2H2 | 5.0 | 160/15.0 | 6.0/2.0 |
| Cobalt | 240.7 | 0.2 | 12.5 | Standard type | Air-C2H2 | 7.5 | 160/15.0 | 7.0/2.2 |
| Copper | 324.8 | 1.3 | 7.5 | Standard type | Air-C2H2 | 7.5 | 160/15.0 | 7.0/2.2 |
| Chromium | 357.9 | 1.3 | 7.5 | Standard type | Air-C2H2 | 7.5 | 160/15.0 | 12/3.0 |
| Nickel | 232.0 | 0.2 | 10.0 | Standard type | Air-C2H2 | 7.5 | 160/15.0 | 7.0/2.2 |
| Zinc | 213.9 | 1.3 | 10.0 | Standard type | Air-C2H2 | 7.5 | 160/15.0 | 6.0/2.0 |

Standards Preparation

Calibrated standards were prepared from the commercially available stock solution (Applichem®) in the form of an aqueous solution (1000 ppm). Highly purified de-ionized water was used for the preparation of working standards. All the glass apparatus used throughout the process of analytical work were immersed in 8N HNO3 overnight and washed with several changes of de-ionized water prior to use.