Supplementary Table 1. Primers for HBV RNA amplification and detection

| Round | Sense | Primer name | Sequence (5'-3') | Position |
| --- | --- | --- | --- | --- |
| RT | - | HBV-3.5RNA-RT | ATTCTCAGACCGTAGCACACGACACCGAGATTGAGATCTTCTGCGAC | nt 2436-2415 |
| PCR | + | 3.5pgRNA_PCR-F | CCTACTGTTCAAGCCTCCAAGC | nt 1856-1877 |
|  | - | HBV RNA_PCR-R | ATTCTCAGACCGTAGCACACGACAC | random sequence |
| QPCR | + | 3.5RNA_QPCR-F | AYAGACCATCAAATGCCC | nt 2295-2312 |
|  | - | HBV RNA_PCR-R | ATTCTCAGACCGTAGCACACGACAC | random sequence |

Supplementary Table 2. The limit of detection for serum HBV RNA was 1466 copies/mL, as calculated by probit analysis

| Nominal input (HBV RNA copies/mL) | No. replicates | No. positives | Positivity rate (%) |
| --- | --- | --- | --- |
| 19200 | 24 | 24 | 100 |
| 1920 | 60 | 58 | 96.67 |
| 960 | 80 | 77 | 96.25 |
| 660 | 80 | 17 | 21.25 |
| 480 | 80 | 7 | 8.75 |
| 240 | 60 | 4 | 6.67 |
| The concentration of HBV RNA that can be detected with a positivity rate of greater than 95% as determined by probit analysis is 1466 copies/mL |  |  |  |