# Water- and energy-efficient dialysate flow optimisation in haemodialysis using a data-driven porous media model

Ruhit Sinha*, Anne E. Staples
Department of Mechanical Engineering, Virginia Tech

*correspondence: ruhitsinha@vt.edu

## Supplementary Notes

### Ergun equation

Ergun (1952)1 introduced a semi-empirical relation for pressure loss gradient based on fluid flow data in a packed column, which was later modified by Mehta and Hawley (1969)2 to include a wall-effect parameter, given by:

|  | ∆PDLeffzz=150M2μDp21-ϵ2ϵ3uDsup+1.75MρDp1-ϵϵ3uDsup2 | (S1) |
| --- | --- | --- |

where ΔPD is the absolute pressure difference between two points along the axis of the packed column, Leff is the effective distance between the two points considered, ϵ is the void fraction or bundle porosity, μ is the viscosity of the fluid, ρ is the density of the fluid, Dp is the characteristic diameter of the porous media, M is the wall effect parameter (Equation S13), uDsup is the superficial velocity of the fluid flowing through the packed column (Equation 13).

The Ergun equation (Equation S1) is also represented in its dimensionless form, given by

|  | f'=αRe'+γ | (S2) |
| --- | --- | --- |

where α is called the Blake-Kozeny constant and typically takes the value 150 and γ is called the Burke-Plummer constant and typically takes the value of 1.75. Another set of values for these constant widely used are α = 180 and γ = 1.8-4 depending particle roughness, as suggested by Macdonald et al. (1979)3. f’ (Equation 9) and Re’ (Equation 11) are called modified friction factor and modified Reynold number respectively.

### Darcy permeability (κ) and Forchheimer coefficient (CF)

#### Module-scale 3D Navier-Stokes equations solution

After module-scale 3D Navier-Stokes equations (3D N-S) solution is obtained (Methods), we extracted the pressure drop data for the range of dialysate flow rate (10-4 to 103 mLmin-1) for all the dialysers listed in Supplementary Table S1. It is well established that pressure drop versus flow rate follows a quadratic behaviour4 given by:

|  | ∆PD=auDsup+buDsup2 | (S3) |
| --- | --- | --- |

We can obtain the values ‘a’ and ‘b’ by regression. Now, according to the Darcy-Forchheimer equation, pressure loss gradient (say along the axis) is given by

|  | ∆PDLeffzz=μκzzuDsup+CFρκzzuDsup2 | (S4) |
| --- | --- | --- |

where κzz is the axial Darcy permeability (also called viscous permeability; unit: m2) and CF is the Forchheimer coefficient (dimensionless). The term √κzz/CF is also called inertial permeability (unit: m), Leff = L-2X = 0.8L

Comparing equations S3 and S4, we get:

|  | κzz=μLeffa | (S5) |
| --- | --- | --- |

|  | CF=bκzzρLeff | (S6) |
| --- | --- | --- |

Notably, it is difficult to directly measure the radial Darcy permeability (κrr) because the dialysate inlet and outlet aren’t collinear and hence difficult to calculate a pressure gradient along the radius of the dialyser. We, therefore, find κrr by a parametric study of κzz and κrr until the pressure profile generated at a very low Reynolds number (say at QDi = 10-4 mLmin-1) matches with the one obtained from 3D N-S equations. All the porous media parameters (κzz, κrr, and CF) calculated for 20 dialyzers can be found in Supplementary Table S2.

#### Ergun equation

According to packed-column theory, κzz can be approximated using Ergun equation (Equation S1). Comparing equation S1 with equation S4, we get:

|  | κzz=Dp2150M2ϵ31-ϵ2 | (S7) |
| --- | --- | --- |

|  | CF=κzz1.75MDp1-ϵϵ3 | (S8) |
| --- | --- | --- |

Note that this theory yields κrr to be same as κzz. Dp takes the value of the diameter of spherical particle for a packed bed. For its implementation in hollow fibre membrane module and shell-side flow to be modelled as a packed column, Dp varies across studies. For example: Suzuki et al. (2001)5 and Papathanasious et al. (2001)6 considered Dp = do, Yoshikawa et al. (1994)7 and Yamamoto et al. (2009)8 considered Dp = dh while Kyan et al. (1970)3 considered Dp = 1.5 do.

#### Single-fibre models

Single-fibre models are idealised representations of flow in hollow-fibre membrane modules in which the complex fibre bundle is reduced to a periodic unit cell containing a single fibre. The surrounding fluid domain is defined based on an assumed packing geometry (e.g., triangular or square array) with a certain packing density 	φ (Supplementary Figure S1) in such a way that the following relation holds:

φ=A1A1+A2=N×12A1Amodule (for triangular array)N×8A1Amodule        (for square array)

|  | ⇒A1+A2=Amodule12N=πdo2412φAmodule8N=πdo248φ | (S9) |
| --- | --- | --- |

The flow profile is then approximated within this simplified domain by the Navier-Stokes equations under the assumption of viscous flow (Stokes flow). As a result, axial Darcy permeability can be given by

|  | κzz=μ∆PDLeffzzuDsup | (S10) |
| --- | --- | --- |

|  | uDsup=QDiAmodule=QdiA1+A2 | (S11) |
| --- | --- | --- |

where QDi is the dialysate flow rate through the entire dialyser module cross-section of area Amodule, Qdi is the dialysate flow rate through the elemental cross-section of area A1 + A2. Decades ago, Sparrow and Loeffler (1959)9 worked on the analytical solution for the longitudinal fully developed laminar flow between cylinders arranged in triangular and square array. Their solution can be interpreted as the pressure drop along the axis versus dialysate flow rate (viscous regime) in the single-fiber model. They expressed the quantity 16Qdiμ ∆PDLeffzzdo2 in terms of pitch-to-outer fiber diameter (do) ratio. We further derived the expression for Darcy permeability for axial flow (κzz) for both triangular and square array as shown in Supplementary Table S3. In the same year, Happel (1959)10 introduced free-surface annulus model for flow around cylinders. Their solution can be interpreted as another single-fiber model based on which Darcy permeability for axial flow (κzz) and radial flow (κrr) has been derived by the author. Radial flow for triangular and square array is a bit complicated. Nonetheless, we derived their κzz and κrr from the findings of Drummond and Tahir (1984)9 for triangular array, and a combination of Keller(1964)11, Sangani and Acrivos (1982)12 and Sangani and Acrivos (1982)13 for square array in the porosity range 0.2146 to 0.4, 0.4 to 0.95 and 0.95 to 1 respectively, as shown in Supplementary Table S3.

### Hydraulic diameter and the Kozeny constant

When the flow through porous structures such as the shell region of a hollow-fibre dialyser is described using Darcy’s law, tube-bundle or capillary models are commonly employed. In this framework, the porous medium is approximated as a bundle of parallel capillaries with a common hydraulic diameter, dh (Equation S12). For a packed column, Mehta and Hawley (1969)2 pointed out that the hydraulic diameter not only depends on the particle diameter but also on the column diameter. They introduced a wall effect parameter M and suggested a modification to the classical Ergun equation to include the wall effect. Following their methodology, we derived the hydraulic diameter dh for a dialyser module of diameter dmodule packed with cylindrical fibres of outer diameter do and length L.

dh=4volume of voidswetted surface=4volume of voidsvolume of modulewetted surface of fibersvolume of module+wetted surface of wallvolume of module

|  | =4ϵ4NπdoLπdmodule2L+4πdmoduleLπdmodule2L=ϵdoφ+dodmodule=ϵdoφM | (S12) |
| --- | --- | --- |

|  | where M= 1+doφdmodule=1+1φN | (S13) |
| --- | --- | --- |

It is noteworthy that M values of all the modules studied here are almost equal to 1 (Supplementary Table S1). The wall effect only starts to be significant when M > 1. Nevertheless, it is always good practice to include M in the subsequent calculations for accuracy and generalisation.

Consequently, it is important to include the wall in the calculation of specific surface s.

|  | s =wetted surfacevolume of fibers=NπdoL+πdmoduleLφ14πdmodule2L=4Mdo=4ϵφdh | (S14) |
| --- | --- | --- |

Combining Darcy permeability with specific surface, porosity, and packing density yields the dimensionless Kozeny constant k (Equation S15), which is often assumed to have a nearly universal value for random porous media, independent of particle size and porosity14. However, capillary theory becomes unreliable at very low (<0.4) or very high (>0.8) porosities6,14,15. In the ideal case of laminar flow in a circular pipe, k = 2 (page 403 in Happel and Brenner (1963)14) , whereas experiments for packed beds typically report k ≈ 514.

|  | k= ϵ3φ21s21κ=ϵ3φ2do216M21κ=ϵdh216κ | (S15) |
| --- | --- | --- |

### 4. Dimensionless pressure-drop metrics

To characterise the fluid flow in a packed column, Ergun (1952)1 defined a modified friction factor fErgun given by

|  | fErgun=∆PDLeffzzdoρuDsup2ϵ3φM | (S16) |
| --- | --- | --- |

and a modified Reynold number

|  | ReErgun=ρuDsupdoμϕM | (S17) |
| --- | --- | --- |

Note that both the terms written above are further modified version to include the wall effect parameter (M) as per Metha and Hawley (1969)2. Also, fErgun = f’ (when Dp = do, Equation 9) and ReErgun = Re’ (when Dp = do, Equation 11)

Hence,

|  | fErgunReErgun=∆PDLeffzzdoρuDsup2ϵ3φM×ρusupdoμϕM=16k=α (ReErgun≪1) | (S18) |
| --- | --- | --- |

This term is also called Blake-Kozeny constant and takes the value of 150 when k ≈ 9.4 for fibrous media6.  However, this term remains almost constant, just like k, when porosity is varied. We would, therefore, consider the following metrics to characterize the fluid flow.

#### i) Fanning friction factor times Reynolds number (fFanningRe)

Following Bird et al. (2002)16’s and Erdim et al. (2015)17’s presentation of defining Fanning friction factor, we find that shear force Fk exerted by the dialysate flow on the fibres and the wall of the module can be expressed as

Fk= fFanningAextK

where,

Aext = total external surface area = specific surface x volume of fibers = sφAmoduleL

Amodule = cross-section area of the module = πdmodule2/4

K = characteristic kinetic energy per unit volume = 12ρum2

um = mean dialysate velocity

A force balance on the dialysate gives:

Fk=∆PϵAmodule

Equating Fk, we get:

∆Pϵ=fFannings φL12ρum2

⇒fFanning=2∆PDLeffzz1ρum2ϵsφ=2∆PDLeffzz1ρum2ϵdo4Mφ=fErgun2

If we define, Reynolds number  Re=ρumdoμM, then

|  | fFanningRe=8φϵk | (S19) |
| --- | --- | --- |

which is basically the already known formula(Equation 8-4.48 in Happel and Brenner (1963)14). This same quantity was also calculated by Sparrow and Loeffler (1959)9, of course for triangular and square arrays. In the ideal case, when the flow region is a circular pipe,  fFanningRe = 16.

#### ii) Darcy friction factor times superficial Reynolds number (fDarcyResup)

Following Cancilla et al. (2021)18’s presentation of defining Darcy friction factor and superficial Reynolds number, we have

fDarcy=2∆PDLeffzzρusup2dh=2∆PDLeffzzρusup2ϵdoφM=4fFanningϵ2
Resup=ρusupdhμ=ρumϵμϵdoφM=Reϵ2φM

|  | ⇒fDarcyResup=4fFanningReφM=32kϵM | (S20) |
| --- | --- | --- |

In the ideal case, when the flow region is a circular pipe, fDarcyResup= 64.

#### iii) Dimensionless drag (F/μusup)

Following previous studies10,12,19, we define drag per unit length of each fibre, F as

F=∆PDLeffzz1NAmodule

|  | ⇒Fμusup=∆PDLeffzzAmoduleμusupN= AmoduleκzzN=16kMdo2ϕ2ϵ3πdmodule24N=4πkϕϵ3M | (S21) |
| --- | --- | --- |

### 5. Experimental conditions for measuring pressure-drop in dialyser modules

#### PHYLTER® HF 17SD and 15SD

For this two dialysers, dialysate-side axial pressure drop is defined as the pressure difference between the tappings separated by 15 cm, which is then extrapolated to the nominal distance between inlet and outlet ports of 23 cm, as reported in the supplementary document of Cancilla et al. (2022)20 or in the PhD dissertation of Cancilla (2022)21. Note that this pressure drop is for the axial flow and doesn’t include the pressure drops associated with the header. Unfortunately, the authors did not report the temperature (which affects the viscosity) at which the hydraulic resistance experiment was conducted. It was through personal communication that we learned that the experiment was conducted in room temperature ranging from 18-28° C and the viscosity value considered was corresponding to 25° C. An error of ± 6% was added to the experimental values due to the experimental uncertainties ascribed as explained20,21.

#### UNIPA test modules

For these modules, dialysate-side axial pressure drop is defined as the pressure difference between the tappings separated by 14.1 cm as reported in the in the PhD dissertation of Cancilla (2022)21. Note that the pressure difference wasn’t extrapolated for these modules. We assumed that the hydraulic experiment was performed at 25° C.

#### CT190G

For this dialyser, we couldn’t find a proper definition of the dialysate-side pressure drop and the temperature at which the experiment was performed in the PhD dissertation of Liao (2002)22. Thus, we assumed it as the pressure difference between dialysate inlet and outlet at 37 °C.

#### F6HPS

Eloot et al. (2005)23 performed a dialysate flow visualization experiment at 37° C where they were more interested on determining radial velocity distribution rather than the pressure drop versus flow rate of dialysate. Hence, we don’t have experimental pressure drop data for this dialyzer. Nonetheless, they did report the axial and radial Darcy permeability (κzz = 6.25 x 10-10 m2, κrr = 2.88 x 10-11 m2).

#### Dialyser A and B

These two dialysers were studied by Suzuki et al. (2001)5 where they reported the experimental pressure drop over a range of dialysate flow rate at 37° C  but no Darcy permeability value. Also, the definition of dialysate-side pressure drop is not clearly defined. Thus, we assumed it to be the dialysate-side axial pressure drop.

## Supplementary Tables

### Table S1

Table S1 | Geometry specifications of the dialyser modules, from published literature, computationally studied in this work.

|  |  |  | Module |  | din [μm] | tM [μm] | do [μm] | N [103] | L [cm] | A [m2] | dmodule [mm] | ϵ | dh [μm] | M |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Training cases |  | 1 | Fresenius | F6024–27 | 200 | 40 | 280 | 9 | 23 | 1.30 | 40 | 0.559 | 349 | 1.016 |
|  |  | 2 | Gambro | Polyflux 210H28–30 | 215 | 50 | 315 | 11.52 | 27 | 2.1 | 50 | 0.543 | 369 | 1.014 |
|  |  | 3 | Gambro | M6031 | 240 | 50 | 340 | 4.4 | 18.1 | 0.60 | 33 | 0.533 | 380 | 1.022 |
|  |  | 4 | Nipro | FB-130U32 | 200 | 15 | 230 | 9.71 | 21.3 | 1.3 | 32 | 0.498 | 225 | 1.014 |
|  |  | 5 | Gambro | GFE-1533 | 230 | 9 | 248 | 8.13 | 23.8 | 1.4 | 31.5 | 0.496 | 240 | 1.016 |
|  |  | 6 | Baxter | Syntra 16034 | 200 | 30 | 260 | 11.58 | 22 | 1.60 | 39 | 0.485 | 242 | 1.013 |
|  |  | 7 | Hospal | Filtral 1224,25,34–36 | 220 | 45 | 310 | 8.32 | 20 | 1.15 | 39 | 0.474 | 263 | 1.015 |
|  |  | 8 | Fresenius | Optiflux 160 NR37 | 200 | 40 | 280 | 10.85 | 22 | 1.50 | 39 | 0.441 | 218 | 1.013 |
|  |  | 9 | Asahi-Kasei | APS-08SA8 | 185 | 45 | 275 | 5.195 | 26.5 | 0.80 | 26.5 | 0.441 | 212 | 1.019 |
|  |  | 10 | Asahi-Kasei | APS-15SA8 | 185 | 45 | 275 | 9.71 | 26.5 | 1.50 | 36.4 | 0.446 | 218 | 1.014 |
|  |  | 11 | Asahi-Kasei | APS-25SA8 | 185 | 45 | 275 | 16.2 | 26.5 | 2.50 | 45.8 | 0.416 | 194 | 1.010 |
|  | Validation cases | 1 | Medtronic | PHYLTER HF 15 SD20 | 200 | 30 | 260 | 9.78 | 24.4 | 1.50 | 37 | 0.517 | 274 | 1.015 |
|  |  | 2 |  | PHYLTER HF 17 SD20 | 200 | 30 | 260 | 11.09 | 24.4 | 1.70 | 40 | 0.531 | 291 | 1.014 |
|  |  | 3 | Special test modules from UNIPA lab21 |  | 200 | 30 | 260 | 6.255 | 24.4 | 0.96 | 27 | 0.42 | 185 | 1.017 |
|  |  | 4 |  |  |  |  |  | 5.176 |  | 0.79 |  | 0.52 | 276 | 1.020 |
|  |  | 5 |  |  |  |  |  | 4.310 |  | 0.66 |  | 0.6 | 381 | 1.024 |
|  |  | 6 |  |  |  |  |  | 3.240 |  | 0.50 |  | 0.7 | 587 | 1.032 |
|  |  | 7 |  |  |  |  |  | 2.160 |  | 0.33 |  | 0.8 | 996 | 1.048 |
|  |  | 8 | Baxter | CT190G22 | 200 | 15 | 230 | 12.1 | 25 | 1.90 | 35 | 0.477 | 208 | 1.013 |
|  |  | 9 | Fresenius | F6HPS23 | 200 | 40 | 280 | 9.2 | 22.5 | 1.30 | 40 | 0.549 | 336 | 1.016 |
| Test cases |  | 10 | - | Dialyzer A5 | 188 | 38.9 | 265.8 | 8.84 | 24 | 1.5 | 36 | 0.518 | 281 | 1.015 |
|  |  | 11 | - | Dialyzer B5 | 214 | 35.3 | 284.6 | 7.26 | 26.5 | 1.5 | 36 | 0.546 | 336 | 1.017 |

Note: din = inner fibre diameter, tM = thickness of hollow fibre membrane, do = outer fibre diameter, N = total number of fibres in the dialyser module, L = length of a fibre or the module, 
A = total internal surface area of membranes, dmodule = diameter of the module, ϵ = porosity of the fibre bundle, dh = hydraulic diameter of the dialysate (shell) domain, M = wall-effect parameter.

### Table S2

Table S2 | The porous media parameters of twenty modules computed from the module-scale three-dimensional Navier-Stokes solution found in this study.

| S.No |  | Module |  | ϵ | κzz [10-10 m2] | κrr [10-10 m2] | κrr/κzz | CF |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Commercial dialysers | Fresenius | F60 | 0.559 | 4.9249 | 2.3428 | 0.476 | 0.0024 |
| 2 |  | Fresenius | F6HPS | 0.549 | 5.6394 | 2.9150 | 0.517 | 0.0073 |
| 3 |  | Gambro | Polyflux 210H | 0.543 | 5.5627 | 2.6413 | 0.475 | 0.0019 |
| 4 |  | Gambro | M60 | 0.533 | 6.2982 | 3.1100 | 0.494 | 0.0070 |
| 5 |  | Medtronic | PHYLTER HF 17SD | 0.531 | 3.5569 | 1.7511 | 0.492 | 0.0029 |
| 6 |  | Medtronic | PHYLTER HF 15SD | 0.517 | 3.1732 | 1.5770 | 0.497 | 0.0052 |
| 7 |  | Nipro | FB-130U | 0.498 | 2.0783 | 0.9060 | 0.436 | 0.0082 |
| 8 |  | Gambro | GFE-15 | 0.496 | 2.2142 | 0.9555 | 0.432 | 0.0064 |
| 9 |  | Baxter | Syntra 160 | 0.485 | 2.4403 | 1.1846 | 0.485 | 0.0132 |
| 10 |  | Baxter | CT190G | 0.477 | 1.6079 | 0.7590 | 0.472 | 0.0056 |
| 11 |  | Hospal | Filtral 12 | 0.476 | 2.9778 | 1.3751 | 0.462 | 0.0181 |
| 12 |  | Asahi-Kasei | APS-15SA | 0.446 | 1.7998 | 0.7965 | 0.443 | 0.0121 |
| 13 |  | Fresenius | Optiflux 160 NR | 0.441 | 1.6747 | 0.6273 | 0.375 | 0.0134 |
| 14 |  | Asahi-Kasei | APS-08SA | 0.441 | 1.7884 | 0.6227 | 0.348 | 0.0168 |
| 15 |  | Asahi-Kasei | APS-25SA | 0.416 | 1.3244 | 0.4857 | 0.367 | 0.0146 |
| 16 | UNIPA test modules |  | 80% porosity | 0.800 | 60.4750 | 30.8250 | 0.510 | 0.0002 |
| 17 |  |  | 70% porosity | 0.700 | 19.830 | 10.3785 | 0.523 | 0.0009 |
| 18 |  |  | 60% porosity | 0.600 | 7.2087 | 3.4921 | 0.484 | 0.0020 |
| 19 |  |  | 52% porosity | 0.520 | 3.1291 | 1.5485 | 0.495 | 0.0046 |
| 20 |  |  | 42% porosity | 0.420 | 1.3439 | 0.4911 | 0.365 | 0.0161 |

Note: ϵ = porosity of the fibre bundle, κzz = axial Darcy permeability, calculated using Supplementary Equation S5, κrr = radial Darcy permeability, calculated using a parametric study, CF = Forchheimer coefficient, calculated using Supplementary Equation S6, UNIPA = University of Palermo

### Table S3

Table S3 | The axial Darcy permeability calculated for 20 modules using module-scale three-dimensional Navier-Stokes equation (3D N-S), both versions of modified Ergun equation developed in this study, and experimental values reported in published literature.

|  |  |  |  |  | Axial Darcy permeability, κzz [10-10 m2] |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| S.No |  | Module |  | ϵ | 3D N-S | Modified Ergun (Dp = do) | Modified Ergun (Dp = dh) | Experimental |
| 1 | Commercial dialysers | Fresenius | F60 | 0.559 | 4.9249 | 5.5723 | 5.3418 | - |
| 2 |  | Fresenius | F6HPS23 | 0.549 | 5.6394 | 5.0506 | 4.8816 | 6.25 |
| 3 |  | Gambro | Polyflux 210H | 0.543 | 5.5627 | 6.0442 | 5.8894 | - |
| 4 |  | Gambro | M60 | 0.533 | 6.2982 | 6.2747 | 6.0637 | - |
| 5 |  | Medtronic | PHYLTER HF 17SD20 | 0.531 | 3.5569 | 3.6557 | 3.5961 | 3.40 |
| 6 |  | Medtronic | PHYLTER HF 15SD20 | 0.517 | 3.1732 | 3.1770 | 3.1573 | 4.50 |
| 7 |  | Nipro | FB-130U | 0.498 | 2.0783 | 2.0579 | 2.0792 | - |
| 8 |  | Gambro | GFE-15 | 0.496 | 2.2142 | 2.3392 | 2.3613 | - |
| 9 |  | Baxter | Syntra 160 | 0.485 | 2.4403 | 2.3143 | 2.3705 | - |
| 10 |  | Baxter | CT190G22 | 0.477 | 1.6079 | 1.6719 | 1.7252 | 12.90 |
| 11 |  | Hospal | Filtral 12 | 0.476 | 2.9778 | 2.9912 | 3.0733 | - |
| 12 |  | Asahi-Kasei | APS-15SA | 0.446 | 1.7998 | 1.7375 | 1.8339 | - |
| 13 |  | Fresenius | Optiflux 160 NR | 0.441 | 1.6747 | 1.7130 | 1.8176 | - |
| 14 |  | Asahi-Kasei | APS-08SA | 0.441 | 1.7884 | 1.6339 | 1.7142 | - |
| 15 |  | Asahi-Kasei | APS-25SA | 0.416 | 1.3244 | 1.2773 | 1.3854 | - |
| 16 | UNIPA test modules |  | 80% porosity21 | 0.800 | 60.4750 | 64.3210 | 68.3961 | 67.22 |
| 17 |  |  | 70% porosity21 | 0.700 | 19.830 | 19.7514 | 17.8540 | 29.50 |
| 18 |  |  | 60% porosity21 | 0.600 | 7.2087 | 7.1066 | 6.5318 | 11.51 |
| 19 |  |  | 52% porosity21 | 0.520 | 3.1291 | 3.2379 | 3.1752 | 5.11 |
| 20 |  |  | 42% porosity21 | 0.420 | 1.3439 | 1.1765 | 1.2571 | 1.69 |

### Table S4

Table S4 | Formulas for Darcy permeabilities based on different single-fibre models used in this study.

| Single-fibre models | Darcy permeability (κ) |  |
| --- | --- | --- |
|  | Axial flow | Radial flow |
| Triangular array9,19 | 4Qdiμ -∆PDLzzdo212ϕπ | do232ϕ-lnϕ-1.4975+2ϕ-0.5ϕ2-0.7391ϕ4+2.5342ϕ51+1.2758ϕ |
| Square array9,11–13 | 4Qdiμ -∆PDLzzdo28ϕπ | do292ϕ1-4ϕπ5/2 (low ϵ)do232ϕ-lnϕ-1.476+2ϕ-1.774ϕ2+4.076ϕ3 (high ϵ) |
| Free-surface annulus10 | do216ϕ-lnϕ-32+2ϕ-ϕ22 | do232ϕ-lnϕ+ϕ2-1ϕ2+1 |

Note: The expansion of the term 16Qdiμ -∆PLzzdo2 for triangular and square array can be found on the Appendix of Sparrow and Loeffler (1959)9.

### Table S5

Table S5 | Mesh parameters considered for solving the dialysate flow field by the module-scale three-dimensional Navier Stokes equations and the continuity equation using the finite-element method in COMSOL 6.1.

|  |  |  | Module |  | Porosity ϵ | Mesh resolution | Maximum element size | Minimum element size | Mesh volume (mm3) | No. of elements |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Training cases | Validation cases | 1 | Medtronic | PHYLTER HF 15 SD | 0.517 | Finer+ | 0.552 | 0.0506 | 9,386 | 4,499,040 |
|  |  | 2 |  | PHYLTER HF 17 SD | 0.531 | Finer+ | 0.5815 | 0.0533 | 11,400 | 5,591,735 |
|  |  | 3 | Special test modules from UNIPA lab |  | 0.42 | Extra fine | 0.343 | 0.0223 | 3553 | 3,534,090 |
|  |  | 4 |  |  | 0.52 | Finer+ | 0.447 | 0.04095 | 5004 | 2,421,615 |
|  |  | 5 |  |  | 0.6 | Finer+ | 0.447 | 0.04095 | 6128 | 3,316,110 |
|  |  | 6 |  |  | 0.7 | Finer | 0.551 | 0.0596 | 7488 | 2,092,860 |
|  |  | 7 |  |  | 0.8 | Finer | 0.551 | 0.0596 | 8851 | 2,721,075 |
|  |  | 8 | Baxter | CT190G | 0.477 | Finer+ | 0.536 | 0.04915 | 7,658 | 5,373,060 |
|  |  | 9 | Fresenius | F6HPS | 0.549 | Finer+ | 0.566 | 0.05185 | 11,070 | 5,201,490 |
|  |  | 1 | Fresenius | F60 | 0.559 | Finer | 0.703 | 0.076 | 11760 | 2,924,670 |
|  |  | 2 | Gambro | Polyflux 210H | 0.543 | Finer | 0.861 | 0.0931 | 20780 | 3,725,925 |
|  |  | 3 | Gambro | M60 | 0.533 | Finer+ | 0.4625 | 0.0424 | 5752 | 2,331,840 |
|  |  | 4 | Nipro | FB-130U | 0.498 | Extra fine | 0.367 | 0.0239 | 5701 | 8,663,382 |
|  |  | 5 | Gambro | GFE-15 | 0.496 | Extra fine | 0.377 | 0.0246 | 6101 | 7,567,382 |
|  |  | 6 | Baxter | Syntra 160 | 0.485 | Extra fine | 0.423 | 0.0276 | 8430 | 8,759,933 |
|  |  | 7 | Hospal | Filtral 12 | 0.474 | Extra fine | 0.41 | 0.0267 | 7358 | 6,023,115 |
|  |  | 8 | Fresenius | Optiflux 160 NR | 0.441 | Extra fine | 0.423 | 0.0276 | 7235 | 6,352,920 |
|  |  | 9 | Asahi-Kasei | APS-08SA | 0.441 | Extra fine | 0.348 | 0.0227 | 4001 | 3,000,270 |
|  |  | 10 | Asahi-Kasei | APS-15SA | 0.446 | Extra fine | 0.43 | 0.0281 | 7745 | 6,445,215 |
|  |  | 11 | Asahi-Kasei | APS-25SA | 0.416 | Extra fine | 0.502 | 0.0327 | 10830 | 9,203,786 |

### Table S6

Table S6 | List of countries/regions whose Dialysis Water Savings Potential score fell in the upper quartile of all countries considered.

|  | Country | Dialysis Water Savings Potential |
| --- | --- | --- |
| 1 | Kuwait | 21.4065 |
| 2 | United Arab Emirates | 6.7569 |
| 3 | Qatar | 5.0826 |
| 4 | Maldives | 1.7427 |
| 5 | Bahrain | 1.2842 |
| 6 | Saudi Arabia | 1.183 |
| 7 | Yemen | 1.057 |
| 8 | Singapore | 0.9326 |
| 9 | Libya | 0.9063 |
| 10 | Jordan | 0.8679 |
| 11 | Malta | 0.6102 |
| 12 | Palestinian Authority | 0.41 |
| 13 | Barbados | 0.3449 |
| 14 | Algeria | 0.3097 |
| 15 | Oman | 0.2979 |
| 16 | Israel | 0.2417 |
| 17 | Tunisia | 0.2212 |
| 18 | Djibouti | 0.2095 |
| 19 | Antigua and Barbuda | 0.1796 |
| 20 | Saint Kitts and Nevis | 0.1755 |
| 21 | Egypt | 0.1504 |
| 22 | Cape Verde | 0.1222 |
| 23 | South Africa | 0.1099 |
| 24 | Lebanon | 0.1045 |
| 25 | Cyprus | 0.0994 |
| 26 | Syria | 0.0967 |
| 27 | Saint Vincent and the Grenadines | 0.0967 |
| 28 | Morocco | 0.0961 |
| 29 | Kenya | 0.0842 |
| 30 | Burkina Faso | 0.0827 |
| 31 | Sudan | 0.0824 |
| 32 | Haiti | 0.0789 |
| 33 | Iran | 0.0776 |
| 34 | Nauru | 0.0771 |
| 35 | Grenada | 0.06 |
| 36 | Saint Lucia | 0.059 |
| 37 | Pakistan | 0.0586 |
| 38 | India | 0.0586 |
| 39 | The Bahamas | 0.056 |
| 40 | Nigeria | 0.0557 |
| 41 | Lesotho | 0.055 |
| 42 | Denmark | 0.055 |
| 43 | Uzbekistan | 0.0537 |
| 44 | Mauritius | 0.0516 |
| 45 | Malawi | 0.0513 |
| 46 | Zimbabwe | 0.0507 |

## Supplementary Figures

### Figure S1

Figure S1 | Schematic of widely used single-fibre models for flow in hollow-fibre membrane modules such as haemodialyser. a, triangular array, b, square array, and c, free-surface annulus. These models are characterised by an envelope of fluid (dialysate) around a single fibre in the shape of hexagon, square, and circle respectively of appropriate size such that the porosity is conserved with respect to the module. do is the outer fibre diameter, pitch is the distance between the centres of two adjacent fibres, dD is the Krogh cylinder diameter, L is the fibre length, and A1, A2 are area outside and inside the envelope, respectively. You can find more details of a,b in Sparrow and Loeffler(1959)9 and c, in Happel (1959)10.

### Figure S2

Figure S2 | Schematic of the computational domain for solving the dialysate flow field using the module-scale three-dimensional Navier Stokes equations and the continuity equation. We consider the following assumptions while modelling this: 1. Fibres are arranged in the most compact regular order i.e. triangular array. Accordingly, we generated a dialysate flow domain representing the shell-region of a dialyser. 2. Distribution tabs are perfect. i.e. the dialysate entering and exiting from the dialysate inlet and outlet ports respectively are distributed uniformly radially towards and away from the axis respectively. Accordingly, we created the inlet and outlet on the peripheral ends of the module. 3. The inlet and outlet were one-tenth of the module length L based on similar calculations of typical distributed tab lengths (also called raised collars) of dialysers38,i.e. X = 0.1L. 4. As a result of all these assumptions, we can assume 1/12th  cross-section with the total length L with side walls being slip wall as the working domain for computation. 5. Outer fibre membrane walls are assigned to no-slip wall boundary conditions on the basis that there is little effect of typical ultrafiltration on the hydrodynamics of dialysate flow20. 6. Module wall is assigned no-slip wall boundary condition. 7. Dialysate is considered an incompressible fluid with density and dynamic viscosity 998.2 kgm-3 and 1.0093 mPas respectively (like water at 20 °C). 8. Pressure at the outlet is assigned to zero while a given flow rate (divided by 12 because of axisymmetry) is assigned at the inlet.

### Figure S3

Figure S3 | Dimensionless parameters that incorporate the radial-to-axial Darcy permeability (κrr/κzz) of the shell region of the dialyser. a, Module aspect ratio defined as the ratio of length of module to the diameter of module (L/dmodule) is another factor incorporated in these parameters. b. Osuga’s parameter39 is one such dimensionless parameter whose value has been found to be between 1 and 10. c. Module geometry number40 is another such dimensionless parameter whose typical value has been found to be in the range 10-100.

### Figure S4

Figure S4 | Schematic of the mesh required to solve the dialysate flow field by the module-scale three-dimensional Navier Stokes equations and the continuity equation using the finite-element method in COMSOL 6.1. a, Representation of triangular meshes in the cross-section of the dialyser with suggested pre-defined mesh sizes. b, Representation of the mesh when “Extra-fine” mesh is selected for the cross-section and then swept along the axis for certain number of elements.

### Figure S5

Figure S5 | Mesh independence study of dialyser comparing the dialysate-side pressure drop (∆PD) as mesh size or parameter is altered. The dialyser module analysed were a, APS-25SA with the lowest porosity, and b, UNIPA 0.8 porosity with the highest porosity among the 20 dailyser tested in this study. We find that any mesh size smaller than “Finer” were within 5% error for both diayser modules. c, APS-25SA after selecting “Finer” mesh in the cross-section. We find that when the number of swept elements was more than 15, ∆PD increased by only approximately 0.01%. Hence, we chose the parameter to be 35 keeping in mind that increasing it further increases the computational cost. Same result was found to be for UNIPA 0.8 porosity module. “Fine+” mesh size is the mesh size calculated by averaging the mesh sizes in  “Fine’” and “Finer” as pre-defined by COMSOL 6.1. The complete list of mesh sizes considered for the module-scale three-dimensional Navier-Stokes simulations are tabulated in Table S5.

### Figure S6

Figure S6 | Schematic of the dialyser module showing the dialysate and blood inlets and outlets in the countercurrent flow configuration. Fundamental flow field quantities are QBi blood flow rate at inlet, QF net ultrafiltration flow rate, QBo blood flow rate at outlet, QDi dialysate flow rate at inlet, QDo dialysate flow rate at outlet, PBin pressure at the blood inlet, PBout pressure at the blood outlet, PDin pressure at the dialysate inlet, PDout pressure at the dialysate outlet.

## Supplementary References

1.	Ergun, S. Fluid Flow through Packed Columns. Chem. Eng. Prog. 48, 89–94 (1952).

2.	Mehta, D. & Hawley, M. C. Wall effect in packed columns. Ind. Eng. Chem. Process Des. Dev. 8, 280–282 (1969).

3.	Macdonald, I. F., El-Sayed, M. S., Mow, K. & Dullien, F. A. L. Flow through Porous Media—the Ergun Equation Revisited. Ind. Eng. Chem. Fundam. 18, 199–208 (1979).

4.	Irmay, S. On the theoretical derivation of Darcy and Forchheimer formulas. Trans. Am. Geophys. Union 39, 702–707 (1958).

5.	Suzuki, Y., Kohori, F. & Sakai, K. Computer-aided design of hollow-fiber dialyzers. J. Artif. Organs 4, 326–330 (2001).

6.	Papathanasiou, T. D., Markicevic, B. & Dendy, E. D. A computational evaluation of the Ergun and Forchheimer equations for fibrous porous media. Phys. Fluids 13, 2795–2804 (2001).

7.	Yoshikawa, S., Ogawa, K. & Minegishi, S. Distributions of pressure and flow rate in a hollow-fiber membrane module for plasma separation. J. Chem. Eng. Japan 27, 385–390 (1994).

8.	Yamamoto, K. I. et al. Computational evaluation of dialysis fluid flow in dialyzers with variously designed jackets. Artif. Organs 33, 481–486 (2009).

9.	Sparrow, E. M. & Loeffler, A. L. Longitudinal laminar flow between cylinders arranged in regular array. AIChE J. 5, 325–330 (1959).

10.	Happel, J. Viscous flow relative to arrays of cylinders. AIChE J. 5, 174–177 (1959).

11.	Keller, J. B. Viscous flow through a grating or lattice of cylinders. J. Fluid Mech. 18, 94–96 (1964).

12.	Sangani, A. S. & Acrivos, A. Slow flow past periodic arrays of cylinders with application to heat transfer. Int. J. Multiph. Flow 8, 193–206 (1982).

13.	Sangani, A. S. & Acrivos, A. Slow flow through a periodic array of spheres. Int. J. Multiph. Flow 8, 343–360 (1982).

14.	Happel, J. & Brenner, H. Low Reynolds number hydrodynamics. vol. 1 (1963).

15.	Skartsis, L. ., Kardos, J. L. . & Khomami, B. Resin flow through fiber beds during composite manufacturing processes. part I: review of Newtonian flow through fiber beds. Polym. Eng. Sci. 32, 221–230 (1992).

16.	Bird, R. B., Stewart, W. E. & Lightfoot, E. N. Transport Phenomena. (2002).

17.	Erdim, E., Akgiray, Ö. & Demir, I. A revisit of pressure drop-flow rate correlations for packed beds of spheres. Powder Technol. 283, 488–504 (2015).

18.	Cancilla, N. et al. CFD prediction of shell-side flow and mass transfer in regular fiber arrays. Int. J. Heat Mass Transf. 168, 120855 (2021).

19.	Drummond, J. E. & Tahir, M. . Laminar viscous flow through regular arrays of parallel solid cylinders. Int. J. Multiph. Flow 10, 515–540 (1984).

20.	Cancilla, N. et al. A porous media CFD model for the simulation of hemodialysis in hollow fiber membrane modules. J. Memb. Sci. 646, 120219 (2022).

21.	Cancilla, N. Experimental and computational investigation of hollow-fiber modules for biomedical applications. (Università degli Studi di Palermo, 2022).

22.	Liao, Z. Numerical and Experimental studies of mass transfer in artificial kidney and hemodialysis. PhD Dissertation (University of Kentucky, 2002).

23.	Eloot, S., D’Asseler, Y., De Bondt, P. & Verdonck, P. Combining SPECT medical imaging and computational fluid dynamics for analyzing blood and dialysate flow in hemodialyzers. Int. J. Artif. Organs 28, 739–749 (2005).

24.	Legallais, C., Catapano, G., Harten, B. von & Baurmeister, U. A theoretical model to predict the in vitro performance of hemodiafilters. J. Memb. Sci. 168, 3–15 (2000).

25.	Abaci, H. E. & Altinkaya, S. A. Modeling of hemodialysis operation. Ann. Biomed. Eng. 38, 3347–3362 (2010).

26.	Eloot, S., De Wachter, D., Van Tricht, I. & Verdonck, P. Computational flow modeling in hollow-fiber dialyzers. Artif. Organs 26, 590–599 (2002).

27.	Eloot, S., Vierendeels, J. & Verdonck, P. Optimisation of solute transport in dialysers using a three-dimensional finite volume model. Comput. Methods Biomech. Biomed. Engin. 9, 363–370 (2006).

28.	Islam, S. & Szpunar, J. Study of dialyzer membrane ( Polyflux 210H ) and effects of different parameters on dialysis performance. Open J. Nephrol. 3, 161–167 (2013).

29.	Yaqoob, T., Ahsan, M., Hussain, A. & Ahmad, I. Computational fluid dynamics (CFD) modeling and simulation of flow regulatory mechanism in artificial kidney using finite element method. Membranes (Basel). 10, 139 (2020).

30.	Donato, D. et al. Optimization of dialyzer design to maximize solute removal with a two-dimensional transport model. J. Memb. Sci. 541, 519–528 (2017).

31.	Goodarzi, M. R. & Mohebbi-Kalhori, D. Numerical modeling and parametric study of a hollow fiber dialyzer using double porous media approach. South African J. Chem. Eng. 42, 1–11 (2022).

32.	Li, W. et al. Simulation and experimental study on the effect of channeling flows on the transport of toxins in hemodialyzers. J. Memb. Sci. 501, 123–133 (2016).

33.	Łabȩcki, M., Piret, J. M. & Bowen, B. D. Two-dimensional analysis of fluid flow in hollow-fibre module. Chem. Eng. Sci. 50, 3369–3384 (1995).

34.	Ding, W. et al. Double porous media model for mass transfer of hemodialyzers. Int. J. Heat Mass Transf. 47, 4849–4855 (2004).

35.	Yang, C., Qinglian, W., Guo, J., Nakayama, A. & Qiu, T. Upscaling solute concentration transport equations of countercurrent dialyzer systems. Chem. Eng. Sci. 134, 108–118 (2015).

36.	Sano, Y. & Nakayama, A. A porous media approach for analyzing a countercurrent dialyzer system. J. Heat Transfer 134, 072602 (2012).

37.	Ding, W. et al. Three-dimensional simulation of mass transfer in artificial kidneys. Artif. Organs 39, E79–E89 (2015).

38.	Annan, K. Mathematical modeling of the dynamic exchange of solutes during bicarbonate dialysis. Math. Comput. Model. 55, 1691–1704 (2012).

39.	Osuga, T. et al. Dialysate pressure isobars in a hollow-fiber dialyzer determined from magnetic resonance imaging and numerical simulation of dialysate flow. Artif. Organs 22, 907–909 (1998).

40.	Lemanski, J. & Lipscomb, G. G. Effect of shell-side flows on hollow-fiber membrane device performance. AIChE J. 41, 2322–2326 (1995).