Table 3. Agreement and correlation of Keratometry and Astigmatism in cataract patients Obtained with four devices
	Parameter
	CCC (95%CI)
	Mean Diff. ± SD
	95%LoA
	P value

	Kf, D
	
	
	
	

	A vs. B
	0.960 (0.944, 0.977)
	-0.15 ± 0.43
	-0.982, 0.688
	<0.001

	A vs. C
	0.973 (0.962, 0.984)
	-0.13 ± 0.36
	-0.838, 0.575
	<0.001

	A vs. D
	0.843 (0.774, 0.912)
	-0.01 ± 1.05
	-2.061, 2.051
	<0.001

	B vs. C
	0.968 (0.955, 0.981)
	0.02 ± 0.41
	-0.787, 0.828
	<0.001

	B vs. D
	0.796 (0.710, 0.882)
	0.17 ± 1.17
	-2.123, 2.460
	<0.001

	C vs. D
	0.820 (0.745, 0.895)
	0.10 ± 1.15
	-2.151, 2.345
	<0.001

	Ks, D
	
	
	
	

	A vs. B
	0.948 (0.927, 0.969)
	-0.24 ± 0.47
	-1.172, 0.684
	<0.001

	A vs. C
	0.971 (0.959, 0.982)
	-0.12 ± 0.38
	-0.866, 0.628
	<0.001

	A vs. D
	0.853 (0.788, 0.918)
	-0.12 ± 1.04
	-2.157, 1.924
	<0.001

	B vs. C
	0.966 (0.952, 0.979)
	0.11 ± 0.42
	-0.716, 0.925
	<0.001

	B vs. D
	0.784 (0.691, 0.877)
	0.20 ± 1.26
	-2.266, 2.658
	<0.001

	C vs. D
	0.820 (0.747, 0.893)
	0.04 ± 1.16
	-2.231, 2.319
	<0.001

	Km, D
	
	
	
	

	A vs. B
	0.970 (0.957, 0.982)
	-0.20 ± 0.34
	-0.866, 0.475
	<0.001

	A vs. C
	0.982 (0.975, 0.989)
	-0.13 ± 0.28
	-0.677, 0.427
	<0.001

	A vs. D
	0.852 (0.786, 0.917)
	-0.06 ± 1.03
	-2.076, 1.954
	<0.001

	B vs. C
	0.982 (0.975, 0.989)
	0.06 ± 0.30
	-0.522, 0.647
	<0.001

	B vs. D
	0.797 (0.711, 0.884)
	0.18 ± 1.19
	-2.134, 2.499
	<0.001

	C vs. D
	0.824 (0.752, 0.897)
	0.07 ± 1.14
	-2.154, 2.295
	<0.001

	J0, D
	
	
	
	
	

	A vs. B
	0.102 (-0.105, 0.309)
	-0.05 ± 0.43
	-0.897, 0.792
	0.333

	A vs. C
	0.135 (-0.340, 0.070)
	0.01 ± 0.44
	-0.850, 0.861
	0.196

	A vs. D
	-0.088 (-0.336, 0.159)
	-0.01 ± 0.44
	-0.862, 0.842
	0.485

	B vs. C
	-0.046 (-0.244, 0.152)
	0.05 ± 0.46
	-0.850, 0.952
	0.650

	B vs. D
	-0.033 (-0.278, 0.212)
	0.07 ± 0.48
	-0.874, 1.009
	0.793

	C vs. D
	0.089 (-0.147, 0.325)
	-0.03 ± 0.39
	-0.806, 0.739
	0.460

	J45, D
	
	
	
	

	A vs. B
	-0.167 (-0.039, 0.373)
	-0.05 ± 0.41
	-0.845, 0.755
	0.112

	A vs. C
	-0.140 (-0.064, 0.345)
	-0.02 ± 0.39
	-0.776, 0.745
	0.177

	A vs. D
	0.044 (-0.206, 0.294)
	0.01 ± 0.38
	-0.743, 0.758
	0.731

	B vs. C
	0.065 (-0.137, 0.266)
	0.04 ± 0.43
	-0.804, 0.881
	0.528

	B vs. D
	-0.012 (-0.259, 0.235)
	0.00 ± 0.43
	-0.831, 0.838
	0.925

	C vs. D
	-0.025 (-0.265, 0.214)
	-0.05 ± 0.41
	-0.849, 0.755
	0.836


CCC, Lin's concordance correlation coefficient.; CI, confidence interval; D, diopter; Diff, Difference; J0, Jackson cross-cylinder at 0 and 90 degrees; J45, Jackson cross-cylinder at 45 and 135 degrees; Kf, flat keratometry, Ks, steep keratometry; Km, mean keratometry; LoA, limits of agreement; SD, standard error.
A=Pentacam, B=Sirius, C=IOLMaster 700, D= iTrace, 
