Table 1. ICP-MS settings

| Parameter | Setting / Value |
| --- | --- |
| Plasma Conditions |  |
| RF power | 1550 W |
| Sampling depth | 10.0 mm |
| Sample Introduction |  |
| Nebulizer gas flow | 1.05 L/min |
| Collision/Reaction Cell Conditions |  |
| Helium (He) mode |  |
| Cell gas flow | He: 5.0 mL/min |
| KED* | 5.0 V |
| Hydrogen (H₂) mode |  |
| Cell gas flow | H₂: 6.0 mL/min |
| KED | 3.0 V |
| Data Acquisition |  |
| Sweeps per reading | 100 |
| Replicates | 3 |
| Integration time per mass | 0.1 – 1.0 s (Isotope dependent) |

*Kinetic energy discrimination

|  | HNO3 vol. (mL) | H2O2 Vol. (mL) | H2O Vol. (mL) | HNO3 conc. (%) | H2O2 conc. (%) |
| --- | --- | --- | --- | --- | --- |
| Sample-1 | 4.0 | 1.00 | 0.00 | 56.2 | 7.1 |
| Sample-2 | 4.0 | 0.55 | 0.45 | 56.2 | 3.9 |
| Sample-3 | 4.0 | 0.10 | 0.90 | 56.2 | 0.7 |
| Sample-4 | 2.3 | 1.00 | 1.70 | 32.3 | 7.1 |
| Sample-5 | 2.3 | 0.55 | 2.15 | 32.3 | 3.9 |
| Sample-6 | 2.3 | 0.10 | 2.60 | 32.3 | 0.7 |
| Sample-7 | 0.6 | 1.00 | 3.40 | 8.4 | 7.1 |
| Sample-8 | 0.6 | 0.55 | 3.85 | 8.4 | 3.9 |
| Sample-9 | 0.6 | 0.10 | 4.30 | 8.4 | 0.7 |

Table 2. Sample preparation

n=2

|  | ICP-MS analytical parameters |  |  |  | Vitamin tablet |  |  |  |  |  |  |  | Chocolate |  |  |  | Cheese |  |  |  | Nori (edible seaweed) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | ICP-MS Gas mode | Detected isotope | Internal standard |  | Certified value ± expanded uncertainty |  |  | Mean*1 | RSD*2 (%) | Max value | Max/certified value (%) |  | Mean | RSD (%) | Max value |  | Mean | RSD (%) | Max value |  | Mean | RSD (%) | Max value |
| Na (μg/g) | H2 | 23 | Sc |  | 330 | ± | 20 | 322.1 | 2.7 | 337.7 | 102.3 |  | 18.1 | 3.9 | 20.0 |  | 9120 | 2.5 | 9534 |  | 6161 | 3.1 | 6442 |
| Mg (mg/g) | He | 24 | Sc |  | 67.8 | ± | 4 | 68.1 | 1.4 | 69.9 | 103.1 |  | 2.12 | 2.4 | 2.2 |  | 0.59 | 1.7 | 0.61 |  | 3.27 | 3.1 | 3.42 |
| K (mg/g) | He | 39 | Sc |  | 53.1 | ± | 7 | 55.4 | 0.8 | 56.0 | 105.5 |  | 9.4 | 1.1 | 9.7 |  | 1.14 | 1.4 | 1.17 |  | 27.9 | 3.2 | 29.1 |
| Ca (mg/g) | H2 | 40 | Sc |  | 110.7 | ± | 5.3 | 107.2 | 1.5 | 111.1 | 100.4 |  | 0.53 | 3.9 | 0.55 |  | 11.4 | 2.1 | 11.8 |  | 1.05 | 3.9 | 1.12 |
| Cr (μg/g) | H2 | 52 | Sc |  | 93.7 | ± | 2.7 | 82.1 | 32.4 | 104.2 | 111.2 |  | 2.6 | 47.3 | 4.13 |  | 0.36 | 27.7 | 0.61 |  | 0.60 | 50.5 | 1.27 |
| Mn (mg/g) | H2 | 55 | Sc |  | 1.44 | ± | 0.11 | 1.46 | 3.6 | 1.57 | 109.1 |  | 0.023 | 0.9 | 0.023 |  | 0.00056 | 2.2 | 0.00059 |  | 0.028 | 2.8 | 0.029 |
| Fe (mg/g) | H2 | 56 | Sc |  | 12.4 | ± | 0.91 | 5.89 | 81.6 | 12.6 | 101.8 |  | 0.069 | 59.2 | 0.099 |  | 0.0027 | 15.6 | 0.0037 |  | 0.065 | 36.6 | 0.077 |
| Co (μg/g) | He | 59 | Sc |  | 0.81 | ± | 0.01 | 0.87 | 2.7 | 0.90 | 111.7 |  | 0.453 | 7.3 | 0.486 |  | 0.013 | 8.4 | 0.015 |  | 0.14 | 4.0 | 0.15 |
| Cu (mg/g) | He | 63 | Ge |  | 1.40 | ± | 0.17 | 1.39 | 2.6 | 1.47 | 105.1 |  | 0.0213 | 5.2 | 0.0227 |  | 0.00047 | 9.2 | 0.00059 |  | 0.0048 | 4.9 | 0.0052 |
| Zn (mg/g) | He | 66 | Ge |  | 10.2 | ± | 0.81 | 10.3 | 0.8 | 10.4 | 102.1 |  | 0.0377 | 2.0 | 0.0389 |  | 0.066 | 1.8 | 0.068 |  | 0.025 | 3.0 | 0.026 |
| Se (μg/g) | H2 | 78 | Ge |  | 17.4 | ± | 0.45 | 15.9 | 7.0 | 18.4 | 105.4 |  | 0.0806 | 22.3 | 0.114 |  | 0.20 | 17.9 | 0.24 |  | 0.036 | 17.0 | 0.044 |
| Mo (μg/g) | He | 95 | Y |  | 70.7 | ± | 4.5 | 54.2 | 23.4 | 70.4 | 99.6 |  | 0.207 | 30.2 | 0.258 |  | 0.097 | 28.6 | 0.13 |  | 0.55 | 40.8 | 0.71 |
| Pb (μg/g) | He | 208 | Y |  | 0.27 | ± | 0.0024 | 0.27 | 3.5 | 0.29 | 105.8 |  | 0.086 | 32.9 | 0.159 |  | 0.043 | 5.6 | 0.048 |  | 0.065 | 4.0 | 0.07 |

Table 3. Quantified minerals in food samples

*1 Mean value from a parallel analysis of 9 conditions
*2 RSD: Relative standard deviation

|  | Mineral | R2 *1 | Adj-R2 *2 | DF*3 | Model F-value | Model p-value | LoF p-value*4 | NRMSE*5 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Vitamin tablet | Cr | 0.922 | 0.890 | 5 | 28.4 | < 0.001 | < 0.001 | 0.1045 |
|  | Fe | 0.970 | 0.957 | 5 | 77.0 | < 0.001 | 0.658 | 0.0819 |
|  | Co | 0.647 | 0.499 | 5 | 4.4 | 0.017 | 0.133 | 0.1847 |
|  | Cu | 0.201 | -0.133 | 5 | 0.6 | 0.700 | 0.752 | 0.2534 |
|  | Se | 0.388 | 0.132 | 5 | 1.5 | 0.256 | 0.791 | 0.249 |
|  | Mo | 0.962 | 0.946 | 5 | 60.2 | < 0.001 | 0.864 | 0.0881 |
|  | Pb | 0.567 | 0.386 | 5 | 3.1 | 0.049 | 0.415 | 0.1686 |
| Chocolate | Cr | 0.958 | 0.940 | 5 | 54.6 | < 0.001 | 0.450 | 0.0839 |
|  | Fe | 0.991 | 0.987 | 5 | 265.0 | < 0.001 | < 0.001 | 0.0474 |
|  | Co | 0.917 | 0.882 | 5 | 26.4 | < 0.001 | 0.019 | 0.119 |
|  | Cu | 0.874 | 0.821 | 5 | 16.6 | < 0.001 | 0.002 | 0.1131 |
|  | Se | 0.849 | 0.785 | 5 | 13.4 | < 0.001 | 0.674 | 0.1369 |
|  | Mo | 0.894 | 0.849 | 5 | 20.2 | < 0.001 | < 0.001 | 0.129 |
|  | Pb | 0.944 | 0.921 | 5 | 40.8 | < 0.001 | 0.609 | 0.0763 |
| Cheese | Cr | 0.651 | 0.506 | 5 | 4.5 | 0.015 | 0.680 | 0.1604 |
|  | Fe | 0.577 | 0.400 | 5 | 3.3 | 0.043 | 0.477 | 0.2047 |
|  | Co | 0.435 | 0.199 | 5 | 1.8 | 0.178 | 0.101 | 0.2292 |
|  | Cu | 0.855 | 0.794 | 5 | 14.1 | < 0.001 | 0.195 | 0.1157 |
|  | Se | 0.924 | 0.893 | 5 | 29.3 | < 0.001 | 0.316 | 0.1106 |
|  | Mo | 0.872 | 0.819 | 5 | 16.4 | < 0.001 | 0.191 | 0.1292 |
|  | Pb | 0.736 | 0.626 | 5 | 6.7 | 0.003 | 0.151 | 0.1528 |
| Nori | Cr | 0.901 | 0.860 | 5 | 21.9 | < 0.001 | 0.295 | 0.1064 |
|  | Fe | 0.809 | 0.729 | 5 | 10.1 | < 0.001 | < 0.001 | 0.1605 |
|  | Co | 0.732 | 0.620 | 5 | 6.6 | 0.004 | 0.113 | 0.1917 |
|  | Cu | 0.797 | 0.713 | 5 | 9.4 | < 0.001 | 0.666 | 0.1458 |
|  | Se | 0.692 | 0.563 | 5 | 5.4 | 0.008 | 0.470 | 0.1906 |
|  | Mo | 0.903 | 0.862 | 5 | 22.3 | < 0.001 | < 0.001 | 0.1239 |
|  | Pb | 0.481 | 0.265 | 5 | 2.2 | 0.119 | 0.052 | 0.2098 |

Table 4. Analysis of variance (ANOVA) and adequacy of the quadratic models for mineral extraction

*1 Coefficient of determination.

*2 Adjusted coefficient of determination.

*3 Degrees of freedom for the regression model.

*4  p -value of the lack-of-fit test (a value > 0.05 indicates no significant lack of fit).

*5 Normalized root mean square error, representing the predictive accuracy of the model.

|  | Optimal HNO3 (%) | Optimal H2O2 (%) | Desirability (Geometric Mean) | Lowest extraction rate (%)* | Lowest extraction Mineral | Average extraction rate (%) |
| --- | --- | --- | --- | --- | --- | --- |
| Vitamin tablet | 38.7 | 4.1 | 98.7 | 97.1 | Pb | 98.7 |
| Chocolate | 56.2 | 7.1 | 98.7 | 93.3 | Fe | 98.7 |
| Cheese | 51.0 | 0.7 | 96.7 | 91.6 | Pb | 98.7 |
| Nori | 56.2 | 1.3 | 95.4 | 84.4 | Fe | 95.5 |
| All data | 47.4 | 5.8 | 93.5 | 77.8 | Cr (Nori) | 93.8 |

Table 5. Optimal extraction conditions for the simultaneous analysis of multiple minerals using the desirability function

* Estimated extraction value/ Max extracted value