Supplementary Material

Robot-Tracked Multi-View 3D Ultrasound Reconstruction for Quantitative Assessment of Simulated Abdominal Haemorrhage in a Tissue-Mimicking Phantom

Supplementary Table S1. Experimental dataset and acquisition summary.

| Target injected fluid volume (mL) | No. of 3D volumes | Acquisition geometry | Volume dimensions (voxels) | Native voxel spacing (mm) | Frames per volume | Approximate FOV (mm) |
| --- | --- | --- | --- | --- | --- | --- |
| 70 | 21 | Group 1 | 512 × 503 × 256 | 0.11918 × 0.10265 × 0.17431 | 256 | 38.3 × 60.9 |
| 100 | 29 | Group 1 | 512 × 503 × 256 | 0.11918 × 0.10265 × 0.17431 | 256 | 38.3 × 60.9 |
| 120 | 22 | Group 2 | 512 × 510 × 256 | 0.12818 × 0.12042 × 0.18762 | 256 | 38.3 × 65.5 |
| 150 | 25 | Group 2 | 512 × 510 × 256 | 0.12818 × 0.12042 × 0.18762 | 256 | 38.3 × 65.5 |
| 190 | 25 | Group 2 | 512 × 510 × 256 | 0.12818 × 0.12042 × 0.18762 | 256 | 38.3 × 65.5 |
| Total | 122 | — | — | — | — | — |

Notes. All acquisitions used the VL13-5 mechanical 3D transducer. The acquisition mode was Free Form / native 3D Cartesian. Each volume comprised 256 frames at 30 Hz (33.33-ms frame interval) and was stored in DICOM format. Two acquisition geometries were used because imaging depth was increased for the higher-volume conditions to maintain coverage of the expanded free-fluid reservoir.

Supplementary Table S2. US–robot hand–eye calibration and held-out cross-validation performance.

- Final calibration performance

| Metric | Result |
| --- | --- |
| Number of calibration acquisitions | 64 |
| RMS point-to-point reconstruction error | 1.23 mm |
| SD of point-to-point reconstruction error | 0.53 mm |
| Mean absolute translation discrepancy, x-axis | 0.50 mm |
| Mean absolute translation discrepancy, y-axis | 0.60 mm |
| Mean absolute translation discrepancy, z-axis | 0.62 mm |
| Mean absolute rotation discrepancy, x-axis | 1.16° |
| Mean absolute rotation discrepancy, y-axis | 1.34° |
| Mean absolute rotation discrepancy, z-axis | 1.17° |

- Thirteen-fold held-out pose cross-validation

| Fold | Held-out poses | RMS error (mm) |
| --- | --- | --- |
| 1 | 5 | 1.17 |
| 2 | 5 | 1.26 |
| 3 | 5 | 1.17 |
| 4 | 4 | 0.97 |
| 5 | 5 | 2.06 |
| 6 | 5 | 1.33 |
| 7 | 5 | 0.61 |
| 8 | 5 | 0.86 |
| 9 | 5 | 1.14 |
| 10 | 4 | 1.57 |
| 11 | 5 | 1.53 |
| 12 | 5 | 1.15 |
| 13 | 5 | 1.13 |
| Mean ± SD | — | 1.23 ± 0.36 |
| Pooled held-out | 63 | 1.27 (SD 0.54) |

Note. The reference pose used to define the relative pose pairs in the hand–eye formulation was retained in the training set for every cross-validation fold. The remaining 63 poses were partitioned across 13 folds and each evaluated once as a held-out observation. The mean ± SD row summarises the 13 fold-wise RMS errors, whereas the pooled held-out result was calculated across all 63 held-out pose errors. RMS, root-mean-square; SD, standard deviation.