Supplementary Material

Table S1. Parameters of DL-HASTE and BLADE sequences

| Variable | DL-HASTE | BLADE |
| --- | --- | --- |
| TR/TE (ms) | 587/103 | 3000/114 |
| Slice thickness (mm) | 6 | 6 |
| Slice gap (mm) | 1.2 | 1.2 |
| FOV (mm2) | 380 × 380 | 380 × 380 |
| Scan matrix Flip angle (degrees) | 320 × 320 160 | 320 × 320 180 |
| Acceleration factor | 3 | 2 (GRAPPA) |
| Partial Fourier factor | 6/8 | N.A. |
| Motion control scheme | Breath-hold | Respiratory triggered |
| Total scan time | 12 s × 2 | 2~3mins |

Note—DL-HASTE: Deep learning accelerated Half-Fourier Acquisition Single-shot Turbo spin Echo imaging. FOV: Field of View. GRAPPA: GeneRalized Autocalibrating Partial Parallel Acquisition.

Table S2. Likert scale scoring for qualitative imaging features

| Variable | Scoring system |
| --- | --- |
| Overall image quality | 1. Unacceptable; 2. Poor; 3. Fair; 4. Good; 5. Excellent |
| Artifact | 1. Unacceptable; 2. Extreme Blur; 3. Moderate Blur; 4. Mild Blur; 5. No Blur |
| Sharpness of liver margin | 1. Unacceptable; 2. Extreme Blur; 3. Moderate Blur; 4. Mild Blur; 5. No Blur |
| Hepatic vessel margin | 1. Unacceptable; 2. Extreme Blur; 3. Moderate Blur; 4. Mild Blur; 5. No Blur |
| Pancreatic duct margin | 1. Unacceptable; 2. Extreme Blur; 3. Moderate Blur; 4. Mild Blur; 5. No Blur |
| Homogeneity of fat suppression | 1. Unacceptable; 2. Poor; 3. Fair; 4. Good; 5. Excellent |
| Strength of fat suppression | 1. Unacceptable; 2. Poor; 3. Fair; 4. Good; 5. Excellent |
| Conspicuity of the liver smallest lesion | 1. Unreadable; 2. Poor; 3. Moderate; 4. Good; 5. Excellent |

Fig. S1—Characterization of liver cysts in a participant. The T2-weighted DL HASTE axial image (A) shows better conspicuity for small cysts than the BLADE axial image (B); C, enhanced image shows no enhancement of the cysts. The largest cyst (arrow) CNR of the DL HASTE sequence and BLADE sequence are 82.42 and 46.74, respectively. Both readers classified the lesion as markedly hyperintense on both sequences. DL HASTE, half-Fourier acquisition single-shot turbo spin echo using deep learning; CNR, contrast-to-noise ratio.