# Tables

Ventilator Circuits as a Potential Source of Airway Microplastics in Mechanically Ventilated Neonates

Table 1. Participant and sampling characteristics by cumulative invasive ventilation duration.

| Characteristic | Overall (n=20) | Short term (n=7) | Intermediate (n=7) | Long term (n=6) |
| --- | --- | --- | --- | --- |
| Gestational age, weeks | 33.4 (30.8, 34.6) | 35.1 (34.6, 35.9) | 33.2 (32.6, 33.8) | 29.6 (29.0, 30.3) |
| Birth weight, g | 1985 (1462, 2152) | 2270 (2165, 2435) | 1960 (1840, 2030) | 1225 (1138, 1372) |
| Male sex, n (%) | 10 (50.0) | 4 (57.1) | 3 (42.9) | 3 (50.0) |
| Invasive ventilation duration, h | 118.5 (68.8, 198.0) | 60.0 (46.5, 67.5) | 122.0 (109.0, 138.0) | 270.0 (225.0, 324.0) |
| Airway specimen volume, mL | 0.30 (0.28, 0.31) | 0.29 (0.28, 0.30) | 0.31 (0.29, 0.33) | 0.29 (0.28, 0.31) |

Values are median (IQR) unless otherwise indicated. Groups were defined by cumulative invasive ventilation duration.

Table 2. Paired microplastic findings in airway secretion and ventilator circuit rinse samples.

| Outcome | Overall (n=20) | Short term (n=7) | Intermediate (n=7) | Long term (n=6) | P value |
| --- | --- | --- | --- | --- | --- |
| Airway secretion samples |  |  |  |  |  |
| Microplastic-positive samples, n (%) | 15 (75.0) | 4 (57.1) | 6 (85.7) | 5 (83.3) | 0.545 |
| Particle count, n | 2.0 (0.8, 4.2) | 1.0 (0.0, 1.0) | 3.0 (1.5, 4.0) | 8.0 (4.0, 10.5) | 0.021 |
| Abundance, particles/mL | 6.15 (2.50, 13.35) | 3.33 (0.00, 3.73) | 10.00 (4.85, 12.16) | 26.45 (12.97, 34.37) | 0.020 |
| Mean particle diameter, micrometers | 27.3 (22.8, 33.6) | 20.5 (19.4, 22.8) | 26.6 (24.6, 29.0) | 38.7 (35.1, 41.3) | 0.008 |
| Predominant polymer | PE | PE | PE | PE | - |
| Ventilator circuit rinse samples |  |  |  |  |  |
| Microplastic-positive samples, n (%) | 17 (85.0) | 5 (71.4) | 6 (85.7) | 6 (100.0) | 0.742 |
| Particle count, n | 2.5 (1.0, 5.2) | 1.0 (0.5, 1.5) | 3.0 (2.5, 5.0) | 10.5 (6.0, 13.5) | 0.003 |
| Abundance, particles/mL | 0.25 (0.10, 0.53) | 0.10 (0.05, 0.15) | 0.30 (0.25, 0.50) | 1.05 (0.60, 1.35) | 0.003 |
| Mean particle diameter, micrometers | 27.4 (22.5, 34.2) | 20.7 (20.1, 22.4) | 28.6 (24.5, 30.8) | 38.5 (34.9, 42.3) | 0.004 |
| Predominant polymer | PE | PE | PE | PE | - |

Values are median (IQR) unless otherwise indicated. P values are from Kruskal-Wallis tests for continuous outcomes and Fisher exact tests for categorical outcomes. Mean particle diameter is reported for positive samples only. PE, polyethylene.

Table 3. Individual paired microplastic findings and recorded specimen volumes.

| Participant ID | Ventilation duration, h | Airway specimen volume, mL | Airway count, n | Airway abundance, particles/mL | Circuit rinse volume, mL | Circuit count, n | Circuit abundance, particles/mL | Predominant polymer |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Num.1 | 29 | 0.28 | 0 | 0.00 | 10.0 | 0 | 0.00 | Not detected |
| Num.2 | 43 | 0.31 | 0 | 0.00 | 10.0 | 1 | 0.10 | PE |
| Num.3 | 50 | 0.25 | 1 | 4.00 | 10.0 | 1 | 0.10 | PE |
| Num.4 | 60 | 0.30 | 1 | 3.33 | 10.0 | 0 | 0.00 | PE |
| Num.5 | 65 | 0.27 | 0 | 0.00 | 10.0 | 2 | 0.20 | PE |
| Num.6 | 70 | 0.33 | 2 | 6.06 | 10.0 | 2 | 0.20 | PE |
| Num.7 | 72 | 0.29 | 1 | 3.45 | 10.0 | 1 | 0.10 | PE |
| Num.8 | 96 | 0.28 | 0 | 0.00 | 10.0 | 2 | 0.20 | PE |
| Num.9 | 103 | 0.32 | 2 | 6.25 | 10.0 | 0 | 0.00 | PE |
| Num.10 | 115 | 0.30 | 3 | 10.00 | 10.0 | 3 | 0.30 | PE |
| Num.11 | 122 | 0.29 | 1 | 3.45 | 10.0 | 3 | 0.30 | PE |
| Num.12 | 132 | 0.35 | 4 | 11.43 | 10.0 | 5 | 0.50 | PE |
| Num.13 | 144 | 0.31 | 4 | 12.90 | 10.0 | 5 | 0.50 | PE |
| Num.14 | 168 | 0.34 | 5 | 14.71 | 10.0 | 6 | 0.60 | PE |
| Num.15 | 192 | 0.30 | 3 | 10.00 | 10.0 | 5 | 0.50 | PE |
| Num.16 | 216 | 0.28 | 0 | 0.00 | 10.0 | 2 | 0.20 | PE |
| Num.17 | 252 | 0.32 | 7 | 21.88 | 10.0 | 9 | 0.90 | PE |
| Num.18 | 288 | 0.29 | 9 | 31.03 | 10.0 | 12 | 1.20 | PE |
| Num.19 | 336 | 0.31 | 11 | 35.48 | 10.0 | 14 | 1.40 | PE |
| Num.20 | 384 | 0.27 | 13 | 48.15 | 10.0 | 16 | 1.60 | PE |

Abundance is particles per mL. PE, polyethylene.

Table 4. Quality assurance and quality control findings.

| Control | n | Particle count | Finding |
| --- | --- | --- | --- |
| Reagent blank | 1 | 0 | Not detected |
| Transport blank | 1 | 0 | Not detected |
| Equipment blank | 1 | 0 | Not detected |
| Procedural blank | 1 | 0 | Not detected |
| Airborne blank | 1 | 1 | Mean diameter 32 micrometers |

Table 5. In vitro polyethylene particle counts in ventilator-circuit release experiments.

| Comparison | Control | After 72 h | Exact P value |
| --- | --- | --- | --- |
| Patient-end gas output, particles/m3 | System blank (n=6) 1.0 (1.0, 1.8) | Patient-end output (n=9) 12.0 (8.0, 17.0) | 0.0004 |
| Circuit eluate, particles/10 mL | Unused circuit (n=9) 2.0 (2.0, 3.0) | Post-run circuit (n=9) 11.0 (9.0, 13.0) | 0.000041 |

Values are median (IQR). Counts were normalized to the stated gas or eluate volume. All reported particles were classified as polyethylene by confocal micro-Raman spectroscopy.