## Guide
Additional file 1 - Search strategy and evidence map
Purpose: Search strategy, final study-selection flow, report-level evidence map, funding/manufacturer involvement, methodological characteristics, device taxonomy and reporting framework for the scoping review.
Bibliographic search | A single structured high-sensitivity PubMed search was run on 1 September 2026 for publications from 1 January 2010 through 1 September 2026. The search returned 643 records.
Study selection | The PubMed search returned 643 records. Review-type records were identified and removed in PubMed using the Review publication-type filter (n=437), leaving 206 records for reviewer screening. Two reviewers completed title/abstract screening and eligibility assessment between 1 and 4 September 2026. Seventy-one reports representing 70 distinct studies met the eligibility criteria.
Protocol | No prospective protocol was written, registered, or publicly posted.
Data charting / extraction | After study selection, data charting and extraction were performed by one reviewer using a structured evidence map. No independent duplicate data extraction was performed.
Critical appraisal | No formal critical appraisal or risk-of-bias assessment was performed. Methodological characteristics relevant to interpretation are reported descriptively.
Final evidence map | 70 distinct included studies represented by 71 reports: 43 R1 studies (44 reports), 17 R2 studies, and 10 R3 studies. Anderson 2018 is a companion secondary analysis of BETR Disinfection and is not counted as a separate study.
Funding | Supported by the KNOWLEDGE IN CLEANING project (Bando a Cascata, PNC ANTHEM; project code PNC0000003). The funding body had no role in manuscript preparation, interpretation, or writing and did not influence its content.
Contents | Search_Strategy; PubMed_Search_Query; Study_Selection; Evidence_Map_Summary; Included_Studies; Methodological_Characteristics; Device_Taxonomy.

## Search_Strategy
PubMed search strategy and selection process
The search was designed as a targeted search for a focused scoping review.
Element | Recorded method | Status / note
Database / platform | PubMed / National Library of Medicine (NIH) | No other database was used.
Search date | 1 September 2026 | Single bibliographic search used for the final review.
Publication-date interval | 2010-01-01 to 2026-09-01 | The 2010 lower limit focused the review on contemporary deployable UV technologies and healthcare implementation.
Search construction | Structured high-sensitivity Boolean query combining controlled vocabulary and title/abstract terms for UV radiation, healthcare settings, surfaces/rooms and disinfection. | The exact reproducible query is provided in PubMed_Search_Query.
Publication-type / language filters | No publication-type or language filters were applied within the Boolean query. | After retrieval, the PubMed Review publication-type filter was used to identify and remove 437 review-type records before reviewer screening.
Records retrieved | 643 | Records returned by the structured PubMed query before publication-type filtering.
Selection reviewers | Two reviewers | Study selection decisions were made by the reviewers; disagreements were resolved by discussion.
Selection period | 1-4 September 2026 | Title/abstract screening and full-text eligibility assessment were completed during this period.
Review-type records removed in PubMed | 437 | Identified and removed using the PubMed Review publication-type filter before reviewer screening.
Records screened by reviewers | 206 | Title/abstract screening and subsequent eligibility assessment by two reviewers.
Records not included after reviewer screening / eligibility | 135 | Calculated from 206 records screened minus 71 included reports.
Full-text primary eligibility | Portable/mobile automated systems or operator-positioned UV sources used for a defined treatment cycle in routine healthcare, genuine healthcare rooms or full-scale simulated rooms; eligible outcomes included measured microbiological surface outcomes, linked patient outcomes or directly related operational outcomes. | Permanently fixed wall/ceiling systems, handheld spot-disinfection devices, air-only interventions, high-level reprocessing of critical/semicritical devices, and dosimetry/model-only studies without a directly measured microbiological surface endpoint were excluded.
Review framework | Scoping review using PCC and PRISMA-ScR; descriptive evidence mapping without statistical pooling. | R1, R2 and R3 are evidence strata, not a formal hierarchy of quality.
Final included corpus | 70 distinct studies represented by 71 reports: 43 R1, 17 R2 and 10 R3 studies. | One included report is a companion secondary analysis of BETR Disinfection and is not counted as an independent study.
Protocol / registration | No prospective protocol was written, registered, or publicly posted. | Reported transparently in the manuscript.
Data charting / extraction | One reviewer performed data charting and extraction for all included studies using the structured evidence map. | No independent duplicate data extraction was performed.
Critical appraisal | No formal critical appraisal or risk-of-bias assessment was performed. | Methodological characteristics were charted descriptively for interpretation.
Review funding | KNOWLEDGE IN CLEANING project, Bando a Cascata PNC ANTHEM, project code PNC0000003. | Funding body had no role in manuscript preparation, interpretation, or writing and did not influence its content.

## PubMed_Search_Query
Structured PubMed search query
Purpose | Bibliographic search for the scoping review of deployable UV systems for environmental surface disinfection in healthcare.
Search date | 2026-09-01
Publication-date interval | 2010-01-01 to 2026-09-01
Records retrieved | 643
Filters | No publication-type or language filters were applied within the Boolean query. After retrieval, the PubMed Review publication-type filter was used to identify and remove 437 review-type records.
Study selection | After removal of 437 review-type records in PubMed, 206 records remained. Two reviewers screened these records and assessed eligibility.
Selection outcome | 71 included reports representing 70 distinct studies; 135 of the 206 records entering reviewer screening were not included after screening/eligibility assessment.
Eligibility boundary | Include portable/mobile automated systems or UV sources positioned by an operator for a defined treatment period; exclude permanently wall/ceiling-mounted installations and handheld spot-disinfection devices. Require a measured microbiological surface endpoint, linked patient outcome, or operational outcome tied to an eligible surface intervention.
Exact query | ( ( ( "Ultraviolet Rays"[Mesh] OR ultraviolet[tiab] OR "ultraviolet C"[tiab] OR "ultraviolet-C"[tiab] OR UVC[tiab] OR "UV-C"[tiab] OR UVGI[tiab] OR "ultraviolet germicidal irradiation"[tiab] OR "germicidal ultraviolet"[tiab] OR "pulsed xenon"[tiab] OR "pulsed-xenon"[tiab] OR "far-UVC"[tiab] OR "far UVC"[tiab] OR ("UV-C"[tiab] AND 222[tiab]) OR (UVC[tiab] AND 222[tiab]) OR ("far-UVC"[tiab] AND 222[tiab]) OR ("far UVC"[tiab] AND 222[tiab]) OR "222 nm"[tiab] OR "222-nm"[tiab] OR "254 nm"[tiab] OR "254-nm"[tiab] OR "265 nm"[tiab] OR "265-nm"[tiab] ) AND ( "Hospitals"[Mesh] OR "Health Facilities"[Mesh] OR hospital*[tiab] OR healthcare[tiab] OR "health care"[tiab] OR "patient room"[tiab] OR "patient rooms"[tiab] OR "isolation room"[tiab] OR "isolation rooms"[tiab] OR ward*[tiab] OR ICU[tiab] OR "intensive care"[tiab] OR "operating room"[tiab] OR "operating rooms"[tiab] OR "operating theatre"[tiab] OR "operating theatres"[tiab] OR "long-term care"[tiab] OR "nursing home"[tiab] OR "nursing homes"[tiab] OR "clinical laboratory"[tiab] OR "clinical laboratories"[tiab] OR radiolog*[tiab] OR endoscop*[tiab] ) AND ( "Fomites"[Mesh] OR "Equipment Contamination"[Mesh] OR surface*[tiab] OR fomite*[tiab] OR "high-touch"[tiab] OR "high touch"[tiab] OR room*[tiab] OR environment*[tiab] OR equipment[tiab] ) AND ( "Disinfection"[Mesh] OR disinfect*[tiab] OR decontaminat*[tiab] OR inactivat*[tiab] OR sanitiz*[tiab] OR sanitise*[tiab] OR sanitization[tiab] OR sanitisation[tiab] OR cleaning[tiab] OR "terminal cleaning"[tiab] OR "terminal disinfection"[tiab] OR "whole-room disinfection"[tiab] OR "whole room disinfection"[tiab] OR "no-touch disinfection"[tiab] ) ) OR ( ( "Ultraviolet Rays"[Mesh] OR ultraviolet[tiab] OR "ultraviolet C"[tiab] OR "ultraviolet-C"[tiab] OR UVC[tiab] OR "UV-C"[tiab] OR UVGI[tiab] OR "ultraviolet germicidal irradiation"[tiab] OR "germicidal ultraviolet"[tiab] OR "pulsed xenon"[tiab] OR "pulsed-xenon"[tiab] OR "far-UVC"[tiab] OR "far UVC"[tiab] OR ("UV-C"[tiab] AND 222[tiab]) OR (UVC[tiab] AND 222[tiab]) OR ("far-UVC"[tiab] AND 222[tiab]) OR ("far UVC"[tiab] AND 222[tiab]) OR "222 nm"[tiab] OR "222-nm"[tiab] OR "254 nm"[tiab] OR "254-nm"[tiab] OR "265 nm"[tiab] OR "265-nm"[tiab] ) AND "terminal disinfection"[tiab] ) ) AND ( "2010/01/01"[Date - Publication] : "2026/09/01"[Date - Publication] )

## Study_Selection
Study selection and evidence classification for the final scoping review
A single structured PubMed search was run on 1 September 2026. Of 643 retrieved records, 437 review-type records were identified and removed in PubMed using the Review publication-type filter, leaving 206 records for screening. Two reviewers completed study selection between 1 and 4 September 2026. Final eligibility includes portable/mobile automated UV systems and operator-positioned sources used for defined treatment cycles; permanently fixed wall/ceiling systems and handheld spot-disinfection devices are outside scope.
Database searched | PubMed | Reported | National Library of Medicine / NIH.
Search date | 2026-09-01 | Reported | Single structured bibliographic search.
Publication-date interval | 2010-01-01 to 2026-09-01 | Reported | Defined in the PubMed query.
PubMed records retrieved | 643 | Identification | Records returned by the structured PubMed query before publication-type filtering.
Selection reviewers | Two reviewers | Reported | Screening and eligibility assessment.
Review-type records removed in PubMed | 437 | Removed before reviewer screening | Identified and removed using the PubMed Review publication-type filter.
Records entering reviewer screening | 206 | Screened | Title/abstract screening and eligibility assessment by two reviewers.
Records not included after reviewer screening / eligibility | 135 | Not included | 206 records screened minus 71 included reports.
Final included reports | 71 | Reports | Represent 70 distinct studies.
Final studies included in scoping review | 70 | Distinct studies | 71 included reports; one is a companion BETR secondary analysis.
R1 routine-care surface effectiveness | 43 | Included studies | 44 R1 reports because Anderson 2018 is a companion BETR report.
R2 healthcare field efficacy | 17 | Included studies | 17 reports.
R3 healthcare-room / full-scale evidence | 10 | Included studies | 10 reports.
Companion report | Anderson 2018 | Included report | Secondary analysis of the same BETR trial as Anderson 2017; not counted as a separate study.
Protocol / registration | None | Reported | No prospective protocol was written, registered, or publicly posted.
Data charting / extraction reviewer | One reviewer | Reported | Structured evidence-map extraction; no independent duplicate data extraction.
Formal critical appraisal | Not performed | Reported | Methodological characteristics were charted descriptively instead.

## Evidence_Map_Summary
Evidence map summary for 70 distinct included studies represented by 71 reports
Stratum | Distinct studies, n | Reports, n | Typical setting/design | Primary question answered | Interpretive emphasis
R1 - Routine-care surface effectiveness | 43 | 44 | Routine healthcare; natural contamination, implementation programs, acquisition/infection studies | Does the deployed intervention reduce environmental burden or patient/program outcomes under routine workflow? | Closest to operational effectiveness; clinical inference varies from randomized trials to uncontrolled before-after programs.
R2 - Healthcare field efficacy | 17 | 17 | Genuine healthcare environments with controlled carriers, inoculated surfaces or mixed laboratory components | Which organism, soil, distance, device or positioning factors modify performance in real healthcare geometry? | Bridges mechanism and field use; strong for efficacy modifiers but generally not patient-level effectiveness.
R3 - Healthcare-room / full-scale performance | 10 | 10 | Unoccupied genuine rooms or full-scale simulated rooms using mobile/operator-positioned sources | How do room geometry, shadowing, distance and positioning determine delivered dose and room-scale microbial reduction? | Translational room-scale evidence; not equivalent to routine-care clinical effectiveness.
Total | 70 | 71 | - | - | -
Additional quantitative descriptors |  |  | Study-level counts; device/source categories may overlap.
R1 studies with clinical/acquisition/program-level HAI outcomes | 19 | 27.1% of all studies; 44.2% of R1 | Primary outcome-domain classification
R1 studies primarily environmental/operational | 24 | 34.3% of all studies; 55.8% of R1 | Primary outcome-domain classification
Studies evaluating pulsed-xenon systems | 23 | 32.9% of all studies | At least one pulsed-xenon system evaluated
Studies using robotic/autonomous configurations | 12 | 17.1% of all studies | Robot/autonomous terminology in device or deployment configuration
Operator-positioned 222-nm far-UVC studies | 1 | 1.4% of all studies | Room-scale far-UVC evidence
Near-UV/UV-A LED robotic studies (340–385 nm) | 1 | 1.4% of all studies | Interpreted separately from germicidal UV-C evidence
Reports with documented direct commercial involvement (minimum) | 38 | 53.5% of 71 reports | Manufacturer/company funding; provision/loan of devices, materials or data; company technical support; or company/developer authorship. Generic unrelated industry disclosures alone were not counted. | Report-level descriptor; “at least” reflects incomplete/unclear funding reporting in some publications. | Interpret commercial involvement as a contextual characteristic, not a formal risk-of-bias score.

## Included_Studies
Included evidence: 70 distinct studies represented by 71 reports
43 R1, 17 R2 and 10 R3 distinct studies. Anderson 2018 is a companion secondary analysis of BETR disinfection and is not counted as a separate study.
First author | Year | Evidence stratum | Healthcare / room setting | Study design | Contamination model | Device / radiation class | Source / wavelength | Configuration | Comparator | Cycle / dose / positioning | Microorganisms | Outcome domains | Key findings | Operational findings | Major limitations | Funding / manufacturer involvement | Report relationship | Main manuscript ref.
Rutala | 2010 | R2 - Field-based surface efficacy | Standard hospital room plus rooms previously occupied by MRSA/VRE patients | Prospective microbiological field study with inoculated carriers and natural contamination | Mixed: inoculated Formica carriers and natural room contamination | Mobile continuous whole-room UV-C | Low-pressure mercury UV-C, approximately 254 nm; Tru-D | Mobile automated unit; reflected-dose endpoint | Baseline and routine terminal cleaning | Vegetative setting about 15 min; sporicidal setting about 50 min | MRSA, VRE, Clostridioides difficile spores | Environmental microbiology; cycle time | Vegetative bacteria decreased by >99.9% in about 15 min; C. difficile spores by 99.8% in about 50 min. In patient rooms, MRSA-positive samples fell from 20.3% to 0.5% and mean total CFU from 384 to 19. | Automated room treatment; device loaned by manufacturer. | Small field component; microbiological surrogate outcomes; no patient infection endpoint. | Research support reported; Tru-D device loaned; authors disclosed consulting relationships. | Primary/independent report | 1
Wong | 2016 | R2 - Field-based surface efficacy | Post-discharge MRSA, VRE and C. difficile isolation rooms | Paired before/after environmental cultures plus parallel carrier comparison of two devices | Mixed: natural contamination and inoculated stainless-steel carriers in challenging locations | Two mobile automated continuous UV-C devices | Steriliz unit and Tru-D SmartUVC; germicidal UV-C | Mobile whole-room units | Before cleaning, after manual cleaning, after UV-C | Manufacturer-defined cycles; sequential evaluation of two devices | MRSA, VRE, C. difficile | Environmental microbiology; device comparison | After UV-C, contaminated rooms decreased to 3.3% for MRSA, 4.9% for VRE and 0% for C. difficile; mean CFU excluding floors fell from 19.6 after manual cleaning to 1.3 after UV-C. Carrier performance differed between devices. | Highlights device-to-device variation under protein load and difficult placement. | Non-random field sampling; carrier findings not equivalent to natural contamination; no clinical endpoint. | No specific commercial funding was reported in the extracted article information. | Primary/independent report | 2
Nerandzic | 2010 | R2 - Field-based surface efficacy | Rooms of hospitalized patients | Laboratory dose study plus paired hospital-room cultures | Mixed: inoculated surfaces and natural contamination | Mobile automated continuous UV-C | Tru-D UV-C system; reflected dose 12,000 or 22,000 µWs/cm² | Mobile whole-room unit | Before versus after UV-C; lower versus higher dose | About 20 min for 12,000 µWs/cm² and 45 min for 22,000 µWs/cm² | C. difficile spores, MRSA, VRE | Environmental microbiology; setup time | Hospital-room positive cultures decreased by 93% for MRSA/VRE and 80% for C. difficile. Higher reflected dose was needed for reliable spore reduction. | Setup required <5 min and no continuous monitoring. | Small field study; no randomized comparator or clinical outcome. | VA grant; Tru-D unit loaned free of charge; company did not fund or write the study. | Primary/independent report | 3
Ali | 2017 | R2 - Field-based surface efficacy | Contaminated hospital patient rooms after terminal cleaning | In-use comparison of two whole-room systems with natural sampling and inoculated indicators | Mixed natural contamination and inoculated indicators with low/high soil | Whole-room continuous UV-C systems | Surfacide Helios triple-emitter 254 nm and a relocatable 265-nm emitter system | Multi-emitter versus single relocatable emitter | Two systems; low versus high soil; direct versus shadowed placement | Manufacturer cycles; longer cycles recommended for C. difficile risk | MRSA, Klebsiella pneumoniae, C. difficile spores; aerobic colony counts | Environmental microbiology; positioning; soil effect | Both systems produced large reductions for MRSA and K. pneumoniae under low soil, but efficacy fell with heavy soil and was inconsistent for C. difficile spores. Careful positioning reduced shadowing. | In-house validation of room layouts and cycle duration was recommended. | Small number of rooms; manufacturer-funded microbiological testing; artificial challenge loads. | Manufacturers funded microbiological testing; one author reported an advisory role with 3M. | Primary/independent report | 4
Yui | 2025 | R2 - Field-based surface efficacy | Hospital room and bathroom | In-vitro challenge plus in-use sequential sampling | Mixed: inoculated indicators with low/high soil and natural aerobic colony counts | Rapid targeted continuous UV-C system | 254-nm emitters mounted on articulating arms | Targeted mobile device; multiple positions | Before cleaning, after manual cleaning, after UV-C; low versus high soil | 90–600 s per position | MRSA, K. pneumoniae, C. difficile spores; aerobic flora | Environmental microbiology; rapid cycle | MRSA/K. pneumoniae reductions were 2.97–4.87 log under low soil and 0.53–3.63 log under high soil. C. difficile reduction was limited, especially with synthetic faeces. Natural contamination persisted on 30% of sampled sites after UV-C. | Designed for rapid targeted use; performance depended on positioning and soil. | Small in-use sample; mixed natural and artificial outcomes; online publication in 2024 / issue year 2025. | No funding; manufacturer supplied devices free of charge and was not involved in study conduct. | Primary/independent report | 49
Kovach | 2017 | R1 - Routine-care surface effectiveness | Single 160-bed nursing home / long-term care facility | Uncontrolled before-after study with environmental sampling and clinical surveillance | Natural environmental contamination | Pulsed-xenon UV whole-room disinfection | Broad-spectrum pulsed xenon UV | Mobile no-touch unit after discharge | 36-month baseline versus 12-month intervention; manual cleaning versus manual + UV | Post-discharge protocol; manufacturer cycle | Environmental bacteria; facility-acquired infections | Environmental microbiology; infection rates; hospitalizations | Positive environmental cultures fell from 33% at baseline to 7.1% after cleaning and 4.4% after UV. Total infections, urinary, respiratory and skin/soft-tissue infections and infection-related hospitalizations decreased during intervention. | Applied after resident discharge in one facility. | Single-site uncontrolled pre-post design; temporal confounding and regression to the mean possible. | Jewish Home Foundation; funder reported no role; no competing interests declared. | Primary/independent report | 36
Boyce | 2016 | R3 - Healthcare-room or full-scale simulation | Unoccupied ICU room and surgical ward room | Prospective dosimetry and inoculated-carrier observational study | Artificial inoculated stainless-steel disks | Mobile continuous UV-C | Spectra 254 UV-C device | Single mobile emitter | Direct versus shaded locations; distances; 5- versus 15-min cycles | 5 and 15 min with six UV sensors | MRSA, VRE, C. difficile spores | Dose distribution; environmental microbiology | Irradiance varied by orders of magnitude with distance and orientation; log reductions ranged from >4 in directly exposed sites to minimal or no reduction at some distant/shaded sites, especially for C. difficile. | Demonstrates need for dose measurement and multiple placements. | Empty-room simulation with artificial carriers; not routine clinical deployment. | Device provided by Spectra 254; consulting/research relationships were disclosed. | Primary/independent report | 5
Carlisle | 2023 | R1 - Routine-care surface effectiveness | Forty hospital rooms after manual cleaning | Randomized room-level comparison | Residual natural healthcare-associated pathogens | Mobile continuous UV-C versus electrostatic hypochlorite spray | UVDI-360 Room Sanitizer; germicidal UV-C | Mobile whole-room device | UV-C versus sodium hypochlorite via electrostatic sprayer | Manufacturer cycle; room-ready time recorded | C. difficile, MRSA and other healthcare-associated pathogens | Environmental microbiology; room turnaround | Both adjunctive methods significantly reduced residual pathogens, with no statistically significant efficacy difference. Median/mean room-ready time was longer for UV-C (23.8 vs 20.4 min). | Direct comparative evidence on workflow and alternative no-touch technology. | Pilot size and insufficient power to exclude meaningful differences; environmental rather than clinical outcome. | Veterans Affairs and gift funding; relevant research grants/consulting relationships disclosed. | Primary/independent report | 70
Jelden | 2017 | R3 - Healthcare-room or full-scale simulation | Hospital biocontainment room | Controlled room experiment | Artificial inoculated coupons | UVGI generator plus UV-reflective paint | TORCH continuous UV-C; UVC Max reflective coating | Portable UVGI generator; standard versus reflective walls | Standard paint versus UV-reflective paint; direct versus indirect sites | 10-min cycle | Nosocomial bacteria including MRSA and VRE | Environmental microbiology; reflected dose | Reflective paint improved decontamination at indirect/shadowed locations, with reductions exceeding 4 log at some sites versus <2 log without reflective coating. | Passive room-surface modification may improve dose distribution. | Controlled challenge, not natural contamination. | University support; TORCH device loaned and reflective paint donated; suppliers were not involved in study conduct. | Primary/independent report | 58
Napolitano | 2015 | R1 - Routine-care surface effectiveness | Community hospital, facility-wide | Pre-post intervention pilot | Clinical HAI surveillance | Mobile continuous UV-C delivered as a contracted service | 254-nm UV-C | Mobile units with trained dedicated personnel and coverage monitoring | Pre-intervention versus intervention period | Frequent monitored cycles; coverage increased and maintained | Healthcare-associated infections, multiple pathogens | Clinical HAI rates; implementation coverage | HAI incidence decreased by 1.3 per 1000 patient-days (34.2%); total number/incidence proportion decreased after higher treatment coverage. | Dedicated service model linked treatment coverage to outcomes. | Uncontrolled pre-post design; no concurrent control; possible co-interventions and secular trends. | Hollywood Community Hospital/PHCMC support; W. Tang reported funding from Clean Sweep Group Inc.; the funding body reported no role in study design, data collection or analysis. | Primary/independent report | 33
Casini | 2023 | R1 - Routine-care surface effectiveness | Operating theatres, ICU isolation room and ward rooms | Before-after high-touch-surface sampling during an enhanced cleaning protocol | Natural environmental contamination | Mobile UV-C disinfection robot | Continuous UV-C; Hyper Light platform | Robot used after standard operating procedure | Before SOP, after SOP and after UV-C | Short terminal cycle; dosimeters at sampled sites | Total viable counts / environmental flora | Environmental microbiology; compliance with hygiene thresholds; traceability | Positive sites decreased from 64.3% after SOP to 17.5% after UV-C; non-compliant sites decreased from 9.3% to 1.2%. | Automatic recording of disinfection parameters enabled process traceability; staff training was emphasized. | No concurrent control or patient outcome; setting-specific thresholds; before-after sampling. | No external funding; authors declared no conflicts of interest. | Primary/independent report | 67
Casini | 2019 | R1 - Routine-care surface effectiveness | Hospital operating theatres and other critical areas | Paired before/after environmental surface study | Natural high-touch surface contamination | Pulsed-xenon UV-C | Pulsed xenon broad-spectrum UV including UVC | Mobile no-touch device after SOP | After SOP versus SOP plus PX-UV | Protocol-defined short cycles | Total viable counts, KPC-producing organisms and C. difficile | Environmental microbiology; hygiene compliance | Across 345 sites, positive samples fell from about 63% after SOP to 18% after PX-UV; operating-room samples met the standard after adjunctive treatment. | Short adjunctive step in critical areas. | Before-after environmental study; no clinical endpoint; same research group as implementation study. | Supported by AB Medica s.r.l.; the company made the pulsed-xenon devices available free of charge and was reported not to be involved in study design, analysis or presentation; authors declared no conflicts. | Primary/independent report | 71
Kitagawa | 2021 | R1 - Routine-care surface effectiveness | Hospital rooms with C. difficile contamination | Paired environmental sampling | Natural C. difficile surface contamination | Pulsed-xenon UV | PX-UV mobile device | Adjunct to manual terminal cleaning | Before cleaning, after manual cleaning, after PX-UV | Standard PX-UV cycle | Clostridioides difficile | Environmental microbiology | Both bleach cleaning and PX-UV reduced contamination; PX-UV provided additional reduction when added to non-sporicidal or routine terminal cleaning. | Useful as an adjunct when chemical cleaning is not reliably sporicidal. | Environmental surrogate; small single-centre study; contamination prevalence may limit power. | Partly funded by Terumo Co., Ltd.; Terumo was reported not to have a role in design/data/manuscript; Hiroki Ohge reported Terumo research funding. | Primary/independent report | 72
Füszl | 2021 | R2 - Field-based surface efficacy | Waiting areas of two hospital outpatient clinics | Field usability/effectiveness study plus controlled C. auris carriers | Mixed: natural surface contamination and inoculated C. auris carriers | Autonomous UV-C disinfection robot | Continuous UV-C | Autonomous mobile robot after routine cleaning | Before and after robot; growth phase and load of C. auris | Manufacturer standard cycle | Natural flora and Candida auris | Environmental microbiology; usability; autonomy | Natural contamination decreased, but the standard cycle did not reliably eliminate stationary-phase/high-load C. auris or shielded/rim locations. | Interface was not intuitive, required technical training and frequent intervention; routine autonomy was limited. | Small field study; mixed natural and artificial outcomes; usability was device-specific. | No specific grant funding; Clean Room Solutions did not fund the study; competing interests were reported as none. | Primary/independent report | 50
Casini | 2025 | R2 - Field-based surface efficacy | Real hospital high-touch surfaces plus controlled laboratory tests | Mixed in-vitro and hospital before/after study | Mixed artificial inocula and natural total viable counts | Experimental autonomous UV-C robot | Low-pressure mercury 254-nm UV-C | Mobile robotic platform | Dirty condition, SOP alone and SOP plus UV-C | Mean dose on sampled hospital surfaces 29.31 mJ/cm² | S. aureus, P. aeruginosa, adenovirus, coronavirus 229E and natural flora | Environmental microbiology; delivered dose | In hospital use, average total viable-count reduction was 67% after SOP and 96% after SOP plus UV-C. Higher dose was required in the hospital than in vitro. | Directly demonstrates the laboratory-to-field performance gap. | Small number of surfaces/rooms; experimental platform; no clinical outcome. | No funding source reported; authors declared no conflict of interest in the article. | Primary/independent report | 51
Villacís | 2019 | R2 - Field-based surface efficacy | Patient rooms and operating-room environments in an Ecuadorian hospital | In-vitro validation plus paired environmental sampling | Mixed artificial challenge and natural high-touch contamination | Pulsed-xenon UV | PX-UV mobile device | Mobile unit after terminal manual cleaning | After manual cleaning versus after PX-UV | Standard cycle; 5-min laboratory exposure | Environmental bacteria | Environmental microbiology | Environmental CFU decreased from 3569 after manual cleaning to 889 after PX-UV (75% reduction); reductions were about 87% in operating rooms and 76% in patient rooms. | Demonstrated applicability in a resource setting outside the USA. | Before-after study without clinical outcomes; manufacturer-affiliated author present. | One author was affiliated with the device company; commercial involvement was disclosed. | Primary/independent report | 55
Raggi | 2018 | R1 - Routine-care surface effectiveness | 377-bed community hospital, facility-wide | Pre-post clinical cohort with operational and cost analysis | Clinical HAI surveillance | Mobile continuous UV-C terminal disinfection | 254-nm UV-C | Facility-wide terminal disinfection service | 12-month pre-intervention versus 12-month intervention | 8449 room cycles; mean cycle about 10.1 min | Five MDRO-related HAIs | Clinical HAI rate; ED throughput; cost | HAI incidence decreased from 4.87 to 3.94 per 1000 patient-days (19.2%; p=.006). Reported direct cost savings were USD 1.22 million over 12 months. | No significant adverse change in emergency-department admission time. | Uncontrolled pre-post design; cost estimate did not fully subtract intervention costs; unmeasured confounding. | Supported by PHCMC; commercial affiliations/relationships were disclosed in the article. | Primary/independent report | 38
Morikane | 2020 | R1 - Routine-care surface effectiveness | Open-style ICU in a Japanese tertiary hospital | Before-after clinical and environmental study | Natural environmental contamination and MDRO acquisition/infection surveillance | Pulsed-xenon UV | Portable PX-UV device | Mobile unit added to terminal cleaning | Pre-intervention versus intervention | Added about 20 min to bed turnaround | MRSA and drug-resistant Acinetobacter; total colony counts | Clinical MDRO rates; environmental microbiology; workflow | MRSA incidence decreased from 13.8 to 9.9 per 10,000 patient-days (IRR 0.71); drug-resistant Acinetobacter from 48.5 to 18.1 (IRR 0.37). Colony counts fell further after PX-UV. | Light/sound initially disturbed staff and equipment; blackout curtains, goggles and earplugs enabled workflow integration. | Single-centre before-after design; co-interventions and time trends possible. | No research funding except provision of Rodac plates; authors reported no competing interests. | Primary/independent report | 41
Kitagawa | 2020 | R1 - Routine-care surface effectiveness | High-touch surfaces in a Japanese hospital | Paired before/after environmental study | Natural MRSA and Acinetobacter contamination | Pulsed-xenon UV | PX-UV device | Mobile unit after manual cleaning | Before versus after PX-UV | Median in-room time about 20 min | MRSA and Acinetobacter baumannii | Environmental microbiology; cycle time | A. baumannii CFU decreased from 14.4 to 1.7 and MRSA from 1.1 to 0.3; positivity also fell substantially. | Provides routine cycle-time data. | Single-centre environmental outcome; paired sites but no clinical endpoint. | Partly funded by Terumo; Hiroki Ohge reported a Terumo research grant. | Primary/independent report | 73
Jelden | 2016 | R3 - Healthcare-room or full-scale simulation | Unoccupied hospital/biocontainment room | Controlled hospital-room challenge experiment | Artificial inoculated coupons on multiple materials/locations | Portable UVGI generator | TORCH continuous UV-C | Single portable generator | Direct versus indirect sites, material and distance | 10-min cycle with measured dose 0–688 mJ/cm² | MRSA and VRE | Log reduction; dose distribution | Average reductions were about 4.6 log for MRSA and 3.9 log for VRE, with better results in line-of-sight locations and material-dependent variation. | Maps performance across a furnished room. | Artificial high-load carriers in an empty room; not routine care. | Academic support; device loaned by supplier; supplier reported not to be involved in design/results. | Primary/independent report | 59
McGinn | 2022 | R1 - Routine-care surface effectiveness | CT treatment room in a hospital | Comparative field study using natural contamination | Natural surface contamination | Autonomous UVGI robot | Akara robotic UV-C platform | Autonomous multi-position mobile robot | Robot protocol versus current manual disinfection | <10-min procedure; mean delivered dose 13.01 ± 4.36 mJ/cm² | Natural environmental flora | Environmental microbiology; procedure time | Both manual and robotic protocols significantly reduced contamination; the robot showed a larger effect size in this study and completed a protocol in under 10 min. | Robot procedure was substantially faster than the existing manual protocol in the evaluated CT room. | Small single-room study; natural counts were low/variable; developer affiliation with Akara Robotics. | Science Foundation Ireland grant; several authors were affiliated with/shareholders in Akara Robotics. | Primary/independent report | 68
McGinn | 2021 | R1 - Routine-care surface effectiveness | CT and X-ray rooms in two Irish hospitals | User-centred robotic feasibility and field microbiology study | Natural surface contamination | Robot-assisted UVGI | Akara autonomous UV-C platform | Mobile robot with planned irradiation positions | Pre/post UV and workflow observations | <15-min deployment | Natural environmental flora | Environmental microbiology; feasibility; human factors | All measurable microbial load was inactivated on 22 of 24 tested surfaces in one deployment series, supporting rapid radiology-room use. | Detailed robot design, navigation, task planning and user-centred deployment considerations. | Small feasibility study; low baseline bioburden; developer involvement. | Science Foundation Ireland grant; two authors reported involvement with a robotics company developing disinfection systems. | Primary/independent report | 69
Różańska | 2023 | R2 - Field-based surface efficacy | Six hospital rooms/wards plus laboratory conditions | Mixed laboratory carrier and hospital before/after study | Artificial carriers plus natural touch-surface contamination | Mobile UV-C robots | OCTA UVC robots, 253.7 nm | Mobile robotic units; 10-min irradiation | Before versus after cleaning/UV; material and location | 10 min | ESKAPE organisms and Candida auris; natural hospital flora | Log reduction; environmental microbiology; material effect | Laboratory reductions ranged from about 1 to 6 log and varied by organism/material/location. Hospital sampling confirmed a marked decrease in bacterial contamination after UV-C. | Supports location-specific validation and adjunctive use. | Mixed designs and outcome metrics; hospital sample size limited; some results based on artificial high loads. | Research cooperation agreement between Jagiellonian University Medical College and Eco-Light Bio Safety; OCTA robots were provided under the collaboration; authors disclosed the funding relationship. | Primary/independent report | 52
Sathitakorn | 2023 | R1 - Routine-care surface effectiveness | Four intensive care units in Thailand | Prospective environmental protocol evaluation with intervention and comparison rooms | Natural Gram-negative contamination | Pulsed-xenon UV plus real-time cleaning feedback | PX-UV mobile device | Adjunct to optimized environmental cleaning | Cleaning protocol with versus without PX-UV | Mean in-room time 15.6 min | Gram-negative microorganisms | Environmental microbiology; workflow feasibility | In intervention rooms, Gram-negative-positive rooms decreased from 40% at baseline to 20% after manual cleaning and 0% after PX-UV; mean contamination fell from 24.1 to 7.5 CFU/cm² after manual cleaning and to 0 after PX-UV. | Judged feasible in a resource-limited setting without compromising room turnover. | Environmental outcome; bundled real-time feedback makes the isolated PX-UV contribution difficult to quantify. | Supported by Terumo and the Thailand Science Research and Innovation Fundamental Fund; authors declared no conflicts. | Primary/independent report | 74
Jinadatha | 2015 | R1 - Routine-care surface effectiveness | Fourteen rooms immediately after discharge of MRSA-positive patients | Paired before/after study without manual cleaning | Natural MRSA contamination | Pulsed-xenon UV | PX-UV mobile no-touch system | Mobile room unit | Before versus after PX-UV alone | Standard device cycle | MRSA | Environmental microbiology | MRSA colonies decreased from 393 before PX-UV to 100 after treatment across 70 positive surface samples (p<.01). | Isolates the no-touch effect from manual cleaning, but does not represent recommended routine practice. | Only 14 rooms; no manual-cleaning comparator; residual contamination remained. | VA Merit Review/Central Texas VA support; laboratory activity supported by a Xenex grant; device owned by the VA. | Primary/independent report | 75
Simmons | 2018 | R1 - Routine-care surface effectiveness | Operating rooms in 23 US hospitals | Multicentre pooled before/after environmental validation | Natural OR surface contamination after manual cleaning | Pulsed-xenon UV | Xenex PX-UV | Mobile no-touch device | After manual cleaning versus after PX-UV | Site protocols | Aerobic environmental flora | Environmental microbiology | Mean CFU after manual cleaning ranged from 5.8 to 34.37 across sites and fell to 0.69–6.43 after PX-UV; positive surfaces decreased from 67% to 38%. | Large multicentre operational dataset. | Pooled validation without standardized control; all authors were affiliated with the device company. | Device-company funded study; authors were affiliated with Xenex Disinfection Services. | Primary/independent report | 76
Su | 2021 | R1 - Routine-care surface effectiveness | COVID-19 isolation ward and ICU rooms with infected patients | Natural environmental sampling followed by controlled in-room UVC exposure | Natural SARS-CoV-2 environmental contamination | Mobile continuous UV-C | 254-nm UVC | Mobile room irradiation | Different irradiation durations at positive sites | 5–30 min; 126 mJ/cm² at 15 min and 252 mJ/cm² at 30 min | SARS-CoV-2 RNA/infectivity proxy | Environmental viral detection; dose/time | Among 12 positive sites, positivity fell from 41.7% pre-UV to 25% after 5 min, 8.3% after 10 min and 0% after 15 min; heavily contaminated washbasins required longer exposure. | Provides in-hospital dose/time evidence during COVID care. | Small number of positive sites; PCR-based endpoints may not equal infectivity; no patient outcome. | Taipei Tzu Chi Hospital/Buddhist Tzu Chi Medical Foundation grant; funder reported no role; authors declared no conflicts. | Primary/independent report | 77
Yang | 2019 | R2 - Field-based surface efficacy | Hospital rooms plus laboratory pathogen tests | Mixed laboratory and in-use before/after study | Artificial pathogens and natural hospital-room contamination | Mobile automatic UV-C robot | Hyper Light P3, 254 nm | Mobile robot | Before/after UV; distance and organism | Typically 15-min in-use cycle | MDR bacteria, M. abscessus, Aspergillus and natural flora | Log reduction; environmental microbiology | The system produced >3-log reduction for most vegetative organisms at practical distances, but VRE/M. abscessus and Aspergillus were less susceptible. Hospital-room contamination decreased after a 15-min cycle. | Performance decreased with distance and organism resistance. | Mixed laboratory and field outcomes; no clinical endpoint. | Funded by Mediland Enterprise Corporation; authors declared no conflicts of interest. | Primary/independent report | 56
Bosco | 2022 | R1 - Routine-care surface effectiveness | Three hospital operating theatres | Before-after model-driven environmental study | Natural environmental contamination | Mobile UVC device | Continuous UV-C | Mobile bedside positions | Contamination before and after 3- or 5-min UVC exposure under different initial cleaning levels | 3–5 min per bedside/position | Total aerobic colony counts | Environmental microbiology; protocol standardization | Mean CFU reduction was about 72%; 3- and 5-min exposure did not differ materially in the evaluated conditions. | Model used to standardize a rapid adjunctive protocol between procedures. | Non-random before-after design; total counts rather than target pathogens; limited generalizability. | Funded by UVDI through the University of Siena; two authors had an ownership relationship with a company receiving UVDI funds; UVDI reported no role in design/data/manuscript. | Primary/independent report | 78
Pavia | 2018 | R1 - Routine-care surface effectiveness | Pediatric long-term care facility | Before-after incidence study | Clinical viral infection surveillance | Continuous UV-C room disinfection | Clorox Healthcare Optimum-UV system | Single mobile device rotated among rooms | Pre-intervention versus 12-month UV-C period | Rooms treated approximately three times weekly according to program | Viral infections | Clinical infection incidence | UV-C implementation was associated with a 44% reduction in viral infection incidence (IRR 0.56; 95% CI 0.37–0.84). | One device was deployed across selected rooms on a scheduled basis. | Uncontrolled before-after design; company authors and device donation; no environmental sampling. | Clorox provided the device; two authors were employees of The Clorox Company. | Primary/independent report | 37
Tholany | 2023 | R3 - Healthcare-room or full-scale simulation | Empty hospital rooms | Controlled room-level challenge pilot | Artificial inoculated agar plates with Gram-negative rods | Multi-emitter continuous UV-C | Surfacide Helios three-tower system, 254 nm | Three mobile emitters | UV-exposed versus unexposed control plates; distance/location | Median runtime about 29 min | Pathogenic enteric and non-fermenting Gram-negative rods | Growth arrest / log reduction; distance failures | Mean recovery was 0.4 CFU on exposed plates versus 87.4 CFU on controls (about 2.35-log reduction); 7.1% of exposed plates showed growth, more often at distant/shielded sites. | Identifies shower rails and distant locations as potential failures. | Artificial plates in empty rooms; pilot sample; no natural contamination. | Department of Veterans Affairs Quality Enhancement Research Initiative grant; no competing interests reported. | Primary/independent report | 64
de la Rosa-Zamboni | 2025 | R1 - Routine-care surface effectiveness | Neonatal intensive care unit with high ESBL burden | Four-period before-after bundle study | Clinical HCAI surveillance | UV-C disinfection plus subsequent copper surface plating | UV-C device; detailed source parameters not central in article | Terminal environmental intervention within a multi-component bundle | Sequential periods: hand hygiene/bundles, +UV-C, +copper, maintenance | Protocol-defined terminal cycles | HCAIs, including ESBL-producing bacteria | Clinical HCAI rates | Across 552 neonates and 15,467 patient-days, HCAIs decreased from 11.03 to 5.35 per 1000 patient-days; bacterial-isolate HCAIs from 5.39 to 1.79. Combined UV-C/copper association RR was 0.49. | Implemented sequentially in a high-risk NICU. | UV-C cannot be separated from hand-hygiene bundles and copper; non-random sequential periods. | Hospital-funded work; UV-360 equipment and copper materials were donated by commercial suppliers, which were reported not to influence the study. | Primary/independent report | 43
Arkusz | 2024 | R3 - Healthcare-room or full-scale simulation | Laboratory and a furnished 1:1 simulated hospital room | Laboratory and full-scale room experiment | Artificial E. coli and S. epidermidis inocula | Direct-type commercial UV-C lamps | Wall/direct UV-C radiators with varied power | Stationary operator-positioned open-type UV-C source in a full-scale simulated room | Power, distance, exposure time, line of sight and furnishing | Up to 60 min; distances 10–300 cm | E. coli and S. epidermidis | Dose and microbial reduction | In the furnished room, reduction declined with distance and shadowing; >90% required very short distance or high power/long exposure. Areas under the bed received little or no effective radiation. | Supports mobile/self-steering systems that reduce source-to-surface distance. | Authors call conditions “real-world,” but the room was a 1:1 simulation rather than routine patient care. | Polish National Centre for Research and Development grant; funder reported no role; authors declared no competing interests. | Primary/independent report | 62
Ishikawa | 2025 | R1 - Routine-care surface effectiveness | NICU/GCU during an MRSA outbreak | Single-centre before-after enhanced-cleaning bundle study | Clinical MRSA incidence and natural environmental cultures | Pulsed-xenon UV | PX-UV terminal disinfection | Mobile device within daily cleaning/hand-hygiene bundle | Before versus after enhanced bundle | Protocol-defined terminal use | MRSA | Clinical infection incidence; environmental cultures; safety | MRSA infection decreased from 2.81 to 0.90 per 1000 patient-days and environmental positives from 12/137 to 0; no neonatal adverse events were reported. | Implemented safely in NICU/GCU. | Single hospital under outbreak conditions; simultaneous daily cleaning and improved hand hygiene confound UV effect. | No specific external funding; authors declared no conflicts of interest. | Primary/independent report | 44
Kitagawa | 2021 | R1 - Routine-care surface effectiveness | Hospital intervention units in Japan | Quasi-experimental before-after study with Poisson regression | Clinical MRSA acquisition surveillance | Pulsed-xenon UV | Xenex PX-UV | Mobile terminal disinfection device | Pre-intervention versus PX-UV intervention units/period | Routine terminal cycle; about 86% eligible-room coverage | MRSA | Clinical acquisition rate; coverage | MRSA acquisition decreased from 3.56 to 2.21 per 1000 patient-days; adjusted IRR 0.556 (95% CI 0.309–0.999). | High but incomplete room coverage was documented. | Non-randomized intervention; borderline confidence interval; possible secular/co-intervention effects. | Hiroki Ohge reported a Terumo research grant; other authors reported no relevant conflicts. | Primary/independent report | 42
Kitagawa | 2021 | R1 - Routine-care surface effectiveness | Hospital rooms and bathrooms in Japan | Paired before/after environmental study | Natural VRE surface contamination | Pulsed-xenon UV | PX-UV mobile device | Adjunct to manual terminal cleaning | Before cleaning, after manual cleaning and after PX-UV | Standard terminal cycle | Vancomycin-resistant Enterococci | Environmental microbiology | Positive samples decreased from 22/60 before cleaning to 14/60 after manual cleaning and 0/60 after PX-UV; bidet toilet nozzles were notable reservoirs before adjunctive treatment. | Identifies bathrooms and bidet nozzles as priority targets. | Small single-centre environmental study; no clinical endpoint. | No specific grant funding reported; authors declared no competing interests. | Primary/independent report | 79
Schaffzin | 2020 | R1 - Routine-care surface effectiveness | High-occupancy pediatric hospital isolation rooms | Implementation time-series / program evaluation | Room-coverage metrics and HAI surveillance | Pulsed-xenon UV robots | Xenex LightStrike | Two-robot fleet with dedicated technicians | Program phases before and after workflow redesign | Shift-based room targets; coverage monitored monthly | HAIs including MDRO/C. difficile context | Coverage, workflow and HAI rates | Six-month averages reached 85.7% of isolation rooms and 87.7% of priority rooms; HAI rates decreased 16.2% after implementation. | A dedicated team, environmental-services ownership and shift goals improved coverage without expanding the robot fleet. | Program evaluation cannot establish causality between UV and HAI reduction; other cleaning initiatives occurred. | No explicit commercial study funding reported; device brand and operational deployment were disclosed. | Primary/independent report | 40
Russo | 2021 | R2 - Field-based surface efficacy | Two outpatient rooms in Perugia Hospital plus laboratory tests | Mixed field and in-vitro study | Natural high-touch contamination plus artificial MDR inocula | Mobile UV-C robotic system | R2S continuous UV-C | Last-generation mobile robot | Manual sanitization versus R2S; surface materials and photoreactivation | Protocol-defined static/dynamic irradiation | S. aureus, MRSA, P. aeruginosa, KPC K. pneumoniae; bacteria, molds and yeasts | Environmental microbiology; log reduction; photoreactivation | R2S reduced bacteria, molds and yeasts at each high-touch site compared with manual sanitization; laboratory reductions reached about 5–6 log but varied by material. | Dynamic robotic use intended to reduce human error and chemical use. | Only two clinical rooms; environmental and laboratory findings combined; device supplier provided equipment/technical support. | Public foundation grant; R2S device supplied by Bazzica Engineering; authors declared no conflicts. | Primary/independent report | 53
Chen | 2020 | R2 - Field-based surface efficacy | Central, microbiology and immunology laboratories in a hospital | Mixed laboratory biofilm and real clinical-laboratory environmental study | Artificial biofilm-forming isolates plus natural air/surface bioburden | Pulsed-xenon UV | PX-UV device | Mobile unit | Before manual cleaning, after manual cleaning and after 6-min PX-UV | 6 min in clinical laboratories | P. aeruginosa, S. aureus, K. pneumoniae and natural laboratory flora | Biofilm-former reduction; air/surface microbiology | PX-UV significantly reduced biofilm-forming pathogens, reduced air bioburden and eliminated detectable bioburden on sampled high-touch surfaces after treatment. | Short cycle in clinical laboratory areas. | No laboratory-acquired infection endpoint despite conclusion; field sample size and control details limited. | Regional/public scientific grants; no commercial funding reported. | Primary/independent report | 54
Lorca-Oró | 2022 | R3 - Healthcare-room or full-scale simulation | 11-m² simulated hospital ICU room | Controlled viral room-simulation study | Artificial desiccated SARS-CoV-2 on surfaces | Mobile autonomous UV-C robot | ASSUM robot, conventional UV-C | Autonomous mobile unit | 12- versus 20-min cycle and direct versus shielded sites | 12 and 20 min | SARS-CoV-2 | Infectious viral reduction; location effect | Direct wall sites achieved ≥99.99% inactivation after 12–20 min; floor and under-table sites received lower dose and showed less consistent reduction. | Autonomous platform and room mapping were demonstrated. | Simulated room with artificial virus; developer/shareholder conflicts disclosed; not routine clinical care. | Academic/public project support; autonomous-system developers included authors and one author was a company shareholder. | Primary/independent report | 63
Anderson | 2013 | R1 - Routine-care surface effectiveness | 39 post-discharge patient rooms in two tertiary hospitals | Prospective multicentre before-after environmental surface study | Natural VRE, C. difficile and Acinetobacter contamination | Mobile automated continuous UV-C | Tru-D SmartUVC, 254 nm | Mobile whole-room reflected-dose system | Before versus after UV-C before routine terminal cleaning | 12,000 or 22,000 µWs/cm² reflected dose; median 25/45 min | VRE, C. difficile, Acinetobacter spp. | Environmental microbiology; line of sight; cycle duration | Target-pathogen CFU fell 1.35 log10 overall; VRE 1.68 log10; C. difficile 1.16 log10; reductions occurred at all nine sampled sites. | Median vegetative/spore cycles 25/45 min. | Convenience sample; small Acinetobacter sample; no patient outcome. | Public/academic support; Tru-D device used; authors disclosed relevant consulting relationships. | Primary/independent report | 20
Anderson | 2017 | R1 - Routine-care surface effectiveness | Nine hospitals; targeted post-discharge patient rooms | Pragmatic cluster-randomized multicentre crossover trial | Routine clinical exposure of subsequent room occupants | Mobile continuous UV-C adjunct to terminal cleaning | Mercury UV-C | Mobile whole-room terminal-disinfection device | Reference vs UV; bleach vs bleach+UV | Four 7-month strategy periods per hospital | MRSA, VRE, C. difficile, multidrug-resistant Acinetobacter | Acquisition/infection; clinical effectiveness | UV plus standard cleaning reduced composite target-organism acquisition/infection (RR 0.70, 95% CI 0.50–0.98); UV added to bleach did not reduce C. difficile infection (RR 1.00, 95% CI 0.57–1.75). | Large pragmatic multicentre deployment. | No simultaneous room-level dosimetry/environmental mediator measurement; effect varied by strategy. | US Centers for Disease Control and Prevention; investigators disclosed Clorox consulting/advisory relationships; companies provided materials but had no role in trial funding/design/analysis. | Primary BETR Disinfection trial report | 26
Boyce | 2011 | R2 - Field-based surface efficacy | 25 post-discharge patient rooms | Prospective patient-room before/after study with natural surface cultures plus a C. difficile carrier test | Natural aerobic surface contamination plus controlled C. difficile spore carriers | Mobile automated continuous UV-C | Germicidal UV-C | Mobile unit; one- or two-stage positioning | Before versus after UV-C; one- versus two-stage procedure | One-stage 34.2-100.1 min; two-stage 72.1-146.3 min | Aerobic flora; Clostridioides difficile spores | Environmental microbiology; line of sight; cycle time; ozone | All five high-touch surfaces showed significant reductions. One-stage C. difficile reductions were 1.8-2.9 log10; the two-stage procedure reduced mean aerobic counts to 0 CFU on all five sites. | Two-stage positioning improved coverage but lengthened room downtime. | No patient outcome; mixed natural and carrier endpoints; long cycle times. | Funding/manufacturer support was not clearly specified in the available article metadata. | Primary/independent report | 45
Havill | 2012 | R2 - Field-based surface efficacy | 15 patient rooms in a 500-bed teaching hospital | Prospective comparative room study of hydrogen peroxide vapor versus UV-C | Natural aerobic surface contamination plus C. difficile and biological-indicator carriers | Mobile whole-room UV-C | Germicidal UV-C | Whole-room mobile unit | Hydrogen peroxide vapor versus UV-C | Mean cycle 73 min for UV-C versus 153 min for HPV | Aerobic flora; C. difficile spores; Geobacillus stearothermophilus indicators | Environmental microbiology; carrier reduction; line of sight; cycle time | UV-C reduced room contamination but was less effective than hydrogen peroxide vapor, particularly for spores and sites outside direct line of sight; mean C. difficile reduction was about 2 log10 with UV-C. | UV-C was faster than hydrogen peroxide vapor. | Comparative environmental study without patient outcomes; carrier challenge may not reflect natural burden. | Funding/manufacturer support was not clearly specified in the available article metadata. | Primary/independent report | 46
Rutala | 2013 | R3 - Healthcare-room or full-scale simulation | Hospital patient-room decontamination experiment | Controlled room-scale evaluation of UV-reflective wall coating | Controlled microbial room challenge | Mobile continuous UV-C plus UV-reflective coating | Tru-D SmartUVC; germicidal UV-C | Mobile room unit; standard versus reflective walls | Standard wall surface versus UV-reflective coating | Time-to-target-dose comparison | MRSA; Clostridioides difficile spores | Room-scale decontamination time; microbial endpoint | Reflective coating shortened MRSA decontamination from 25 min 13 s to 5 min 3 s and C. difficile spore decontamination from 43 min 42 s to 9 min 24 s. | Passive modification of room reflectivity markedly reduced cycle duration. | Controlled room-scale study; coating-specific generalizability. | Funding/manufacturer support was not clearly specified in the available article metadata. | Primary/independent report | 65
Haas | 2014 | R1 - Routine-care surface effectiveness | 643-bed tertiary acute-care hospital | Retrospective implementation and before/during clinical outcome study | Routine clinical exposure; hospital-acquired MDRO and C. difficile surveillance | Portable UV room disinfection | UV-C environmental disinfection | Mobile units deployed after discharge cleaning | 30-month pre-UV versus 22-month UV implementation period | Average 51 min; 11,389 uses; 76% of contact-precaution discharges treated | MDROs; Clostridioides difficile | HAI/MDRO incidence; treatment coverage; workflow | Hospital-acquired MDRO plus C. difficile rate decreased from 2.67 to 2.14 per 1000 patient-days (rate ratio 0.80, 95% CI 0.73-0.88). | Coverage reached 76% of contact-precaution discharges; machines were used 30% of available time. | Historical comparison; residual confounding and concurrent changes possible. | Funding/manufacturer support was not clearly specified in the available article metadata; authors reported no conflicts in the article. | Primary/independent report | 30
Ghantoji | 2015 | R1 - Routine-care surface effectiveness | 30 C. difficile isolation rooms in a comprehensive cancer center | Prospective comparative environmental study | Natural C. difficile contamination on high-touch surfaces | Pulsed-xenon UV | Pulsed xenon germicidal UV | Mobile room unit positioned for a 15-min treatment | Bleach cleaning versus non-bleach cleaning followed by PX-UV | 15 min PX-UV | Clostridioides difficile | Environmental culture; comparative efficacy | PX-UV reduced C. difficile burden by 83-95% depending on analysis; final contamination did not differ significantly from bleach-treated rooms. | Explored a non-bleach adjunctive strategy. | Small study; baseline contamination differed substantially between arms; company authors involved. | Institutional/NIH support plus Xenex support for laboratory analysis/editorial assistance; Xenex employees were coauthors. | Primary/independent report | 80
Nerandzic | 2015 | R2 - Field-based surface efficacy | Two acute-care hospitals | Mixed carrier and hospital-room surface study | Controlled carriers plus native high-touch surface contamination | Pulsed-xenon UV | Pulsed xenon UV; compared with continuous UV-C | Mobile room device | Distance, organic load, shading and PX-UV versus continuous UV-C; before/after room cultures | 10 min at 4 ft in comparative tests; 10-min room exposure | C. difficile spores; MRSA; VRE | Carrier log reduction; natural surface positivity; distance effects | At 4 ft for 10 min, carrier reductions were 0.55 log10 for C. difficile, 1.85 for MRSA and 0.60 for VRE. Native pathogen positivity on high-touch surfaces also decreased; distance strongly reduced efficacy. | Short room cycle, but performance was distance-sensitive. | Mixed natural and artificial endpoints; no patient outcome. | Funding/manufacturer support was not clearly specified in the available article metadata. | Primary/independent report | 47
Nagaraja | 2015 | R1 - Routine-care surface effectiveness | Acute-care hospital using pulsed-xenon room disinfection | Before/during implementation clinical analysis | Routine clinical CDI surveillance | Pulsed-xenon UV | Pulsed xenon UV | Mobile room disinfection after cleaning | One year before versus first year of UV deployment | Routine program deployment | Clostridioides difficile | Hospital-acquired CDI incidence; occupancy and room exposure | Hospital-acquired CDI was 22% lower during UV use, but the difference did not reach conventional statistical significance (P = 0.06). | Explored how deployment and room occupancy related to CDI. | Historical before-after design; limited causal attribution. | Funding source not specified in the available article metadata; authors reported no conflicts. | Primary/independent report | 34
Vianna | 2016 | R1 - Routine-care surface effectiveness | Community hospital; ICU plus non-ICU areas | Historical before-after clinical implementation study | Routine HAI surveillance | Pulsed-xenon UV | Pulsed xenon UV | Mobile device; three 5-min positions after cleaning | Baseline versus intervention periods | Three 5-min positions; >4,400 rooms treated | C. difficile; MRSA; VRE | Facility and ICU HAI rates; implementation | Combined facility rate for C. difficile, MRSA and VRE decreased 29% (P = 0.01); ICU combined rate decreased 61% (P = 0.01), but organism-specific effects were heterogeneous and non-ICU MRSA increased. | More than 200 rooms per month were treated. | Historical comparison; industry authors; uncontrolled changes in infection prevention. | Industry involvement: multiple authors were Xenex employees and disclosed financial/intellectual interests. | Primary/independent report | 35
Hosein | 2016 | R2 - Field-based surface efficacy | 40 isolation rooms in a 700-bed UK hospital | Patient-room before/after study plus controlled MDRO plate challenge and staff survey | Natural aerobic contamination plus inoculated MDRO plates | Pulsed-xenon UV | Pulsed xenon germicidal UV | Mobile room unit | After discharge, after manual cleaning, after PX-UV | Average device turnaround about 1 h | Aerobic flora; MDRO challenge organisms | Environmental microbiology; controlled log reduction; workflow and acceptability | PX-UV produced a further 78.4% reduction after manual cleaning (91% from initial burden) and about 5-log reduction on spiked MDRO plates. | Minimal effect on patient flow; staff and operators reported good acceptability. | No clinical endpoint; controlled plate test occurred in a non-patient-care area. | Laboratory processing supported by a Sodexo grant; Xenex loaned devices; authors reported no relevant conflicts. | Primary/independent report | 57
Pegues | 2017 | R1 - Routine-care surface effectiveness | Three adult hematology-oncology units in a tertiary hospital | Interrupted time-series analysis with comparison arm | Routine healthcare-onset CDI surveillance | Mobile UV-C germicidal irradiation | UV-C | No-touch terminal room device targeted by electronic patient-flow system | Baseline versus intervention in study and non-study units | 542/2,569 discharges treated (21.1%) over 52 weeks | Clostridioides difficile | Healthcare-onset CDI; coverage; room turnaround | CDI incidence decreased on study units; mixed-effects analysis showed IRR 0.49 (95% CI 0.26-0.94; P = 0.03), with no corresponding decrease on non-study units. | Targeted deployment did not materially increase room turnaround time. | Unit-level quasi-experimental design; only 21.1% of discharges received UV. | Funding/manufacturer support was not clearly specified in the available article metadata. | Primary/independent report | 31
Zeber | 2018 | R1 - Routine-care surface effectiveness | Four Veterans Affairs hospitals | Multicentre implementation comparison of PX-UV sites versus manual-only sites | Natural MRSA and aerobic surface contamination | Portable pulsed-xenon UV | Pulsed xenon UV | Mobile units after terminal cleaning | Two hospitals with PX-UV versus two manual-disinfection-only hospitals | Site protocols after terminal cleaning | MRSA; aerobic bacterial colonies | Environmental microbiology; multicentre comparative effectiveness | PX-UV sites reduced MRSA and aerobic counts by 75.3% and 84.1%, respectively, versus roughly 25-30% reductions at manual-only sites; adjusted models favored PX-UV. | Demonstrated multicentre deployment in VA hospitals. | Nonrandom facility allocation; modest environmental sample size. | Veterans Affairs Health Services Research & Development/public support; multicentre VA study. | Primary/independent report | 81
Brite | 2018 | R1 - Routine-care surface effectiveness | Bone marrow transplant unit at a tertiary cancer center | Interrupted time-series analysis with active and clinical surveillance | Routine VRE and toxigenic C. difficile acquisition surveillance | UV environmental disinfection | UV-C room disinfection | Mobile terminal-room UV for all rooms | Pre- versus post-implementation trend and level | Terminal disinfection of all rooms | VRE; Clostridioides difficile | Acquisition and clinical infection | No significant change in VRE or C. difficile acquisition/incidence was detected after UV implementation. | Universal terminal-room deployment on the BMT unit. | Single specialized unit; interrupted time series; no environmental mediator measurement. | New York State Department of Health HAI Prevention Project and MSKCC Cancer Center support; authors reported no relevant conflicts. | Primary/independent report | 39
Zeber | 2019 | R1 - Routine-care surface effectiveness | Patient rooms with five sampled high-touch surfaces | Repeated environmental sampling before cleaning, after cleaning and after UV | Natural aerobic bacterial and MRSA contamination | Supplemental UV room irradiation | Germicidal UV | Mobile no-touch adjunct after manual cleaning | Four manual products followed by UV | Routine post-cleaning UV cycle | Aerobic bacteria; MRSA | Incremental environmental bioburden reduction | After adjustment for cleaner/disinfectant and EVS staff, UV provided an additional estimated 56% reduction in aerobic counts and 93% reduction in MRSA compared with manual cleaning alone. | Incremental value was greatest when residual burden after manual cleaning was higher. | Environmental surrogate; practical incremental benefit depends on baseline/manual-cleaning performance. | VA Merit Review plus Xenex support for laboratory activity; authors disclosed prior Xenex-funded research roles. | Primary/independent report | 82
Hodges | 2022 | R1 - Routine-care surface effectiveness | Single-patient rooms at six hospitals in an academic health system | Retrospective cohort of prior-room-occupant pathogen transfer | Routine MDRO exposure and acquisition data | Adjunct UV-C terminal room disinfection | UV-C | Mobile UV-C after chlorine-based terminal cleaning | Rooms treated with adjunct UV-C versus standard chlorine cleaning only | 5,802/33,771 prior rooms received UV-C | Hospital-acquired multidrug-resistant pathogens | Pathogen transfer/acquisition from prior room occupant | After adjustment, UV-C-treated rooms had comparable risk of pathogen transfer (OR 1.06, 95% CI 0.84-1.32; P = 0.64). | Large multicentre real-world implementation dataset. | Retrospective nonrandom treatment allocation; facility-level deployment differences. | No financial support relevant to the article; authors reported no conflicts. | Primary/independent report | 29
Rock | 2022 | R1 - Routine-care surface effectiveness | Four cancer and one solid-organ transplant inpatient units | Cluster-randomized crossover controlled trial | Routine clinical VRE and C. difficile surveillance | UV-C adjunct to daily and discharge room cleaning | UV-C | Mobile room disinfection integrated into daily and terminal cleaning | UV-C plus standard cleaning versus standard cleaning; crossover by unit | Two study years with 5-week washout | VRE; Clostridioides difficile | Healthcare-associated infection/acquisition | UV-C did not reduce VRE (IRR 0.98, 95% CI 0.78-1.22) or C. difficile (IRR 1.43, 95% CI 0.93-2.21) rates. | Daily plus discharge deployment in high-risk units. | Only five clusters; high-risk immunocompromised population; no linked room dosimetry/environmental mediator. | Research grant from The Clorox Company to Johns Hopkins University; Clorox UV-C devices used. | Primary/independent report | 27
Fourman | 2026 | R1 - Routine-care surface effectiveness | 33 internal-medicine rooms occupied by CDI patients or asymptomatic carriers | Prospective interventional before-after environmental study | Natural C. difficile contamination on ten high-touch sites per room | Mobile UV-C or pulsed-xenon UV devices | UV-C and pulsed-xenon UV | Portable room devices used after bleach cleaning | After routine bleach cleaning, before versus after UV application | Device-specific protocol; 16 PX-UV and 17 UV-C rooms | Clostridioides difficile | Environmental CFU burden; contaminated sites; contamination score | No statistically significant incremental reduction after UV: mean contamination score 1.94 to 1.58 (P = 0.08), CFU 29.61 to 17.97 (P = 0.10), and contaminated sites 2.24 to 1.91 (P = 0.07). | Directly tests added value after daily bleach cleaning. | Single centre; 33 rooms; relatively low power and strong preceding sporicidal cleaning. | No funding reported in the article. | Primary/independent report | 66
Vincent | 2026 | R3 - Healthcare-room or full-scale simulation | Two unoccupied two-bed hospital patient rooms | Partially blinded crossover room-scale carrier study | Inoculated stainless-steel disc carriers at low and high soil levels | Operator-positioned far-UVC 222-nm luminaires | Filtered UV-C222 | Luminaires positioned behind beds and in bathroom for defined exposure | Active UV versus no-UV control; low versus high soil | 24-h exposure; measured dose 5-637 mJ/cm2 | Candida auris; MRSA; T1 bacteriophage | Room-scale microbial log reduction; soil and dose effects | In hospital rooms, UV-C222 reduced C. auris by 2.6 log10 at low soil and 1.0 log10 at high soil, MRSA by 1.0 log10 at low soil, and T1 by 0.6 log10 at low soil. | Extends room-scale evidence to operator-positioned far-UVC; occupied-space benefit remains unproven. | Only two unoccupied rooms; 24-h exposure; controlled carriers rather than natural contamination. | Ushio sponsored the research; Myna Life Technology loaned precommercial luminaires; two authors reported a UV Smart grant. | Primary/independent report | 17
Allen | 2019 | R1 - Routine-care surface effectiveness | Specialist cystic-fibrosis outpatient clinic; clinic rooms after patient encounters | Sequential before/after environmental surface study during routine outpatient care | Natural environmental contamination | Mobile/operator-positioned UV-C room-disinfection system | UV-C Ultra-V system (Hygiene Solutions); wavelength not specified in the article | Portable room system monitored by six perimeter sensors; repositioned when required | After patient encounter; after chlorine-containing manual cleaning; after UV-C | Cumulative exposure monitored in real time; treatment completed within 10–20 min, with repositioning in some rooms | Aerobic environmental bacteria; patients carried P. aeruginosa, MRSA or ESBL E. coli | Environmental microbiology; cycle time; positioning | Manual cleaning produced inconsistent changes, whereas UV-C produced a significant reduction in CFU across sampled surfaces and often reduced counts to undetectable levels. | Variable room size/shape sometimes required repositioning; all treatments completed within 10–20 min. | Small single-centre study; environmental surrogate outcome; no acquisition/infection endpoint. | Hygiene Solutions provided the UV-C unit and contact plates/covered microbiological processing; one author was an employee of Hygiene Solutions. | Primary/independent report | 21
Anderson | 2018 | R1 - Routine-care surface effectiveness | Nine hospitals in the southeastern USA; hospital-wide outcomes during BETR Disinfection | Prespecified secondary analysis of the pragmatic multicentre cluster-randomized crossover BETR Disinfection trial | Routine clinical acquisition/infection surveillance | Mobile continuous UV-C adjunct to terminal cleaning | UV-C; same BETR deployment as primary 2017 report | Targeted terminal room disinfection; four crossover cleaning strategies | Reference vs UV; bleach vs bleach+UV strategy periods | Same 7-month crossover strategy periods as BETR Disinfection | C. difficile, MRSA, VRE, multidrug-resistant Acinetobacter spp. | Hospital-wide acquisition/infection incidence; pathogen-specific clinical outcomes | Hospital-wide risk for all target organisms was not significantly different in the UV period versus reference (RR 0.89, 95% CI 0.79–1.00; P=0.052); decreases were observed for C. difficile (RR 0.89, 95% CI 0.80–0.99) and VRE (RR 0.56, 95% CI 0.31–0.996). | Extends the BETR analysis from next-room exposure to hospital-wide incidence. | Companion secondary analysis of the same BETR trial as ref. 19; not an independent study and should not be double-counted. | US Centers for Disease Control and Prevention; authors reported advisory-panel relationships with Clorox; Tru-D SmartUVC, Ecolab and Clorox contributed materials but had no role in funding, design, analysis or manuscript preparation. | Companion secondary analysis of BETR Disinfection (same trial as ref. 19); not counted as an independent study | 19
Armellino | 2019 | R1 - Routine-care surface effectiveness | Operating-room suite and patient-care equipment inside/outside the OR environment | Three-phase nonrandomized observational study with multisided pre/post surface sampling | Natural microbial contamination on patient-care equipment | Focused multivector ultraviolet (FMUV) system with enclosed shadowless delivery | Focused multivector germicidal UV; wavelength not specified in the article | Modular panels surrounding the target zone; equipment placed inside treatment enclosure | Manual-chemical disinfection in phase 1; pre/post FMUV without chemical disinfection in phases 2–3 | 90-second FMUV cycle | Aerobic environmental microbial burden | Environmental microbiology; rapid cycle; multisided equipment coverage | FMUV reduced active microbial burden by 92.0–97.7% in OR phases and 96.3–99.6% on equipment outside the OR; overall reductions were highly significant. | A 90-second enclosed treatment was designed for rapid between-case/equipment decontamination. | Nonrandomized study; environmental outcomes only; proprietary enclosed system limits generalizability. | Supported by Northwell Ventures; PurpleSun was a Northwell Ventures portfolio company and received external grant support; two authors were paid advisors to PurpleSun. | Primary/independent report | 22
Choi | 2022 | R1 - Routine-care surface effectiveness | COVID-19 ICU airborne-infection isolation rooms and community treatment-center isolation rooms | Before/after environmental surface study after patient discharge without preceding manual terminal cleaning | Natural SARS-CoV-2 RNA surface contamination | Articulated UV-LED irradiation robot | UV-A/near-UV LED, 340–385 nm with peak approximately 365 nm | Robot with multi-jointed arm positioned to irradiate room surfaces | Environmental sampling before versus after robot irradiation | Surface irradiation programmed by robot; overall room procedures approximately 30–70 min depending on room | SARS-CoV-2 RNA; culture attempted on PCR-positive samples | Environmental viral RNA detection; feasibility; room-size/shadowing effects | Surface PCR positivity decreased after UV-LED irradiation, with complete conversion of positive samples in one ICU room but residual RNA in smaller treatment-center rooms; viable virus was not recovered from PCR-positive samples. | Performance varied with room size, shadowing and possible wheel-mediated recontamination. | Near-UV rather than UV-C; PCR endpoint does not establish infectivity; small room sample and no patient outcome. | Supported by Korea University Medical Center and the National Research Foundation of Korea; funders reported no role; authors declared no related conflicts. | Primary/independent report | 18
Dexter | 2022 | R1 - Routine-care surface effectiveness | Operating rooms and hospital rooms with routine 3-tower UV-C deployments | Large operational analysis of completed UV-C treatment logs | Not applicable; operational implementation dataset | Three-tower multi-emitter UV-C room-disinfection system | UV-C; Surfacide multi-tower system | Three mobile emitters deployed by room | Room-specific treatment-duration distributions | 133,927 completed treatments across 700 rooms; pooled mean treatment duration 18.3 min | Not applicable | Operational treatment time; scheduling variability; prediction limits | Treatment duration varied substantially both between rooms and across repeated treatments of the same room; room-specific, distribution-free prediction limits were recommended rather than generic UV-C scheduling assumptions. | Provides large-scale workflow evidence for room-specific scheduling of UV-C treatment. | Operational dataset supplied by the manufacturer; no microbiological or patient-effectiveness endpoint. | Unrestricted research grant from Surfacide to the University of Iowa Department of Anesthesia; Surfacide also supplied the treatment-time data; author disclosures included other industry relationships. | Primary/independent report | 25
Dhar | 2024 | R1 - Routine-care surface effectiveness | Two hospitals and 15 inpatient wards | Cluster-randomized, controlled, double-blinded interventional crossover trial | Routine clinical HAI surveillance | Pulsed-xenon UV room-disinfection devices with visually indistinguishable sham units | Pulsed-xenon broad-spectrum UV, approximately 200–320 nm | Up to four device positions after standard terminal cleaning; 5-min light cycle per position | Standard terminal cleaning + PX-UV versus standard terminal cleaning + sham | 145,230 intervention/sham cycles; trained technicians; overall discharge-level compliance 84.2% | VRE, ESBL-producing E. coli/K. pneumoniae, MRSA, A. baumannii, C. difficile | Environmentally implicated HAI incidence; randomized clinical effectiveness; implementation compliance | No reduction in environmentally implicated HAIs was demonstrated: 3.49 versus 3.17 infections per 1000 patient-days (RR 1.10, 95% CI 0.94–1.29; P=0.23), with similar findings across pathogen, hospital and unit strata. | Large blinded sham-controlled deployment with high documented use and trained technicians. | Clinical endpoint without linked room-level dosimetry/environmental mediator; technical support provided by manufacturer. | AHRQ grant R01HS024709; Xenex provided equipment and technician support. Several authors disclosed industry relationships, and one author was employed by Xenex during the study. | Primary/independent report | 28
Jinadatha | 2014 | R1 - Routine-care surface effectiveness | Twenty rooms vacated by MRSA-positive patients | Comparative environmental study of standard manual cleaning versus pulsed-xenon UV treatment | Natural MRSA and heterotrophic bacterial contamination on high-touch surfaces | Portable pulsed-xenon UV room-disinfection device | Pulsed xenon UV | Portable device repositioned for multiple 5-min cycles; approximately 15 min UV exposure | Standard manual room cleaning versus PPX-UV | Approximately 15 min UV exposure plus setup | MRSA and heterotrophic plate counts | Environmental microbiology; room-cleaning time | PPX-UV produced larger reductions than manual cleaning for both MRSA and heterotrophic counts; reported reductions were approximately 99.4% for MRSA and 98.1% for heterotrophic counts in the PPX-UV arm. | Mean room time was not significantly longer than manual cleaning in this small comparison. | Only 20 rooms; baseline burden differed between groups; environmental endpoint only. | Supported by Central Texas Veterans Health Care System/VA-related support; laboratory activity and the PPX-UV machine were supported by a Xenex Healthcare Services grant. | Primary/independent report | 23
Kanamori | 2016 | R3 - Healthcare-room or full-scale simulation | Unoccupied hospital patient room and bathroom | Controlled room-scale challenge comparing one versus two source positions | Artificial high-load MRSA and carbapenem-resistant K. pneumoniae on Formica carriers | Portable fixed-cycle-time UV-C device | Clorox Healthcare Optimum-UV, 254 nm | One central/foot-of-bed position versus sequential positions on both sides of bed | Single 5-min cycle versus two 5-min cycles; direct versus indirect sites | 5 min versus 10 min total exposure | MRSA and carbapenem-resistant K. pneumoniae | Room-scale log reduction; distance; line-of-sight effects | Overall reductions were 5.01 versus 5.87 log10 for carbapenem-resistant K. pneumoniae and 4.61 versus 5.06 log10 for MRSA for one versus two positions; distance strongly correlated with lower reduction. | Additional positioning improved coverage but increased treatment time. | Controlled inoculated carriers in an empty room; no natural contamination or patient outcome. | Internal funding from UNC Health Care; Kanamori received JSPS postdoctoral support; Rutala and Weber reported consultancy relationships with Clorox. | Primary/independent report | 60
Mustapha | 2018 | R2 - Healthcare field efficacy | Post-discharge hospital rooms of patients under contact precautions for MRSA | Natural floor sampling before/after manual cleaning and adjunctive UV-C plus controlled MRSA carrier experiment | Mixed natural floor contamination and inoculated MRSA carriers | Mobile UV-C room-decontamination device | Clorox Healthcare Optimum-UV, 254 nm | Cycles on both sides of bed and in bathroom | Before cleaning; after manual cleaning; after adjunctive UV-C; controlled exposed versus unexposed carriers | 3–5 min per position | MRSA, C. difficile, Candida spp.; MRSA carriers | Environmental floor microbiology; controlled carrier log reduction | Manual cleaning markedly reduced floor contamination; residual MRSA positivity decreased further after UV-C (from 9% after manual cleaning to 1%), although this incremental difference was not statistically significant. Controlled carriers confirmed UV-C activity at floor and elevated locations. | Demonstrates floor exposure using standard multi-position room cycles. | Small convenience sample; floor contamination may not represent high-touch transmission; mixed natural and controlled components. | Department of Veterans Affairs Merit Review grant; Donskey disclosed research grants/advisory relationships with several companies, not reported as direct study funding. | Primary/independent report | 48
Rock | 2016 | R3 - Healthcare-room or full-scale simulation | Empty patient room and bathroom at Johns Hopkins Hospital | Controlled room-scale carrier study on multiple high-touch sites | Artificial CRE inocula on Formica carriers | Portable UV-C room-disinfection device | Clorox Healthcare Optimum-UV, 254 nm | 5 min on each side of the bed plus 5 min in bathroom | UV-exposed carriers versus shielded negative controls; direct and indirect sites | 15 min total exposure | Carbapenem-resistant E. coli, K. pneumoniae and Enterobacter cloacae | Room-scale microbial recovery; high-touch surface coverage; shadowing | Only 2 of 133 high-touch carrier plates yielded CRE after UV-C exposure, corresponding to near-complete multi-log reduction; residual growth occurred at difficult/shadowed locations. | Short multi-position cycle covered patient room and bathroom. | Artificial high-load carriers in an empty room; no natural contamination or patient outcome. | Financial support provided by Clorox; authors reported no relevant conflicts of interest. | Primary/independent report | 61
Sitzlar | 2012 | R1 - Routine-care surface effectiveness | Shared room in a long-term care facility occupied by two residents with recurrent CDI | Two-patient case report/letter with environmental cultures before and after automated UV decontamination | Natural C. difficile room contamination | Automated UV radiation room-disinfection device | Tru-D Rapid Room Disinfection; germicidal UV-C | Automated treatment of room and bathroom | Environmental cultures before versus after UV decontamination within ongoing cleaning program | Reflected dose 22,000 µWs/cm²; approximately 90 min for room/bathroom treatment | Clostridium difficile | Environmental culture; recurrent CDI clinical course | Environmental C. difficile contamination persisted despite routine measures; after UV treatment only one bedrail colony was recovered. Both residents subsequently had negative rectal swabs and no further CDI during the reported follow-up. | Long cycle in a long-term care shared room. | Only two patients; uncontrolled case report; recurrence cannot be causally attributed to UV. | VA Merit Review and T. Franklin Williams Scholarship support; Donskey and Jump disclosed unrelated industry relationships. | Primary/independent report | 32
Uwayo | 2025 | R1 - Routine-care surface effectiveness | Marburg virus treatment environments in Rwanda: ward, ICU, laboratory areas and medical equipment | Sequential outbreak-environment sampling before cleaning, after chemical disinfection and after UV-C robot deployment | Natural MARV RNA contamination | Autonomous UV-C robot | THOR UV system, 254 nm | Mobile robot for room/equipment decontamination | Baseline; after chemical disinfection; after UV-C robot | Robot rated for approximately 16 m²/min | Marburg virus RNA by RT-PCR | Environmental viral RNA detection; outbreak implementation | Thirteen samples were positive at baseline. Chemical treatment cleared most sites, while two equipment-associated samples remained positive and became negative after UV-C robot deployment. | Illustrates robotic deployment during a viral-hemorrhagic-fever outbreak in a resource-limited context. | RT-PCR detects viral RNA rather than infectivity; very small number of residual positive sites; sequential intervention prevents isolation of UV effect for most surfaces. | No study funding statement was identified in the article; authors reported no conflicts of interest. | Primary/independent report | 24

## Methodological_Characteristics
Descriptive methodological characteristics for 71 included reports representing 70 distinct studies
No formal critical appraisal or risk-of-bias assessment was performed; the table describes design and interpretation characteristics only. Data charting/extraction was performed by one reviewer.
Ref. | Study | Evidence stratum | Design / comparator profile | Outcome type / directness | Implementation / technical features | Main interpretive limitation
1 | Rutala (2010) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Automated room treatment; device loaned by manufacturer. | Small field component; microbiological surrogate outcomes; no patient infection endpoint.
2 | Wong (2016) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Highlights device-to-device variation under protein load and difficult placement. | Non-random field sampling; carrier findings not equivalent to natural contamination; no clinical endpoint.
3 | Nerandzic (2010) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Setup required <5 min and no continuous monitoring. | Small field study; no randomized comparator or clinical outcome.
4 | Ali (2017) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | In-house validation of room layouts and cycle duration was recommended. | Small number of rooms; manufacturer-funded microbiological testing; artificial challenge loads.
49 | Yui (2025) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Designed for rapid targeted use; performance depended on positioning and soil. | Small in-use sample; mixed natural and artificial outcomes; online publication in 2024 / issue year 2025.
36 | Kovach (2017) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Patient-level clinical/acquisition outcome | Applied after resident discharge in one facility. | Single-site uncontrolled pre-post design; temporal confounding and regression to the mean possible.
5 | Boyce (2016) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room or full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Demonstrates need for dose measurement and multiple placements. | Empty-room simulation with artificial carriers; not routine clinical deployment.
70 | Carlisle (2023) | R1 - Routine-care surface effectiveness | Randomized room-level comparator | Natural surface contamination and/or operational outcome | Direct comparative evidence on workflow and alternative no-touch technology. | Pilot size and insufficient power to exclude meaningful differences; environmental rather than clinical outcome.
58 | Jelden (2017) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room or full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Passive room-surface modification may improve dose distribution. | Controlled challenge, not natural contamination.
33 | Napolitano (2015) | R1 - Routine-care surface effectiveness | Historical pre/post, feasibility or program evaluation | Patient-level clinical/acquisition outcome | Dedicated service model linked treatment coverage to outcomes. | Uncontrolled pre-post design; no concurrent control; possible co-interventions and underlying temporal trends.
67 | Casini (2023) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Automatic recording of disinfection parameters enabled process traceability; staff training was emphasized. | No concurrent control or patient outcome; setting-specific thresholds; before-after sampling.
71 | Casini (2019) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Short adjunctive step in critical areas. | Before-after environmental study; no clinical endpoint; same research group as implementation study.
72 | Kitagawa (2021) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Useful as an adjunct when chemical cleaning is not reliably sporicidal. | Environmental surrogate; small single-centre study; contamination prevalence may limit power.
50 | Füszl (2021) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Interface was not intuitive, required technical training and frequent intervention; routine autonomy was limited. | Small field study; mixed natural and artificial outcomes; usability was device-specific.
51 | Casini (2025) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Directly demonstrates the laboratory-to-field performance gap. | Small number of surfaces/rooms; experimental platform; no clinical outcome.
55 | Villacís (2019) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Demonstrated applicability in a resource setting outside the USA. | Before-after study without clinical outcomes; manufacturer-affiliated author present.
38 | Raggi (2018) | R1 - Routine-care surface effectiveness | Historical pre/post, feasibility or program evaluation | Patient-level clinical/acquisition outcome | No significant adverse change in emergency-department admission time. | Uncontrolled pre-post design; cost estimate did not fully subtract intervention costs; unmeasured confounding.
41 | Morikane (2020) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Patient-level clinical/acquisition outcome | Light/sound initially disturbed staff and equipment; blackout curtains, goggles and earplugs enabled workflow integration. | Single-centre before-after design; co-interventions and time trends possible.
73 | Kitagawa (2020) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Provides routine cycle-time data. | Single-centre environmental outcome; paired sites but no clinical endpoint.
59 | Jelden (2016) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room or full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Maps performance across a furnished room. | Artificial high-load carriers in an empty room; not routine care.
68 | McGinn (2022) | R1 - Routine-care surface effectiveness | Concurrent/comparative element without randomization | Natural surface contamination and/or operational outcome | Robot procedure was substantially faster than the existing manual protocol in the evaluated CT room. | Small single-room study; natural counts were low/variable; developer affiliation with Akara Robotics.
69 | McGinn (2021) | R1 - Routine-care surface effectiveness | Historical pre/post, feasibility or program evaluation | Natural surface contamination and/or operational outcome | Detailed robot design, navigation, task planning and user-centred deployment considerations. | Small feasibility study; low baseline bioburden; developer involvement.
52 | Różańska (2023) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Supports location-specific validation and adjunctive use. | Mixed designs and outcome metrics; hospital sample size limited; some results based on artificial high loads.
74 | Sathitakorn (2023) | R1 - Routine-care surface effectiveness | Concurrent/comparative element without randomization | Natural surface contamination and/or operational outcome | Judged feasible in a resource-limited setting without compromising room turnover. | Environmental outcome; bundled real-time feedback makes the isolated PX-UV contribution difficult to quantify.
75 | Jinadatha (2015) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Isolates the no-touch effect from manual cleaning, but does not represent recommended routine practice. | Only 14 rooms; no manual-cleaning comparator; residual contamination remained.
76 | Simmons (2018) | R1 - Routine-care surface effectiveness | Historical pre/post, feasibility or program evaluation | Natural surface contamination and/or operational outcome | Large multicentre operational dataset. | Pooled validation without standardized control; all authors were affiliated with the device company.
77 | Su (2021) | R1 - Routine-care surface effectiveness | Historical pre/post, feasibility or program evaluation | Natural surface contamination and/or operational outcome | Provides in-hospital dose/time evidence during COVID care. | Small number of positive sites; PCR-based endpoints may not equal infectivity; no patient outcome.
56 | Yang (2019) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Performance decreased with distance and organism resistance. | Mixed laboratory and field outcomes; no clinical endpoint.
78 | Bosco (2022) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Model used to standardize a rapid adjunctive protocol between procedures. | Non-random before-after design; total counts rather than target pathogens; limited generalizability.
37 | Pavia (2018) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Patient-level clinical/acquisition outcome | One device was deployed across selected rooms on a scheduled basis. | Uncontrolled before-after design; company authors and device donation; no environmental sampling.
64 | Tholany (2023) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room or full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Identifies shower rails and distant locations as potential failures. | Artificial plates in empty rooms; pilot sample; no natural contamination.
43 | de la Rosa-Zamboni (2025) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Patient-level clinical/acquisition outcome | Implemented sequentially in a high-risk NICU. | UV-C cannot be separated from hand-hygiene bundles and copper; non-random sequential periods.
62 | Arkusz (2024) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room or full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Supports mobile/self-steering systems that reduce source-to-surface distance. | Authors call conditions “real-world,” but the room was a 1:1 simulation rather than routine patient care.
44 | Ishikawa (2025) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Patient-level clinical/acquisition outcome | Implemented safely in NICU/GCU. | Single hospital under outbreak conditions; simultaneous daily cleaning and improved hand hygiene confound UV effect.
42 | Kitagawa (2021) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Patient-level clinical/acquisition outcome | High but incomplete room coverage was documented. | Non-randomized intervention; borderline confidence interval; possible secular/co-intervention effects.
79 | Kitagawa (2021) | R1 - Routine-care surface effectiveness | Within-site pre/post comparison | Natural surface contamination and/or operational outcome | Identifies bathrooms and bidet nozzles as priority targets. | Small single-centre environmental study; no clinical endpoint.
40 | Schaffzin (2020) | R1 - Routine-care surface effectiveness | Historical pre/post, feasibility or program evaluation | Patient-level clinical/acquisition outcome | A dedicated team, environmental-services ownership and shift goals improved coverage without expanding the robot fleet. | Program evaluation cannot establish causality between UV and HAI reduction; other cleaning initiatives occurred.
53 | Russo (2021) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Dynamic robotic use intended to reduce human error and chemical use. | Only two clinical rooms; environmental and laboratory findings combined; device supplier provided equipment/technical support.
54 | Chen (2020) | R2 - Field-based surface efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Short cycle in clinical laboratory areas. | No laboratory-acquired infection endpoint despite conclusion; field sample size and control details limited.
63 | Lorca-Oró (2022) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room or full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Autonomous platform and room mapping were demonstrated. | Simulated room with artificial virus; developer/shareholder conflicts disclosed; not routine clinical care.
20 | Anderson (2013) | R1 - Routine-care surface effectiveness | Prospective multicentre before-after environmental study | Natural surface contamination outcome | Cycle duration, reflected-dose protocol and line of sight reported | Natural contamination in 39 patient rooms at two hospitals; no patient endpoint.
26 | Anderson (2017) – BETR Disinfection | R1 - Routine-care surface effectiveness | Pragmatic cluster-randomized multicentre crossover trial | Patient-level acquisition/infection outcome | Large implementation program; room-level dose not linked to outcome | Randomization strengthens causal inference for the UV-vs-reference comparison; lack of C. difficile benefit with bleach+UV demonstrates context dependence.
45 | Boyce (2011) | R2 - Field-based surface efficacy | Prospective patient-room before/after study with natural surface cultures plus a C. difficile carrier test | Healthcare-relevant microbiological performance | Two-stage positioning improved coverage but lengthened room downtime. | No patient outcome; mixed natural and carrier endpoints; long cycle times.
46 | Havill (2012) | R2 - Field-based surface efficacy | Prospective comparative room study of hydrogen peroxide vapor versus UV-C | Healthcare-relevant microbiological performance | UV-C was faster than hydrogen peroxide vapor. | Comparative environmental study without patient outcomes; carrier challenge may not reflect natural burden.
65 | Rutala (2013) | R3 - Healthcare-room or full-scale simulation | Controlled room-scale evaluation of UV-reflective wall coating | Healthcare-relevant microbiological performance | Passive modification of room reflectivity markedly reduced cycle duration. | Controlled room-scale study; coating-specific generalizability.
30 | Haas (2014) | R1 - Routine-care surface effectiveness | Retrospective implementation and before/during clinical outcome study | Patient-level or program outcome | Coverage reached 76% of contact-precaution discharges; machines were used 30% of available time. | Historical comparison; residual confounding and concurrent changes possible.
80 | Ghantoji (2015) | R1 - Routine-care surface effectiveness | Prospective comparative environmental study | Natural surface microbiology / operational outcome | Explored a non-bleach adjunctive strategy. | Small study; baseline contamination differed substantially between arms; company authors involved.
47 | Nerandzic (2015) | R2 - Field-based surface efficacy | Mixed carrier and hospital-room surface study | Healthcare-relevant microbiological performance | Short room cycle, but performance was distance-sensitive. | Mixed natural and artificial endpoints; no patient outcome.
34 | Nagaraja (2015) | R1 - Routine-care surface effectiveness | Before/during implementation clinical analysis | Patient-level clinical/acquisition outcome | Explored how deployment and room occupancy related to CDI. | Historical before-after design; limited causal attribution.
35 | Vianna (2016) | R1 - Routine-care surface effectiveness | Historical before-after clinical implementation study | Patient-level or program outcome | More than 200 rooms per month were treated. | Historical comparison; industry authors; uncontrolled changes in infection prevention.
57 | Hosein (2016) | R2 - Field-based surface efficacy | Patient-room before/after study plus controlled MDRO plate challenge and staff survey | Healthcare-relevant microbiological performance | Minimal effect on patient flow; staff and operators reported good acceptability. | No clinical endpoint; controlled plate test occurred in a non-patient-care area.
31 | Pegues (2017) | R1 - Routine-care surface effectiveness | Interrupted time-series analysis with comparison arm | Patient-level clinical/acquisition outcome | Targeted deployment did not materially increase room turnaround time. | Unit-level quasi-experimental design; only 21.1% of discharges received UV.
81 | Zeber (2018) | R1 - Routine-care surface effectiveness | Multicentre implementation comparison of PX-UV sites versus manual-only sites | Natural surface microbiology / operational outcome | Demonstrated multicentre deployment in VA hospitals. | Nonrandom facility allocation; modest environmental sample size.
39 | Brite (2018) | R1 - Routine-care surface effectiveness | Interrupted time-series analysis with active and clinical surveillance | Patient-level or program outcome | Universal terminal-room deployment on the BMT unit. | Single specialized unit; interrupted time series; no environmental mediator measurement.
82 | Zeber (2019) | R1 - Routine-care surface effectiveness | Repeated environmental sampling before cleaning, after cleaning and after UV | Natural surface microbiology / operational outcome | Incremental value was greatest when residual burden after manual cleaning was higher. | Environmental surrogate; practical incremental benefit depends on baseline/manual-cleaning performance.
29 | Hodges (2022) | R1 - Routine-care surface effectiveness | Retrospective cohort of prior-room-occupant pathogen transfer | Patient-level or program outcome | Large multicentre real-world implementation dataset. | Retrospective nonrandom treatment allocation; facility-level deployment differences.
27 | Rock (2022) | R1 - Routine-care surface effectiveness | Cluster-randomized crossover controlled trial | Patient-level or program outcome | Daily plus discharge deployment in high-risk units. | Only five clusters; high-risk immunocompromised population; no linked room dosimetry/environmental mediator.
66 | Fourman (2026) | R1 - Routine-care surface effectiveness | Prospective interventional before-after environmental study | Natural surface microbiology / operational outcome | Directly tests added value after daily bleach cleaning. | Single centre; 33 rooms; relatively low power and strong preceding sporicidal cleaning.
17 | Vincent (2026) | R3 - Healthcare-room or full-scale simulation | Partially blinded crossover room-scale carrier study | Healthcare-relevant microbiological performance | Extends room-scale evidence to operator-positioned far-UVC; occupied-space benefit remains unproven. | Only two unoccupied rooms; 24-h exposure; controlled carriers rather than natural contamination.
21 | Allen (2019) | R1 - Routine-care surface effectiveness | Sequential before/after environmental surface study during routine outpatient care | Natural surface contamination and/or operational outcome | Variable room size/shape sometimes required repositioning; all treatments completed within 10–20 min. | Small single-centre study; environmental surrogate outcome; no acquisition/infection endpoint.
19 | Anderson (2018) | R1 - Routine-care surface effectiveness | Companion prespecified secondary analysis of cluster-randomized BETR trial | Patient-level/hospital-wide acquisition/infection outcome | Extends the BETR analysis from next-room exposure to hospital-wide incidence. | Same underlying BETR trial as ref. 19; retained as a companion report and not counted as an independent study.
22 | Armellino (2019) | R1 - Routine-care surface effectiveness | Three-phase nonrandomized observational study with multisided pre/post surface sampling | Natural surface contamination and/or operational outcome | A 90-second enclosed treatment was designed for rapid between-case/equipment decontamination. | Nonrandomized study; environmental outcomes only; proprietary enclosed system limits generalizability.
18 | Choi (2022) | R1 - Routine-care surface effectiveness | Before/after environmental surface study after patient discharge without preceding manual terminal cleaning | Natural surface contamination and/or operational outcome | Performance varied with room size, shadowing and possible wheel-mediated recontamination. | Near-UV rather than UV-C; PCR endpoint does not establish infectivity; small room sample and no patient outcome.
25 | Dexter (2022) | R1 - Routine-care surface effectiveness | Large operational analysis of completed UV-C treatment logs | Natural surface contamination and/or operational outcome | Provides large-scale workflow evidence for room-specific scheduling of UV-C treatment. | Operational dataset supplied by the manufacturer; no microbiological or patient-effectiveness endpoint.
28 | Dhar (2024) | R1 - Routine-care surface effectiveness | Cluster-randomized, blinded/sham-controlled crossover trial | Patient-level clinical/acquisition outcome | Large blinded sham-controlled deployment with high documented use and trained technicians. | Clinical endpoint without linked room-level dosimetry/environmental mediator; technical support provided by manufacturer.
23 | Jinadatha (2014) | R1 - Routine-care surface effectiveness | Comparative environmental study of standard manual cleaning versus pulsed-xenon UV treatment | Natural surface contamination and/or operational outcome | Mean room time was not significantly longer than manual cleaning in this small comparison. | Only 20 rooms; baseline burden differed between groups; environmental endpoint only.
60 | Kanamori (2016) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room/full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Additional positioning improved coverage but increased treatment time. | Controlled inoculated carriers in an empty room; no natural contamination or patient outcome.
48 | Mustapha (2018) | R2 - Healthcare field efficacy | Field pre/post plus controlled challenge or mixed laboratory component | Field-based surface efficacy; not patient-level effectiveness | Demonstrates floor exposure using standard multi-position room cycles. | Small convenience sample; floor contamination may not represent high-touch transmission; mixed natural and controlled components.
61 | Rock (2016) | R3 - Healthcare-room or full-scale simulation | Controlled healthcare-room/full-scale experiment | Room-scale surface performance; not routine-care effectiveness | Short multi-position cycle covered patient room and bathroom. | Artificial high-load carriers in an empty room; no natural contamination or patient outcome.
32 | Sitzlar (2012) | R1 - Routine-care surface effectiveness | Two-patient case report/letter with environmental cultures before and after automated UV decontamination | Case-level environmental and clinical outcome | Long cycle in a long-term care shared room. | Only two patients; uncontrolled case report; recurrence cannot be causally attributed to UV.
24 | Uwayo (2025) | R1 - Routine-care surface effectiveness | Sequential outbreak-environment sampling before cleaning, after chemical disinfection and after UV-C robot deployment | Natural surface contamination and/or operational outcome | Illustrates robotic deployment during a viral-hemorrhagic-fever outbreak in a resource-limited context. | RT-PCR detects viral RNA rather than infectivity; very small number of residual positive sites; sequential intervention prevents isolation of UV effect for most surfaces.

## Device_Taxonomy
Multiaxial taxonomy for the 70 distinct included studies (71 reports)
Technology is described on independent axes; laboratory-only emerging-source evidence is outside the included evidence map.
Axis | Categories represented in included studies | Key extracted variables | Interpretive use | Representative references
Source and spectrum | Low-pressure mercury UV-C; pulsed xenon; far-UVC 222 nm; other continuous germicidal UV-C; one near-UV/UV-A LED system (340–385 nm, peak ~365 nm) | Wavelength/spectral output; continuous or pulsed emission; radiant output | Mechanism, organism sensitivity, dose and safety; near-UV/UV-A and far-UVC evidence remain sparse and require wavelength-specific interpretation. | [1,2,4,17,18,20–22,24,26,28,48–50,60,63,67,68]
Deployment configuration | Mobile stationary whole-room; targeted/articulating; multi-emitter; autonomous/semi-autonomous robot; operator-positioned luminaires | Number/orientation of emitters; movement; placements; interface; navigation; whether the source is portable or operator-positioned | Coverage, shadow management, setup and repeatability. Permanently fixed wall/ceiling systems and handheld spot-disinfection devices are outside scope. | [4,17,18,21–25,28,32,48–53,60,61,63,64,67–69]
Experimental setting | Routine healthcare; genuine healthcare rooms; unoccupied healthcare rooms; full-scale simulated hospital rooms | Occupancy, furnishing, room geometry, natural vs controlled contamination | Separates effectiveness, field efficacy and room-scale performance | R1, R2 and R3 strata
Surface target | High-touch surfaces; whole-room surfaces; selected equipment; direct and indirect sites | Surface type, room area, natural/artificial contamination | Prevents broad whole-room claims from exceeding sampled targets | [1–5,9,13,14,17–24,26,27,29–32,34–69]
Dose-control strategy | Fixed time; reflected-dose endpoint; radiometer/sensor mapping; multi-position/navigation algorithm | Irradiance, dose, distance, time, sensor positions and calibration | Whether intended dose reaches direct and shaded sites | [1–5,17,18,20,21,28,45–49,51,52,58–65,68,69]