Supplemental table 1. Analytical performance (within-run, between-run, and within-laboratory imprecision) of the pTau217, Aβ40 and Aβ42 assays on the Sysmex Automated Immunoassay System HISCL determined by measuring pTau217 (RUO), Aβ40 en Aβ42 QC-solutions in a 5x5 protocol (CLSI E15-A3)
	
	Plasma p-tau217
	Plasma Aβ1-40
	Aβ1-42

	 
	Level 1
	Level 2
	Level 1
	Level 2
	Level 1
	 Level 2

	Mean concentration (pg/mL)
	0.53
	4.00
	92.0
	174.5
	14.6
	 30.7

	Target (pg/mL)
	0.50
	4.13
	95
	182
	15.3
	 32.4

	Imprecision
	 
	 
	 
	
	
	 

	SDwr (pg/mL)
	0.02
	0.06
	1.98
	2.60
	0.30
	 0.55

	CVwr (%)
	2.8
	1.4
	2.2
	1.5
	2.0
	 1.8

	SDbr (pg/mL)
	0.001
	0.04
	1.63
	4.39
	0.60
	 0.55

	CVbr (%)
	0.3
	1.0
	1.8
	2.5
	4.1
	 1.8

	SDwl (pg/mL)
	0.02
	0.07
	2.6
	5.11
	0.67
	 0.78

	CVwl (%)
	2.8
	1.7
	2.8
	2.9
	4.6
	 2.5


Abbreviations: WR: within-run; BR: between-run, WL: within-lab, SD: standard deviation, CV: coefficient of variation, CVwr: within-run coefficient
