Supplemental table 1. Analytical performance (within-run, between-run, and within-laboratory imprecision) of the pTau217, Aβ40 and Aβ42 assays on the Sysmex Automated Immunoassay System HISCL determined by measuring pTau217 (RUO), Aβ40 en Aβ42 QC-solutions in a 5x5 protocol (CLSI E15-A3)

|  | Plasma p-tau217 |  | Plasma Aβ1-40 |  | Aβ1-42 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Level 1 | Level 2 | Level 1 | Level 2 | Level 1 | Level 2 |
| Mean concentration (pg/mL) | 0.53 | 4.00 | 92.0 | 174.5 | 14.6 | 30.7 |
| Target (pg/mL) | 0.50 | 4.13 | 95 | 182 | 15.3 | 32.4 |
| Imprecision |  |  |  |  |  |  |
| SDwr (pg/mL) | 0.02 | 0.06 | 1.98 | 2.60 | 0.30 | 0.55 |
| CVwr (%) | 2.8 | 1.4 | 2.2 | 1.5 | 2.0 | 1.8 |
| SDbr (pg/mL) | 0.001 | 0.04 | 1.63 | 4.39 | 0.60 | 0.55 |
| CVbr (%) | 0.3 | 1.0 | 1.8 | 2.5 | 4.1 | 1.8 |
| SDwl (pg/mL) | 0.02 | 0.07 | 2.6 | 5.11 | 0.67 | 0.78 |
| CVwl (%) | 2.8 | 1.7 | 2.8 | 2.9 | 4.6 | 2.5 |

Abbreviations: WR: within-run; BR: between-run, WL: within-lab, SD: standard deviation, CV: coefficient of variation, CVwr: within-run coefficient