Comparative Clinical and Biomechanical Assessment of Patient-Specific 3D-Printed Monoblock Versus Commercial Modular Acetabular Prostheses in Hip Arthroplasty for Reconstruction of Severe Acetabular Bone Defects

| Table S1 Intra- and inter-observer reliability of radiographic measurements |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Parameter | Mean Intra-observer ICC | 95% CI | Inter-observer ICC | 95% CI | SEM (mm) | SDD (mm) |
| Pre-operative LLD | 0.95 | 0.897–0.973 | 0.89 | 0.795–0.943 | 4.7 | 13.1 |
| Post-operative LLD | 0.94 | 0.874–0.967 | 0.85 | 0.715–0.921 | 1.0 | 2.9 |
| Pre-operative COR horizontal deviation | 0.94 | 0.874–0.967 | 0.88 | 0.777–0.937 | 2.0 | 5.5 |
| Post-operative COR horizontal deviation | 0.98 | 0.958–0.989 | 0.80 | 0.649–0.893 | 0.5 | 1.3 |
| Pre-operative COR vertical deviation | 0.95 | 0.893–0.972 | 0.89 | 0.794–0.943 | 2.9 | 8.0 |
| Post-operative COR vertical deviation | 0.98 | 0.960–0.990 | 0.89 | 0.791–0.942 | 0.4 | 1.2 |
| Pre-operative FOD | 0.94 | 0.874–0.967 | 0.88 | 0.777–0.937 | 2.2 | 6.1 |
| Post-operative FOD | 0.98 | 0.954–0.988 | 0.81 | 0.662–0.898 | 0.7 | 1.8 |
| ICC, intraclass correlation coefficient; CI, confidence interval; SEM, standard error of measurement; SDD, smallest detectable difference; LLD, limb length discrepancy; COR, center of rotation; FOD, femoral offset distance. Intra-observer ICC was calculated using a two-way random-effects model with absolute agreement (ICC(A,1)). SEM = pooled SD × √(1− mean intra-observer ICC); SDD = 1.96 × √2 × SEM |  |  |  |  |  |  |

| Table S2 General data of the cadaveric specimens |  |  |  |  |
| --- | --- | --- | --- | --- |
| Specimen No. and Site | Gender | Age (years) | BMI (kg/m2) | BMD (g/cm2) |
| 1 Right | Male | 51 | 18.17 | 1.192 |
| 1 Left |  |  |  | 1.136 |
| 2 Right | Male | 48 | 23.83 | 1.158 |
| 2 Left |  |  |  | 1.181 |
| 3 Right | Female | 52 | 20.76 | 1.005 |
| 3 Left |  |  |  | 1.012 |
| 4 Right | Female | 59 | 22.06 | 0.814 |
| 4 Left |  |  |  | 0.821 |
| 5 Right | Female | 49 | 23.03 | 1.022 |
| 5 Left |  |  |  | 1.031 |
| BMI, body mass index; BMD, bone mineral density |  |  |  |  |

| Table S3 Specifications of the stereo-3D-DIC optical measurement system |  |  |
| --- | --- | --- |
| Parameter | Specification | Reference / Note |
| SYSTEM |  |  |
| Developer | Southeast University, China | Camera motion OL&E, 2021 DOI:10.1016/j.optlaseng.2021.106809 |
| Measurement principle | Three-dimensional digital image correlation (3D-DIC) | Sutton et al., 1986 DOI:10.1016/0262-8856(86)90057-5; Chao et al., 1993 DOI:10.1007/BF0232248 |
| Error compensation | Speckle-based camera motion correction | Camera motion OL&E, 2021 DOI:10.1016/j.optlaseng.2021.106809 |
| HARDWARE |  |  |
| Camera model | EoSens84CXP high-speed camera (×2) | Mikrotron GmbH, Germany |
| Lens model | Schneider JADE 2.825C | Schneider Kreuznach, Germany |
| Focal length | 25 mm | — |
| Image sensor | 4/3" CMOS | — |
| Sensor resolution | 2336×1728 pixels | — |
| Full-frame frame rate | 563 fps | — |
| Speckle pattern | Random black-and-white spray-painted pattern | — |
| METROLOGICAL VALIDATION |  |  |
| Calibration institution | Jiangsu Provincial Institute of Metrology | Certificate No. L2023-3179122* |
| Expanded uncertainty | U = 3.3 μm | Coverage factor k = 2 (95% confidence) |
| Reported precision | 1 μm (rounded) | Consistent with uncertainty budget |
| *Calibration certificate available upon request. All DIC-derived displacement values (micromotion and migration) were rounded to the nearest integer micron to ensure that reported precision reflects the instrument's metrological capability rather than software output resolution |  |  |

| Table S4 Material properties defined in the finite element (FE) models |  |  |  |
| --- | --- | --- | --- |
| Components | Materials | Elastic Modulus (MPa) | Poisson’s ratio (v) |
| Cortical bone | Cortical bone | 17,300 | 0.265 |
| Cancellous bone | Cancellous bone | 400 | 0.2 |
| Screws | Titanium alloy | 110,600 | 0.326 |
| Acetabular cup |  |  |  |
| Metal augment |  |  |  |
| Monoblock prosthesis |  |  |  |
| Ceramic femoral head | Ceramic | 350,000 | 0.22 |
| Liner | Polyethylene | 800 | 0.45 |

| Table S5 Permanent migration at the bone–prosthesis interface after cyclic loading and results of the linear mixed-effects model |  |  |  |
| --- | --- | --- | --- |
| Cycles | 3D-printed cup–bone | Modular cup–bone | Modular augment–bone |
| 200 | 34 (31–36) | 48 (48–55) | 38 (35–41) |
| 400 | 35 (33–39) | 107 (104–111) | 57 (51–61) |
| 600 | 34 (32–39) | 140 (136–144) | 81 (76–86) |
| 800 | 86 (82–94) | 181 (178–186) | 135 (125–140) |
| Linear mixed-effects model (3D-printed vs modular cup–bone; specimen as random intercept) |  |  |  |
| LMM effect | β (3D − modular) | 95% CI | p value |
| Group effect | −71.95 | −78.88 to −65.02 | <0.001 |
| Cycle effect | 8.07 | 5.88 to 10.26 | <0.001 |
| Group × cycle interaction | −13.21 | −16.31 to −10.11 | <0.001 |
| EMM difference at 800 cycles (μm) | −111.6 | — | d = 10.0 |
| Wilcoxon signed-rank, 800 cycles | — | — | p = 0.0625 |
| Values are median (interquartile range) of five paired cadaveric specimens. Between-group comparisons were performed using linear mixed-effects models with specimen as a random intercept ( group × cycle); β denotes the estimated difference (3D-printed minus modular cup–bone). EMM, estimated marginal mean; d, standardized effect size (EMM difference divided by the residual standard deviation). The modular augment–bone interface is reported descriptively. —, not applicable |  |  |  |

| Table S6 Maximal inducible micromotion at the bone–prosthesis interface during cyclic loading and results of the linear mixed-effects model |  |  |  |
| --- | --- | --- | --- |
| Cycles | 3D-printed cup–bone | Modular cup–bone | Modular augment–bone |
| 50 | 89 (88–90) | 226 (192–253) | 110 (106–129) |
| 100 | 93 (90–96) | 187 (166–213) | 110 (102–128) |
| 150 | 90 (87–94) | 234 (197–273) | 122 (120–142) |
| 200 | 90 (88–93) | 199 (138–244) | 144 (131–165) |
| 250 | 79 (78–85) | 207 (175–262) | 106 (97–122) |
| 300 | 86 (84–90) | 179 (166–207) | 95 (88–111) |
| 350 | 84 (83–85) | 236 (199–278) | 119 (114–140) |
| 400 | 87 (87–93) | 207 (140–255) | 118 (114–140) |
| 450 | 81 (78–82) | 178 (151–273) | 104 (101–120) |
| 500 | 93 (88–95) | 165 (149–175) | 99 (93–118) |
| 550 | 95 (91–100) | 182 (154–240) | 119 (111–139) |
| 600 | 95 (87–98) | 189 (173–206) | 175 (164–200) |
| 650 | 111 (107–112) | 187 (158–246) | 167 (163–188) |
| 700 | 116 (113–121) | 175 (158–190) | 180 (166–203) |
| 750 | 125 (117–127) | 334 (282–340) | 240 (229–261) |
| 800 | 126 (119–128) | 289 (248–296) | 229 (216–253) |
| Linear mixed-effects model (3D-printed vs modular cup–bone; specimen as random intercept) |  |  |  |
| LMM effect | β (3D − modular) | 95% CI | p value |
| Group effect | −113.75 | −126.88 to −100.61 | <0.001 |
| Load stage effect | 33.67 | 9.83 to 57.50 | 0.006 |
| Group × load interaction | −9.00 | −35.25 to 17.25 | 0.502 |
| EMM difference, high-load stage (μm) | −122.8 | — | d = 3.35 |
| Wilcoxon signed-rank, 800 cycles | — | — | p = 0.0625 |
| Values are median (interquartile range) of five paired cadaveric specimens. Between-group comparisons were performed using linear mixed-effects models with specimen as a random intercept (group × load stage); β denotes the estimated difference (3D-printed minus modular cup–bone). EMM, estimated marginal mean; d, standardized effect size (EMM difference divided by the residual standard deviation). The modular augment–bone interface is reported descriptively. —, not applicable |  |  |  |

| Table S7 Peak strain on the acetabular bone |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Measurement (με) | Point 1 Strain 1 | Point 1 Strain 2 | Point 2 Strain 1 | Point 2 Strain 2 | Point 3 Strain 1 | Point 3 Strain 2 |
| 3D-printed prostheses 50-500N | 184.57 (177.96–194.95) | 103.89 (97.62–107.20) | 301.23 (292.67–322.33) | 214.56 (204.79–220.18) | 452.32 (430.18–481.62) | 88.68 (82.22–90.16) |
| Modular prostheses 50-500N | 38.23 (37.02–39.85) | 76.23 (73.48–79.23) | 231.69 (222.01–236.34) | 186.79 (181.13–190.73) | 276.22 (270.23–281.18) | 119.72 (116.01–123.23) |
| 3D-printed prostheses 84-840N | 334.76 (330.23–339.96) | 196.50 (183.90–203.70) | 589.06 (581.40–598.73) | 375.55 (371.28–383.73) | 908.87 (883.40–917.90) | 166.08 (164.51–171.01) |
| Modular prostheses 84-840N | 47.96 (45.28–50.68) | 139.89 (136.39–141.80) | 369.24 (360.62–373.85) | 303.32 (299.85–307.18) | 428.82 (417.12–436.96) | 190.06 (187.57–192.46) |
| Data are given as median (interquartile range) [M (Q1, Q3)]. Strain data are reported descriptively; Strain 1, Compressive; Strain 2. Tensile no formal statistical testing was performed for strain outcomes given the small number of paired specimens (n = 5) |  |  |  |  |  |  |

Fig. S1 Preoperative design and long-term radiographic outcomes of 3D-printed monoblock acetabular reconstruction in a representative case. A representative case with a Paprosky type IIIB acetabular defect. a Preoperative radiograph demonstrating the severe acetabular bone loss and compromised hip center. b Immediate postoperative radiograph showing the 3D-printed monoblock prosthesis in situ with satisfactory restoration of the hip center of rotation. c Radiograph at the final follow-up, confirming stable prosthesis fixation with no signs of loosening, migration, or osteolysis. d Preoperative planning and design schematic

Fig. S2 Preoperative design and long-term radiographic outcomes of 3D-printed monoblock acetabular reconstruction in a representative case. A representative case with a Paprosky type IIIB acetabular defect. a Preoperative radiograph demonstrating the severe acetabular bone loss and compromised hip center. b Immediate postoperative radiograph showing the 3D-printed monoblock prosthesis in situ with satisfactory restoration of the hip center of rotation. c Radiograph at the final follow-up, confirming stable prosthesis fixation with no signs of loosening, migration, or osteolysis. d Preoperative planning and design schematic

Fig. S3 Preoperative design and long-term radiographic outcomes of 3D-printed monoblock acetabular reconstruction in a representative case. A representative case with a Paprosky type IIC acetabular defect. a Preoperative radiograph demonstrating the severe acetabular bone loss and compromised hip center. b Immediate postoperative radiograph showing the 3D-printed monoblock prosthesis in situ with satisfactory restoration of the hip center of rotation. c Radiograph at the final follow-up, confirming stable prosthesis fixation with no signs of loosening, migration, or osteolysis. d Preoperative planning and design schematic

Fig. S4 Comparison of periacetabular bone strain under the two loading protocols. a, b Strain distribution under the low-magnitude load (50-500 N) for the modular (a) and 3D-printed (b) groups. c, d Strain distribution under the high-magnitude load (84-840 N) for the modular (c) and 3D-printed (d) groups.