## Crossover MV Data
[5 empty] | Respiratory Mechanics | [7 empty] | ABG |  |  |  | Pulse Oximeter |  | EtCO2 Monitor | BP monitor
| Order of Ventilation | Timepoint (min) | VT (mL) | VT (mL/kg) | Target PIP (cm H2O) | PIP (cm H2O) | Delta PIP (cm H2O) | Target PEEP (cm H2O) | PEEP (cm H2O) | Delta PEEP (cm H2O) | RR (breaths/min) | I:E Ratio | PAO2 (mm Hg) | PACO2 (mm Hg) | pH | cLac (mmol/L) | SpO2 | Pulse Rate (beats/min) | Peak CO2 (mmHg) | MAP (mm Hg)
Rabbit MV1 | Carestation | 0 | 17.7794 | 6.3725 | 11 | 11.3653 | 0.36526 | 3 | 2.9558 | -0.044211 | 39.9726 | 1.0066 | NA |  |  |  | 97.4 | 142.4 | 34.5963 | 59.1589
|  | 10 | 17.6338 | 6.3204 |  | 11.4574 | 0.45744 |  | 3.0197 | 0.019744 | 39.9766 | 1.0112
|  | 20 | 18.1502 | 6.5054 |  | 11.3408 | 0.34077 |  | 2.8956 | -0.10436 | 39.9785 | 1.009
|  | 30 | 17.6817 | 6.3375 |  | 11.1621 | 0.16205 |  | 2.699 | -0.30103 | 39.9818 | 1.0092 | 108 | 34.1 | 7.563 | 1.5
|  | 40 | 18.0479 | 6.4688 |  | 11.1654 | 0.16538 |  | 2.6767 | -0.32333 | 39.9741 | 1.0164 | NA
|  | 50 | 17.8963 | 6.4144 |  | 11.0326 | 0.032632 |  | 2.6063 | -0.39368 | 39.9867 | 1.0169
|  | 60 | 18.2896 | 6.5554 |  | 10.9853 | -0.014737 |  | 2.4976 | -0.50237 | 39.9989 | 0.95558 | 103 | 31.6 | 7.57 | 1.6
| FALCON | 0 | 16.8413 | 6.0363 |  | 11.2723 | 0.27231 |  | 3.1251 | 0.12513 | 39.8219 | 0.70291 | NA |  |  |  | 98.2 | 137.4 | 42.4466 | 60.864
|  | 10 | 16.1121 | 5.775 |  | 11.0995 | 0.099474 |  | 3.0497 | 0.049737 | 39.8437 | 0.6935
|  | 20 | 16.3831 | 5.8721 |  | 11.6382 | 0.63821 |  | 3.0977 | 0.097692 | 39.8356 | 0.67846
|  | 30 | 14.646 | 5.2495 |  | 11.321 | 0.32103 |  | 3.019 | 0.018974 | 39.8424 | 0.66906 | 89.2 | 43 | 7.461 | 1.3
|  | 40 | 12.9006 | 4.6239 |  | 9.7771 | -1.2229 |  | 2.6297 | -0.37026 | 39.8397 | 0.68559 | NA
|  | 50 | 15.0604 | 5.398 |  | 11.309 | 0.30897 |  | 2.8638 | -0.13615 | 39.8234 | 0.67494
|  | 60 | 16.8968 | 6.0562 |  | 12.1471 | 1.1471 |  | 2.6416 | -0.35842 | 39.8437 | 0.68307 | 107 | 40.1 | 7.474 | 1.2
Rabbit MV2 | Carestation | 0 | 23.623 | 7.4286 | 11 | 11.5108 | 0.51077 | 3 | 3.0331 | 0.033077 | 39.7439 | 1.0011 | NA |  |  |  | 95.5 | 164.1 | 31.2088 | 54.9569
|  | 10 | 23.1839 | 7.2905 |  | 11.4803 | 0.48026 |  | 2.9268 | -0.073158 | 39.9744 | 1.0106
|  | 20 | 22.6295 | 7.1162 |  | 11.2641 | 0.2641 |  | 2.6477 | -0.35231 | 39.9805 | 1.005
|  | 30 | 22.0259 | 6.9264 |  | 11.2576 | 0.25763 |  | 2.6276 | -0.37237 | 39.6533 | 0.9929 | 74.1 | 39 | 7.624 | 1.7
|  | 40 | 21.8287 | 6.8644 |  | 11.0859 | 0.085897 |  | 2.4831 | -0.51692 | 39.9796 | 1.009 | NA
|  | 50 | 21.7131 | 6.828 |  | 10.9292 | -0.070789 |  | 2.3421 | -0.65789 | 39.9995 | 1.0173
|  | 60 | 21.2516 | 6.6829 |  | 10.8262 | -0.17379 |  | 2.2317 | -0.76828 | 39.9797 | 1.0154 | 73 | 37.7 | 7.62 | 2
| FALCON | 0 | 20.1088 | 6.3235 |  | 12.1059 | 1.1059 |  | 3.0695 | 0.069487 | 39.8273 | 0.70756 | NA |  |  |  | 88.5 | 164.4 | 46.3848 | 53.0151
|  | 10 | 18.5355 | 5.8288 |  | 11.2961 | 0.29605 |  | 2.9011 | -0.098947 | 39.8175 | 0.83099
|  | 20 | 19.0764 | 5.9989 |  | 11.825 | 0.825 |  | 2.7697 | -0.23028 | 38.0874 | 0.6339
|  | 30 | 18.2148 | 5.7279 |  | 11.792 | 0.792 |  | 2.5926 | -0.40743 | 37.5302 | 0.62014 | 69.1 | 52.4 | 7.48 | 2.2
|  | 40 | 17.0797 | 5.371 |  | 11.7 | 0.7 |  | 2.6426 | -0.35743 | 38.0606 | 0.62252 | NA
|  | 50 | 16.8882 | 5.3108 |  | 11.2908 | 0.29083 |  | 2.2081 | -0.79194 | 38.5079 | 0.62163
|  | 60 | 17.1652 | 5.3978 |  | 11.7682 | 0.76821 |  | 2.2515 | -0.74846 | 39.8493 | 0.67976 | 71.2 | 56.1 | 7.448 | 2.3
Rabbit MV3 | FALCON | 0 | 18.6815 | 6.0069 | 12 | 12.0685 | 0.068462 | 3 | 3.4149 | 0.41487 | 39.9371 | 0.69648 | NA |  |  |  | 94.7 | 176.9 | 52.1381 | 67.9578
|  | 10 | 24.4388 | 7.8581 |  | 13.0013 | 1.0013 |  | 2.9631 | -0.036923 | 39.9604 | 0.62659
|  | 20 | 21.8037 | 7.0108 |  | 13.4626 | 1.4626 |  | 2.8738 | -0.12615 | 39.9383 | 0.69379
|  | 30 | 20.8143 | 6.6927 |  | 12.9253 | 0.92526 |  | 2.8268 | -0.17316 | 39.9825 | 0.71807 | 76.8 | 49.4 | 7.448 | 2.8
|  | 40 | 23.5088 | 7.5591 |  | 12.8508 | 0.85077 |  | 2.6864 | -0.31359 | 40.019 | 0.67882 | NA
|  | 50 | 23.5056 | 7.5581 |  | 12.7918 | 0.79179 |  | 2.5062 | -0.49385 | 39.9396 | 0.68994
|  | 60 | 21.2708 | 6.8395 |  | 12.9741 | 0.9741 |  | 2.9803 | -0.019744 | 39.9504 | 0.65008 | 78 | 47.8 | 7.458 | 3.5
| Carestation | 0 | 21.2289 | 6.826 |  | 12.721 | 0.72103 |  | 3.0677 | 0.067692 | 39.9885 | 1.0335 | NA |  |  |  | 99 | 167.3 | 36.7169 | 76.9087
|  | 10 | 21.4196 | 6.8873 |  | 12.5863 | 0.58632 |  | 2.9011 | -0.098947 | 39.4674 | 0.99093
|  | 20 | 21.0066 | 6.7545 |  | 12.5397 | 0.53974 |  | 2.71 | -0.29 | 39.9805 | 1.0254
|  | 30 | 20.3403 | 6.5403 |  | 12.4285 | 0.42846 |  | 2.5818 | -0.41821 | 39.9801 | 1.0204 | 100 | 32.2 | 7.591 | 3.5
|  | 40 | 20.1465 | 6.478 |  | 12.3016 | 0.30158 |  | 2.4566 | -0.54342 | 39.6509 | 1.0303 | NA
|  | 50 | 19.8975 | 6.3979 |  | 12.2672 | 0.26718 |  | 2.3359 | -0.6641 | 39.9955 | 1.0362
|  | 60 | 19.6009 | 6.3025 |  | 12.1303 | 0.13026 |  | 2.1949 | -0.80513 | 39.9942 | 1.0354 | 96.1 | 31 | 7.587 | 3.5
Rabbit MV4 | FALCON | 0 | 22.6759 | 7.5086 | 13 | 11.9889 | -1.0111 | 2 | 3.0176 | 1.0176 | 39.4652 | 0.69128 | NA |  |  |  | 90.8 | 201.7 | 37.8596 | 61.7427
|  | 10 | 29.9034 | 9.9018 |  | 12.5632 | -0.43676 |  | 1.3943 | -0.60568 | 39.258 | 0.71064
|  | 20 | 30.8949 | 10.2301 |  | 13.15 | 0.15 |  | 1.2433 | -0.75667 | 39.9407 | 0.7477
|  | 30 | 31.2274 | 10.3402 |  | 13.2611 | 0.26105 |  | 1.0516 | -0.94842 | 39.415 | 0.7251 | 65.6 | 41.5 | 7.562 | 1.2
|  | 40 | 29.9709 | 9.9241 |  | 12.9453 | -0.054737 |  | 1.0095 | -0.99053 | 39.9426 | 0.7379 | NA
|  | 50 | 30.6546 | 10.1505 |  | 13.1328 | 0.13282 |  | 0.93615 | -1.0638 | 39.9637 | 0.7345
|  | 60 | 29.5116 | 9.7721 |  | 12.7795 | -0.22051 |  | 0.79282 | -1.2072 | 39.9346 | 0.74547 | 73.2 | 36 | 7.592 | 1.4
| Carestation | 0 | 22.3276 | 7.3932 |  | 13.1831 | 0.18308 |  | 2.6031 | 0.60308 | 40.02 | 1.0813 | NA |  |  |  | 93.4 | 197.6 | 27.7464 | 57.6683
|  | 10 | 23.6021 | 7.8153 |  | 13.2792 | 0.27923 |  | 2.0928 | 0.092821 | 40.0103 | 1.0744
|  | 20 | 23.2916 | 7.7125 |  | 13.1566 | 0.15658 |  | 1.9924 | -0.0076316 | 40.0642 | 1.0862
|  | 30 | 22.8236 | 7.5575 |  | 12.9908 | -0.0092105 |  | 1.8042 | -0.19579 | 40.0077 | 1.0902 | 67.5 | 31.2 | 7.625 | 1.8
|  | 40 | 22.5852 | 7.4785 |  | 12.9726 | -0.027368 |  | 1.7682 | -0.23184 | 39.9978 | 1.0825 | NA
|  | 50 | 22.6581 | 7.5027 |  | 12.8262 | -0.17385 |  | 1.6472 | -0.35282 | 39.9289 | 1.0921
|  | 60 | 23.0366 | 7.628 |  | 12.8854 | -0.11462 |  | 1.6479 | -0.35205 | 40.0034 | 1.0822 | NA
Rabbit MV5 | Carestation | 0 | 25.2019 | 8.457 | 13 | 13.6716 | 0.67158 | 5 | 5.1913 | 0.19132 | 40.0271 | 1.0385 | NA |  |  |  | 96.9 | 166.4 | 30.6781 | 41.1548
|  | 10 | 24.124 | 8.0953 |  | 13.5379 | 0.53789 |  | 4.9868 | -0.013158 | 39.9947 | 1.0243
|  | 20 | 23.6295 | 7.9294 |  | 13.4808 | 0.48079 |  | 4.8697 | -0.13026 | 39.9939 | 1.0341
|  | 30 | 22.6939 | 7.6154 |  | 13.5003 | 0.50026 |  | 4.8915 | -0.10846 | 40.0021 | 1.0443 | 104 | 30.4 | 7.457 | 4.8
|  | 40 | 21.8925 | 7.3465 |  | 13.4539 | 0.45395 |  | 4.7776 | -0.22237 | 39.9667 | 1.0234 | NA
|  | 50 | 21.1959 | 7.1127 |  | 13.3041 | 0.3041 |  | 4.6192 | -0.38077 | 39.9964 | 1.0272
|  | 60 | 20.4045 | 6.8471 |  | 13.2239 | 0.22395 |  | 4.4992 | -0.50079 | 39.9888 | 1.027 | 115 | 27.9 | 7.477 | 4.6
| FALCON | 0 | 22.8949 | 7.6828 |  | 14.1997 | 1.1997 |  | 4.8962 | -0.10385 | 39.9336 | 0.65904 | NA |  |  |  | 92.1 | 170.9 | 44.5513 | 44.2708
|  | 10 | 21.4909 | 7.2117 |  | 14.5056 | 1.5056 |  | 5.2597 | 0.25974 | 39.935 | 0.59149
|  | 20 | 19.3015 | 6.477 |  | 13.7918 | 0.79184 |  | 4.9245 | -0.075526 | 39.6334 | 0.56032
|  | 30 | 18.1437 | 6.0885 |  | 14.0142 | 1.0142 |  | 4.5447 | -0.45526 | 39.8762 | 0.59566 | 84.2 | 39.6 | 7.355 | 4.5
|  | 40 | 15.7771 | 5.2943 |  | 14.0347 | 1.0347 |  | 5.0853 | 0.085263 | 39.9221 | 0.5547 | NA
|  | 50 | 14.6783 | 4.9256 |  | 13.8077 | 0.80769 |  | 5.229 | 0.22897 | 40.2298 | 0.55386
|  | 60 | 16.0336 | 5.3804 |  | 13.9256 | 0.92564 |  | 4.4433 | -0.55667 | 39.9138 | 0.60887 | 74.6 | 45.7 | 7.278 | 4.9

## Alter RR Data
| RR (breaths/min) |  |  |  | VT (mL) |  |  |  | VT (mL/kg) |  |  |  | Minute Ventilation (mL/kg/min) |  |  |  | PIP (cm H2O) |  |  |  | Delta PIP (cm H2O) |  |  |  | PEEP (cm H2O) |  |  |  | Delta PEEP (cm H2O) |  |  |  | EtCO2 (mm Hg)
Target RR (breaths/min) | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120 | 30 | 40 | 60 | 120
| 30.0646 | 39.6852 | 60.048 | 115.429 | 16.801 | 15.5435 | 14.165 | 9.2042 | 6.0219 | 5.5712 | 5.0771 | 3.299 | 181.0449 | 221.0926 | 304.8678 | 380.8017 | 11.83 | 11.17 | 12.7 | 12.51 | 0.83 | 0.17 | 1.7 | 1.51 | 2.49 | 2.5 | 2.86 | 3.81 | -0.51 | -0.5 | -0.14 | 0.81 | 47 | 43 | 36 | 32
| 29.7841 | 40.0534 | 59.2008 | 122.1498 | 19.3311 | 15.8377 | 14.0701 | 8.7197 | 6.9287 | 5.6766 | 5.0431 | 3.1253 | 206.3655 | 227.3668 | 298.5533 | 381.7587 | 11.77 | 11.21 | 12.61 | 12.65 | 0.77 | 0.21 | 1.61 | 1.65 | 1.71 | 2.53 | 3.13 | 3.97 | -1.29 | -0.47 | 0.13 | 0.97 | 47 | 43 | 36 | 32
| 29.8611 | 39.6694 | 60.1745 | 121.0165 | 19.0282 | 15.4034 | 12.6616 | 8.4146 | 6.8201 | 5.5209 | 4.5382 | 3.016 | 203.657 | 219.0121 | 273.0841 | 364.9843 | 11.86 | 11.05 | 11.75 | 11.89 | 0.86 | 0.05 | 0.75 | 0.89 | 1.7 | 2.55 | 3.12 | 4 | -1.3 | -0.45 | 0.12 | 1 | 47 | 43 | 36 | 32
| 30.0045 | 39.9334 | 60.3318 | 118.5771 | 19.2444 | 15.4389 | 12.6951 | 9.7857 | 6.8976 | 5.5337 | 4.5502 | 3.5074 | 206.9603 | 220.9786 | 274.523 | 415.8992 | 11.77 | 11.07 | 11.64 | 12.05 | 0.77 | 0.07 | 0.64 | 1.05 | 1.63 | 2.47 | 3.06 | 3.95 | -1.37 | -0.53 | 0.06 | 0.95 | 47 | 43 | 36 | 32
| 29.5217 | 39.772 | 59.2944 | 122.7998 | 19.2092 | 15.297 | 13.1504 | 9.6615 | 6.885 | 5.4828 | 4.7134 | 3.4629 | 203.2575 | 218.0622 | 279.4777 | 425.2421 | 11.8 | 11.18 | 12.83 | 13.09 | 0.8 | 0.18 | 1.83 | 2.09 | 1.75 | 2.64 | 3.48 | 4.11 | -1.25 | -0.36 | 0.48 | 1.11 | 47 | 43 | 36 | 32
| 29.5698 | 39.9228 | 59.0261 | 117.9477 | 17.7758 | 15.5056 | 13.3201 | 8.6256 | 6.3713 | 5.5576 | 4.7742 | 3.0916 | 188.3966 | 221.8742 | 281.8048 | 364.6505 | 11.87 | 11.21 | 12.35 | 12.66 | 0.87 | 0.21 | 1.35 | 1.66 | 2.21 | 2.49 | 2.99 | 4.05 | -0.79 | -0.51 | -0.01 | 1.05 | 47 | 43 | 36 | 32
| 29.9461 | 39.7456 | 61.0438 | 116.0093 | 16.2292 | 15.8259 | 12.5569 | 9.6204 | 5.8169 | 5.6724 | 4.5007 | 3.4482 | 174.1943 | 225.4517 | 274.74 | 400.0215 | 11.81 | 11.15 | 11.59 | 11.75 | 0.81 | 0.15 | 0.59 | 0.75 | 2.59 | 2.58 | 2.99 | 3.94 | -0.41 | -0.42 | -0.01 | 0.94 | 47 | 43 | 36 | 32
| 29.8656 | 39.6747 | 59.2242 | 124.818 | 16.2756 | 15.7026 | 14.2424 | 9.4521 | 5.8335 | 5.6282 | 5.1048 | 3.3878 | 174.2223 | 223.2961 | 302.3267 | 422.8646 | 11.84 | 11.15 | 12.78 | 12.3 | 0.84 | 0.15 | 1.78 | 1.3 | 2.62 | 2.54 | 2.89 | 3.94 | -0.38 | -0.46 | -0.11 | 0.94 | 47 | 43 | 36 | 32
| 29.8092 | 39.984 | 59.952 | 116.9819 | 16.0295 | 15.5882 | 14.6158 | 9.6469 | 5.7453 | 5.5872 | 5.2386 | 3.4577 | 171.2637 | 223.3978 | 314.0675 | 404.4864 | 11.8 | 11.14 | 12.47 | 13 | 0.8 | 0.14 | 1.47 | 2 | 2.64 | 2.6 | 2.98 | 4.05 | -0.36 | -0.4 | -0.02 | 1.05 | 47 | 43 | 36 | 32
| 29.8686 | 39.976 | 59.1191 | 117.0503 | 16.0003 | 15.6106 | 13.9551 | 9.4131 | 5.7349 | 5.5952 | 5.0018 | 3.3739 | 171.2928 | 223.6739 | 295.7036 | 394.9138 | 11.75 | 11.13 | 12.81 | 12.96 | 0.75 | 0.13 | 1.81 | 1.96 | 2.65 | 2.44 | 2.96 | 4.02 | -0.35 | -0.56 | -0.04 | 1.02 | 47 | 43 | 36 | 32
| 30.0873 | 39.7825 | 60.6183 | 125.9182 | 16.198 | 15.5732 | 13.7239 | 8.2553 | 5.8057 | 5.5818 | 4.919 | 2.9589 | 174.6782 | 222.0583 | 298.1798 | 372.5785 | 11.83 | 11.16 | 12.45 | 11.89 | 0.83 | 0.16 | 1.45 | 0.89 | 2.76 | 2.71 | 3.05 | 3.95 | -0.24 | -0.29 | 0.05 | 0.95 | 47 | 43 | 36 | 32
| 29.5567 | 39.7035 | 59.5948 | 114.7008 | 16.0367 | 15.409 | 14.0805 | 10.383 | 5.7479 | 5.5229 | 5.0468 | 3.7215 | 169.8893 | 219.2801 | 300.7605 | 426.8602 | 11.73 | 11.16 | 12.82 | 12.3 | 0.73 | 0.16 | 1.82 | 1.3 | 2.61 | 2.56 | 3.07 | 3.92 | -0.39 | -0.44 | 0.07 | 0.92 | 47 | 43 | 36 | 32
| 9.2002 | 39.883 | 59.5415 | 121.6298 | 13.6891 | 15.5879 | 14.3489 | 9.8182 | 4.9065 | 5.5871 | 5.143 | 3.5191 | 45.1406 | 222.8294 | 306.2207 | 428.0233 | 11.97 | 11.24 | 12.51 | 12.85 | 0.97 | 0.24 | 1.51 | 1.85 | 3.52 | 2.47 | 2.95 | 3.86 | 0.52 | -0.53 | -0.05 | 0.86 | 47 | 43 | 36 | 32
| 41.8206 | 39.7246 | 59.1424 | 120.6515 | 10.653 | 15.0818 | 14.4994 | 8.9045 | 3.8183 | 5.4057 | 5.1969 | 3.1916 | 159.6827 | 214.738 | 307.3578 | 385.0697 | 7.35 | 11.1 | 12.91 | 12.85 | -3.65 | 0.1 | 1.91 | 1.85 | 0.28 | 2.64 | 2.97 | 4.08 | -2.72 | -0.36 | -0.03 | 1.08 | 47 | 43 | 36 | 32
| 28.9729 | 39.9202 | 60.1323 | 117.0047 | 22.4611 | 15.3188 | 14.3706 | 9.733 | 8.0506 | 5.4906 | 5.1508 | 3.4885 | 233.2483 | 219.1864 | 309.7274 | 408.1754 | 12.01 | 11.13 | 12.54 | 11.88 | 1.01 | 0.13 | 1.54 | 0.88 | 0.8 | 2.53 | 3.08 | 3.91 | -2.2 | -0.47 | 0.08 | 0.91 | 47 | 43 | 36 | 32
| 29.8671 | 39.992 | 60.423 | 120.7729 | 16.4975 | 15.0685 | 13.8754 | 9.2971 | 5.9131 | 5.4009 | 4.9733 | 3.3323 | 176.6064 | 215.9928 | 300.5 | 402.4528 | 11.83 | 11.15 | 12.9 | 12.4 | 0.83 | 0.15 | 1.9 | 1.4 | 2.69 | 2.83 | 3.05 | 3.62 | -0.31 | -0.17 | 0.05 | 0.62 | 47 | 43 | 36 | 32
| 29.8626 | 39.6773 | 58.9971 | 119.8083 | 16.4306 | 15.2121 | 14.8465 | 9.1588 | 5.8891 | 5.4524 | 5.3213 | 3.2827 | 175.8641 | 216.3352 | 313.9424 | 393.2961 | 12 | 11.14 | 12.56 | 12.73 | 1 | 0.14 | 1.56 | 1.73 | 2.62 | 2.57 | 3.03 | 3.93 | -0.38 | -0.43 | 0.03 | 0.93 | 47 | 43 | 36 | 32
| 29.8998 | 40.0614 | 60.1383 | 119.2132 | 16.0384 | 15.0904 | 14.5053 | 9.409 | 5.7485 | 5.4088 | 5.199 | 3.3724 | 171.8805 | 216.6827 | 312.6618 | 402.0334 | 11.73 | 11.1 | 12.92 | 12.35 | 0.73 | 0.1 | 1.92 | 1.35 | 2.74 | 2.81 | 3.05 | 3.86 | -0.26 | -0.19 | 0.05 | 0.86 | 47 | 43 | 36 | 31
| 29.8894 | 39.5179 | 60.006 | 119.2606 | 15.7231 | 15.6064 | 13.7273 | 9.3863 | 5.6355 | 5.5937 | 4.9202 | 3.3643 | 168.4423 | 221.0513 | 295.2407 | 401.2229 | 11.88 | 11.11 | 12.65 | 12.54 | 0.88 | 0.11 | 1.65 | 1.54 | 2.86 | 2.51 | 3.27 | 4.04 | -0.14 | -0.49 | 0.27 | 1.04 | 47 | 43 | 36 | 31
| 29.9118 | 40.016 | 59.512 | 115.8301 | 16.2121 | 15.393 | 14.1969 | 9.0423 | 5.8108 | 5.5172 | 5.0885 | 3.241 | 173.8106 | 220.7771 | 302.8272 | 375.4016 | 11.86 | 11.12 | 12.81 | 12.56 | 0.86 | 0.12 | 1.81 | 1.56 | 2.68 | 2.76 | 3.06 | 3.97 | -0.32 | -0.24 | 0.06 | 0.97 | 47 | 43 | 36 | 31
| 29.8389 | 39.6275 | 59.9401 | 124.1979 | 16.1139 | 15.1416 | 13.6401 | 8.8583 | 5.7756 | 5.4271 | 4.8889 | 3.175 | 172.3376 | 215.0625 | 293.0427 | 394.3326 | 11.91 | 11.03 | 12.64 | 12.8 | 0.91 | 0.03 | 1.64 | 1.8 | 2.71 | 2.64 | 2.97 | 3.85 | -0.29 | -0.36 | -0.03 | 0.85 | 47 | 43 | 36 | 31
| 29.9297 | 40.1391 | 59.4001 | 117.7163 | 16.1283 | 15.5414 | 14.1969 | 9.415 | 5.7808 | 5.5704 | 5.0885 | 3.3746 | 173.0162 | 223.591 | 302.2567 | 397.2407 | 11.83 | 11.08 | 12.86 | 12.14 | 0.83 | 0.08 | 1.86 | 1.14 | 2.72 | 2.57 | 2.97 | 3.94 | -0.28 | -0.43 | -0.03 | 0.94 | 47 | 43 | 36 | 31
| 29.8122 | 39.5987 | 59.6896 | 118.6944 | 16.1094 | 15.4069 | 14.542 | 9.9006 | 5.774 | 5.5222 | 5.2122 | 3.5486 | 172.1351 | 218.6712 | 311.1136 | 421.1998 | 11.93 | 10.99 | 12.68 | 12.6 | 0.93 | -0.01 | 1.68 | 1.6 | 2.59 | 2.46 | 3.09 | 4.04 | -0.41 | -0.54 | 0.09 | 1.04 | 47 | 43 | 36 | 31
| 29.882 | 39.9787 | 60.1564 | 121.1632 | 16.3367 | 15.6109 | 13.1208 | 9.1251 | 5.8554 | 5.5953 | 4.7028 | 3.2706 | 174.9721 | 223.6934 | 282.9034 | 396.2817 | 11.86 | 11.16 | 12.9 | 12.18 | 0.86 | 0.16 | 1.9 | 1.18 | 2.57 | 2.53 | 3.13 | 3.7 | -0.43 | -0.47 | 0.13 | 0.7 | 47 | 43 | 36 | 31
| 29.8374 | 39.96 | 58.9912 | 123.7879 | 16.1539 | 15.2653 | 13.6048 | 7.8068 | 5.7899 | 5.4714 | 4.8763 | 2.7982 | 172.7565 | 218.6389 | 287.6569 | 346.3772 | 11.9 | 11.15 | 12.84 | 12.7 | 0.9 | 0.15 | 1.84 | 1.7 | 2.74 | 2.61 | 3.26 | 3.96 | -0.26 | -0.39 | 0.26 | 0.96 | 47 | 43 | 36 | 31
| 30.0105 | 39.838 | 60.4108 | 119.1895 | 15.4662 | 15.4308 | 13.7839 | 9.9963 | 5.5435 | 5.5308 | 4.9405 | 3.5829 | 166.362 | 220.3344 | 298.457 | 427.0441 | 11.86 | 11.23 | 12.75 | 12.41 | 0.86 | 0.23 | 1.75 | 1.41 | 2.91 | 2.45 | 2.96 | 3.95 | -0.09 | -0.55 | -0.04 | 0.95 | 47 | 42 | 36 | 31
| 29.8433 | 39.672 | 59.7134 | 113.5074 | 15.9769 | 15.4948 | 13.8226 | 10.0949 | 5.7265 | 5.5537 | 4.9543 | 3.6182 | 170.8978 | 220.3263 | 295.8404 | 410.6954 | 11.93 | 11.24 | 12.73 | 12.8 | 0.93 | 0.24 | 1.73 | 1.8 | 2.57 | 2.51 | 3.15 | 4.17 | -0.43 | -0.49 | 0.15 | 1.17 | 47 | 42 | 36 | 32
| 29.8463 | 39.9574 | 59.4766 | 121.6545 | 16.0712 | 15.0127 | 13.843 | 8.2681 | 5.7603 | 5.3809 | 4.9616 | 2.9635 | 171.9228 | 215.0071 | 295.1011 | 360.5223 | 11.93 | 11.16 | 12.67 | 12.47 | 0.93 | 0.16 | 1.67 | 1.47 | 2.7 | 2.64 | 3.05 | 3.88 | -0.3 | -0.36 | 0.05 | 0.88 | 47 | 42 | 36 | 32
| 29.9133 | 39.7931 | 60.2289 | 119.6411 | 16.1289 | 15.2418 | 13.6093 | 9.831 | 5.781 | 5.463 | 4.8779 | 3.5236 | 172.9274 | 217.3903 | 293.7903 | 421.5726 | 11.9 | 11 | 12.63 | 13.02 | 0.9 | 0 | 1.63 | 2.02 | 2.55 | 2.57 | 2.91 | 3.98 | -0.45 | -0.43 | -0.09 | 0.98 | 47 | 42 | 36 | 32
| 29.9745 | 39.9361 | 59.755 | 119.5457 | 16.0972 | 14.7416 | 13.6702 | 8.4873 | 5.7696 | 5.2837 | 4.8997 | 3.0421 | 172.9407 | 211.0118 | 292.7832 | 363.6647 | 11.92 | 11.16 | 12.72 | 12.07 | 0.92 | 0.16 | 1.72 | 1.07 | 2.62 | 2.6 | 3.24 | 4.05 | -0.38 | -0.4 | 0.24 | 1.05 | 47 | 42 | 36 | 32
| 29.8389 | 39.4971 | 59.3413 | 120.2646 | 16.3706 | 15.0557 | 13.8223 | 9.0149 | 5.8676 | 5.3963 | 4.9542 | 3.2311 | 175.0825 | 213.1388 | 293.9907 | 388.5908 | 11.91 | 11.2 | 12.69 | 12.04 | 0.91 | 0.2 | 1.69 | 1.04 | 2.73 | 2.49 | 3.13 | 4 | -0.27 | -0.51 | 0.13 | 1 | 47 | 42 | 36 | 32
| 29.8522 | 40.0748 | 60.1202 | 118.0405 | 16.2443 | 14.9366 | 13.2769 | 9.1555 | 5.8223 | 5.3536 | 4.7587 | 3.2815 | 173.8097 | 214.5453 | 286.0967 | 387.3529 | 11.92 | 10.94 | 12.44 | 12.24 | 0.92 | -0.06 | 1.44 | 1.24 | 2.61 | 2.48 | 3.23 | 3.67 | -0.39 | -0.52 | 0.23 | 0.67 | 47 | 41 | 36 | 32
| 29.8344 | 39.8248 | 59.0319 | 117.7625 | 16.2091 | 15.3739 | 14.0532 | 9.3029 | 5.8097 | 5.5103 | 5.037 | 3.3344 | 173.329 | 219.4483 | 297.3426 | 392.6639 | 11.84 | 11.11 | 12.81 | 13.01 | 0.84 | 0.11 | 1.81 | 2.01 | 2.66 | 2.62 | 2.95 | 4.05 | -0.34 | -0.38 | -0.05 | 1.05 | 47 | 41 | 36 | 32
| 29.9312 | 39.6799 | 60.8026 | 121.6792 | 16.0868 | 15.6573 | 13.8658 | 8.8743 | 5.7659 | 5.6119 | 4.9698 | 3.1807 | 172.5797 | 222.6806 | 302.1788 | 387.0307 | 11.8 | 11.1 | 12.54 | 11.89 | 0.8 | 0.1 | 1.54 | 0.89 | 2.64 | 2.42 | 3.11 | 3.96 | -0.36 | -0.58 | 0.11 | 0.96 | 47 | 41 | 36 | 32
| 29.8864 | 39.9281 | 59.1774 | 123.6858 | 16.3845 | 15.4554 | 13.886 | 8.6958 | 5.8726 | 5.5396 | 4.9771 | 3.1168 | 175.5107 | 221.1849 | 294.5296 | 385.5013 | 11.9 | 11.3 | 13.03 | 12.39 | 0.9 | 0.3 | 2.03 | 1.39 | 2.73 | 2.58 | 3.35 | 3.97 | -0.27 | -0.42 | 0.35 | 0.97 | 47 | 40 | 36 | 32
| 29.882 | 39.8036 | 59.7134 | 113.8304 | 16.0565 | 15.5186 | 13.9048 | 10.0086 | 5.755 | 5.5622 | 4.9838 | 3.5873 | 171.9714 | 221.3965 | 297.6002 | 408.3471 | 11.89 | 11.15 | 12.38 | 13.11 | 0.89 | 0.15 | 1.38 | 2.11 | 2.59 | 2.52 | 3.09 | 4.06 | -0.41 | -0.48 | 0.09 | 1.06 | 47 | 40 | 36 | 32
| 29.8537 | 39.9547 | 59.8922 | 125.7862 | 16.1538 | 15.7232 | 14.697 | 8.7259 | 5.7899 | 5.6356 | 5.2678 | 3.1276 | 172.85 | 225.1672 | 315.4974 | 393.4028 | 12.02 | 11.09 | 12.51 | 12.33 | 1.02 | 0.09 | 1.51 | 1.33 | 2.67 | 2.6 | 2.97 | 3.94 | -0.33 | -0.4 | -0.03 | 0.94 | 47 | 40 | 36 | 32
| 29.9715 | 39.6511 | 59.7372 | 116.5275 | 16.0781 | 15.7281 | 13.569 | 10.2123 | 5.7628 | 5.6373 | 4.8635 | 3.6603 | 172.719 | 223.5261 | 290.5288 | 426.5283 | 11.85 | 11.16 | 12.55 | 11.99 | 0.85 | 0.16 | 1.55 | 0.99 | 2.69 | 2.49 | 3.12 | 3.89 | -0.31 | -0.51 | 0.12 | 0.89 | 47 | 40 | 36 | 32
| 29.8656 | 39.8089 | 59.8444 | 118.6474 | 16.1486 | 15.8773 | 13.8261 | 9.049 | 5.788 | 5.6908 | 4.9556 | 3.2434 | 172.8629 | 226.5436 | 296.5643 | 384.8153 | 12.03 | 11.24 | 12.85 | 12.46 | 1.03 | 0.24 | 1.85 | 1.46 | 2.57 | 2.62 | 3.01 | 3.97 | -0.43 | -0.38 | 0.01 | 0.97 | 47 | 39 | 36 | 32
| 29.8805 | 39.7693 | 59.3296 | 122.2494 | 16.335 | 15.4202 | 14.057 | 8.9499 | 5.8548 | 5.5269 | 5.0384 | 3.2078 | 174.945 | 219.8028 | 298.9236 | 392.1559 | 11.97 | 11.15 | 12.51 | 12.8 | 0.97 | 0.15 | 1.51 | 1.8 | 2.65 | 2.51 | 2.97 | 3.93 | -0.35 | -0.49 | -0.03 | 0.93 | 47 | 39 | 36 | 32
| 29.8003 | 40.0909 | 60.3258 | 116.6861 | 16.2795 | 15.6312 | 14.1053 | 9.3936 | 5.835 | 5.6026 | 5.0556 | 3.3669 | 173.8839 | 224.6118 | 304.9858 | 392.8698 | 12.03 | 11.17 | 12.93 | 12.5 | 1.03 | 0.17 | 1.93 | 1.5 | 2.68 | 2.55 | 3.15 | 4.03 | -0.32 | -0.45 | 0.15 | 1.03 | 47 | 39 | 36 | 32
| 29.9506 | 39.4867 | 59.4825 | 123.3553 | 16.4119 | 15.4346 | 13.9595 | 9.0458 | 5.8824 | 5.5321 | 5.0034 | 3.2422 | 176.1813 | 218.4449 | 297.6146 | 399.9437 | 11.93 | 11.12 | 12.41 | 12.42 | 0.93 | 0.12 | 1.41 | 1.42 | 2.71 | 2.51 | 3.18 | 3.95 | -0.29 | -0.49 | 0.18 | 0.95 | 47 | 39 | 36 | 32
| 29.8507 | 39.9627 | 59.512 | 119.4743 | 16.2323 | 15.6003 | 14.0017 | 9.1577 | 5.818 | 5.5915 | 5.0185 | 3.2823 | 173.6721 | 223.4522 | 298.6635 | 392.1528 | 11.84 | 11.15 | 12.86 | 13.04 | 0.84 | 0.15 | 1.86 | 2.04 | 2.59 | 2.55 | 3.15 | 4.11 | -0.41 | -0.45 | 0.15 | 1.11 | 47 | 39 | 36 | 32
| 29.8418 | 39.6799 | 60.1745 | 113.4859 | 16.2908 | 15.7274 | 13.5914 | 9.6194 | 5.839 | 5.6371 | 4.8715 | 3.4478 | 174.2466 | 223.6788 | 293.1383 | 391.28 | 11.92 | 11.22 | 12.57 | 12.86 | 0.92 | 0.22 | 1.57 | 1.86 | 2.7 | 2.5 | 2.98 | 4.07 | -0.3 | -0.5 | -0.02 | 1.07 | 47 | 39 | 36 | 32
| 29.9745 | 40.1204 | 59.9221 | 122.6994 | 16.3652 | 15.7014 | 14.2639 | 8.0123 | 5.8657 | 5.6278 | 5.1125 | 2.8718 | 175.8205 | 225.7874 | 306.3528 | 352.3692 | 12.03 | 11.18 | 12.85 | 11.93 | 1.03 | 0.18 | 1.85 | 0.93 | 2.66 | 2.54 | 3.08 | 3.93 | -0.34 | -0.46 | 0.08 | 0.93 | 47 | 40 | 36 | 33
| 29.8463 | 39.6327 | 59.8623 | 116.0766 | 16.2154 | 15.578 | 14.2923 | 8.7965 | 5.812 | 5.5835 | 5.1227 | 3.1529 | 173.4658 | 221.2902 | 306.6569 | 365.9726 | 11.89 | 11.04 | 12.66 | 11.83 | 0.89 | 0.04 | 1.66 | 0.83 | 2.53 | 2.61 | 3.19 | 3.62 | -0.47 | -0.39 | 0.19 | 0.62 | 47 | 40 | 36 | 33
| 29.7855 | 40 | 59.4884 | 124.6624 | 16.0312 | 15.5306 | 13.8441 | 8.346 | 5.7459 | 5.5665 | 4.9621 | 2.9914 | 171.146 | 222.6609 | 295.1849 | 372.915 | 11.95 | 11.05 | 12.68 | 12.12 | 0.95 | 0.05 | 1.68 | 1.12 | 2.53 | 2.44 | 3.06 | 3.95 | -0.47 | -0.56 | 0.06 | 0.95 | 47 | 40 | 36 | 33
| 29.9371 | 39.7957 | 59.2768 | 122.2245 | 16.1391 | 15.6537 | 13.7409 | 9.4065 | 5.7846 | 5.6106 | 4.925 | 3.3715 | 173.1753 | 223.2794 | 291.9406 | 412.0823 | 11.88 | 11.18 | 12.61 | 12.92 | 0.88 | 0.18 | 1.61 | 1.92 | 2.64 | 2.56 | 2.99 | 3.96 | -0.36 | -0.44 | -0.01 | 0.96 | 47 | 40 | 36 | 33
| 29.9356 | 39.6563 | 60.3865 | 112.6338 | 16.2401 | 15.3801 | 14.093 | 9.2727 | 5.8208 | 5.5126 | 5.0513 | 3.3236 | 174.2505 | 218.6079 | 305.0279 | 374.3444 | 11.99 | 11.23 | 12.84 | 12.17 | 0.99 | 0.23 | 1.84 | 1.17 | 2.72 | 2.5 | 3.11 | 3.98 | -0.28 | -0.5 | 0.11 | 0.98 | 47 | 40 | 36 | 33
| 29.7354 | 40.1123 | 59.4059 | 121.9512 | 16.385 | 15.4192 | 14.0052 | 7.4651 | 5.8728 | 5.5266 | 5.0198 | 2.6757 | 174.6287 | 221.6838 | 298.2058 | 326.3021 | 11.97 | 11.11 | 12.54 | 11.81 | 0.97 | 0.11 | 1.54 | 0.81 | 2.62 | 2.59 | 3.11 | 3.71 | -0.38 | -0.41 | 0.11 | 0.71 | 47 | 40 | 36 | 33
| 29.8998 | 39.7799 | 60.1685 | 117.9014 | 16.538 | 15.6799 | 14.4342 | 8.3126 | 5.9276 | 5.62 | 5.1735 | 2.9794 | 177.2347 | 223.5648 | 311.2837 | 351.2786 | 11.95 | 11.09 | 12.82 | 11.75 | 0.95 | 0.09 | 1.82 | 0.75 | 2.64 | 2.48 | 2.98 | 3.83 | -0.36 | -0.52 | -0.02 | 0.83 | 47 | 40 | 36 | 33
| 29.9207 | 39.8406 | 59.976 | 121.2366 | 16.2432 | 15.8399 | 14.2667 | 8.4117 | 5.822 | 5.6774 | 5.1135 | 3.0149 | 174.197 | 226.1912 | 306.6873 | 365.5199 | 12.03 | 11.24 | 12.54 | 12.28 | 1.03 | 0.24 | 1.54 | 1.28 | 2.72 | 2.55 | 3.05 | 3.8 | -0.28 | -0.45 | 0.05 | 0.8 | 47 | 40 | 36 | 33
| 29.7855 | 39.9361 | 59.6244 | 121.9017 | 16.2469 | 15.5625 | 13.6974 | 8.7859 | 5.8233 | 5.5779 | 4.9095 | 3.1491 | 173.4487 | 222.7615 | 292.7245 | 383.8753 | 12 | 11.28 | 12.83 | 11.69 | 1 | 0.28 | 1.83 | 0.69 | 2.54 | 2.46 | 3.09 | 3.92 | -0.46 | -0.54 | 0.09 | 0.92 | 47 | 41 | 36 | 33
| 29.8701 | 39.6301 | 59.5948 | 118.7884 | 16.2143 | 15.8368 | 13.9416 | 9.3192 | 5.8116 | 5.6763 | 4.997 | 3.3402 | 173.5924 | 224.9514 | 297.7944 | 396.7793 | 11.99 | 11.11 | 12.56 | 12.45 | 0.99 | 0.11 | 1.56 | 1.45 | 2.58 | 2.5 | 2.89 | 4.05 | -0.42 | -0.5 | -0.11 | 1.05 | 47 | 41 | 36 | 32
| 30.0135 | 40.0187 | 59.8384 | 118.0638 | 16.3351 | 15.5298 | 15.332 | 9.176 | 5.8549 | 5.5662 | 5.4953 | 3.2889 | 175.7257 | 222.753 | 328.8317 | 388.2981 | 11.91 | 11.18 | 12.5 | 12.7 | 0.91 | 0.18 | 1.5 | 1.7 | 2.6 | 2.61 | 2.96 | 4.01 | -0.4 | -0.39 | -0.04 | 1.01 | 47 | 41 | 36 | 32
| 29.8864 | 39.6799 | 59.577 | 123.0012 | 16.3164 | 15.8383 | 14.9624 | 8.6634 | 5.8482 | 5.6768 | 5.3629 | 3.1052 | 174.7814 | 225.256 | 319.504 | 381.9378 | 12.05 | 11.22 | 12.62 | 12.43 | 1.05 | 0.22 | 1.62 | 1.43 | 2.75 | 2.57 | 2.92 | 3.98 | -0.25 | -0.43 | -0.08 | 0.98 | 47 | 41 | 36 | 32
| 29.7442 | 40.0107 | 60.006 | 115.7407 | 16.1392 | 15.5717 | 14.7576 | 9.2102 | 5.7847 | 5.5812 | 5.2894 | 3.3012 | 172.0606 | 223.3094 | 317.3985 | 382.0779 | 12.01 | 11.22 | 12.52 | 12.12 | 1.01 | 0.22 | 1.52 | 1.12 | 2.68 | 2.61 | 3.03 | 4 | -0.32 | -0.39 | 0.03 | 1 | 47 | 41 | 36 | 32
| 29.9506 | 39.7009 | 59.6303 | 123.5585 | 16.083 | 15.7123 | 15.112 | 9.6993 | 5.7645 | 5.6316 | 5.4165 | 3.4764 | 172.6509 | 223.581 | 322.9876 | 429.544 | 11.94 | 11.14 | 12.56 | 12.35 | 0.94 | 0.14 | 1.56 | 1.35 | 2.59 | 2.47 | 2.97 | 4.05 | -0.41 | -0.53 | -0.03 | 1.05 | 47 | 41 | 36 | 32
| 29.9028 | 39.7588 | 59.3061 | 116.5275 | 16.3701 | 15.6034 | 15.3402 | 9.9692 | 5.8674 | 5.5926 | 5.4983 | 3.5732 | 175.4524 | 222.356 | 326.0808 | 416.3754 | 12.06 | 11.09 | 12.54 | 12.84 | 1.06 | 0.09 | 1.54 | 1.84 | 2.71 | 2.68 | 2.91 | 4.04 | -0.29 | -0.32 | -0.09 | 1.04 | 47 | 41 | 36 | 32
| 29.8894 | 39.8804 | 60.3865 | 121.4575 | 16.2663 | 15.6287 | 15.2107 | 9.1126 | 5.8302 | 5.6017 | 5.4519 | 3.2662 | 174.2622 | 223.3967 | 329.2197 | 396.6994 | 11.98 | 11.14 | 12.88 | 12.43 | 0.98 | 0.14 | 1.88 | 1.43 | 2.58 | 2.47 | 3.01 | 3.95 | -0.42 | -0.53 | 0.01 | 0.95 | 47 | 41 | 36 | 32
| 29.8418 | 40.0641 | 60.036 | 119.2369 | 16.1247 | 15.7953 | 15.4045 | 9.7795 | 5.7795 | 5.6614 | 5.5213 | 3.5052 | 172.47 | 226.8185 | 331.4781 | 417.9498 | 11.93 | 11.21 | 12.49 | 12.19 | 0.93 | 0.21 | 1.49 | 1.19 | 2.53 | 2.56 | 3.04 | 3.99 | -0.47 | -0.44 | 0.04 | 0.99
| 30.021 | 39.4763 | 58.9565 | 116.8452 | 16.5581 | 15.5342 | 14.4587 | 8.5834 | 5.9348 | 5.5678 | 5.1823 | 3.0765 | 178.1684 | 219.7966 | 305.5324 | 359.472 | 12.01 | 11.06 | 12.78 | 12.02 | 1.01 | 0.06 | 1.78 | 1.02 | 2.73 | 2.59 | 2.98 | 3.82 | -0.27 | -0.41 | -0.02 | 0.82
| 29.8671 | 40.2387 | 60.1745 | 121.3838 | 16.1516 | 15.3681 | 14.4998 | 9.2361 | 5.7891 | 5.5083 | 5.197 | 3.3104 | 172.9033 | 221.6457 | 312.7296 | 401.8347 | 12.05 | 11.01 | 12.6 | 12.57 | 1.05 | 0.01 | 1.6 | 1.57 | 2.6 | 2.5 | 3.15 | 3.86 | -0.4 | -0.5 | 0.15 | 0.86
| 29.7471 | 39.6066 | 59.6303 | 118.5302 | 16.2965 | 15.4401 | 14.4526 | 8.7609 | 5.8411 | 5.5341 | 5.1801 | 3.1401 | 173.7548 | 219.1864 | 308.8933 | 372.1978 | 11.95 | 11.11 | 12.71 | 12.71 | 0.95 | 0.11 | 1.71 | 1.71 | 2.6 | 2.63 | 3.1 | 3.99 | -0.4 | -0.37 | 0.1 | 0.99
| 29.9118 | 39.6458 | 60.2349 | 120.5788 | 16.2975 | 15.4675 | 15.0559 | 9.7781 | 5.8414 | 5.5439 | 5.3964 | 3.5047 | 174.727 | 219.793 | 325.0499 | 422.5897 | 12.14 | 11.16 | 12.41 | 12.17 | 1.14 | 0.16 | 1.41 | 1.17 | 2.68 | 2.46 | 2.92 | 4.02 | -0.32 | -0.54 | -0.08 | 1.02
| 29.8611 | 40.0828 | 59.3531 | 119.6172 | 16.4085 | 15.5942 | 14.9478 | 9.5921 | 5.8812 | 5.5893 | 5.3576 | 3.438 | 175.6184 | 224.0363 | 317.9927 | 411.2456 | 12.02 | 11.15 | 12.64 | 12.39 | 1.02 | 0.15 | 1.64 | 1.39 | 2.65 | 2.58 | 2.99 | 4 | -0.35 | -0.42 | -0.01 | 1
| 29.8478 | 39.6223 | 59.9101 | 116.9135 | 16.4027 | 15.3127 | 15.1672 | 9.144 | 5.8791 | 5.4884 | 5.4363 | 3.2774 | 175.4782 | 217.4642 | 325.6878 | 383.1735 | 12.06 | 11.15 | 12.69 | 12.78 | 1.06 | 0.15 | 1.69 | 1.78 | 2.65 | 2.45 | 2.99 | 4.03 | -0.35 | -0.55 | -0.01 | 1.03
| 29.8404 | 39.8883 | 59.7372 | 122.976 | 16.2552 | 15.3996 | 15.7655 | 7.4269 | 5.8262 | 5.5196 | 5.6507 | 2.662 | 173.8567 | 220.1664 | 337.5583 | 327.3579 | 12.01 | 11.16 | 12.74 | 12.02 | 1.01 | 0.16 | 1.74 | 1.02 | 2.6 | 2.52 | 2.95 | 3.82 | -0.4 | -0.48 | -0.05 | 0.82
| 29.9073 | 39.8036 | 59.3354 | 115.5402 | 16.4983 | 15.4741 | 15.0826 | 9.2078 | 5.9134 | 5.5463 | 5.406 | 3.3003 | 176.8523 | 220.7622 | 320.7654 | 381.3142 | 11.97 | 11.22 | 12.56 | 11.62 | 0.97 | 0.22 | 1.56 | 0.62 | 2.66 | 2.5 | 2.9 | 3.77 | -0.34 | -0.5 | -0.1 | 0.77
| 29.9625 | 40.0695 | 60.2531 | 123.9157 | 16.3353 | 15.6288 | 15.2309 | 8.3695 | 5.8549 | 5.6017 | 5.4591 | 2.9998 | 175.4289 | 224.4574 | 328.9285 | 371.725 | 12.06 | 11.21 | 12.97 | 11.9 | 1.06 | 0.21 | 1.97 | 0.9 | 2.62 | 2.54 | 3 | 3.84 | -0.38 | -0.46 | 0 | 0.84
| 29.8582 | 39.8406 | 60.3926 | 119.1422 | 16.2719 | 15.3024 | 15.3337 | 9.135 | 5.8322 | 5.4847 | 5.4959 | 3.2742 | 174.1397 | 218.5145 | 331.9138 | 390.0931 | 12.03 | 11.13 | 12.65 | 12.54 | 1.03 | 0.13 | 1.65 | 1.54 | 2.57 | 2.49 | 3.13 | 3.85 | -0.43 | -0.51 | 0.13 | 0.85
| 29.8626 | 39.6904 | 59.5061 | 118.8825 | 16.1437 | 15.8128 | 15.6338 | 8.9977 | 5.7863 | 5.6677 | 5.6035 | 3.225 | 172.7935 | 224.9522 | 333.4423 | 383.3936 | 12.08 | 11.1 | 12.92 | 12.35 | 1.08 | 0.1 | 1.92 | 1.35 | 2.64 | 2.5 | 3.01 | 4 | -0.36 | -0.5 | 0.01 | 1
| 29.9491 | 39.7852 | 59.3648 | 120.919 | 16.1328 | 15.3168 | 14.9248 | 8.8337 | 5.7824 | 5.4899 | 5.3494 | 3.1662 | 173.177 | 218.4164 | 317.5647 | 382.8559 | 12.1 | 11.09 | 12.65 | 12.15 | 1.1 | 0.09 | 1.65 | 1.15 | 2.7 | 2.56 | 2.96 | 3.93 | -0.3 | -0.44 | -0.04 | 0.93
| 29.753 | 39.6983 | 59.8504 | 117.4398 | 16.2177 | 15.3384 | 14.8786 | 8.5679 | 5.8128 | 5.4976 | 5.3328 | 3.0709 | 172.9489 | 218.2466 | 319.1729 | 360.6497 | 12.07 | 11.27 | 12.95 | 12.17 | 1.07 | 0.27 | 1.95 | 1.17 | 2.62 | 2.45 | 3.11 | 3.87 | -0.38 | -0.55 | 0.11 | 0.87
| 29.9745 | 40.115 | 59.8205 | 123.3046 | 16.4344 | 15.3158 | 15.5421 | 9.0631 | 5.8905 | 5.4895 | 5.5706 | 3.2484 | 176.564 | 220.213 | 333.2385 | 400.5472 | 11.98 | 11.18 | 12.66 | 13 | 0.98 | 0.18 | 1.66 | 2 | 2.62 | 2.67 | 3.02 | 3.98 | -0.38 | -0.33 | 0.02 | 0.98
| 29.8641 | 39.4477 | 59.5061 | 117.0503 | 16.4273 | 15.5533 | 14.9683 | 9.671 | 5.8879 | 5.5747 | 5.365 | 3.4663 | 175.8375 | 219.9082 | 319.2489 | 405.734 | 11.99 | 11.13 | 12.98 | 12.28 | 0.99 | 0.13 | 1.98 | 1.28 | 2.68 | 2.5 | 3.24 | 4.03 | -0.32 | -0.5 | 0.24 | 1.03
| 29.8493 | 39.9494 | 59.6184 | 118.3665 | 16.4519 | 15.9658 | 14.4418 | 9.1389 | 5.8968 | 5.7225 | 5.1763 | 3.2756 | 176.0138 | 228.6112 | 308.6013 | 387.7185 | 12.04 | 11.14 | 12.66 | 12.25 | 1.04 | 0.14 | 1.66 | 1.25 | 2.55 | 2.56 | 3.02 | 4.01 | -0.45 | -0.44 | 0.02 | 1.01
| 29.8879 | 40.0133 | 60.012 | 119.7366 | 16.1962 | 15.5092 | 14.8992 | 9.7299 | 5.8051 | 5.5589 | 5.3402 | 3.4874 | 173.5016 | 222.4286 | 320.4763 | 417.5712 | 12.01 | 11.21 | 12.88 | 12.67 | 1.01 | 0.21 | 1.88 | 1.67 | 2.64 | 2.54 | 3.2 | 3.95 | -0.36 | -0.46 | 0.2 | 0.95
| 29.8849 | 39.9255 | 59.8563 | 120.3128 | 16.1723 | 15.7716 | 15.578 | 8.7456 | 5.7965 | 5.6529 | 5.5835 | 3.1346 | 173.2287 | 225.6953 | 334.2095 | 377.1334 | 11.91 | 11.24 | 12.86 | 12.75 | 0.91 | 0.24 | 1.86 | 1.75 | 2.68 | 2.54 | 2.94 | 4.03 | -0.32 | -0.46 | -0.06 | 1.03
| 29.8745 | 39.7483 | 59.6837 | 118.7178 | 16.2628 | 15.3202 | 14.4601 | 9.3292 | 5.829 | 5.4911 | 5.1828 | 3.3438 | 174.1378 | 218.2623 | 309.3306 | 396.9707 | 12.01 | 11.07 | 12.97 | 12.1 | 1.01 | 0.07 | 1.97 | 1.1 | 2.63 | 2.57 | 3.09 | 3.99 | -0.37 | -0.43 | 0.09 | 0.99
| 29.8151 | 40.1284 | 60.4595 | 121.4083 | 16.253 | 15.5331 | 14.8375 | 9.1194 | 5.8255 | 5.5674 | 5.3181 | 3.2686 | 173.6871 | 223.4123 | 321.5287 | 396.8356 | 12.04 | 11.02 | 12.6 | 12.27 | 1.04 | 0.02 | 1.6 | 1.27 | 2.59 | 2.52 | 3.12 | 3.96 | -0.41 | -0.48 | 0.12 | 0.96
| 29.9611 | 39.6327 | 58.7314 | 119.1658 | 16.2052 | 15.6513 | 14.52 | 9.1988 | 5.8083 | 5.6098 | 5.2043 | 3.2971 | 174.0234 | 222.3308 | 305.6568 | 392.8982 | 12.04 | 11.09 | 13.07 | 12.51 | 1.04 | 0.09 | 2.07 | 1.51 | 2.62 | 2.6 | 3.62 | 3.9 | -0.38 | -0.4 | 0.62 | 0.9
| 29.9327 | 39.8248 | 59.4589 | 116.3919 | 16.4014 | 15.5097 | 14.5922 | 9.8214 | 5.8786 | 5.559 | 5.2302 | 3.5202 | 175.9629 | 221.3875 | 310.981 | 409.7232 | 12.09 | 11.18 | 12.71 | 12.56 | 1.09 | 0.18 | 1.71 | 1.56 | 2.57 | 2.52 | 3.05 | 3.87 | -0.43 | -0.48 | 0.05 | 0.87
| 29.83 | 39.672 | 60.4839 | 119.0476 | 16.3982 | 15.7033 | 15.381 | 9.538 | 5.8775 | 5.6284 | 5.5129 | 3.4186 | 175.3253 | 223.2916 | 333.4415 | 406.9801 | 12.02 | 11.25 | 12.69 | 12.38 | 1.02 | 0.25 | 1.69 | 1.38 | 2.5 | 2.61 | 3.01 | 4.02 | -0.5 | -0.39 | 0.01 | 1.02
| 29.9476 | 39.6616 | 59.2534 | 121.8522 | 16.1155 | 15.2533 | 14.6635 | 8.7534 | 5.7762 | 5.4671 | 5.2557 | 3.1374 | 172.9827 | 216.8356 | 311.4202 | 382.3033 | 12.17 | 11.26 | 12.85 | 12.52 | 1.17 | 0.26 | 1.85 | 1.52 | 2.53 | 2.6 | 3.13 | 4.04 | -0.47 | -0.4 | 0.13 | 1.04
| 29.8226 | 40.2226 | 59.9341 | 121.0165 | 16.155 | 15.6789 | 14.6924 | 8.374 | 5.7903 | 5.6197 | 5.2661 | 3.0014 | 172.6827 | 226.0371 | 315.6187 | 363.2216 | 12.08 | 11.29 | 12.72 | 12.58 | 1.08 | 0.29 | 1.72 | 1.58 | 2.66 | 2.54 | 3.11 | 3.99 | -0.34 | -0.46 | 0.11 | 0.99
| 29.8255 | 39.7272 | 59.512 | 115.4512 | 16.1497 | 15.2835 | 14.9592 | 9.1558 | 5.7884 | 5.478 | 5.3617 | 3.2816 | 172.6432 | 217.6245 | 319.0863 | 378.8702 | 12 | 11.3 | 12.73 | 12.59 | 1 | 0.3 | 1.73 | 1.59 | 2.55 | 2.56 | 3.07 | 3.97 | -0.45 | -0.44 | 0.07 | 0.97
| 29.9133 | 39.9521 | 59.5061 | 125.6808 | 16.3596 | 15.694 | 14.9893 | 9.1372 | 5.8637 | 5.6251 | 5.3725 | 3.275 | 175.4009 | 224.7336 | 319.6973 | 411.6034 | 12.16 | 11.05 | 12.51 | 11.76 | 1.16 | 0.05 | 1.51 | 0.76 | 2.59 | 2.52 | 3.07 | 3.86 | -0.41 | -0.48 | 0.07 | 0.86
| 29.8641 | 39.6668 | 60.1263 | 113.9168 | 16.3003 | 15.4876 | 14.8381 | 9.7138 | 5.8424 | 5.5511 | 5.3183 | 3.4817 | 174.4779 | 220.1948 | 319.7714 | 396.6204 | 12.09 | 11.09 | 12.74 | 12.4 | 1.09 | 0.09 | 1.74 | 1.4 | 2.58 | 2.51 | 2.86 | 4.09 | -0.42 | -0.49 | -0.14 | 1.09
| 29.97 | 39.8857 | 59.8982 | 124.9479 | 15.9671 | 15.5189 | 15.0624 | 8.2636 | 5.723 | 5.5623 | 5.3987 | 2.9619 | 171.5178 | 221.8572 | 323.3722 | 370.081 | 12.08 | 11.1 | 12.47 | 12.48 | 1.08 | 0.1 | 1.47 | 1.48 | 2.53 | 2.62 | 2.95 | 3.91 | -0.47 | -0.38 | -0.05 | 0.91
| 29.8969 | 39.891 | 60.0721 | 117.3709 | 16.0997 | 15.6401 | 14.8627 | 9.0441 | 5.7705 | 5.6058 | 5.3271 | 3.2416 | 172.52 | 223.6198 | 320.0126 | 380.4706 | 12.03 | 11.35 | 12.65 | 12.11 | 1.03 | 0.35 | 1.65 | 1.11 | 2.65 | 2.57 | 3.06 | 3.38 | -0.35 | -0.43 | 0.06 | 0.38
| 29.7885 | 39.6118 | 59.7491 | 116.1665 | 15.9439 | 15.7172 | 15.0537 | 9.2113 | 5.7147 | 5.6334 | 5.3956 | 3.3015 | 170.2311 | 223.1495 | 322.3804 | 383.5281 | 12.15 | 11.25 | 12.54 | 11.66 | 1.15 | 0.25 | 1.54 | 0.66 | 2.57 | 2.49 | 3 | 3.94 | -0.43 | -0.51 | 0 | 0.94
| 29.9252 | 39.7483 | 59.6481 | 126.4489 | 16.3952 | 15.8474 | 14.7613 | 8.5615 | 5.8764 | 5.6801 | 5.2908 | 3.0686 | 175.8529 | 225.7727 | 315.5848 | 388.024 | 12.21 | 11.25 | 12.76 | 12.15 | 1.21 | 0.25 | 1.76 | 1.15 | 2.56 | 2.61 | 2.99 | 3.91 | -0.44 | -0.39 | -0.01 | 0.91
| 29.8077 | 40.0481 | 60.0661 | 117.2333 | 16.132 | 15.2307 | 15.6707 | 8.6423 | 5.7821 | 5.459 | 5.6168 | 3.0976 | 172.3505 | 218.623 | 337.3763 | 363.1408 | 12.02 | 11.1 | 12.51 | 12.15 | 1.02 | 0.1 | 1.51 | 1.15 | 2.56 | 2.57 | 3.11 | 3.36 | -0.44 | -0.43 | 0.11 | 0.36
| 29.7974 | 39.8168 | 59.453 | 119.9041 | 16.2105 | 15.8726 | 14.6427 | 9.3031 | 5.8102 | 5.6891 | 5.2483 | 3.3344 | 173.1295 | 226.5221 | 312.0253 | 399.8115 | 12.09 | 11.2 | 12.57 | 12.47 | 1.09 | 0.2 | 1.57 | 1.47 | 2.59 | 2.56 | 3.02 | 3.95 | -0.41 | -0.44 | 0.02 | 0.95
| 29.9521 | 39.7588 | 59.97 | 117.0047 | 12.8384 | 15.3397 | 15.17 | 9.0559 | 4.6016 | 5.4981 | 5.4373 | 3.2458 | 137.8268 | 218.5979 | 326.0742 | 379.778 | 12.21 | 11.21 | 12.47 | 12.23 | 1.21 | 0.21 | 1.47 | 1.23 | 3.83 | 2.43 | 3.04 | 4.03 | 0.83 | -0.57 | 0.04 | 1.03
| 29.9715 | 40.0855 | 58.9623 | 119.6888 | 16.0964 | 15.6305 | 14.8071 | 9.0768 | 5.7693 | 5.6023 | 5.3072 | 3.2533 | 172.9158 | 224.5723 | 312.924 | 389.3884 | 12.26 | 11.11 | 12.7 | 12.45 | 1.26 | 0.11 | 1.7 | 1.45 | 2.57 | 2.52 | 2.91 | 4.01 | -0.43 | -0.48 | -0.09 | 1.01
| 29.8641 | 39.5596 | 60.3197 | 117.3709 | 16.0084 | 15.3061 | 15.2596 | 8.8789 | 5.7378 | 5.486 | 5.4694 | 3.1824 | 171.3531 | 217.0255 | 329.9129 | 373.5214 | 12.15 | 11.13 | 12.57 | 12.48 | 1.15 | 0.13 | 1.57 | 1.48 | 2.54 | 2.58 | 2.99 | 3.94 | -0.46 | -0.42 | -0.01 | 0.94
| 29.8641 | 39.8195 | 60.1383 | 124.2236 | 16.131 | 15.4193 | 14.1781 | 8.1396 | 5.7817 | 5.5266 | 5.0817 | 2.9174 | 172.6657 | 220.0677 | 305.6074 | 362.4144 | 12.13 | 11.12 | 12.74 | 11.96 | 1.13 | 0.12 | 1.74 | 0.96 | 2.5 | 2.44 | 2.93 | 3.96 | -0.5 | -0.56 | -0.07 | 0.96
| 30.024 | 39.8459 | 59.312 | 115.3624 | 16.1923 | 15.6676 | 14.6689 | 9.213 | 5.8037 | 5.6156 | 5.2577 | 3.3021 | 174.2498 | 223.7599 | 311.8432 | 380.9432 | 12.25 | 11.1 | 12.25 | 11.73 | 1.25 | 0.1 | 1.25 | 0.73 | 2.64 | 2.56 | 2.92 | 3.84 | -0.36 | -0.44 | -0.08 | 0.84
| 29.6501 | 39.8301 | 60.1504 | 122.2743 | 16.1346 | 15.3152 | 14.9807 | 8.0148 | 5.783 | 5.4893 | 5.3694 | 2.8727 | 171.4672 | 218.6402 | 322.9722 | 351.2569 | 12.26 | 11.14 | 12.84 | 12.18 | 1.26 | 0.14 | 1.84 | 1.18 | 2.57 | 2.47 | 3.03 | 3.68 | -0.43 | -0.53 | 0.03 | 0.68
| 30.0481 | 39.9281 | 59.7669 | 121.7039 | 16.1876 | 15.5613 | 14.9686 | 8.5764 | 5.802 | 5.5775 | 5.3651 | 3.074 | 174.3391 | 222.6999 | 320.6551 | 374.1141 | 12.14 | 11.1 | 12.45 | 12.73 | 1.14 | 0.1 | 1.45 | 1.73 | 2.63 | 2.54 | 3.02 | 3.83 | -0.37 | -0.46 | 0.02 | 0.83
| 29.6971 | 39.6066 | 59.4354 | 116.8224 | 16.0633 | 15.4622 | 15.3073 | 9.1335 | 5.7575 | 5.542 | 5.4865 | 3.2737 | 170.9799 | 219.5 | 326.0906 | 382.4366 | 12.15 | 11.09 | 12.99 | 11.48 | 1.15 | 0.09 | 1.99 | 0.48 | 2.64 | 2.44 | 3.02 | 3.73 | -0.36 | -0.56 | 0.02 | 0.73
| 30.125 | 40.123 | 59.8205 | 120.7486 | 13.3551 | 15.679 | 14.9449 | 8.4696 | 4.7868 | 5.6197 | 5.3566 | 3.0357 | 144.202 | 225.48 | 320.4348 | 366.5575 | 12.51 | 11.34 | 12.69 | 11.96 | 1.51 | 0.34 | 1.69 | 0.96 | 3.65 | 2.51 | 3 | 3.89 | 0.65 | -0.49 | 0 | 0.89
| 29.7782 | 39.7325 | 59.1716 | 117.6932 | 16.3045 | 15.6361 | 14.7318 | 9.1031 | 5.8439 | 5.6043 | 5.2802 | 3.2628 | 174.0208 | 222.6738 | 312.4386 | 384.0061 | 12.32 | 11.21 | 12.85 | 12.23 | 1.32 | 0.21 | 1.85 | 1.23 | 2.72 | 2.56 | 3.05 | 3.87 | -0.28 | -0.44 | 0.05 | 0.87
| 29.8359 | 39.8671 | 60.429 | 119.928 | 16.2425 | 15.6157 | 14.3937 | 8.6725 | 5.8217 | 5.597 | 5.159 | 3.1084 | 173.6956 | 223.1377 | 311.7554 | 372.7891 | 12.27 | 11.02 | 12.51 | 12.87 | 1.27 | 0.02 | 1.51 | 1.87 | 2.63 | 2.51 | 2.96 | 3.84 | -0.37 | -0.49 | -0.04 | 0.84
| 29.876 | 39.9494 | 60.1866 | 119.7127 | 16.2046 | 15.608 | 14.6584 | 8.6334 | 5.8081 | 5.5943 | 5.2539 | 3.0944 | 173.5234 | 223.4876 | 316.215 | 370.4384 | 12.14 | 10.98 | 12.71 | 11.96 | 1.14 | -0.02 | 1.71 | 0.96 | 2.67 | 2.62 | 2.96 | 4.02 | -0.33 | -0.38 | -0.04 | 1.02
| 29.9103 | 39.4867 | 59.4236 | 122.3242 | 16.2547 | 15.7676 | 15.0947 | 8.1397 | 5.8261 | 5.6515 | 5.4103 | 2.9174 | 174.2592 | 223.1573 | 321.4996 | 356.8739 | 12.12 | 11.09 | 12.58 | 11.82 | 1.12 | 0.09 | 1.58 | 0.82 | 2.56 | 2.46 | 3.04 | 3.94 | -0.44 | -0.54 | 0.04 | 0.94
| 29.8805 | 39.9947 | 59.5948 | 116.3693 | 16.1437 | 15.5473 | 14.4557 | 8.809 | 5.7863 | 5.5725 | 5.1813 | 3.1573 | 172.8968 | 222.8713 | 308.7755 | 367.4178 | 12.19 | 11.16 | 12.83 | 12.26 | 1.19 | 0.16 | 1.83 | 1.26 | 2.52 | 2.57 | 3.04 | 4 | -0.48 | -0.43 | 0.04 | 1
| 29.9267 | 39.7404 | 59.5297 | 120.7729 | 16.4134 | 15.6461 | 15.2315 | 9.1667 | 5.8829 | 5.6079 | 5.4593 | 3.2856 | 176.0566 | 222.8605 | 324.9927 | 396.8073 | 12.18 | 11.09 | 12.52 | 12.39 | 1.18 | 0.09 | 1.52 | 1.39 | 2.67 | 2.47 | 2.88 | 3.84 | -0.33 | -0.53 | -0.12 | 0.84
| 29.8745 | 39.9574 | 59.6184 | 118.859 | 16.0366 | 15.5999 | 14.9902 | 9.0879 | 5.7479 | 5.5914 | 5.3728 | 3.2573 | 171.7155 | 223.4168 | 320.3195 | 387.1591 | 12.09 | 11 | 12.66 | 12 | 1.09 | 0 | 1.66 | 1 | 2.55 | 2.5 | 2.99 | 3.98 | -0.45 | -0.5 | -0.01 | 0.98
| 29.876 | 39.7614 | 60.2531 | 118.7884 | 16.3546 | 15.2993 | 15.1888 | 8.8531 | 5.8618 | 5.4836 | 5.444 | 3.1731 | 175.1286 | 218.0367 | 328.0189 | 376.9318 | 12.19 | 11.2 | 12.57 | 11.78 | 1.19 | 0.2 | 1.57 | 0.78 | 2.52 | 2.45 | 3.02 | 3.93 | -0.48 | -0.55 | 0.02 | 0.93
| 29.8344 | 39.9122 | 59.8802 | 122.2992 | 16.0307 | 15.5326 | 15.3049 | 9.1354 | 5.7458 | 5.5672 | 5.4856 | 3.2744 | 171.422 | 222.201 | 328.4807 | 400.4505 | 12.13 | 11.09 | 12.75 | 12.61 | 1.13 | 0.09 | 1.75 | 1.61 | 2.69 | 2.42 | 3.01 | 3.88 | -0.31 | -0.58 | 0.01 | 0.88
| 29.827 | 39.6852 | 59.4648 | 117.7625 | 16.2784 | 15.776 | 14.5083 | 9.1418 | 5.8346 | 5.6545 | 5.2001 | 3.2766 | 174.0272 | 224.399 | 309.224 | 385.8659 | 12.2 | 11.16 | 12.6 | 12.74 | 1.2 | 0.16 | 1.6 | 1.74 | 2.54 | 2.59 | 2.95 | 3.99 | -0.46 | -0.41 | -0.05 | 0.99
| 29.9805 | 40.0187 | 59.4648 | 121.3347 | 16.1796 | 15.6751 | 14.231 | 9.3939 | 5.7991 | 5.6183 | 5.1007 | 3.367 | 173.8613 | 224.8381 | 303.3135 | 408.5317 | 12.2 | 11.19 | 12.74 | 12.01 | 1.2 | 0.19 | 1.74 | 1.01 | 2.6 | 2.47 | 3.16 | 3.99 | -0.4 | -0.53 | 0.16 | 0.99
| 29.8567 | 39.6197 | 60.0901 | 116.6181 | 16.3444 | 15.7514 | 14.8386 | 9.0344 | 5.8582 | 5.6457 | 5.3185 | 3.2382 | 174.9066 | 223.6791 | 319.5881 | 377.6272 | 12.3 | 11.24 | 12.71 | 12.07 | 1.3 | 0.24 | 1.71 | 1.07 | 2.71 | 2.55 | 3.04 | 3.92 | -0.29 | -0.45 | 0.04 | 0.92
| 29.8389 | 39.7772 | 60.0601 | 118.2965 | 16.1914 | 15.0986 | 14.8972 | 9.1759 | 5.8034 | 5.4117 | 5.3395 | 3.2889 | 173.1662 | 215.2623 | 320.6913 | 389.0609 | 12.17 | 11.22 | 12.64 | 12.52 | 1.17 | 0.22 | 1.64 | 1.52 | 2.58 | 2.69 | 3.02 | 4.01 | -0.42 | -0.31 | 0.02 | 1.01
| 29.9327 | 40.1016 | 59.3472 | 124.5589 | 16.2255 | 15.6865 | 15.384 | 7.6761 | 5.8156 | 5.6224 | 5.514 | 2.7513 | 174.0763 | 225.4679 | 327.2397 | 342.698 | 12.14 | 11.17 | 12.76 | 12.31 | 1.14 | 0.17 | 1.76 | 1.31 | 2.59 | 2.57 | 2.97 | 3.68 | -0.41 | -0.43 | -0.03 | 0.68
| 29.7989 | 39.604 | 59.9341 | 113.9601 | 16.3096 | 15.4221 | 14.5735 | 8.7182 | 5.8457 | 5.5276 | 5.2235 | 3.1248 | 174.1961 | 218.9158 | 313.0651 | 356.1014 | 12.18 | 11.09 | 12.51 | 11.54 | 1.18 | 0.09 | 1.51 | 0.54 | 2.58 | 2.44 | 3.08 | 3.82 | -0.42 | -0.56 | 0.08 | 0.82
| 29.8374 | 39.9787 | 59.512 | 125.7598 | 14.5483 | 15.349 | 14.6401 | 7.8613 | 5.2145 | 5.5014 | 5.2474 | 2.8177 | 155.5858 | 219.9403 | 312.2812 | 354.3481 | 12.25 | 11.17 | 12.74 | 11.71 | 1.25 | 0.17 | 1.74 | 0.71 | 3.1 | 2.48 | 2.91 | 3.93 | 0.1 | -0.52 | -0.09 | 0.93
| 29.9745 | 39.6904 | 60.1323 | 119.928 | 15.5223 | 15.4718 | 14.4462 | 8.441 | 5.5636 | 5.5454 | 5.1778 | 3.0255 | 166.7651 | 220.1006 | 311.3558 | 362.8374 | 12.2 | 11.22 | 12.59 | 12.42 | 1.2 | 0.22 | 1.59 | 1.42 | 3.03 | 2.6 | 2.97 | 3.95 | 0.03 | -0.4 | -0.03 | 0.95
| 29.8611 | 39.7957 | 59.7253 | 116.3693 | 16.0053 | 15.1808 | 15.2616 | 9.2193 | 5.7367 | 5.4411 | 5.4701 | 3.3044 | 171.3031 | 216.5337 | 326.704 | 384.5328 | 12.3 | 11.23 | 12.79 | 12.58 | 1.3 | 0.23 | 1.79 | 1.58 | 2.62 | 2.55 | 3.03 | 3.96 | -0.38 | -0.45 | 0.03 | 0.96
| 29.9312 | 39.8883 | 59.5297 | 120.2646 | 16.076 | 15.6859 | 13.9559 | 8.6145 | 5.762 | 5.6222 | 5.0021 | 3.0876 | 172.4631 | 224.2595 | 297.774 | 371.3335 | 12.27 | 11.17 | 12.49 | 12.16 | 1.27 | 0.17 | 1.49 | 1.16 | 2.6 | 2.58 | 3.03 | 4.03 | -0.4 | -0.42 | 0.03 | 1.03
| 29.8181 | 39.7483 | 60.3926 | 118.8825 | 16.1556 | 15.6014 | 15.0038 | 9.0801 | 5.7905 | 5.5919 | 5.3777 | 3.2545 | 172.6629 | 222.269 | 324.7728 | 386.9055 | 12.31 | 11.12 | 12.72 | 12.18 | 1.31 | 0.12 | 1.72 | 1.18 | 2.54 | 2.5 | 2.94 | 3.9 | -0.46 | -0.5 | -0.06 | 0.9
| 29.9581 | 39.9867 | 58.8293 | 123.1274 | 16.1705 | 15.295 | 15.1488 | 8.5498 | 5.7959 | 5.4821 | 5.4297 | 3.0645 | 173.6336 | 219.2097 | 319.4236 | 377.3186 | 12.32 | 11.21 | 12.52 | 12.48 | 1.32 | 0.21 | 1.52 | 1.48 | 2.53 | 2.69 | 2.92 | 3.88 | -0.47 | -0.31 | -0.08 | 0.88
| 29.7855 | 39.9016 | 60.5877 | 115.0969 | 16.342 | 15.3722 | 12.7913 | 9.4824 | 5.8573 | 5.5098 | 4.5847 | 3.3987 | 174.4643 | 219.8482 | 277.7762 | 391.1807 | 12.23 | 11.22 | 11.78 | 12.68 | 1.23 | 0.22 | 0.78 | 1.68 | 2.69 | 2.51 | 3.08 | 4.08 | -0.31 | -0.49 | 0.08 | 1.08
| 29.9446 | 39.8618 | 59.4707 | 121.8027 | 16.1355 | 15.0962 | 15.058 | 8.8911 | 5.7833 | 5.4108 | 5.3971 | 3.1868 | 173.1795 | 215.6853 | 320.9706 | 388.1555 | 12.19 | 11.21 | 12.26 | 11.74 | 1.19 | 0.21 | 1.26 | 0.74 | 2.55 | 2.62 | 2.96 | 3.95 | -0.45 | -0.38 | -0.04 | 0.95
| 29.7929 | 39.7667 | 59.7193 | 119.9041 | 16.1875 | 15.4862 | 15.3834 | 8.7684 | 5.802 | 5.5506 | 5.5137 | 3.1428 | 172.8577 | 220.7299 | 329.2772 | 376.8331 | 12.27 | 11.18 | 12.76 | 12.25 | 1.27 | 0.18 | 1.76 | 1.25 | 2.52 | 2.58 | 2.89 | 3.84 | -0.48 | -0.42 | -0.11 | 0.84
| 30.0135 | 39.7509 | 59.9101 | 120.8703 | 16.2838 | 15.7163 | 15.3664 | 8.9552 | 5.8365 | 5.6331 | 5.5077 | 3.2097 | 175.1732 | 223.9206 | 329.9654 | 387.9621 | 12.19 | 11.18 | 12.51 | 12.33 | 1.19 | 0.18 | 1.51 | 1.33 | 2.65 | 2.47 | 2.99 | 3.97 | -0.35 | -0.53 | -0.01 | 0.97
| 29.8374 | 39.8671 | 59.5356 | 114.0251 | 16.0762 | 15.7185 | 14.9536 | 9.4386 | 5.7621 | 5.6339 | 5.3597 | 3.383 | 171.9254 | 224.6064 | 319.0949 | 385.7479 | 12.26 | 11.13 | 12.92 | 12.23 | 1.26 | 0.13 | 1.92 | 1.23 | 2.65 | 2.57 | 3.02 | 3.94 | -0.35 | -0.43 | 0.02 | 0.94
| 29.9416 | 39.6747 | 59.577 | 120.3852 | 16.1293 | 15.1564 | 14.996 | 7.5465 | 5.7811 | 5.4324 | 5.3749 | 2.7048 | 173.0957 | 215.5291 | 320.2217 | 325.6227 | 12.35 | 11.22 | 12.53 | 11.71 | 1.35 | 0.22 | 1.53 | 0.71 | 2.61 | 2.5 | 2.97 | 3.87 | -0.39 | -0.5 | -0.03 | 0.87
| 29.7501 | 40.04 | 59.4413 | 122.7496 | 15.9363 | 15.5724 | 15.3077 | 8.1148 | 5.7119 | 5.5815 | 5.4866 | 2.9085 | 169.9305 | 223.4843 | 326.1315 | 357.0223 | 12.26 | 11.01 | 12.82 | 11.89 | 1.26 | 0.01 | 1.82 | 0.89 | 2.64 | 2.46 | 2.94 | 3.88 | -0.36 | -0.54 | -0.06 | 0.88
| 29.9103 | 39.7852 | 60.5144 | 117.5779 | 15.8741 | 15.5385 | 15.236 | 8.4744 | 5.6897 | 5.5694 | 5.4609 | 3.0374 | 170.1792 | 221.5775 | 330.4649 | 357.132 | 12.25 | 11.1 | 12.37 | 11.69 | 1.25 | 0.1 | 1.37 | 0.69 | 2.71 | 2.5 | 3.02 | 3.96 | -0.29 | -0.5 | 0.02 | 0.96
| 29.9491 | 39.7483 | 59.1075 | 118.6709 | 16.2437 | 15.5382 | 14.9231 | 8.1721 | 5.8221 | 5.5692 | 5.3488 | 2.9291 | 174.3668 | 221.3679 | 316.1534 | 347.5968 | 12.2 | 11.1 | 12.86 | 11.37 | 1.2 | 0.1 | 1.86 | 0.37 | 2.67 | 2.48 | 2.99 | 3.77 | -0.33 | -0.52 | -0.01 | 0.77
| 29.8389 | 39.8036 | 59.755 | 117.855 | 16.0778 | 15.5817 | 15.2288 | 7.7146 | 5.7627 | 5.5848 | 5.4584 | 2.7651 | 171.9513 | 222.2967 | 326.1641 | 325.8811 | 12.33 | 11.21 | 12.51 | 11.32 | 1.33 | 0.21 | 1.51 | 0.32 | 2.67 | 2.55 | 2.91 | 3.85 | -0.33 | -0.45 | -0.09 | 0.85
| 29.9222 | 39.801 | 59.23 | 119.6649 | 15.9045 | 15.6308 | 14.341 | 7.3372 | 5.7005 | 5.6025 | 5.1402 | 2.6298 | 170.573 | 222.9832 | 304.4514 | 314.6963 | 12.23 | 11.2 | 12.87 | 11.18 | 1.23 | 0.2 | 1.87 | 0.18 | 2.49 | 2.44 | 3.02 | 3.81 | -0.51 | -0.56 | 0.02 | 0.81
| 29.8701 | 40.0507 | 60.8581 | 119.8083 | 16.5187 | 15.5399 | 15.3234 | 8.0298 | 5.9207 | 5.5698 | 5.4923 | 2.878 | 176.8515 | 223.0765 | 334.249 | 344.8143 | 12.14 | 11.2 | 12.61 | 11.48 | 1.14 | 0.2 | 1.61 | 0.48 | 2.67 | 2.44 | 3.05 | 3.9 | -0.33 | -0.56 | 0.05 | 0.9
| 29.8389 | 39.8195 | 59.1833 | 119.6172 | 16.4251 | 15.3239 | 15.2851 | 7.8657 | 5.8871 | 5.4924 | 5.4785 | 2.8193 | 175.6657 | 218.7054 | 324.2382 | 337.2313 | 12.15 | 11.07 | 12.97 | 10.84 | 1.15 | 0.07 | 1.97 | -0.16 | 2.58 | 2.56 | 3.04 | 3.95 | -0.42 | -0.44 | 0.04 | 0.95
| 29.8463 | 39.6852 | 30.2084 | 127.6053 | 15.9551 | 15.4697 | 14.7904 | 7.9781 | 5.7187 | 5.5447 | 5.3012 | 2.8595 | 170.6811 | 220.0416 | 160.1413 | 364.8905 | 12.1 | 11.18 | 12.55 | 10.92 | 1.1 | 0.18 | 1.55 | -0.08 | 2.6 | 2.44 | 3.1 | 3.89 | -0.4 | -0.56 | 0.1 | 0.89
| 30.036 | 40.0374 | 59.2476 | 113.2075 | 16.5733 | 15.6652 | 18.2549 | 8.6804 | 5.9403 | 5.6148 | 6.543 | 3.1112 | 178.4219 | 224.801 | 387.6546 | 352.2162 | 12.09 | 11.22 | 11.65 | 11.25 | 1.09 | 0.22 | 0.65 | 0.25 | 2.71 | 2.53 | 0.98 | 3.96 | -0.29 | -0.47 | -2.02 | 0.96
| 29.6399 | 39.6668 | 58.9912 | 122.549 | 16.2879 | 15.6515 | 15.9566 | 7.7062 | 5.838 | 5.6099 | 5.7192 | 2.7621 | 173.0366 | 222.525 | 337.3831 | 338.4917 | 12.29 | 11.07 | 12.63 | 11.42 | 1.29 | 0.07 | 1.63 | 0.42 | 2.54 | 2.46 | 2.8 | 3.84 | -0.46 | -0.54 | -0.2 | 0.84
| 29.8716 | 40.1016 | 59.9401 | 119.2369 | 14.2943 | 15.3455 | 14.7503 | 8.054 | 5.1234 | 5.5002 | 5.2868 | 2.8867 | 153.0435 | 220.5653 | 316.8929 | 344.2059 | 12.26 | 11.02 | 12.23 | 10.85 | 1.26 | 0.02 | 1.23 | -0.15 | 3.33 | 2.49 | 2.99 | 3.92 | 0.33 | -0.51 | -0.01 | 0.92
| 29.9252 | 39.5023 | 59.8922 | 117.7856 | 11.9881 | 15.9823 | 13.8536 | 8.4607 | 4.2968 | 5.7284 | 4.9655 | 3.0325 | 128.5832 | 226.2861 | 297.3923 | 357.188 | 12.41 | 11.17 | 12.32 | 11.34 | 1.41 | 0.17 | 1.32 | 0.34 | 4.23 | 2.53 | 2.93 | 3.92 | 1.23 | -0.47 | -0.07 | 0.92
| 30.036 | 39.9494 | 59.9161 | 121.5313 | 16.2313 | 15.4649 | 14.6424 | 8.7343 | 5.8177 | 5.543 | 5.2482 | 3.1306 | 174.7401 | 221.4381 | 314.4491 | 380.4605 | 12.24 | 11.02 | 12.13 | 12.19 | 1.24 | 0.02 | 1.13 | 1.19 | 2.73 | 2.51 | 2.97 | 3.91 | -0.27 | -0.49 | -0.03 | 0.91
| 29.8374 | 39.9547 | 59.3472 | 121.7039 | 15.5789 | 15.493 | 14.5326 | 8.2752 | 5.5838 | 5.5531 | 5.2088 | 2.966 | 166.6069 | 221.8706 | 309.1286 | 360.9765 | 12.25 | 11.1 | 12.52 | 11.71 | 1.25 | 0.1 | 1.52 | 0.71 | 2.66 | 2.67 | 3.17 | 3.96 | -0.34 | -0.33 | 0.17 | 0.96
|  | 39.6301 | 59.4059 | 118.9061 | 16.1243 | 15.4006 | 15.1649 | 8.3717 | 5.7793 | 5.5199 | 5.4355 | 3.0006 |  | 218.7554 | 322.8985 | 356.7908 | 12.21 | 11.08 | 12.51 | 10.65 | 1.21 | 0.08 | 1.51 | -0.35 | 2.49 | 2.39 | 2.91 | 3.82 | -0.51 | -0.61 | -0.09 | 0.82
|  | 39.838 | 59.8683 | 119.7605 |  | 15.2783 | 14.8085 | 8.5233 |  | 5.4761 | 5.3077 | 3.0549 |  | 218.157 | 317.7625 | 365.86 |  | 11.13 | 12.54 | 11.6 |  | 0.13 | 1.54 | 0.6 |  | 2.67 | 2.96 | 3.91 |  | -0.33 | -0.04 | 0.91
|  | 40.0133 | 30.1917 | 117.9477 |  | 15.2539 | 14.8586 | 8.9866 |  | 5.4673 | 5.3257 | 3.221 |  | 218.7665 | 160.7908 | 379.9105 |  | 11.09 | 12.56 | 12.19 |  | 0.09 | 1.56 | 1.19 |  | 2.54 | 3.04 | 4.12 |  | -0.46 | 0.04 | 1.12
|  | 39.8221 | 58.8351 | 119.6411 |  | 15.5488 | 17.7315 | 8.1986 |  | 5.5731 | 6.3554 | 2.9386 |  | 221.9313 | 373.9197 | 351.5731 |  | 11.02 | 11.73 | 11.7 |  | 0.02 | 0.73 | 0.7 |  | 2.45 | 0.78 | 4.06 |  | -0.55 | -2.22 | 1.06
|  | 39.8486 | 59.5415 | 120.1442 |  | 15.7513 | 15.9006 | 8.0924 |  | 5.6456 | 5.6991 | 2.9005 |  | 224.9695 | 339.3348 | 348.4786 |  | 11.17 | 12.22 | 10.9 |  | 0.17 | 1.22 | -0.1 |  | 2.56 | 2.82 | 3.91 |  | -0.44 | -0.18 | 0.91
|  | 39.6223 | 59.6599 | 117.3939 |  | 15.3244 | 14.4599 | 8.9945 |  | 5.4926 | 5.1827 | 3.2238 |  | 217.6293 | 309.2022 | 378.4574 |  | 11.05 | 12.65 | 11.61 |  | 0.05 | 1.65 | 0.61 |  | 2.43 | 3.08 | 3.76 |  | -0.57 | 0.08 | 0.76
|  | 39.9734 | 60.4108 | 117.3939 |  | 15.4872 | 14.9988 | 8.5228 |  | 5.551 | 5.3759 | 3.0548 |  | 221.8913 | 324.7622 | 358.6106 |  | 11.01 | 12.37 | 12.21 |  | 0.01 | 1.37 | 1.21 |  | 2.46 | 3.14 | 3.78 |  | -0.54 | 0.14 | 0.78
|  | 39.638 | 59.1833 | 122.8501 |  | 15.5294 | 14.953 | 8.1418 |  | 5.5661 | 5.3595 | 2.9182 |  | 220.6293 | 317.1917 | 358.5019 |  | 10.94 | 12.44 | 12.26 |  | -0.06 | 1.44 | 1.26 |  | 2.42 | 2.97 | 3.83 |  | -0.58 | -0.03 | 0.83
|  | 39.9893 | 59.3942 | 119.5695 |  | 15.4517 | 14.4872 | 8.5763 |  | 5.5382 | 5.1925 | 3.074 |  | 221.4704 | 308.4068 | 367.5517 |  | 11.01 | 12.21 | 11.36 |  | 0.01 | 1.21 | 0.36 |  | 2.4 | 2.95 | 3.89 |  | -0.6 | -0.05 | 0.89
|  | 39.891 | 59.2125 | 116.144 |  | 15.6139 | 14.2633 | 8.7853 |  | 5.5964 | 5.1123 | 3.1488 |  | 223.2452 | 302.7112 | 365.7199 |  | 11.15 | 12.39 | 11.78 |  | 0.15 | 1.39 | 0.78 |  | 2.51 | 3.05 | 3.86 |  | -0.49 | 0.05 | 0.86
|  | 39.7219 | 61.4565 | 119.8083 |  | 15.4303 | 14.4267 | 8.2381 |  | 5.5306 | 5.1708 | 2.9527 |  | 219.6846 | 317.7823 | 353.7609 |  | 11.09 | 12.15 | 12.16 |  | 0.09 | 1.15 | 1.16 |  | 2.39 | 2.98 | 3.75 |  | -0.61 | -0.02 | 0.75
|  | 39.8883 | 58.8466 | 121.7532 |  | 15.7203 | 14.8425 | 7.9167 |  | 5.6345 | 5.3199 | 2.8375 |  | 224.7514 | 313.0574 | 345.4789 |  | 11 | 12.83 | 11.75 |  | 0 | 1.83 | 0.75 |  | 2.5 | 2.92 | 3.85 |  | -0.5 | -0.08 | 0.85
|  | 39.8142 | 59.7312 | 121.6792 |  | 15.5507 | 13.8245 | 7.9262 |  | 5.5737 | 4.955 | 2.8409 |  | 221.9133 | 295.9692 | 345.6813 |  | 11.03 | 12.06 | 10.93 |  | 0.03 | 1.06 | -0.07 |  | 2.56 | 2.98 | 3.83 |  | -0.44 | -0.02 | 0.83
|  | 39.9787 | 60.4656 | 116.0542 |  | 15.3713 | 14.1886 | 8.1119 |  | 5.5094 | 5.0855 | 2.9075 |  | 220.2592 | 307.4994 | 337.4279 |  | 11.08 | 12.25 | 11.52 |  | 0.08 | 1.25 | 0.52 |  | 2.34 | 2.94 | 4.04 |  | -0.66 | -0.06 | 1.04
|  | 39.8195 | 59.5061 | 121.2366 |  | 15.7939 | 13.9447 | 7.6467 |  | 5.6609 | 4.9981 | 2.7408 |  | 225.4138 | 297.4173 | 332.28 |  | 11.09 | 12.15 | 11.73 |  | 0.09 | 1.15 | 0.73 |  | 2.52 | 3.08 | 3.85 |  | -0.48 | 0.08 | 0.85
|  | 39.6458 | 60.4169 | 119.1658 |  | 15.3456 | 13.9031 | 8.5518 |  | 5.5002 | 4.9832 | 3.0652 |  | 218.061 | 301.0684 | 365.2615 |  | 11.18 | 12.28 | 11.46 |  | 0.18 | 1.28 | 0.46 |  | 2.41 | 2.92 | 3.94 |  | -0.59 | -0.08 | 0.94
|  | 40.0534 | 59.23 | 121.2856 |  | 15.4128 | 15.9854 | 8.1106 |  | 5.5243 | 5.7295 | 2.907 |  | 221.267 | 339.3607 | 352.5802 |  | 10.98 | 12.78 | 11.28 |  | -0.02 | 1.78 | 0.28 |  | 2.43 | 2.9 | 3.98 |  | -0.57 | -0.1 | 0.98
|  | 39.9122 | 60.1564 | 116.8907 |  | 15.435 | 15.2814 | 8.2238 |  | 5.5323 | 5.4772 | 2.9476 |  | 220.8047 | 329.4896 | 344.5482 |  | 11.02 | 12.99 | 11.29 |  | 0.02 | 1.99 | 0.29 |  | 2.43 | 2.94 | 3.85 |  | -0.57 | -0.06 | 0.85
|  | 39.4322 | 59.6481 | 118.087 |  | 15.7498 | 15.4521 | 8.3865 |  | 5.6451 | 5.5384 | 3.0059 |  | 222.598 | 330.3544 | 354.9594 |  | 11.16 | 12.8 | 12.12 |  | 0.16 | 1.8 | 1.12 |  | 2.42 | 3.04 | 3.83 |  | -0.58 | 0.04 | 0.83
|  | 40.2172 | 59.2534 | 119.1185 |  | 15.5246 | 15.1425 | 7.2683 |  | 5.5644 | 5.4274 | 2.6051 |  | 223.7834 | 321.5923 | 310.3203 |  | 11.05 | 13.08 | 10.98 |  | 0.05 | 2.08 | -0.02 |  | 2.5 | 3.04 | 3.56 |  | -0.5 | 0.04 | 0.56
|  | 39.6458 | 59.8384 | 119.0004 |  | 15.3063 | 15.1926 | 7.5673 |  | 5.4861 | 5.4454 | 2.7123 |  | 217.502 | 325.8418 | 322.7653 |  | 11 | 12.77 | 11.14 |  | 0 | 1.77 | 0.14 |  | 2.41 | 2.96 | 3.84 |  | -0.59 | -0.04 | 0.84
|  | 39.8221 | 60.2954 | 120.7001 |  | 15.5623 | 15.2187 | 7.67 |  | 5.5779 | 5.4547 | 2.7491 |  | 222.1229 | 328.8945 | 331.8172 |  | 11.06 | 13.14 | 11.42 |  | 0.06 | 2.14 | 0.42 |  | 2.41 | 3.06 | 3.77 |  | -0.59 | 0.06 | 0.77
|  | 40.107 | 60.1383 | 121.8769 |  | 15.5198 | 15.7206 | 8.0823 |  | 5.5627 | 5.6346 | 2.8969 |  | 223.1013 | 338.8567 | 353.064 |  | 11.09 | 12.97 | 11.48 |  | 0.09 | 1.97 | 0.48 |  | 2.5 | 3.09 | 3.91 |  | -0.5 | 0.09 | 0.91
|  | 39.6642 | 58.548 | 120.1923 |  | 16.0834 | 15.7463 | 8.4399 |  | 5.7647 | 5.6438 | 3.0251 |  | 228.6509 | 330.4356 | 363.5891 |  | 11.19 | 13.14 | 11.05 |  | 0.19 | 2.14 | 0.05 |  | 2.13 | 2.97 | 3.94 |  | -0.87 | -0.03 | 0.94
|  | 40.2901 | 60.1383 | 116.3016 |  | 16.0361 | 15.4803 | 8.7873 |  | 5.7477 | 5.5485 | 3.1496 |  | 231.5759 | 333.6773 | 366.3001 |  | 11.08 | 12.91 | 11.27 |  | 0.08 | 1.91 | 0.27 |  | 2.07 | 3.01 | 3.88 |  | -0.93 | 0.01 | 0.88
|  | 39.5909 | 60.2107 | 119.6411 |  | 16.1991 | 15.1264 | 8.084 |  | 5.8061 | 5.4216 | 2.8975 |  | 229.8693 | 326.4413 | 346.6588 |  | 11.16 | 13 | 11.39 |  | 0.16 | 2 | 0.39 |  | 2.08 | 3.12 | 3.92 |  | -0.92 | 0.12 | 0.92
|  | 39.3959 | 59.6362 | 122.2743 |  | 15.9454 | 14.9376 | 7.6546 |  | 5.7152 | 5.354 | 2.7436 |  | 225.1554 | 319.2901 | 335.4697 |  | 10.98 | 12.96 | 12.01 |  | -0.02 | 1.96 | 1.01 |  | 2.07 | 3.02 | 3.83 |  | -0.93 | 0.02 | 0.83
|  | 39.976 | 59.7193 | 116.8679 |  | 15.8981 | 15.4072 | 8.0477 |  | 5.6983 | 5.5223 | 2.8845 |  | 227.7935 | 329.7884 | 337.1039 |  | 11.06 | 12.91 | 10.75 |  | 0.06 | 1.91 | -0.25 |  | 2.19 | 3.03 | 3.43 |  | -0.81 | 0.03 | 0.43
|  | 41.2712 | 60.0601 | 121.7285 |  | 16.184 | 14.9719 | 7.4883 |  | 5.8007 | 5.3663 | 2.684 |  | 239.4017 | 322.2976 | 326.7163 |  | 11.03 | 12.82 | 11.31 |  | 0.03 | 1.82 | 0.31 |  | 1.77 | 3.03 | 3.86 |  | -1.23 | 0.03 | 0.86
|  | 39.7957 | 59.4943 | 120.2646 |  | 21.7464 | 15.6001 | 8.1325 |  | 7.7944 | 5.5915 | 2.9149 |  | 310.1842 | 332.6595 | 350.5581 |  | 10.99 | 12.82 | 11.53 |  | -0.01 | 1.82 | 0.53 |  | -0.52 | 3.07 | 3.82 |  | -3.52 | 0.07 | 0.82
|  | 40.1849 | 58.9971 | 117.1646 |  | 22.1561 | 14.3543 | 7.8508 |  | 7.9413 | 5.1449 | 2.8139 |  | 319.118 | 303.5344 | 329.6891 |  | 10.99 | 12.79 | 11.63 |  | -0.01 | 1.79 | 0.63 |  | -0.53 | 2.96 | 3.92 |  | -3.53 | -0.04 | 0.92
|  | 39.6249 | 29.8404 | 122.8501 |  | 22.7226 | 15.1472 | 8.0269 |  | 8.1443 | 5.4291 | 2.877 |  | 322.7167 | 162.0068 | 353.4425 |  | 11.1 | 12.81 | 11.09 |  | 0.1 | 1.81 | 0.09 |  | -0.49 | 3 | 3.87 |  | -3.49 | 0 | 0.87
|  | 39.6301 | 61.2745 | 114.2422 |  | 22.5313 | 17.07 | 8.5061 |  | 8.0757 | 6.1183 | 3.0488 |  | 320.0426 | 374.894 | 348.3008 |  | 11.11 | 11.31 | 11.53 |  | 0.11 | 0.31 | 0.53 |  | -0.47 | 0.67 | 3.74 |  | -3.47 | -2.33 | 0.74
|  | 40.4968 | 59.3178 | 120.4577 |  | 22.9427 | 15.5252 | 7.2801 |  | 8.2232 | 5.5646 | 2.6094 |  | 333.0119 | 330.0787 | 314.3189 |  | 11.1 | 12.85 | 11.72 |  | 0.1 | 1.85 | 0.72 |  | -0.55 | 2.78 | 3.59 |  | -3.55 | -0.22 | 0.59
|  | 39.44 | 59.7669 | 120.5788 |  | 22.7388 | 15.2788 | 7.9838 |  | 8.1501 | 5.4763 | 2.8616 |  | 321.4403 | 327.3006 | 345.0451 |  | 10.95 | 12.7 | 11.91 |  | -0.05 | 1.7 | 0.91 |  | -0.5 | 2.93 | 3.77 |  | -3.5 | -0.07 | 0.77
|  | 39.6275 | 59.2359 | 118.7178 |  | 22.4601 | 15.6716 | 7.8887 |  | 8.0502 | 5.6171 | 2.8275 |  | 319.0101 | 332.732 | 335.6726 |  | 10.94 | 12.97 | 11.42 |  | -0.06 | 1.97 | 0.42 |  | -0.53 | 3.06 | 3.84 |  | -3.53 | 0.06 | 0.84
|  | 40.1365 | 30.2221 | 119.7366 |  | 23.1979 | 15.3578 | 7.7994 |  | 8.3147 | 5.5046 | 2.7955 |  | 333.721 | 166.3602 | 334.7211 |  | 11.19 | 12.62 | 11.55 |  | 0.19 | 1.62 | 0.55 |  | -0.54 | 3.05 | 3.6 |  | -3.54 | 0.05 | 0.6
|  | 38.8123 | 58.7544 | 121.9017 |  | 21.9083 | 19.2112 | 7.0212 |  | 7.8524 | 6.8857 | 2.5166 |  | 304.7718 | 404.5668 | 306.7741 |  | 11.08 | 11.86 | 11.46 |  | 0.08 | 0.86 | 0.46 |  | 0.76 | 0.7 | 3.72 |  | -2.24 | -2.3 | 0.72
|  | 40.0481 | 59.9042 | 120.3852 |  | 16.9301 | 16.7699 | 8.4682 |  | 6.0681 | 6.0107 | 3.0352 |  | 243.0173 | 360.0663 | 365.3917 |  | 11.04 | 12.89 | 11.75 |  | 0.04 | 1.89 | 0.75 |  | 1.92 | 2.87 | 3.73 |  | -1.08 | -0.13 | 0.73
|  | 39.4919 | 59.3824 | 118.5068 |  | 16.8107 | 16.2179 | 7.9676 |  | 6.0253 | 5.8129 | 2.8558 |  | 237.952 | 345.1823 | 338.429 |  | 11.01 | 12.72 | 11.02 |  | 0.01 | 1.72 | 0.02 |  | 2.01 | 3.03 | 3.9 |  | -0.99 | 0.03 | 0.9
|  | 39.9441 | 59.8265 | 119.2132 |  | 17.0083 | 16.688 | 8.097 |  | 6.0962 | 5.9814 | 2.9022 |  | 243.506 | 357.8433 | 345.9756 |  | 11.12 | 13.06 | 11.13 |  | 0.12 | 2.06 | 0.13 |  | 2.15 | 3.03 | 3.74 |  | -0.85 | 0.03 | 0.74
|  | 39.9574 | 59.6184 | 117.624 |  | 16.8293 | 16.7374 | 8.6532 |  | 6.032 | 5.9991 | 3.1015 |  | 241.0228 | 357.6549 | 364.8105 |  | 11.02 | 12.92 | 12.16 |  | 0.02 | 1.92 | 1.16 |  | 1.97 | 2.96 | 3.7 |  | -1.03 | -0.04 | 0.7
|  | 39.4192 | 60.2773 | 120.9677 |  | 16.6719 | 16.0032 | 7.2649 |  | 5.9756 | 5.7359 | 2.6039 |  | 235.5536 | 345.7447 | 314.9874 |  | 11.05 | 13.11 | 10.91 |  | 0.05 | 2.11 | -0.09 |  | 2 | 2.96 | 3.73 |  | -1 | -0.04 | 0.73
|  | 40.3741 | 59.2183 | 119.6411 |  | 16.7935 | 15.3709 | 7.3227 |  | 6.0192 | 5.5093 | 2.6246 |  | 243.0194 | 326.2511 | 314.0128 |  | 11.07 | 12.75 | 10.84 |  | 0.07 | 1.75 | -0.16 |  | 2.09 | 3.16 | 3.8 |  | -0.91 | 0.16 | 0.8
|  | 39.3236 | 60.2773 | 120.9921 |  | 16.846 | 15.7071 | 7.3252 |  | 6.038 | 5.6298 | 2.6255 |  | 237.4358 | 339.349 | 317.6681 |  | 11.03 | 12.98 | 11.54 |  | 0.03 | 1.98 | 0.54 |  | 1.96 | 3.01 | 3.36 |  | -1.04 | 0.01 | 0.36
|  | 39.693 | 59.7669 | 119.7366 |  | 17.0621 | 16.4 | 7.6792 |  | 6.1155 | 5.8781 | 2.7524 |  | 242.7411 | 351.3186 | 329.5636 |  | 11.1 | 12.82 | 11.63 |  | 0.1 | 1.82 | 0.63 |  | 2 | 2.96 | 3.94 |  | -1 | -0.04 | 0.94
|  | 40.5899 | 59.6007 | 117.9709 |  | 16.7772 | 17.1081 | 7.3177 |  | 6.0134 | 6.1319 | 2.6228 |  | 244.0813 | 365.4671 | 309.4162 |  | 11.1 | 13.12 | 10.88 |  | 0.1 | 2.12 | -0.12 |  | 2.09 | 2.92 | 3.86 |  | -0.91 | -0.08 | 0.86
|  | 39.1978 | 60.3561 | 119.0476 |  | 16.6741 | 17.0144 | 7.5129 |  | 5.9764 | 6.0983 | 2.6928 |  | 234.2607 | 368.0725 | 320.5713 |  | 11.05 | 12.8 | 11.23 |  | 0.05 | 1.8 | 0.23 |  | 1.98 | 3.01 | 3.46 |  | -1.02 | 0.01 | 0.46
|  | 39.8089 | 59.3003 | 118.1102 |  | 16.5791 | 17.0197 | 7.2937 |  | 5.9423 | 6.1002 | 2.6142 |  | 236.5584 | 361.7461 | 308.7681 |  | 10.98 | 13.04 | 11.26 |  | -0.02 | 2.04 | 0.26 |  | 1.96 | 3.04 | 3.8 |  | -1.04 | 0.04 | 0.8
|  | 39.9414 | 59.761 | 123.3046 |  | 16.7922 | 16.701 | 7.2436 |  | 6.0187 | 5.986 | 2.5963 |  | 240.3953 | 357.7313 | 320.1321 |  | 11.13 | 12.8 | 11.35 |  | 0.13 | 1.8 | 0.35 |  | 2.06 | 2.92 | 3.75 |  | -0.94 | -0.08 | 0.75
|  | 39.9973 | 59.6718 | 119.8083 |  | 16.6284 | 16.8275 | 8.7705 |  | 5.96 | 6.0314 | 3.1436 |  | 238.3847 | 359.9028 | 376.6237 |  | 11.01 | 12.91 | 11.26 |  | 0.01 | 1.91 | 0.26 |  | 1.94 | 3.04 | 3.87 |  | -1.06 | 0.04 | 0.87
|  | 39.6878 | 60.2349 | 117.5549 |  | 17.1854 | 16.3619 | 8.5939 |  | 6.1596 | 5.8645 | 3.0802 |  | 244.4623 | 353.2474 | 362.0976 |  | 11.05 | 12.99 | 11.17 |  | 0.05 | 1.99 | 0.17 |  | 1.96 | 3.05 | 3.87 |  | -1.04 | 0.05 | 0.87
|  |  | 59.5002 | 120.4336 |  | 16.516 | 16.5442 | 8.7487 |  | 5.9197 | 5.9298 | 3.1357 |  |  | 352.8246 | 377.6488 |  | 11.04 | 12.77 | 12.18 |  | 0.04 | 1.77 | 1.18 |  | 2.12 | 2.94 | 3.87 |  | -0.88 | -0.06 | 0.87
|  |  | 59.7491 | 119.7844 |  |  | 15.9285 | 8.6775 |  |  | 5.7091 | 3.1102 |  |  | 341.1152 | 372.5533 |  |  | 12.94 | 12.4 |  |  | 1.94 | 1.4 |  |  | 2.97 | 4.05 |  |  | -0.03 | 1.05
|  |  | 59.8325 | 116.9819 |  |  | 16.4484 | 8.1117 |  |  | 5.8955 | 2.9074 |  |  | 352.7405 | 340.1169 |  |  | 12.89 | 11.24 |  |  | 1.89 | 0.24 |  |  | 3.1 | 3.99 |  |  | 0.1 | 0.99
|  |  | 60.012 | 119.2843 |  |  | 17.0367 | 7.8361 |  |  | 6.1063 | 2.8086 |  |  | 366.4537 | 335.026 |  |  | 13 | 11.62 |  |  | 2 | 0.62 |  |  | 3.08 | 3.34 |  |  | 0.08 | 0.34
|  |  | 58.8986 | 120.6758 |  |  | 15.8658 | 8.2321 |  |  | 5.6867 | 2.9506 |  |  | 334.9376 | 356.0627 |  |  | 12.75 | 12.02 |  |  | 1.75 | 1.02 |  |  | 2.91 | 3.87 |  |  | -0.09 | 0.87
|  |  | 59.8623 | 122.0008 |  |  | 16.4466 | 8.3479 |  |  | 5.8948 | 2.9921 |  |  | 352.8784 | 365.0373 |  |  | 12.9 | 12.06 |  |  | 1.9 | 1.06 |  |  | 2.91 | 3.9 |  |  | -0.09 | 0.9
|  |  | 60.0661 | 115.6738 |  |  | 16.0866 | 8.2489 |  |  | 5.7658 | 2.9566 |  |  | 346.33 | 341.9991 |  |  | 12.74 | 10.76 |  |  | 1.74 | -0.24 |  |  | 3.19 | 3.81 |  |  | 0.19 | 0.81
|  |  | 59.8384 | 120.7001 |  |  | 15.5928 | 7.5721 |  |  | 5.5888 | 2.714 |  |  | 334.4266 | 327.5819 |  |  | 12.93 | 11.3 |  |  | 1.93 | 0.3 |  |  | 2.95 | 3.73 |  |  | -0.05 | 0.73
|  |  | 59.4648 | 121.5067 |  |  | 16.4399 | 8.0869 |  |  | 5.8924 | 2.8985 |  |  | 350.3929 | 352.1924 |  |  | 12.79 | 11.92 |  |  | 1.79 | 0.92 |  |  | 2.96 | 3.83 |  |  | -0.04 | 0.83
|  |  | 60.2712 | 121.3838 |  |  | 16.1813 | 8.4689 |  |  | 5.7997 | 3.0354 |  |  | 349.5569 | 368.453 |  |  | 12.98 | 11.58 |  |  | 1.98 | 0.58 |  |  | 2.99 | 3.87 |  |  | -0.01 | 0.87
|  |  | 59.0958 | 117.4858 |  |  | 16.0677 | 8.6452 |  |  | 5.759 | 3.0986 |  |  | 340.3338 | 364.0441 |  |  | 12.8 | 11.38 |  |  | 1.8 | 0.38 |  |  | 2.9 | 3.63 |  |  | -0.1 | 0.63
|  |  | 60.6306 | 115.6515 |  |  | 16.8007 | 8.3666 |  |  | 6.0218 | 2.9988 |  |  | 365.1024 | 346.8149 |  |  | 13.04 | 11.72 |  |  | 2.04 | 0.72 |  |  | 2.89 | 3.78 |  |  | -0.11 | 0.78
|  |  | 59.1833 | 123.2539 |  |  | 16.75 | 7.5263 |  |  | 6.0036 | 2.6976 |  |  | 355.312 | 332.4889 |  |  | 12.65 | 11.77 |  |  | 1.65 | 0.77 |  |  | 3.01 | 3.54 |  |  | 0.01 | 0.54
|  |  | 60.024 | 117.7163 |  |  | 16.2575 | 7.4335 |  |  | 5.8271 | 2.6643 |  |  | 349.7644 | 313.6338 |  |  | 13.04 | 11.1 |  |  | 2.04 | 0.1 |  |  | 2.9 | 3.9 |  |  | -0.1 | 0.9
|  |  | 59.2008 | 119.7127 |  |  | 16.2941 | 8.178 |  |  | 5.8402 | 2.9312 |  |  | 345.7437 | 350.9011 |  |  | 12.84 | 11.48 |  |  | 1.84 | 0.48 |  |  | 2.89 | 3.81 |  |  | -0.11 | 0.81
|  |  | 60.423 | 121.1632 |  |  | 16.0466 | 7.6111 |  |  | 5.7515 | 2.728 |  |  | 347.5206 | 330.5322 |  |  | 13.11 | 11.39 |  |  | 2.11 | 0.39 |  |  | 2.92 | 3.64 |  |  | -0.08 | 0.64
|  |  | 59.6837 | 119.6649 |  |  | 16.3785 | 7.744 |  |  | 5.8704 | 2.7756 |  |  | 350.3692 | 332.1466 |  |  | 12.92 | 11.44 |  |  | 1.92 | 0.44 |  |  | 2.99 | 3.79 |  |  | -0.01 | 0.79
|  |  | 59.4825 | 116.1665 |  |  | 15.7577 | 7.4618 |  |  | 5.6479 | 2.6745 |  |  | 335.9515 | 310.6834 |  |  | 13.05 | 10.94 |  |  | 2.05 | -0.06 |  |  | 2.99 | 3.74 |  |  | -0.01 | 0.74
|  |  | 59.6481 | 119.952 |  |  | 15.8194 | 7.7199 |  |  | 5.6701 | 2.767 |  |  | 338.2078 | 331.908 |  |  | 12.94 | 10.75 |  |  | 1.94 | -0.25 |  |  | 2.94 | 3.82 |  |  | -0.06 | 0.82
|  |  | 60.1504 | 118.5068 |  |  | 16.1828 | 7.8046 |  |  | 5.8003 | 2.7973 |  |  | 348.8889 | 331.5046 |  |  | 13.05 | 11.4 |  |  | 2.05 | 0.4 |  |  | 2.97 | 3.79 |  |  | -0.03 | 0.79
|  |  | 59.9101 | 120.2164 |  |  | 16.4756 | 7.5385 |  |  | 5.9052 | 2.702 |  |  | 353.7837 | 324.822 |  |  | 12.92 | 11.35 |  |  | 1.92 | 0.35 |  |  | 2.94 | 3.75 |  |  | -0.06 | 0.75
|  |  | 59.23 | 122.125 |  |  | 16.4459 | 7.7405 |  |  | 5.8946 | 2.7744 |  |  | 349.1363 | 338.819 |  |  | 13.09 | 10.76 |  |  | 2.09 | -0.24 |  |  | 2.96 | 3.72 |  |  | -0.04 | 0.72
|  |  | 60.0841 | 120.3611 |  |  | 15.5435 | 8.5216 |  |  | 5.5712 | 3.0543 |  |  | 334.7379 | 367.6239 |  |  | 12.85 | 10.87 |  |  | 1.85 | -0.13 |  |  | 3.03 | 3.9 |  |  | 0.03 | 0.9
|  |  | 59.5297 | 121.1143 |  |  | 16.7013 | 8.1045 |  |  | 5.9861 | 2.9048 |  |  | 356.3529 | 351.8157 |  |  | 13.11 | 11.02 |  |  | 2.11 | 0.02 |  |  | 2.93 | 3.87 |  |  | -0.07 | 0.87
|  |  | 59.9101 | 116.7315 |  |  | 15.8761 | 9.22 |  |  | 5.6904 | 3.3046 |  |  | 340.9102 | 385.7559 |  |  | 12.91 | 11.96 |  |  | 1.91 | 0.96 |  |  | 2.9 | 3.96 |  |  | -0.1 | 0.96
|  |  | 59.7312 | 120.1923 |  |  | 16.7608 | 8.0777 |  |  | 6.0075 | 2.8952 |  |  | 358.8333 | 347.9826 |  |  | 12.97 | 11.63 |  |  | 1.97 | 0.63 |  |  | 3.03 | 3.94 |  |  | 0.03 | 0.94
|  |  | 59.7967 | 121.2366 |  |  | 17.2037 | 8.4819 |  |  | 6.1662 | 3.0401 |  |  | 368.7187 | 368.5733 |  |  | 12.8 | 11.33 |  |  | 1.8 | 0.33 |  |  | 2.85 | 3.61 |  |  | -0.15 | 0.61
|  |  | 59.5002 | 117.3479 |  |  | 15.5934 | 7.6507 |  |  | 5.589 | 2.7422 |  |  | 332.5494 | 321.7891 |  |  | 12.86 | 11.84 |  |  | 1.86 | 0.84 |  |  | 3.03 | 3.89 |  |  | 0.03 | 0.89
|  |  | 60.3682 | 120.1682 |  |  | 16.9768 | 7.6296 |  |  | 6.0849 | 2.7346 |  |  | 367.334 | 328.6146 |  |  | 12.95 | 11.45 |  |  | 1.95 | 0.45 |  |  | 2.94 | 3.8 |  |  | -0.06 | 0.8
|  |  | 59.512 | 121.0654 |  |  | 16.1765 | 7.6179 |  |  | 5.798 | 2.7304 |  |  | 345.0522 | 330.5616 |  |  | 12.93 | 10.82 |  |  | 1.93 | -0.18 |  |  | 3.03 | 3.91 |  |  | 0.03 | 0.91
|  |  | 59.512 | 117.7394 |  |  | 16.5539 | 8.1042 |  |  | 5.9333 | 2.9047 |  |  | 353.1024 | 342.0002 |  |  | 12.95 | 11.13 |  |  | 1.95 | 0.13 |  |  | 2.9 | 3.83 |  |  | -0.1 | 0.83
|  |  | 60.0721 | 117.8088 |  |  | 15.816 | 8.7629 |  |  | 5.6688 | 3.1408 |  |  | 340.5369 | 370.0171 |  |  | 12.69 | 11.75 |  |  | 1.69 | 0.75 |  |  | 3.06 | 3.8 |  |  | 0.06 | 0.8
|  |  | 59.5888 | 118.8354 |  |  | 16.9437 | 7.9781 |  |  | 6.073 | 2.8595 |  |  | 361.8836 | 339.8148 |  |  | 12.84 | 11.99 |  |  | 1.84 | 0.99 |  |  | 3.03 | 3.89 |  |  | 0.03 | 0.89
|  |  | 59.2593 | 118.8119 |  |  | 15.8883 | 8.257 |  |  | 5.6947 | 2.9595 |  |  | 337.4651 | 351.6245 |  |  | 12.58 | 10.96 |  |  | 1.58 | -0.04 |  |  | 2.95 | 3.89 |  |  | -0.05 | 0.89
|  |  | 60.3258 | 121.4329 |  |  | 15.5529 | 7.7752 |  |  | 5.5745 | 2.7868 |  |  | 336.2869 | 338.4123 |  |  | 12.84 | 10.98 |  |  | 1.84 | -0.02 |  |  | 2.92 | 3.74 |  |  | -0.08 | 0.74
|  |  | 59.8146 | 118.7884 |  |  | 15.3259 | 7.2764 |  |  | 5.4932 | 2.608 |  |  | 328.5713 | 309.8017 |  |  | 12.64 | 11.61 |  |  | 1.64 | 0.61 |  |  | 3.01 | 3.59 |  |  | 0.01 | 0.59
|  |  | 59.7907 | 122.2743 |  |  | 16.4325 | 7.4617 |  |  | 5.8898 | 2.6745 |  |  | 352.1541 | 327.0166 |  |  | 12.95 | 11.32 |  |  | 1.95 | 0.32 |  |  | 2.86 | 3.73 |  |  | -0.14 | 0.73
|  |  | 59.8205 | 121.5805 |  |  | 16.7008 | 8.3459 |  |  | 5.9859 | 2.9913 |  |  | 358.0819 | 363.6896 |  |  | 12.74 | 11.25 |  |  | 1.74 | 0.25 |  |  | 2.87 | 3.78 |  |  | -0.13 | 0.78
|  |  | 59.9281 | 118.7178 |  |  | 16.421 | 7.9789 |  |  | 5.8857 | 2.8598 |  |  | 352.7164 | 339.5103 |  |  | 12.92 | 11.25 |  |  | 1.92 | 0.25 |  |  | 3.04 | 3.86 |  |  | 0.04 | 0.86
|  |  | 59.3648 | 118.1102 |  |  | 15.4056 | 8.903 |  |  | 5.5217 | 3.191 |  |  | 327.7966 | 376.8933 |  |  | 12.81 | 12.06 |  |  | 1.81 | 1.06 |  |  | 3.01 | 3.86 |  |  | 0.01 | 0.86
|  |  | 59.4177 | 117.5549 |  |  | 16.0465 | 7.8747 |  |  | 5.7514 | 2.8225 |  |  | 341.7371 | 331.794 |  |  | 12.99 | 11.31 |  |  | 1.99 | 0.31 |  |  | 2.85 | 3.86 |  |  | -0.15 | 0.86
|  |  | 60.1022 | 118.624 |  |  | 16.8712 | 7.4514 |  |  | 6.047 | 2.6708 |  |  | 363.4402 | 316.8172 |  |  | 12.82 | 10.52 |  |  | 1.82 | -0.48 |  |  | 2.94 | 3.8 |  |  | -0.06 | 0.8
|  |  | 60.3865 | 119.952 |  |  | 16.3264 | 8.0016 |  |  | 5.8518 | 2.8679 |  |  | 353.3673 | 344.0159 |  |  | 13.03 | 11.7 |  |  | 2.03 | 0.7 |  |  | 3.04 | 3.77 |  |  | 0.04 | 0.77
|  |  | 59.09 | 122.3491 |  |  | 16.3274 | 8.2066 |  |  | 5.8521 | 2.9414 |  |  | 345.8009 | 359.8822 |  |  | 12.9 | 12.01 |  |  | 1.9 | 1.01 |  |  | 2.95 | 3.83 |  |  | -0.05 | 0.83
|  |  | 59.958 | 116.8224 |  |  | 16.3643 | 8.3974 |  |  | 5.8653 | 3.0098 |  |  | 351.6748 | 351.6144 |  |  | 12.95 | 11.57 |  |  | 1.95 | 0.57 |  |  | 2.98 | 3.91 |  |  | -0.02 | 0.91
|  |  | 59.6718 | 126.4755 |  |  | 16.3938 | 7.7832 |  |  | 5.8759 | 2.7897 |  |  | 350.6259 | 352.8253 |  |  | 12.91 | 11.17 |  |  | 1.91 | 0.17 |  |  | 3.03 | 3.88 |  |  | 0.03 | 0.88
|  |  | 59.4825 | 114.7228 |  |  | 15.9648 | 9.2836 |  |  | 5.7222 | 3.3275 |  |  | 340.3685 | 381.7367 |  |  | 12.96 | 11.77 |  |  | 1.96 | 0.77 |  |  | 3.01 | 3.95 |  |  | 0.01 | 0.95
|  |  | 59.8802 | 120.2164 |  |  | 15.8635 | 7.6644 |  |  | 5.6858 | 2.7471 |  |  | 340.4695 | 330.2442 |  |  | 12.71 | 11.86 |  |  | 1.71 | 0.86 |  |  | 2.95 | 3.79 |  |  | -0.05 | 0.79
|  |  | 59.9281 | 121.041 |  |  | 15.6752 | 7.9829 |  |  | 5.6184 | 2.8612 |  |  | 336.6976 | 346.3272 |  |  | 12.95 | 11.04 |  |  | 1.95 | 0.04 |  |  | 3.03 | 3.88 |  |  | 0.03 | 0.88
|  |  | 59.23 | 120.7486 |  |  | 15.547 | 8.0587 |  |  | 5.5724 | 2.8884 |  |  | 330.0543 | 348.7748 |  |  | 12.78 | 10.65 |  |  | 1.78 | -0.35 |  |  | 3.02 | 3.83 |  |  | 0.02 | 0.83
|  |  | 60.5449 | 116.0093 |  |  | 15.8683 | 7.8177 |  |  | 5.6876 | 2.802 |  |  | 344.354 | 325.0634 |  |  | 12.87 | 11.21 |  |  | 1.87 | 0.21 |  |  | 2.89 | 3.87 |  |  | -0.11 | 0.87
|  |  | 59.4825 | 121.5067 |  |  | 16.2768 | 8.502 |  |  | 5.834 | 3.0473 |  |  | 347.0207 | 370.2673 |  |  | 12.84 | 11.53 |  |  | 1.84 | 0.53 |  |  | 2.92 | 3.71 |  |  | -0.08 | 0.71
|  |  | 59.8922 | 116.9819 |  |  | 16.6148 | 7.7464 |  |  | 5.9551 | 2.7765 |  |  | 356.6656 | 324.7982 |  |  | 12.77 | 10.8 |  |  | 1.77 | -0.2 |  |  | 3.02 | 3.82 |  |  | 0.02 | 0.82
|  |  | 59.976 | 118.2499 |  |  | 16.2981 | 7.7485 |  |  | 5.8416 | 2.7772 |  |  | 350.3561 | 328.4063 |  |  | 12.69 | 11.16 |  |  | 1.69 | 0.16 |  |  | 2.92 | 3.39 |  |  | -0.08 | 0.39
|  |  | 59.2944 | 121.9264 |  |  | 16.7545 | 8.31 |  |  | 6.0052 | 2.9785 |  |  | 356.0745 | 363.1562 |  |  | 12.74 | 11.8 |  |  | 1.74 | 0.8 |  |  | 2.99 | 3.84 |  |  | -0.01 | 0.84
|  |  | 59.8444 | 116.5728 |  |  | 15.1475 | 8.8213 |  |  | 5.4292 | 3.1618 |  |  | 324.9084 | 368.5745 |  |  | 12.92 | 11.91 |  |  | 1.92 | 0.91 |  |  | 3.08 | 3.93 |  |  | 0.08 | 0.93
|  |  | 59.976 | 120.7001 |  |  | 15.8275 | 8.1883 |  |  | 5.6729 | 2.9349 |  |  | 340.2403 | 354.2382 |  |  | 12.58 | 10.86 |  |  | 1.58 | -0.14 |  |  | 3.04 | 3.76 |  |  | 0.04 | 0.76
|  |  | 59.1191 | 122.8753 |  |  | 15.5036 | 8.2889 |  |  | 5.5568 | 2.9709 |  |  | 328.5152 | 365.055 |  |  | 13 | 11.51 |  |  | 2 | 0.51 |  |  | 2.86 | 3.88 |  |  | -0.14 | 0.88
|  |  | 60.2531 | 119.8562 |  |  | 15.8238 | 8.0234 |  |  | 5.6716 | 2.8758 |  |  | 341.733 | 344.6801 |  |  | 12.61 | 11.76 |  |  | 1.61 | 0.76 |  |  | 2.96 | 3.86 |  |  | -0.04 | 0.86
|  |  | 59.9341 | 118.859 |  |  | 15.6101 | 8.3149 |  |  | 5.595 | 2.9803 |  |  | 335.3322 | 354.2308 |  |  | 12.84 | 11.82 |  |  | 1.84 | 0.82 |  |  | 3.02 | 3.33 |  |  | 0.02 | 0.33
|  |  | 59.6718 | 115.5847 |  |  | 16.771 | 8.5843 |  |  | 6.0111 | 3.0768 |  |  | 358.693 | 355.6318 |  |  | 12.62 | 10.81 |  |  | 1.62 | -0.19 |  |  | 2.94 | 3.86 |  |  | -0.06 | 0.86
|  |  | 59.6837 | 121.8769 |  |  | 16.3925 | 7.8583 |  |  | 5.8755 | 2.8166 |  |  | 350.6686 | 343.2794 |  |  | 12.94 | 11.19 |  |  | 1.94 | 0.19 |  |  | 2.89 | 3.8 |  |  | -0.11 | 0.8
|  |  | 59.654 | 120.2646 |  |  | 15.8862 | 7.6767 |  |  | 5.694 | 2.7515 |  |  | 339.6682 | 330.9096 |  |  | 12.72 | 11.79 |  |  | 1.72 | 0.79 |  |  | 3.03 | 3.83 |  |  | 0.03 | 0.83
|  |  | 60.1685 | 117.3479 |  |  | 16.6829 | 7.7409 |  |  | 5.9795 | 2.7745 |  |  | 359.7792 | 325.5823 |  |  | 12.96 | 11.5 |  |  | 1.96 | 0.5 |  |  | 2.95 | 3.95 |  |  | -0.05 | 0.95
|  |  | 59.5652 | 120.0961 |  |  | 15.8357 | 7.2872 |  |  | 5.6759 | 2.6119 |  |  | 338.0841 | 313.6807 |  |  | 12.83 | 10.65 |  |  | 1.83 | -0.35 |  |  | 3.06 | 3.75 |  |  | 0.06 | 0.75
|  |  | 59.9401 | 119.8801 |  |  | 16.046 | 7.4881 |  |  | 5.7513 | 2.6839 |  |  | 344.7316 | 321.7452 |  |  | 13.11 | 10.98 |  |  | 2.11 | -0.02 |  |  | 3.11 | 3.47 |  |  | 0.11 | 0.47
|  |  | 59.7253 | 118.9061 |  |  | 16.3439 | 7.6257 |  |  | 5.858 | 2.7332 |  |  | 349.873 | 324.9979 |  |  | 12.74 | 11.75 |  |  | 1.74 | 0.75 |  |  | 3.05 | 3.8 |  |  | 0.05 | 0.8
|  |  | 59.1949 | 119.2132 |  |  | 15.8144 | 7.9104 |  |  | 5.6682 | 2.8353 |  |  | 335.5316 | 338.0015 |  |  | 12.91 | 11.37 |  |  | 1.91 | 0.37 |  |  | 2.99 | 3.82 |  |  | -0.01 | 0.82
|  |  | 60.49 | 118.7884 |  |  | 16.5075 | 7.2926 |  |  | 5.9167 | 2.6139 |  |  | 357.9 | 310.4951 |  |  | 12.84 | 10.86 |  |  | 1.84 | -0.14 |  |  | 2.99 | 3.71 |  |  | -0.01 | 0.71
|  |  | 59.5061 | 118.8825 |  |  | 16.3671 | 7.8126 |  |  | 5.8663 | 2.8002 |  |  | 349.0826 | 332.8977 |  |  | 12.98 | 11.47 |  |  | 1.98 | 0.47 |  |  | 3.14 | 3.8 |  |  | 0.14 | 0.8
|  |  | 60.4534 | 120.6273 |  |  | 15.8143 | 7.6338 |  |  | 5.6682 | 2.7361 |  |  | 342.6618 | 330.0525 |  |  | 12.86 | 11.58 |  |  | 1.86 | 0.58 |  |  | 2.94 | 3.78 |  |  | -0.06 | 0.78
|  |  | 59.0958 | 120.6273 |  |  | 17.0185 | 8.3495 |  |  | 6.0998 | 2.9927 |  |  | 360.4735 | 360.9953 |  |  | 13.19 | 11.3 |  |  | 2.19 | 0.3 |  |  | 3.09 | 3.74 |  |  | 0.09 | 0.74
|  |  | 59.6007 | 120.6273 |  |  | 16.2054 | 8.0175 |  |  | 5.8084 | 2.8737 |  |  | 346.1842 | 346.6412 |  |  | 12.87 | 11.03 |  |  | 1.87 | 0.03 |  |  | 3.08 | 3.9 |  |  | 0.08 | 0.9
|  |  | 59.8146 | 119.4743 |  |  | 15.8162 | 7.5257 |  |  | 5.6689 | 2.6974 |  |  | 339.0813 | 322.2664 |  |  | 13.06 | 11.11 |  |  | 2.06 | 0.11 |  |  | 3.01 | 3.34 |  |  | 0.01 | 0.34
|  |  | 59.3531 | 119.8562 |  |  | 15.2786 | 8.0306 |  |  | 5.4762 | 2.8783 |  |  | 325.0286 | 344.9865 |  |  | 12.71 | 11.57 |  |  | 1.71 | 0.57 |  |  | 2.87 | 3.83 |  |  | -0.13 | 0.83
|  |  | 61.5385 | 117.4398 |  |  | 15.784 | 8.1806 |  |  | 5.6573 | 2.9321 |  |  | 348.1443 | 344.3491 |  |  | 12.98 | 11.37 |  |  | 1.98 | 0.37 |  |  | 3.17 | 3.86 |  |  | 0.17 | 0.86
|  |  | 59.4589 | 121.778 |  |  | 15.8947 | 8.3859 |  |  | 5.697 | 3.0057 |  |  | 338.7397 | 366.0269 |  |  | 13 | 10.95 |  |  | 2 | -0.05 |  |  | 3.13 | 3.86 |  |  | 0.13 | 0.86
|  |  | 59.1075 | 117.6471 |  |  | 15.7867 | 7.8184 |  |  | 5.6583 | 2.8023 |  |  | 334.4484 | 329.6833 |  |  | 12.81 | 11.16 |  |  | 1.81 | 0.16 |  |  | 2.99 | 3.86 |  |  | -0.01 | 0.86
|  |  | 59.3296 | 121.0165 |  |  | 15.5653 | 8.3806 |  |  | 5.579 | 3.0038 |  |  | 330.9979 | 363.509 |  |  | 12.82 | 12.04 |  |  | 1.82 | 1.04 |  |  | 2.93 | 3.79 |  |  | -0.07 | 0.79
|  |  | 59.9221 | 123.1274 |  |  | 15.5026 | 7.1469 |  |  | 5.5565 | 2.5616 |  |  | 332.9558 | 315.4061 |  |  | 12.73 | 11 |  |  | 1.73 | 0 |  |  | 3.05 | 3.81 |  |  | 0.05 | 0.81
|  |  | 59.6956 | 114.7447 |  |  | 15.5503 | 8.0175 |  |  | 5.5736 | 2.8736 |  |  | 332.7192 | 329.735 |  |  | 12.88 | 11.31 |  |  | 1.88 | 0.31 |  |  | 2.99 | 3.91 |  |  | -0.01 | 0.91
|  |  | 60.344 | 120.9921 |  |  | 15.8662 | 6.8861 |  |  | 5.6868 | 2.4681 |  |  | 343.1655 | 298.6267 |  |  | 12.66 | 11.57 |  |  | 1.66 | 0.57 |  |  | 2.97 | 3.81 |  |  | -0.03 | 0.81
|  |  | 59.1658 | 117.7856 |  |  | 15.7137 | 7.5829 |  |  | 5.6321 | 2.7179 |  |  | 333.2302 | 320.1294 |  |  | 12.85 | 11.6 |  |  | 1.85 | 0.6 |  |  | 3.07 | 3.93 |  |  | 0.07 | 0.93
|  |  | 59.3942 | 118.46 |  |  | 16.9436 | 7.3098 |  |  | 6.073 | 2.62 |  |  | 360.7002 | 310.365 |  |  | 12.56 | 10.73 |  |  | 1.56 | -0.27 |  |  | 2.85 | 3.86 |  |  | -0.15 | 0.86
|  |  | 59.8504 | 120.8703 |  |  | 16.321 | 7.3081 |  |  | 5.8498 | 2.6194 |  |  | 350.1129 | 316.6074 |  |  | 12.88 | 11 |  |  | 1.88 | 0 |  |  | 2.87 | 3.83 |  |  | -0.13 | 0.83
|  |  | 60.6612 | 121.5067 |  |  | 15.7295 | 7.3242 |  |  | 5.6378 | 2.6252 |  |  | 341.9969 | 318.9765 |  |  | 12.66 | 11.55 |  |  | 1.66 | 0.55 |  |  | 3.08 | 3.85 |  |  | 0.08 | 0.85
|  |  | 58.6797 | 118.3665 |  |  | 16.0079 | 8.3684 |  |  | 5.7376 | 2.9994 |  |  | 336.6806 | 355.0302 |  |  | 13.08 | 11.68 |  |  | 2.08 | 0.68 |  |  | 3.1 | 3.87 |  |  | 0.1 | 0.87
|  |  | 60.2228 | 117.9709 |  |  | 15.8379 | 7.8698 |  |  | 5.6767 | 2.8207 |  |  | 341.8653 | 332.7634 |  |  | 12.69 | 10.63 |  |  | 1.69 | -0.37 |  |  | 2.96 | 3.9 |  |  | -0.04 | 0.9
|  |  | 59.6007 | 119.4268 |  |  | 15.5547 | 7.1864 |  |  | 5.5752 | 2.5758 |  |  | 332.2833 | 307.6141 |  |  | 12.99 | 11.03 |  |  | 1.99 | 0.03 |  |  | 3.03 | 3.62 |  |  | 0.03 | 0.62
|  |  | 60 | 121.3347 |  |  | 15.9329 | 6.8397 |  |  | 5.7107 | 2.4515 |  |  | 342.6426 | 297.4506 |  |  | 12.69 | 11.7 |  |  | 1.69 | 0.7 |  |  | 2.94 | 3.8 |  |  | -0.06 | 0.8
|  |  | 59.4589 | 119.5934 |  |  | 16.1061 | 8.1679 |  |  | 5.7728 | 2.9276 |  |  | 343.2442 | 350.1171 |  |  | 12.91 | 11.81 |  |  | 1.91 | 0.81 |  |  | 2.96 | 3.88 |  |  | -0.04 | 0.88
|  |  | 60.1624 | 117.3479 |  |  | 15.7706 | 8.4791 |  |  | 5.6526 | 3.0391 |  |  | 340.0717 | 356.634 |  |  | 12.67 | 11.37 |  |  | 1.67 | 0.37 |  |  | 2.97 | 3.89 |  |  | -0.03 | 0.89
|  |  | 59.4589 | 121.7039 |  |  | 15.3545 | 7.7335 |  |  | 5.5034 | 2.7719 |  |  | 327.2255 | 337.3485 |  |  | 12.94 | 11.5 |  |  | 1.94 | 0.5 |  |  | 3.1 | 3.83 |  |  | 0.1 | 0.83
|  |  | 59.9101 | 119.9041 |  |  | 15.5552 | 7.3849 |  |  | 5.5753 | 2.6469 |  |  | 334.0192 | 317.3781 |  |  | 12.83 | 11.62 |  |  | 1.83 | 0.62 |  |  | 3.03 | 3.62 |  |  | 0.03 | 0.62
|  |  | 59.1541 | 122.474 |  |  | 16.6446 | 8.2464 |  |  | 5.9658 | 2.9557 |  |  | 352.901 | 361.9983 |  |  | 12.97 | 11.48 |  |  | 1.97 | 0.48 |  |  | 2.97 | 3.88 |  |  | -0.03 | 0.88
|  |  | 60.7841 | 115.9196 |  |  | 15.6075 | 8.531 |  |  | 5.5941 | 3.0577 |  |  | 340.0311 | 354.4479 |  |  | 12.73 | 11.27 |  |  | 1.73 | 0.27 |  |  | 3 | 3.9 |  |  | 0 | 0.9
|  |  |  | 122.5741 |  |  | 17.0802 | 7.7126 |  |  | 6.1219 | 2.7644 |  |  |  | 338.842 |  |  | 13.01 | 11.55 |  |  | 2.01 | 0.55 |  |  | 3.06 | 3.92 |  |  | 0.06 | 0.92
|  |  |  | 113.8088 |  |  |  | 7.7338 |  |  |  | 2.772 |  |  |  | 315.4736 |  |  |  | 11.83 |  |  |  | 0.83 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 121.7039 |  |  |  | 7.308 |  |  |  | 2.6194 |  |  |  | 318.7871 |  |  |  | 11.02 |  |  |  | 0.02 |  |  |  | 3.73 |  |  |  | 0.73
|  |  |  | 122.0504 |  |  |  | 7.3824 |  |  |  | 2.646 |  |  |  | 322.9462 |  |  |  | 10.97 |  |  |  | -0.03 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 117.0275 |  |  |  | 8.0369 |  |  |  | 2.8806 |  |  |  | 337.1102 |  |  |  | 11.38 |  |  |  | 0.38 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 118.5068 |  |  |  | 7.5029 |  |  |  | 2.6892 |  |  |  | 318.6919 |  |  |  | 11.83 |  |  |  | 0.83 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 120.6515 |  |  |  | 7.9885 |  |  |  | 2.8633 |  |  |  | 345.4559 |  |  |  | 10.93 |  |  |  | -0.07 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 120.1682 |  |  |  | 8.0513 |  |  |  | 2.8858 |  |  |  | 346.7767 |  |  |  | 11.3 |  |  |  | 0.3 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 122.399 |  |  |  | 8.4456 |  |  |  | 3.0271 |  |  |  | 370.5141 |  |  |  | 11.35 |  |  |  | 0.35 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 119.5219 |  |  |  | 8.9158 |  |  |  | 3.1956 |  |  |  | 381.9483 |  |  |  | 11.99 |  |  |  | 0.99 |  |  |  | 3.99 |  |  |  | 0.99
|  |  |  | 116.1215 |  |  |  | 8.1089 |  |  |  | 2.9064 |  |  |  | 337.4959 |  |  |  | 11.14 |  |  |  | 0.14 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 121.6545 |  |  |  | 6.3644 |  |  |  | 2.2812 |  |  |  | 277.5135 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 120.3128 |  |  |  | 7.4424 |  |  |  | 2.6675 |  |  |  | 320.9376 |  |  |  | 11.32 |  |  |  | 0.32 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 116.0991 |  |  |  | 7.6285 |  |  |  | 2.7342 |  |  |  | 317.4418 |  |  |  | 11.2 |  |  |  | 0.2 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 119.2369 |  |  |  | 7.2465 |  |  |  | 2.5973 |  |  |  | 309.6954 |  |  |  | 10.64 |  |  |  | -0.36 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.5805 |  |  |  | 7.1281 |  |  |  | 2.5549 |  |  |  | 310.623 |  |  |  | 10.82 |  |  |  | -0.18 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 122.1996 |  |  |  | 6.9117 |  |  |  | 2.4773 |  |  |  | 302.7247 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.5 |  |  |  | 0.5
|  |  |  | 118.0173 |  |  |  | 8.2299 |  |  |  | 2.9498 |  |  |  | 348.1253 |  |  |  | 11.47 |  |  |  | 0.47 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 120 |  |  |  | 8.1082 |  |  |  | 2.9062 |  |  |  | 348.7402 |  |  |  | 10.9 |  |  |  | -0.1 |  |  |  | 3.74 |  |  |  | 0.74
|  |  |  | 116.5275 |  |  |  | 7.7879 |  |  |  | 2.7914 |  |  |  | 325.2724 |  |  |  | 11.89 |  |  |  | 0.89 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 120.2405 |  |  |  | 7.7746 |  |  |  | 2.7866 |  |  |  | 335.0607 |  |  |  | 11.69 |  |  |  | 0.69 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 119.4743 |  |  |  | 8.071 |  |  |  | 2.8928 |  |  |  | 345.6174 |  |  |  | 11.96 |  |  |  | 0.96 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 122.3741 |  |  |  | 8.8392 |  |  |  | 3.1682 |  |  |  | 387.7036 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 118.9768 |  |  |  | 8.3096 |  |  |  | 2.9783 |  |  |  | 354.353 |  |  |  | 11.89 |  |  |  | 0.89 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 119.4505 |  |  |  | 8.2667 |  |  |  | 2.963 |  |  |  | 353.9292 |  |  |  | 12.17 |  |  |  | 1.17 |  |  |  | 3.99 |  |  |  | 0.99
|  |  |  | 120.4336 |  |  |  | 8.248 |  |  |  | 2.9563 |  |  |  | 356.0352 |  |  |  | 11.3 |  |  |  | 0.3 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 118.4834 |  |  |  | 8.6085 |  |  |  | 3.0855 |  |  |  | 365.5788 |  |  |  | 11.21 |  |  |  | 0.21 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 119.6649 |  |  |  | 7.7312 |  |  |  | 2.771 |  |  |  | 331.5965 |  |  |  | 11.55 |  |  |  | 0.55 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 120.048 |  |  |  | 8.2235 |  |  |  | 2.9475 |  |  |  | 353.8401 |  |  |  | 11.86 |  |  |  | 0.86 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 115.2959 |  |  |  | 7.9199 |  |  |  | 2.8387 |  |  |  | 327.2892 |  |  |  | 10.87 |  |  |  | -0.13 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 123.7624 |  |  |  | 7.7634 |  |  |  | 2.7826 |  |  |  | 344.3771 |  |  |  | 10.84 |  |  |  | -0.16 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 117.4628 |  |  |  | 7.1935 |  |  |  | 2.5783 |  |  |  | 302.8547 |  |  |  | 11.28 |  |  |  | 0.28 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 118.18 |  |  |  | 7.2493 |  |  |  | 2.5983 |  |  |  | 307.0674 |  |  |  | 11.51 |  |  |  | 0.51 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 120.9434 |  |  |  | 8.5759 |  |  |  | 3.0738 |  |  |  | 371.7577 |  |  |  | 10.86 |  |  |  | -0.14 |  |  |  | 3.68 |  |  |  | 0.68
|  |  |  | 118.5771 |  |  |  | 7.6496 |  |  |  | 2.7418 |  |  |  | 325.1145 |  |  |  | 11.29 |  |  |  | 0.29 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 121.8522 |  |  |  | 7.4311 |  |  |  | 2.6635 |  |  |  | 324.5502 |  |  |  | 11.55 |  |  |  | 0.55 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 116.9819 |  |  |  | 7.46 |  |  |  | 2.6738 |  |  |  | 312.7899 |  |  |  | 11.7 |  |  |  | 0.7 |  |  |  | 4 |  |  |  | 1
|  |  |  | 121.7285 |  |  |  | 7.8832 |  |  |  | 2.8255 |  |  |  | 343.947 |  |  |  | 10.96 |  |  |  | -0.04 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 125.5756 |  |  |  | 8.1325 |  |  |  | 2.9149 |  |  |  | 366.035 |  |  |  | 11.43 |  |  |  | 0.43 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 113.8952 |  |  |  | 8.6924 |  |  |  | 3.1156 |  |  |  | 354.8472 |  |  |  | 13.6 |  |  |  | 2.6 |  |  |  | 4.05 |  |  |  | 1.05
|  |  |  | 119.1658 |  |  |  | 7.5723 |  |  |  | 2.7141 |  |  |  | 323.4247 |  |  |  | 12.1 |  |  |  | 1.1 |  |  |  | 4.07 |  |  |  | 1.07
|  |  |  | 120.7973 |  |  |  | 7.439 |  |  |  | 2.6663 |  |  |  | 322.0811 |  |  |  | 10.85 |  |  |  | -0.15 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 120.0961 |  |  |  | 8.5713 |  |  |  | 3.0721 |  |  |  | 368.9522 |  |  |  | 11.62 |  |  |  | 0.62 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 118.7648 |  |  |  | 8.7047 |  |  |  | 3.12 |  |  |  | 370.5424 |  |  |  | 12.19 |  |  |  | 1.19 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 118.3199 |  |  |  | 8.2814 |  |  |  | 2.9682 |  |  |  | 351.2009 |  |  |  | 11.65 |  |  |  | 0.65 |  |  |  | 3.99 |  |  |  | 0.99
|  |  |  | 117.4168 |  |  |  | 7.5467 |  |  |  | 2.7049 |  |  |  | 317.6016 |  |  |  | 10.7 |  |  |  | -0.3 |  |  |  | 3.95 |  |  |  | 0.95
|  |  |  | 124.5847 |  |  |  | 7.8601 |  |  |  | 2.8173 |  |  |  | 350.987 |  |  |  | 11.68 |  |  |  | 0.68 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 119.7605 |  |  |  | 8.813 |  |  |  | 3.1588 |  |  |  | 378.2967 |  |  |  | 11.79 |  |  |  | 0.79 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 116.6407 |  |  |  | 8.4496 |  |  |  | 3.0285 |  |  |  | 353.2508 |  |  |  | 11.86 |  |  |  | 0.86 |  |  |  | 4.01 |  |  |  | 1.01
|  |  |  | 122.6492 |  |  |  | 7.6306 |  |  |  | 2.735 |  |  |  | 335.4456 |  |  |  | 11.09 |  |  |  | 0.09 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 117.6932 |  |  |  | 7.9406 |  |  |  | 2.8461 |  |  |  | 334.9656 |  |  |  | 11.34 |  |  |  | 0.34 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 123.0265 |  |  |  | 7.6424 |  |  |  | 2.7392 |  |  |  | 336.9939 |  |  |  | 12.12 |  |  |  | 1.12 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 114.9646 |  |  |  | 9.1818 |  |  |  | 3.291 |  |  |  | 378.3436 |  |  |  | 11.41 |  |  |  | 0.41 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 119.6888 |  |  |  | 8.4884 |  |  |  | 3.0424 |  |  |  | 364.1442 |  |  |  | 11.65 |  |  |  | 0.65 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 120.0961 |  |  |  | 8.2645 |  |  |  | 2.9622 |  |  |  | 355.7473 |  |  |  | 11.83 |  |  |  | 0.83 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 121.7039 |  |  |  | 8.8334 |  |  |  | 3.1661 |  |  |  | 385.3272 |  |  |  | 12.37 |  |  |  | 1.37 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 117.9709 |  |  |  | 8.8309 |  |  |  | 3.1652 |  |  |  | 373.3994 |  |  |  | 11.41 |  |  |  | 0.41 |  |  |  | 3.97 |  |  |  | 0.97
|  |  |  | 120.5788 |  |  |  | 8.5449 |  |  |  | 3.0627 |  |  |  | 369.2934 |  |  |  | 11.74 |  |  |  | 0.74 |  |  |  | 4.02 |  |  |  | 1.02
|  |  |  | 120.048 |  |  |  | 8.0604 |  |  |  | 2.889 |  |  |  | 346.822 |  |  |  | 12 |  |  |  | 1 |  |  |  | 4.02 |  |  |  | 1.02
|  |  |  | 119.6172 |  |  |  | 8.3413 |  |  |  | 2.9897 |  |  |  | 357.6211 |  |  |  | 12.03 |  |  |  | 1.03 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 118.6944 |  |  |  | 7.9992 |  |  |  | 2.8671 |  |  |  | 340.308 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 120.603 |  |  |  | 8.458 |  |  |  | 3.0316 |  |  |  | 365.6146 |  |  |  | 11.43 |  |  |  | 0.43 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.7413 |  |  |  | 9.0156 |  |  |  | 3.2314 |  |  |  | 383.6989 |  |  |  | 11.99 |  |  |  | 0.99 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 119.7605 |  |  |  | 7.815 |  |  |  | 2.8011 |  |  |  | 335.4577 |  |  |  | 11.72 |  |  |  | 0.72 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 119.0476 |  |  |  | 7.8334 |  |  |  | 2.8077 |  |  |  | 334.248 |  |  |  | 11.07 |  |  |  | 0.07 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 116.6407 |  |  |  | 8.4653 |  |  |  | 3.0341 |  |  |  | 353.9047 |  |  |  | 11.52 |  |  |  | 0.52 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.1143 |  |  |  | 7.5978 |  |  |  | 2.7232 |  |  |  | 329.8227 |  |  |  | 11.75 |  |  |  | 0.75 |  |  |  | 3.55 |  |  |  | 0.55
|  |  |  | 117.3939 |  |  |  | 7.4462 |  |  |  | 2.6689 |  |  |  | 313.3092 |  |  |  | 11.4 |  |  |  | 0.4 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 122.3491 |  |  |  | 7.6518 |  |  |  | 2.7426 |  |  |  | 335.551 |  |  |  | 11.31 |  |  |  | 0.31 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 120.7729 |  |  |  | 7.7399 |  |  |  | 2.7742 |  |  |  | 335.0438 |  |  |  | 11.14 |  |  |  | 0.14 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 117.096 |  |  |  | 7.9141 |  |  |  | 2.8366 |  |  |  | 332.1553 |  |  |  | 11.75 |  |  |  | 0.75 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 121.1143 |  |  |  | 7.7522 |  |  |  | 2.7786 |  |  |  | 336.525 |  |  |  | 10.97 |  |  |  | -0.03 |  |  |  | 3.35 |  |  |  | 0.35
|  |  |  | 120.3852 |  |  |  | 8.1885 |  |  |  | 2.9349 |  |  |  | 353.3231 |  |  |  | 11.02 |  |  |  | 0.02 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 119.2606 |  |  |  | 8.1271 |  |  |  | 2.9129 |  |  |  | 347.3994 |  |  |  | 11.39 |  |  |  | 0.39 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 117.8782 |  |  |  | 8.1059 |  |  |  | 2.9054 |  |  |  | 342.4779 |  |  |  | 11.68 |  |  |  | 0.68 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 121.0898 |  |  |  | 7.4208 |  |  |  | 2.6598 |  |  |  | 322.0714 |  |  |  | 10.64 |  |  |  | -0.36 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 117.8782 |  |  |  | 7.7856 |  |  |  | 2.7905 |  |  |  | 328.9437 |  |  |  | 11.71 |  |  |  | 0.71 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 121.1143 |  |  |  | 8.0929 |  |  |  | 2.9007 |  |  |  | 351.3131 |  |  |  | 11.88 |  |  |  | 0.88 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 120.024 |  |  |  | 8.6604 |  |  |  | 3.1041 |  |  |  | 372.5652 |  |  |  | 11.95 |  |  |  | 0.95 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 119.7605 |  |  |  | 8.703 |  |  |  | 3.1194 |  |  |  | 373.5774 |  |  |  | 11.26 |  |  |  | 0.26 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 117.7394 |  |  |  | 7.8873 |  |  |  | 2.827 |  |  |  | 332.847 |  |  |  | 11.52 |  |  |  | 0.52 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 121.2121 |  |  |  | 8.1585 |  |  |  | 2.9242 |  |  |  | 354.4467 |  |  |  | 12.21 |  |  |  | 1.21 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.8354 |  |  |  | 7.9781 |  |  |  | 2.8595 |  |  |  | 339.8119 |  |  |  | 12.13 |  |  |  | 1.13 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 118.7884 |  |  |  | 7.626 |  |  |  | 2.7333 |  |  |  | 324.6876 |  |  |  | 11.23 |  |  |  | 0.23 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 120.024 |  |  |  | 7.4806 |  |  |  | 2.6812 |  |  |  | 321.81 |  |  |  | 11.37 |  |  |  | 0.37 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 119.8083 |  |  |  | 7.8746 |  |  |  | 2.8224 |  |  |  | 338.1509 |  |  |  | 11.8 |  |  |  | 0.8 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 118.624 |  |  |  | 8.5601 |  |  |  | 3.0682 |  |  |  | 363.9563 |  |  |  | 11.6 |  |  |  | 0.6 |  |  |  | 3.95 |  |  |  | 0.95
|  |  |  | 118.1102 |  |  |  | 7.4105 |  |  |  | 2.6561 |  |  |  | 313.7107 |  |  |  | 11.07 |  |  |  | 0.07 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 120.4577 |  |  |  | 7.5991 |  |  |  | 2.7237 |  |  |  | 328.0881 |  |  |  | 11.4 |  |  |  | 0.4 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 119.1185 |  |  |  | 7.8375 |  |  |  | 2.8091 |  |  |  | 334.6186 |  |  |  | 11.76 |  |  |  | 0.76 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 122.524 |  |  |  | 8.0166 |  |  |  | 2.8733 |  |  |  | 352.0508 |  |  |  | 11.47 |  |  |  | 0.47 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 117.9245 |  |  |  | 8.3961 |  |  |  | 3.0094 |  |  |  | 354.8776 |  |  |  | 11.02 |  |  |  | 0.02 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 120.5788 |  |  |  | 8.1394 |  |  |  | 2.9174 |  |  |  | 351.7711 |  |  |  | 11.34 |  |  |  | 0.34 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 118.6944 |  |  |  | 8.2492 |  |  |  | 2.9567 |  |  |  | 350.9439 |  |  |  | 11.98 |  |  |  | 0.98 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 117.1875 |  |  |  | 8.1 |  |  |  | 2.9032 |  |  |  | 340.2227 |  |  |  | 11.37 |  |  |  | 0.37 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 119.4505 |  |  |  | 7.3212 |  |  |  | 2.6241 |  |  |  | 313.4492 |  |  |  | 11.16 |  |  |  | 0.16 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 123.533 |  |  |  | 7.8637 |  |  |  | 2.8185 |  |  |  | 348.1829 |  |  |  | 11.57 |  |  |  | 0.57 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 118.7884 |  |  |  | 8.1621 |  |  |  | 2.9255 |  |  |  | 347.5148 |  |  |  | 12.07 |  |  |  | 1.07 |  |  |  | 3.96 |  |  |  | 0.96
|  |  |  | 119.8083 |  |  |  | 7.5649 |  |  |  | 2.7114 |  |  |  | 324.8529 |  |  |  | 10.88 |  |  |  | -0.12 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 117.2104 |  |  |  | 8.9564 |  |  |  | 3.2102 |  |  |  | 376.2679 |  |  |  | 11.22 |  |  |  | 0.22 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 122.8753 |  |  |  | 8.5529 |  |  |  | 3.0655 |  |  |  | 376.6797 |  |  |  | 11.87 |  |  |  | 0.87 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 115.4512 |  |  |  | 7.9896 |  |  |  | 2.8636 |  |  |  | 330.6118 |  |  |  | 11.86 |  |  |  | 0.86 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 123.0517 |  |  |  | 7.3971 |  |  |  | 2.6513 |  |  |  | 326.2439 |  |  |  | 11.02 |  |  |  | 0.02 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.46 |  |  |  | 7.7469 |  |  |  | 2.7767 |  |  |  | 328.9254 |  |  |  | 11.05 |  |  |  | 0.05 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.8119 |  |  |  | 7.8845 |  |  |  | 2.826 |  |  |  | 335.759 |  |  |  | 11.59 |  |  |  | 0.59 |  |  |  | 3.78 |  |  |  | 0.78
|  |  |  | 118.46 |  |  |  | 7.3865 |  |  |  | 2.6475 |  |  |  | 313.6204 |  |  |  | 11.61 |  |  |  | 0.61 |  |  |  | 3.69 |  |  |  | 0.69
|  |  |  | 118.2266 |  |  |  | 7.2948 |  |  |  | 2.6146 |  |  |  | 309.1193 |  |  |  | 10.81 |  |  |  | -0.19 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 119.5457 |  |  |  | 7.5842 |  |  |  | 2.7183 |  |  |  | 324.9658 |  |  |  | 10.92 |  |  |  | -0.08 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 119.7127 |  |  |  | 7.4911 |  |  |  | 2.685 |  |  |  | 321.427 |  |  |  | 11.39 |  |  |  | 0.39 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 119.7605 |  |  |  | 7.7034 |  |  |  | 2.7611 |  |  |  | 330.669 |  |  |  | 11.34 |  |  |  | 0.34 |  |  |  | 3.74 |  |  |  | 0.74
|  |  |  | 120.5303 |  |  |  | 7.5997 |  |  |  | 2.7239 |  |  |  | 328.3143 |  |  |  | 10.89 |  |  |  | -0.11 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 119.4268 |  |  |  | 7.5592 |  |  |  | 2.7094 |  |  |  | 323.5737 |  |  |  | 11.43 |  |  |  | 0.43 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 122.2743 |  |  |  | 7.6331 |  |  |  | 2.7359 |  |  |  | 334.5291 |  |  |  | 11.65 |  |  |  | 0.65 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 119.308 |  |  |  | 7.9978 |  |  |  | 2.8666 |  |  |  | 342.0082 |  |  |  | 11.14 |  |  |  | 0.14 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 119.6888 |  |  |  | 8.3866 |  |  |  | 3.0059 |  |  |  | 359.7783 |  |  |  | 10.9 |  |  |  | -0.1 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 121.4329 |  |  |  | 8.4235 |  |  |  | 3.0192 |  |  |  | 366.6254 |  |  |  | 11.22 |  |  |  | 0.22 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 118.3199 |  |  |  | 8.122 |  |  |  | 2.9111 |  |  |  | 344.442 |  |  |  | 11.87 |  |  |  | 0.87 |  |  |  | 4.05 |  |  |  | 1.05
|  |  |  | 120.1923 |  |  |  | 8.1222 |  |  |  | 2.9112 |  |  |  | 349.9028 |  |  |  | 11.37 |  |  |  | 0.37 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 115.8078 |  |  |  | 7.7117 |  |  |  | 2.7641 |  |  |  | 320.1004 |  |  |  | 10.96 |  |  |  | -0.04 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.8027 |  |  |  | 7.1913 |  |  |  | 2.5775 |  |  |  | 313.9504 |  |  |  | 11.25 |  |  |  | 0.25 |  |  |  | 3.68 |  |  |  | 0.68
|  |  |  | 119.8322 |  |  |  | 8.0714 |  |  |  | 2.893 |  |  |  | 346.6724 |  |  |  | 11.66 |  |  |  | 0.66 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 120.3852 |  |  |  | 8.2609 |  |  |  | 2.9609 |  |  |  | 356.447 |  |  |  | 11.23 |  |  |  | 0.23 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 117.9245 |  |  |  | 8.665 |  |  |  | 3.1057 |  |  |  | 366.2406 |  |  |  | 11.16 |  |  |  | 0.16 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 119.952 |  |  |  | 7.6236 |  |  |  | 2.7325 |  |  |  | 327.7648 |  |  |  | 11.48 |  |  |  | 0.48 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 118.5068 |  |  |  | 7.8503 |  |  |  | 2.8137 |  |  |  | 333.4461 |  |  |  | 11.69 |  |  |  | 0.69 |  |  |  | 3.97 |  |  |  | 0.97
|  |  |  | 117.9709 |  |  |  | 7.2581 |  |  |  | 2.6015 |  |  |  | 306.8958 |  |  |  | 11.01 |  |  |  | 0.01 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 120.9434 |  |  |  | 8.1217 |  |  |  | 2.911 |  |  |  | 352.0664 |  |  |  | 11.08 |  |  |  | 0.08 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 121.3592 |  |  |  | 7.7085 |  |  |  | 2.7629 |  |  |  | 335.305 |  |  |  | 11.75 |  |  |  | 0.75 |  |  |  | 3.98 |  |  |  | 0.98
|  |  |  | 117.096 |  |  |  | 8.0435 |  |  |  | 2.883 |  |  |  | 337.5861 |  |  |  | 11.51 |  |  |  | 0.51 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 117.4628 |  |  |  | 7.5001 |  |  |  | 2.6882 |  |  |  | 315.7644 |  |  |  | 10.78 |  |  |  | -0.22 |  |  |  | 3.74 |  |  |  | 0.74
|  |  |  | 127.6596 |  |  |  | 7.5996 |  |  |  | 2.7239 |  |  |  | 347.7268 |  |  |  | 11.33 |  |  |  | 0.33 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 114.7228 |  |  |  | 7.6802 |  |  |  | 2.7528 |  |  |  | 315.8061 |  |  |  | 11.79 |  |  |  | 0.79 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 121.3347 |  |  |  | 7.6275 |  |  |  | 2.7339 |  |  |  | 331.712 |  |  |  | 11.58 |  |  |  | 0.58 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 116.8679 |  |  |  | 7.6658 |  |  |  | 2.7476 |  |  |  | 321.1076 |  |  |  | 11.67 |  |  |  | 0.67 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 123.406 |  |  |  | 7.4737 |  |  |  | 2.6787 |  |  |  | 330.5719 |  |  |  | 11.57 |  |  |  | 0.57 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 119.308 |  |  |  | 8.3192 |  |  |  | 2.9818 |  |  |  | 355.7514 |  |  |  | 11.9 |  |  |  | 0.9 |  |  |  | 3.99 |  |  |  | 0.99
|  |  |  | 116.2115 |  |  |  | 7.9648 |  |  |  | 2.8548 |  |  |  | 331.7585 |  |  |  | 11.41 |  |  |  | 0.41 |  |  |  | 3.99 |  |  |  | 0.99
|  |  |  | 119.5457 |  |  |  | 7.2628 |  |  |  | 2.6032 |  |  |  | 311.1978 |  |  |  | 10.86 |  |  |  | -0.14 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 118.6709 |  |  |  | 7.6636 |  |  |  | 2.7468 |  |  |  | 325.9678 |  |  |  | 11.48 |  |  |  | 0.48 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 123.2539 |  |  |  | 7.6232 |  |  |  | 2.7323 |  |  |  | 336.771 |  |  |  | 11.75 |  |  |  | 0.75 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 119.5219 |  |  |  | 8.6746 |  |  |  | 3.1092 |  |  |  | 371.6149 |  |  |  | 11.36 |  |  |  | 0.36 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 119.7127 |  |  |  | 8.3441 |  |  |  | 2.9907 |  |  |  | 358.0248 |  |  |  | 11.55 |  |  |  | 0.55 |  |  |  | 3.96 |  |  |  | 0.96
|  |  |  | 120.5546 |  |  |  | 7.9643 |  |  |  | 2.8546 |  |  |  | 344.1338 |  |  |  | 11.94 |  |  |  | 0.94 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 117.3939 |  |  |  | 8.5459 |  |  |  | 3.0631 |  |  |  | 359.5849 |  |  |  | 11.84 |  |  |  | 0.84 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 119.0712 |  |  |  | 7.8176 |  |  |  | 2.802 |  |  |  | 333.637 |  |  |  | 10.81 |  |  |  | -0.19 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 118.0405 |  |  |  | 7.2727 |  |  |  | 2.6067 |  |  |  | 307.6946 |  |  |  | 10.94 |  |  |  | -0.06 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 120.2405 |  |  |  | 7.3292 |  |  |  | 2.627 |  |  |  | 315.868 |  |  |  | 11.55 |  |  |  | 0.55 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 117.8088 |  |  |  | 7.5878 |  |  |  | 2.7197 |  |  |  | 320.3987 |  |  |  | 11.99 |  |  |  | 0.99 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 120.6758 |  |  |  | 7.0817 |  |  |  | 2.5383 |  |  |  | 306.3057 |  |  |  | 10.67 |  |  |  | -0.33 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 123.2033 |  |  |  | 7.1322 |  |  |  | 2.5563 |  |  |  | 314.9481 |  |  |  | 11.17 |  |  |  | 0.17 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 119.5219 |  |  |  | 8.3634 |  |  |  | 2.9976 |  |  |  | 358.2844 |  |  |  | 11.72 |  |  |  | 0.72 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 117.2104 |  |  |  | 8.3184 |  |  |  | 2.9815 |  |  |  | 349.4628 |  |  |  | 11.58 |  |  |  | 0.58 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 124.5589 |  |  |  | 6.9511 |  |  |  | 2.4914 |  |  |  | 310.3287 |  |  |  | 10.83 |  |  |  | -0.17 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 117.0732 |  |  |  | 8.5324 |  |  |  | 3.0582 |  |  |  | 358.0343 |  |  |  | 11.17 |  |  |  | 0.17 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 120 |  |  |  | 8.3606 |  |  |  | 2.9966 |  |  |  | 359.5974 |  |  |  | 11.74 |  |  |  | 0.74 |  |  |  | 3.78 |  |  |  | 0.78
|  |  |  | 117.2791 |  |  |  | 8.186 |  |  |  | 2.9341 |  |  |  | 344.1048 |  |  |  | 11.6 |  |  |  | 0.6 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 119.6888 |  |  |  | 7.6088 |  |  |  | 2.7272 |  |  |  | 326.4103 |  |  |  | 10.99 |  |  |  | -0.01 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 117.2791 |  |  |  | 7.528 |  |  |  | 2.6982 |  |  |  | 316.4434 |  |  |  | 11.42 |  |  |  | 0.42 |  |  |  | 3.98 |  |  |  | 0.98
|  |  |  | 120.3852 |  |  |  | 7.0442 |  |  |  | 2.5248 |  |  |  | 303.9502 |  |  |  | 11.5 |  |  |  | 0.5 |  |  |  | 3.7 |  |  |  | 0.7
|  |  |  | 120.2405 |  |  |  | 7.724 |  |  |  | 2.7684 |  |  |  | 332.8797 |  |  |  | 11.63 |  |  |  | 0.63 |  |  |  | 3.78 |  |  |  | 0.78
|  |  |  | 117.325 |  |  |  | 7.7652 |  |  |  | 2.7832 |  |  |  | 326.5401 |  |  |  | 11.07 |  |  |  | 0.07 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 123.8135 |  |  |  | 7.3878 |  |  |  | 2.648 |  |  |  | 327.8532 |  |  |  | 11.29 |  |  |  | 0.29 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 120.3611 |  |  |  | 8.0501 |  |  |  | 2.8853 |  |  |  | 347.2807 |  |  |  | 11.78 |  |  |  | 0.78 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 118.1568 |  |  |  | 7.9073 |  |  |  | 2.8342 |  |  |  | 334.8758 |  |  |  | 11.29 |  |  |  | 0.29 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 119.4981 |  |  |  | 8.3658 |  |  |  | 2.9985 |  |  |  | 358.3145 |  |  |  | 11.16 |  |  |  | 0.16 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 120.024 |  |  |  | 7.9174 |  |  |  | 2.8378 |  |  |  | 340.6025 |  |  |  | 11.35 |  |  |  | 0.35 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 121.041 |  |  |  | 7.913 |  |  |  | 2.8362 |  |  |  | 343.2949 |  |  |  | 11.9 |  |  |  | 0.9 |  |  |  | 4.01 |  |  |  | 1.01
|  |  |  | 119.3317 |  |  |  | 7.8045 |  |  |  | 2.7973 |  |  |  | 333.8094 |  |  |  | 11.15 |  |  |  | 0.15 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 114.6351 |  |  |  | 8.333 |  |  |  | 2.9867 |  |  |  | 342.3848 |  |  |  | 11.4 |  |  |  | 0.4 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 125.2871 |  |  |  | 7.4687 |  |  |  | 2.6769 |  |  |  | 335.3863 |  |  |  | 11.48 |  |  |  | 0.48 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 118.087 |  |  |  | 8.6401 |  |  |  | 3.0968 |  |  |  | 365.6941 |  |  |  | 11.62 |  |  |  | 0.62 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 120.5303 |  |  |  | 8.0067 |  |  |  | 2.8698 |  |  |  | 345.8953 |  |  |  | 11.21 |  |  |  | 0.21 |  |  |  | 3.95 |  |  |  | 0.95
|  |  |  | 117.2333 |  |  |  | 8.9428 |  |  |  | 3.2053 |  |  |  | 375.7702 |  |  |  | 11.42 |  |  |  | 0.42 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 118.624 |  |  |  | 7.7566 |  |  |  | 2.7802 |  |  |  | 329.7928 |  |  |  | 11.46 |  |  |  | 0.46 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 117.4398 |  |  |  | 7.8635 |  |  |  | 2.8184 |  |  |  | 330.9978 |  |  |  | 11.72 |  |  |  | 0.72 |  |  |  | 3.73 |  |  |  | 0.73
|  |  |  | 121.4329 |  |  |  | 7.6406 |  |  |  | 2.7386 |  |  |  | 332.5527 |  |  |  | 10.98 |  |  |  | -0.02 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 121.4329 |  |  |  | 8.076 |  |  |  | 2.8946 |  |  |  | 351.5011 |  |  |  | 11.47 |  |  |  | 0.47 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 119.4743 |  |  |  | 7.3943 |  |  |  | 2.6503 |  |  |  | 316.6423 |  |  |  | 11.35 |  |  |  | 0.35 |  |  |  | 3.44 |  |  |  | 0.44
|  |  |  | 119.6172 |  |  |  | 7.6132 |  |  |  | 2.7287 |  |  |  | 326.4052 |  |  |  | 11.67 |  |  |  | 0.67 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 118.3665 |  |  |  | 8.3636 |  |  |  | 2.9977 |  |  |  | 354.8289 |  |  |  | 10.73 |  |  |  | -0.27 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.1632 |  |  |  | 7.6453 |  |  |  | 2.7402 |  |  |  | 332.0167 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 120.1682 |  |  |  | 7.8961 |  |  |  | 2.8301 |  |  |  | 340.092 |  |  |  | 11.61 |  |  |  | 0.61 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 118.6005 |  |  |  | 8.4424 |  |  |  | 3.0259 |  |  |  | 358.8783 |  |  |  | 11.72 |  |  |  | 0.72 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 120.6273 |  |  |  | 8.0149 |  |  |  | 2.8727 |  |  |  | 346.5296 |  |  |  | 10.98 |  |  |  | -0.02 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 115.119 |  |  |  | 7.4866 |  |  |  | 2.6834 |  |  |  | 308.9056 |  |  |  | 11.19 |  |  |  | 0.19 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 120.0961 |  |  |  | 6.9881 |  |  |  | 2.5047 |  |  |  | 300.8028 |  |  |  | 11.62 |  |  |  | 0.62 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 120.8459 |  |  |  | 7.1388 |  |  |  | 2.5587 |  |  |  | 309.2089 |  |  |  | 11.32 |  |  |  | 0.32 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 117.5779 |  |  |  | 8.158 |  |  |  | 2.924 |  |  |  | 343.8004 |  |  |  | 10.93 |  |  |  | -0.07 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 121.1632 |  |  |  | 8.1923 |  |  |  | 2.9363 |  |  |  | 355.7714 |  |  |  | 11.68 |  |  |  | 0.68 |  |  |  | 3.99 |  |  |  | 0.99
|  |  |  | 118.624 |  |  |  | 7.1924 |  |  |  | 2.5779 |  |  |  | 305.8016 |  |  |  | 11.92 |  |  |  | 0.92 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 118.1335 |  |  |  | 7.1643 |  |  |  | 2.5678 |  |  |  | 303.3479 |  |  |  | 11.44 |  |  |  | 0.44 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 120.8703 |  |  |  | 7.7269 |  |  |  | 2.7695 |  |  |  | 334.751 |  |  |  | 11.17 |  |  |  | 0.17 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 123.6094 |  |  |  | 7.7913 |  |  |  | 2.7926 |  |  |  | 345.1885 |  |  |  | 11.4 |  |  |  | 0.4 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 117.9709 |  |  |  | 8.4506 |  |  |  | 3.0289 |  |  |  | 357.3206 |  |  |  | 11.78 |  |  |  | 0.78 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.087 |  |  |  | 7.6892 |  |  |  | 2.756 |  |  |  | 325.4447 |  |  |  | 11.42 |  |  |  | 0.42 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 120.2887 |  |  |  | 7.8048 |  |  |  | 2.7974 |  |  |  | 336.4993 |  |  |  | 11.14 |  |  |  | 0.14 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 117.4858 |  |  |  | 7.4102 |  |  |  | 2.656 |  |  |  | 312.0418 |  |  |  | 11.39 |  |  |  | 0.39 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 120.5546 |  |  |  | 7.428 |  |  |  | 2.6624 |  |  |  | 320.9598 |  |  |  | 11.52 |  |  |  | 0.52 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 121.6052 |  |  |  | 7.4426 |  |  |  | 2.6676 |  |  |  | 324.3929 |  |  |  | 11.24 |  |  |  | 0.24 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 119.8801 |  |  |  | 7.5344 |  |  |  | 2.7005 |  |  |  | 323.7382 |  |  |  | 11.08 |  |  |  | 0.08 |  |  |  | 3.42 |  |  |  | 0.42
|  |  |  | 117.7856 |  |  |  | 8.3042 |  |  |  | 2.9764 |  |  |  | 350.5786 |  |  |  | 11.85 |  |  |  | 0.85 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 120.8459 |  |  |  | 7.5785 |  |  |  | 2.7163 |  |  |  | 328.2546 |  |  |  | 11.77 |  |  |  | 0.77 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 116.7315 |  |  |  | 8.0349 |  |  |  | 2.8799 |  |  |  | 336.1747 |  |  |  | 10.72 |  |  |  | -0.28 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 122.549 |  |  |  | 7.8908 |  |  |  | 2.8282 |  |  |  | 346.5987 |  |  |  | 11.19 |  |  |  | 0.19 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 118.8119 |  |  |  | 7.7081 |  |  |  | 2.7628 |  |  |  | 328.2492 |  |  |  | 11.57 |  |  |  | 0.57 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 117.0732 |  |  |  | 7.2012 |  |  |  | 2.5811 |  |  |  | 302.1752 |  |  |  | 11.69 |  |  |  | 0.69 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 123.8902 |  |  |  | 8.1808 |  |  |  | 2.9322 |  |  |  | 363.2699 |  |  |  | 11.06 |  |  |  | 0.06 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 118.4834 |  |  |  | 7.8553 |  |  |  | 2.8155 |  |  |  | 333.5934 |  |  |  | 11.3 |  |  |  | 0.3 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 120.1442 |  |  |  | 8.0535 |  |  |  | 2.8865 |  |  |  | 346.8017 |  |  |  | 11.59 |  |  |  | 0.59 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 119.4743 |  |  |  | 8.2695 |  |  |  | 2.964 |  |  |  | 354.1185 |  |  |  | 11.94 |  |  |  | 0.94 |  |  |  | 3.97 |  |  |  | 0.97
|  |  |  | 120.1442 |  |  |  | 8.0189 |  |  |  | 2.8741 |  |  |  | 345.3112 |  |  |  | 10.99 |  |  |  | -0.01 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 116.3467 |  |  |  | 8.0409 |  |  |  | 2.882 |  |  |  | 335.3163 |  |  |  | 11.65 |  |  |  | 0.65 |  |  |  | 3.87 |  |  |  | 0.87
|  |  |  | 121.1632 |  |  |  | 6.6547 |  |  |  | 2.3852 |  |  |  | 288.9992 |  |  |  | 11.6 |  |  |  | 0.6 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 117.6701 |  |  |  | 7.4759 |  |  |  | 2.6795 |  |  |  | 315.3017 |  |  |  | 11.44 |  |  |  | 0.44 |  |  |  | 3.39 |  |  |  | 0.39
|  |  |  | 120.4336 |  |  |  | 7.0246 |  |  |  | 2.5178 |  |  |  | 303.2257 |  |  |  | 10.9 |  |  |  | -0.1 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 117.1417 |  |  |  | 7.3284 |  |  |  | 2.6267 |  |  |  | 307.693 |  |  |  | 11.42 |  |  |  | 0.42 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 122.1498 |  |  |  | 6.7069 |  |  |  | 2.4039 |  |  |  | 293.6372 |  |  |  | 11.52 |  |  |  | 0.52 |  |  |  | 3.68 |  |  |  | 0.68
|  |  |  | 120.024 |  |  |  | 8.7415 |  |  |  | 3.1332 |  |  |  | 376.0537 |  |  |  | 11.37 |  |  |  | 0.37 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 121.2366 |  |  |  | 8.3875 |  |  |  | 3.0063 |  |  |  | 364.4713 |  |  |  | 11.59 |  |  |  | 0.59 |  |  |  | 4.03 |  |  |  | 1.03
|  |  |  | 119.2132 |  |  |  | 7.4842 |  |  |  | 2.6825 |  |  |  | 319.7919 |  |  |  | 11.57 |  |  |  | 0.57 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 120.024 |  |  |  | 7.3103 |  |  |  | 2.6202 |  |  |  | 314.4827 |  |  |  | 11.97 |  |  |  | 0.97 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 118.1335 |  |  |  | 8.075 |  |  |  | 2.8943 |  |  |  | 341.9081 |  |  |  | 11.21 |  |  |  | 0.21 |  |  |  | 3.84 |  |  |  | 0.84
|  |  |  | 121.0898 |  |  |  | 7.5686 |  |  |  | 2.7128 |  |  |  | 328.4882 |  |  |  | 10.97 |  |  |  | -0.03 |  |  |  | 3.62 |  |  |  | 0.62
|  |  |  | 116.7315 |  |  |  | 7.7382 |  |  |  | 2.7736 |  |  |  | 323.7607 |  |  |  | 11.46 |  |  |  | 0.46 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.1632 |  |  |  | 7.3107 |  |  |  | 2.6203 |  |  |  | 317.4887 |  |  |  | 11.46 |  |  |  | 0.46 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 117.6471 |  |  |  | 8.5277 |  |  |  | 3.0565 |  |  |  | 359.591 |  |  |  | 11.47 |  |  |  | 0.47 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 118.6005 |  |  |  | 7.1527 |  |  |  | 2.5637 |  |  |  | 304.0558 |  |  |  | 11.46 |  |  |  | 0.46 |  |  |  | 3.97 |  |  |  | 0.97
|  |  |  | 119.4268 |  |  |  | 7.7461 |  |  |  | 2.7764 |  |  |  | 331.5735 |  |  |  | 11.63 |  |  |  | 0.63 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 122.0256 |  |  |  | 7.4396 |  |  |  | 2.6665 |  |  |  | 325.384 |  |  |  | 11.8 |  |  |  | 0.8 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 120.8703 |  |  |  | 7.5129 |  |  |  | 2.6928 |  |  |  | 325.4799 |  |  |  | 10.86 |  |  |  | -0.14 |  |  |  | 3.78 |  |  |  | 0.78
|  |  |  | 114.5257 |  |  |  | 8.2116 |  |  |  | 2.9432 |  |  |  | 337.0732 |  |  |  | 11.15 |  |  |  | 0.15 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 121.4575 |  |  |  | 7.8245 |  |  |  | 2.8045 |  |  |  | 340.6252 |  |  |  | 11.62 |  |  |  | 0.62 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 120.7729 |  |  |  | 7.1339 |  |  |  | 2.557 |  |  |  | 308.8128 |  |  |  | 11.55 |  |  |  | 0.55 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 121.5559 |  |  |  | 7.29 |  |  |  | 2.6129 |  |  |  | 317.6154 |  |  |  | 10.82 |  |  |  | -0.18 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 120.072 |  |  |  | 7.4257 |  |  |  | 2.6616 |  |  |  | 319.5788 |  |  |  | 11.05 |  |  |  | 0.05 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 117.9477 |  |  |  | 8.1852 |  |  |  | 2.9338 |  |  |  | 346.0294 |  |  |  | 11.9 |  |  |  | 0.9 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 116.7542 |  |  |  | 7.1339 |  |  |  | 2.557 |  |  |  | 298.5366 |  |  |  | 11.48 |  |  |  | 0.48 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 121.3592 |  |  |  | 7.3877 |  |  |  | 2.6479 |  |  |  | 321.3511 |  |  |  | 10.81 |  |  |  | -0.19 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 118.0638 |  |  |  | 7.3691 |  |  |  | 2.6413 |  |  |  | 311.8378 |  |  |  | 11.54 |  |  |  | 0.54 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 121.8274 |  |  |  | 7.137 |  |  |  | 2.5581 |  |  |  | 311.6406 |  |  |  | 11.52 |  |  |  | 0.52 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 118.7648 |  |  |  | 7.4758 |  |  |  | 2.6795 |  |  |  | 318.2318 |  |  |  | 11.53 |  |  |  | 0.53 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 119.7844 |  |  |  | 7.4012 |  |  |  | 2.6528 |  |  |  | 317.7595 |  |  |  | 10.93 |  |  |  | -0.07 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 119.0712 |  |  |  | 7.0079 |  |  |  | 2.5118 |  |  |  | 299.0809 |  |  |  | 11.85 |  |  |  | 0.85 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 117.3709 |  |  |  | 7.4558 |  |  |  | 2.6723 |  |  |  | 313.6529 |  |  |  | 12.28 |  |  |  | 1.28 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 122.2992 |  |  |  | 7.0295 |  |  |  | 2.5195 |  |  |  | 308.1367 |  |  |  | 11.07 |  |  |  | 0.07 |  |  |  | 3.8 |  |  |  | 0.8
|  |  |  | 118.4834 |  |  |  | 7.3497 |  |  |  | 2.6343 |  |  |  | 312.1223 |  |  |  | 11.09 |  |  |  | 0.09 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 119.024 |  |  |  | 7.1486 |  |  |  | 2.5622 |  |  |  | 304.964 |  |  |  | 11.22 |  |  |  | 0.22 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 119.2606 |  |  |  | 6.992 |  |  |  | 2.5061 |  |  |  | 298.8791 |  |  |  | 11.58 |  |  |  | 0.58 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 121.2366 |  |  |  | 7.1999 |  |  |  | 2.5806 |  |  |  | 312.8644 |  |  |  | 11.15 |  |  |  | 0.15 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 119.2606 |  |  |  | 6.5763 |  |  |  | 2.3571 |  |  |  | 281.1105 |  |  |  | 10.92 |  |  |  | -0.08 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 122.1996 |  |  |  | 6.7745 |  |  |  | 2.4281 |  |  |  | 296.7189 |  |  |  | 11.42 |  |  |  | 0.42 |  |  |  | 3.79 |  |  |  | 0.79
|  |  |  | 116.5728 |  |  |  | 7.5996 |  |  |  | 2.7239 |  |  |  | 317.531 |  |  |  | 12.08 |  |  |  | 1.08 |  |  |  | 4 |  |  |  | 1
|  |  |  | 123.406 |  |  |  | 6.9779 |  |  |  | 2.5011 |  |  |  | 308.6448 |  |  |  | 11.34 |  |  |  | 0.34 |  |  |  | 3.96 |  |  |  | 0.96
|  |  |  | 117.6471 |  |  |  | 8.1273 |  |  |  | 2.913 |  |  |  | 342.7055 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.3432 |  |  |  | 8.1199 |  |  |  | 2.9104 |  |  |  | 344.4218 |  |  |  | 11.92 |  |  |  | 0.92 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 121.976 |  |  |  | 7.9111 |  |  |  | 2.8355 |  |  |  | 345.8643 |  |  |  | 11.76 |  |  |  | 0.76 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 118.18 |  |  |  | 8.169 |  |  |  | 2.9279 |  |  |  | 346.025 |  |  |  | 11.69 |  |  |  | 0.69 |  |  |  | 4.09 |  |  |  | 1.09
|  |  |  | 118.0405 |  |  |  | 7.5192 |  |  |  | 2.695 |  |  |  | 318.1239 |  |  |  | 11.35 |  |  |  | 0.35 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 118.3432 |  |  |  | 7.3241 |  |  |  | 2.6251 |  |  |  | 310.6676 |  |  |  | 11.72 |  |  |  | 0.72 |  |  |  | 3.58 |  |  |  | 0.58
|  |  |  | 118.624 |  |  |  | 7.5018 |  |  |  | 2.6888 |  |  |  | 318.9569 |  |  |  | 11.75 |  |  |  | 0.75 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.7285 |  |  |  | 7.1927 |  |  |  | 2.578 |  |  |  | 313.8216 |  |  |  | 11 |  |  |  | 0 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 119.9041 |  |  |  | 8.082 |  |  |  | 2.8968 |  |  |  | 347.3351 |  |  |  | 11.71 |  |  |  | 0.71 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 118.0173 |  |  |  | 8.118 |  |  |  | 2.9097 |  |  |  | 343.3911 |  |  |  | 11.62 |  |  |  | 0.62 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 125.0782 |  |  |  | 8.5213 |  |  |  | 3.0542 |  |  |  | 382.0191 |  |  |  | 11.84 |  |  |  | 0.84 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 114.6789 |  |  |  | 7.8209 |  |  |  | 2.8032 |  |  |  | 321.4683 |  |  |  | 10.82 |  |  |  | -0.18 |  |  |  | 3.93 |  |  |  | 0.93
|  |  |  | 120.2887 |  |  |  | 7.3067 |  |  |  | 2.6189 |  |  |  | 315.0217 |  |  |  | 11.28 |  |  |  | 0.28 |  |  |  | 3.74 |  |  |  | 0.74
|  |  |  | 117.9014 |  |  |  | 7.3032 |  |  |  | 2.6176 |  |  |  | 308.6211 |  |  |  | 12.01 |  |  |  | 1.01 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 119.5695 |  |  |  | 7.4399 |  |  |  | 2.6666 |  |  |  | 318.8493 |  |  |  | 11.28 |  |  |  | 0.28 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 124.4297 |  |  |  | 7.6617 |  |  |  | 2.7461 |  |  |  | 341.7011 |  |  |  | 11.14 |  |  |  | 0.14 |  |  |  | 3.94 |  |  |  | 0.94
|  |  |  | 115.0527 |  |  |  | 7.9814 |  |  |  | 2.8607 |  |  |  | 329.1323 |  |  |  | 11.52 |  |  |  | 0.52 |  |  |  | 3.9 |  |  |  | 0.9
|  |  |  | 124.1722 |  |  |  | 7.5607 |  |  |  | 2.7099 |  |  |  | 336.4985 |  |  |  | 11.7 |  |  |  | 0.7 |  |  |  | 3.82 |  |  |  | 0.82
|  |  |  | 116.777 |  |  |  | 7.6504 |  |  |  | 2.7421 |  |  |  | 320.2121 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.48 |  |  |  | 0.48
|  |  |  | 119.403 |  |  |  | 7.1632 |  |  |  | 2.5674 |  |  |  | 306.5601 |  |  |  | 11.02 |  |  |  | 0.02 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 118.4834 |  |  |  | 7.4161 |  |  |  | 2.6581 |  |  |  | 314.9415 |  |  |  | 11.23 |  |  |  | 0.23 |  |  |  | 3.75 |  |  |  | 0.75
|  |  |  | 120.5788 |  |  |  | 7.491 |  |  |  | 2.6849 |  |  |  | 323.7466 |  |  |  | 11.83 |  |  |  | 0.83 |  |  |  | 3.78 |  |  |  | 0.78
|  |  |  | 119.024 |  |  |  | 7.5952 |  |  |  | 2.7223 |  |  |  | 324.0186 |  |  |  | 10.94 |  |  |  | -0.06 |  |  |  | 3.91 |  |  |  | 0.91
|  |  |  | 122.0256 |  |  |  | 7.6336 |  |  |  | 2.7361 |  |  |  | 333.8703 |  |  |  | 11.02 |  |  |  | 0.02 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 118.7178 |  |  |  | 8.2537 |  |  |  | 2.9583 |  |  |  | 351.203 |  |  |  | 11.39 |  |  |  | 0.39 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 120.8703 |  |  |  | 7.8715 |  |  |  | 2.8213 |  |  |  | 341.0152 |  |  |  | 11.63 |  |  |  | 0.63 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 117.7625 |  |  |  | 7.7485 |  |  |  | 2.7773 |  |  |  | 327.056 |  |  |  | 11.29 |  |  |  | 0.29 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 120.4094 |  |  |  | 7.377 |  |  |  | 2.6441 |  |  |  | 318.3736 |  |  |  | 11.13 |  |  |  | 0.13 |  |  |  | 3.73 |  |  |  | 0.73
|  |  |  | 118.6944 |  |  |  | 7.4128 |  |  |  | 2.6569 |  |  |  | 315.361 |  |  |  | 11.54 |  |  |  | 0.54 |  |  |  | 3.81 |  |  |  | 0.81
|  |  |  | 121.4083 |  |  |  | 7.7893 |  |  |  | 2.7919 |  |  |  | 338.956 |  |  |  | 11.74 |  |  |  | 0.74 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 117.5549 |  |  |  | 7.8265 |  |  |  | 2.8052 |  |  |  | 329.7652 |  |  |  | 10.88 |  |  |  | -0.12 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 120.4094 |  |  |  | 7.7549 |  |  |  | 2.7795 |  |  |  | 334.6815 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.85 |  |  |  | 0.85
|  |  |  | 116.3467 |  |  |  | 7.9187 |  |  |  | 2.8382 |  |  |  | 330.2207 |  |  |  | 11.85 |  |  |  | 0.85 |  |  |  | 3.77 |  |  |  | 0.77
|  |  |  | 125.1303 |  |  |  | 7.5645 |  |  |  | 2.7113 |  |  |  | 339.2638 |  |  |  | 11.79 |  |  |  | 0.79 |  |  |  | 3.89 |  |  |  | 0.89
|  |  |  | 118.2266 |  |  |  | 8.4118 |  |  |  | 3.015 |  |  |  | 356.4511 |  |  |  | 11.55 |  |  |  | 0.55 |  |  |  | 3.98 |  |  |  | 0.98
|  |  |  | 117.1189 |  |  |  | 7.6455 |  |  |  | 2.7403 |  |  |  | 320.9417 |  |  |  | 11.57 |  |  |  | 0.57 |  |  |  | 3.97 |  |  |  | 0.97
|  |  |  | 121.2121 |  |  |  | 7.3389 |  |  |  | 2.6304 |  |  |  | 318.8399 |  |  |  | 11.6 |  |  |  | 0.6 |  |  |  | 3.83 |  |  |  | 0.83
|  |  |  | 118.8119 |  |  |  | 8.4138 |  |  |  | 3.0157 |  |  |  | 358.2999 |  |  |  | 11.64 |  |  |  | 0.64 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 119.7844 |  |  |  | 7.8354 |  |  |  | 2.8084 |  |  |  | 336.3999 |  |  |  | 11.63 |  |  |  | 0.63 |  |  |  | 4.07 |  |  |  | 1.07
|  |  |  | 118.2033 |  |  |  | 8.4507 |  |  |  | 3.0289 |  |  |  | 358.0297 |  |  |  | 11.58 |  |  |  | 0.58 |  |  |  | 3.88 |  |  |  | 0.88
|  |  |  | 118.3432 |  |  |  | 7.2129 |  |  |  | 2.5853 |  |  |  | 305.9474 |  |  |  | 12.16 |  |  |  | 1.16 |  |  |  | 3.92 |  |  |  | 0.92
|  |  |  | 118.6944 |  |  |  | 7.2207 |  |  |  | 2.5881 |  |  |  | 307.19 |  |  |  | 11.1 |  |  |  | 0.1 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 123.1274 |  |  |  | 7.2891 |  |  |  | 2.6126 |  |  |  | 321.6817 |  |  |  | 10.97 |  |  |  | -0.03 |  |  |  | 3.76 |  |  |  | 0.76
|  |  |  | 120.2646 |  |  |  | 8.0607 |  |  |  | 2.8891 |  |  |  | 347.4604 |  |  |  | 11.45 |  |  |  | 0.45 |  |  |  | 3.72 |  |  |  | 0.72
|  |  |  | 119.2606 |  |  |  | 7.5073 |  |  |  | 2.6908 |  |  |  | 320.9041 |  |  |  | 11.63 |  |  |  | 0.63 |  |  |  | 3.96 |  |  |  | 0.96
|  |  |  | 119.308 |  |  |  | 7.282 |  |  |  | 2.61 |  |  |  | 311.397 |  |  |  | 11.12 |  |  |  | 0.12 |  |  |  | 3.86 |  |  |  | 0.86
|  |  |  | 119.6649 |  |  |  | 8.3703 |  |  |  | 3.0001 |  |  |  | 359.0084 |  |  |  | 11.18 |  |  |  | 0.18 |  |  |  | 3.78 |  |  |  | 0.78
|  |  |  | 120.024 |  |  |  | 8.482 |  |  |  | 3.0401 |  |  |  | 364.8896 |  |  |  | 12.17 |  |  |  | 1.17 |  |  |  | 4.02 |  |  |  | 1.02
|  |  |  | 118.0173 |  |  |  | 8.0618 |  |  |  | 2.8895 |  |  |  | 341.0148 |  |  |  | 11.94 |  |  |  | 0.94 |  |  |  | 3.96 |  |  |  | 0.96
|  |  |  | 122.549 |  |  |  | 6.5802 |  |  |  | 2.3585 |  |  |  | 289.0304 |  |  |  | 11.07 |  |  |  | 0.07 |  |  |  | 3.73 |  |  |  | 0.73
[8 empty] | 7.7355 |  |  |  | 2.7726 | [7 empty] | 11.24 |  |  |  | 0.24 |  |  |  | 3.8 |  |  |  | 0.8

## Poincare Data
MV 1 | [7 empty] | MV2 | [7 empty] | MV3 | [7 empty] | MV4 | [7 empty] | MV5 | [8 empty] | Data Summary
Carestation |  |  |  | FALCON |  |  |  | Carestation |  |  |  | FALCON |  |  |  | Carestation |  |  |  | FALCON |  |  |  | Carestation |  |  |  | FALCON |  |  |  | Carestation |  |  |  | FALCON |  |  |  |  | SD1 PIP |  | SD1 PEEP |  | SD2 PIP |  | SD2 PEEP
PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 | PIP n | PIP n+1 | PEEP n | PEEP n+1 |  | Carestation | FALCON | Carestation | FALCON | Carestation | FALCON | Carestation | FALCON
11.39 | 11.39 | 2.99 | 2.96 | 11.2 | 11.19 | 2.98 | 3.35 | 11.5 | 11.58 | 3.04 | 3.03 | 12.1 | 12.06 | 3.01 | 3.08 | 13.27 | 13.03 | 3.1 | 3.08 | 12 | 12.36 | 3.04 | 3.13 | 13.69 | 13.68 | 5.19 | 5.22 | 14.59 | 14.53 | 5.37 | 5.39 | 12.77 | 12.75 | 3.03 | 3.05 | 11.85 | 11.76 | 3.46 | 3.42 | MV1 | 0.03171 | 0.1022 | 0.03505 | 0.1071 | 0.2417 | 0.9719 | 0.2565 | 0.293
11.39 | 11.36 | 2.96 | 2.94 | 11.19 | 11.1 | 3.35 | 3.06 | 11.58 | 11.62 | 3.03 | 3 | 12.06 | 12.08 | 3.08 | 3.16 | 13.03 | 13.31 | 3.08 | 3.1 | 12.36 | 12.43 | 3.13 | 3.12 | 13.68 | 13.73 | 5.22 | 5.23 | 14.53 | 14.31 | 5.39 | 5.39 | 12.75 | 12.8 | 3.05 | 3.04 | 11.76 | 11.89 | 3.42 | 3.4 | MV2 | 0.06082 | 0.4348 | 0.03374 | 0.5433 | 0.3465 | 0.6028 | 0.3782 | 0.6636
11.36 | 11.43 | 2.94 | 2.95 | 11.1 | 11.25 | 3.06 | 3.28 | 11.62 | 11.54 | 3 | 2.99 | 12.08 | 12.13 | 3.16 | 3.15 | 13.31 | 13.32 | 3.1 | 3.13 | 12.43 | 12.02 | 3.12 | 3.15 | 13.73 | 13.73 | 5.23 | 5.23 | 14.31 | 14.74 | 5.39 | 5.19 | 12.8 | 12.73 | 3.04 | 3.06 | 11.89 | 11.76 | 3.4 | 3.42 | MV3 | 0.0542 | 0.2367 | 0.03584 | 0.1377 | 0.286 | 0.6284 | 0.4106 | 0.4184
11.43 | 11.4 | 2.95 | 2.97 | 11.25 | 11.2 | 3.28 | 3.13 | 11.54 | 11.62 | 2.99 | 3 | 12.13 | 12.1 | 3.15 | 3.03 | 13.32 | 13.26 | 3.13 | 3.15 | 12.02 | 12.14 | 3.15 | 3.13 | 13.73 | 13.72 | 5.23 | 5.19 | 14.74 | 14.55 | 5.19 | 5.28 | 12.73 | 12.77 | 3.06 | 3.03 | 11.76 | 11.87 | 3.42 | 3.49 | MV4 | 0.07983 | 0.2254 | 0.04336 | 0.1292 | 0.2489 | 0.6553 | 0.508 | 1.007
11.4 | 11.45 | 2.97 | 3 | 11.2 | 11.26 | 3.13 | 3.2 | 11.62 | 11.55 | 3 | 3.06 | 12.1 | 12.17 | 3.03 | 3.04 | 13.26 | 13.26 | 3.15 | 3.18 | 12.14 | 11.84 | 3.13 | 3.1 | 13.72 | 13.69 | 5.19 | 5.21 | 14.55 | 14.64 | 5.28 | 5.25 | 12.77 | 12.73 | 3.03 | 3.06 | 11.87 | 12.03 | 3.49 | 3.37 | MV5 | 0.03411 | 0.5547 | 0.0247 | 0.2456 | 0.2084 | 0.748 | 0.3032 | 0.5399
11.45 | 11.44 | 3 | 2.96 | 11.26 | 11.17 | 3.2 | 3.06 | 11.55 | 11.56 | 3.06 | 3.09 | 12.17 | 12.11 | 3.04 | 3.08 | 13.26 | 13.28 | 3.18 | 3.19 | 11.84 | 12.21 | 3.1 | 3.05 | 13.69 | 13.69 | 5.21 | 5.19 | 14.64 | 15.02 | 5.25 | 5.25 | 12.73 | 12.67 | 3.06 | 3.05 | 12.03 | 11.96 | 3.37 | 3.42
11.44 | 11.36 | 2.96 | 2.96 | 11.17 | 11.21 | 3.06 | 3.13 | 11.56 | 11.57 | 3.09 | 3.06 | 12.11 | 12.2 | 3.08 | 3.09 | 13.28 | 13.23 | 3.19 | 3.09 | 12.21 | 11.78 | 3.05 | 3.05 | 13.69 | 13.67 | 5.19 | 5.18 | 15.02 | 14.52 | 5.25 | 5.29 | 12.67 | 12.6 | 3.05 | 3.09 | 11.96 | 11.82 | 3.42 | 3.42
11.36 | 11.36 | 2.96 | 2.97 | 11.21 | 11.32 | 3.13 | 3.08 | 11.57 | 11.55 | 3.06 | 3.06 | 12.2 | 12.12 | 3.09 | 3.11 | 13.23 | 13.1 | 3.09 | 3.12 | 11.78 | 12.21 | 3.05 | 3.06 | 13.67 | 13.67 | 5.18 | 5.17 | 14.52 | 14.27 | 5.29 | 5.35 | 12.6 | 12.59 | 3.09 | 3.1 | 11.82 | 12.02 | 3.42 | 3.41
11.36 | 11.34 | 2.97 | 2.93 | 11.32 | 11.27 | 3.08 | 2.99 | 11.55 | 11.51 | 3.06 | 3.02 | 12.12 | 12.11 | 3.11 | 3.13 | 13.1 | 13.05 | 3.12 | 3.14 | 12.21 | 12.05 | 3.06 | 3.08 | 13.67 | 13.64 | 5.17 | 5.2 | 14.27 | 13.89 | 5.35 | 5.19 | 12.59 | 12.63 | 3.1 | 3.12 | 12.02 | 11.92 | 3.41 | 3.36
11.34 | 11.3 | 2.93 | 2.94 | 11.27 | 11.25 | 2.99 | 3.09 | 11.51 | 11.45 | 3.02 | 2.81 | 12.11 | 12.03 | 3.13 | 3.19 | 13.05 | 13.08 | 3.14 | 3.19 | 12.05 | 12.06 | 3.08 | 3.21 | 13.64 | 13.6 | 5.2 | 5.18 | 13.89 | 13.78 | 5.19 | 5.22 | 12.63 | 12.68 | 3.12 | 3.12 | 11.92 | 12.06 | 3.36 | 3.44
11.3 | 11.35 | 2.94 | 2.97 | 11.25 | 11.23 | 3.09 | 3.11 | 11.45 | 11.45 | 2.81 | 3.02 | 12.03 | 12.17 | 3.19 | 3.11 | 13.08 | 13.14 | 3.19 | 3.22 | 12.06 | 11.92 | 3.21 | 3.06 | 13.6 | 13.64 | 5.18 | 5.18 | 13.78 | 14.15 | 5.22 | 5.25 | 12.68 | 12.74 | 3.12 | 3.13 | 12.06 | 11.85 | 3.44 | 3.35
11.35 | 11.36 | 2.97 | 2.94 | 11.23 | 11.27 | 3.11 | 3.22 | 11.45 | 11.46 | 3.02 | 3.01 | 12.17 | 12.14 | 3.11 | 3.05 | 13.14 | 13.16 | 3.22 | 3.15 | 11.92 | 11.57 | 3.06 | 3.07 | 13.64 | 13.67 | 5.18 | 5.2 | 14.15 | 14.48 | 5.25 | 5.24 | 12.74 | 12.77 | 3.13 | 3.1 | 11.85 | 11.62 | 3.35 | 3.3
11.36 | 11.37 | 2.94 | 2.96 | 11.27 | 11.31 | 3.22 | 3.03 | 11.46 | 11.48 | 3.01 | 3.01 | 12.14 | 12.09 | 3.05 | 3.04 | 13.16 | 13.24 | 3.15 | 3.17 | 11.57 | 11.85 | 3.07 | 3.07 | 13.67 | 13.69 | 5.2 | 5.22 | 14.48 | 14.52 | 5.24 | 5.26 | 12.77 | 12.81 | 3.1 | 3.1 | 11.62 | 11.85 | 3.3 | 3.54
11.37 | 11.4 | 2.96 | 2.96 | 11.31 | 11.35 | 3.03 | 3.06 | 11.48 | 11.54 | 3.01 | 3.03 | 12.09 | 12.08 | 3.04 | 3 | 13.24 | 13.17 | 3.17 | 3.21 | 11.85 | 11.12 | 3.07 | 2.81 | 13.69 | 13.72 | 5.22 | 5.23 | 14.52 | 14.26 | 5.26 | 5.14 | 12.81 | 12.8 | 3.1 | 3.11 | 11.85 | 11.85 | 3.54 | 3.41
11.4 | 11.36 | 2.96 | 2.94 | 11.35 | 11.26 | 3.06 | 3.03 | 11.54 | 11.5 | 3.03 | 3.09 | 12.08 | 12.1 | 3 | 3.07 | 13.17 | 13.23 | 3.21 | 3.16 | 11.12 | 10.05 | 2.81 | 1.24 | 13.72 | 13.7 | 5.23 | 5.21 | 14.26 | 13.95 | 5.14 | 4.73 | 12.8 | 12.82 | 3.11 | 3.09 | 11.85 | 11.65 | 3.41 | 3.35
11.36 | 11.39 | 2.94 | 2.97 | 11.26 | 11.39 | 3.03 | 3.23 | 11.5 | 11.55 | 3.09 | 3.04 | 12.1 | 12.15 | 3.07 | 3.12 | 13.23 | 12.22 | 3.16 | 2.36 | 10.05 | 11.61 | 1.24 | 2.86 | 13.7 | 13.69 | 5.21 | 5.18 | 13.95 | 14.17 | 4.73 | 4.6 | 12.82 | 12.8 | 3.09 | 3.04 | 11.65 | 11.86 | 3.35 | 3.42
11.39 | 11.4 | 2.97 | 2.96 | 11.39 | 11.38 | 3.23 | 3.14 | 11.55 | 11.5 | 3.04 | 3.05 | 12.15 | 12.12 | 3.12 | 3.1 | 12.22 | 13.27 | 2.36 | 2.37 | 11.61 | 11.7 | 2.86 | 3.07 | 13.69 | 13.67 | 5.18 | 5.17 | 14.17 | 14.02 | 4.6 | 4.65 | 12.8 | 12.76 | 3.04 | 3.03 | 11.86 | 11.91 | 3.42 | 3.51
11.4 | 11.42 | 2.96 | 2.95 | 11.38 | 11.2 | 3.14 | 3.35 | 11.5 | 11.54 | 3.05 | 3.01 | 12.12 | 12.2 | 3.1 | 3.06 | 13.27 | 13.16 | 2.37 | 2.26 | 11.7 | 11.66 | 3.07 | 3.14 | 13.67 | 13.68 | 5.17 | 5.16 | 14.02 | 13.94 | 4.65 | 4.65 | 12.76 | 12.68 | 3.03 | 3 | 11.91 | 11.79 | 3.51 | 3.34
11.42 | 11.38 | 2.95 | 2.97 | 11.2 | 11.19 | 3.35 | 3.05 | 11.54 | 11.52 | 3.01 | 3.01 | 12.2 | 12.16 | 3.06 | 3.04 | 13.16 | 13.08 | 2.26 | 2.28 | 11.66 | 11.48 | 3.14 | 3.07 | 13.68 | 13.71 | 5.16 | 5.19 | 13.94 | 14.33 | 4.65 | 4.62 | 12.68 | 12.74 | 3 | 3.02 | 11.79 | 11.89 | 3.34 | 3.42
11.38 | 11.38 | 2.97 | 2.95 | 11.19 | 11.19 | 3.05 | 3.1 | 11.52 | 11.56 | 3.01 | 2.96 | 12.16 | 12.06 | 3.04 | 3.04 | 13.08 | 13.18 | 2.28 | 2.27 | 11.48 | 12.45 | 3.07 | 3.1 | 13.71 | 13.65 | 5.19 | 5.17 | 14.33 | 14.21 | 4.62 | 4.64 | 12.74 | 12.59 | 3.02 | 3.04 | 11.89 | 11.9 | 3.42 | 3.44
11.38 | 11.37 | 2.95 | 2.96 | 11.19 | 11.33 | 3.1 | 3.04 | 11.56 | 11.47 | 2.96 | 2.98 | 12.06 | 12.26 | 3.04 | 3.09 | 13.18 | 13.2 | 2.27 | 2.19 | 12.45 | 11.84 | 3.1 | 3.01 | 13.65 | 13.6 | 5.17 | 5.16 | 14.21 | 13.86 | 4.64 | 4.67 | 12.59 | 12.64 | 3.04 | 3.06 | 11.9 | 11.97 | 3.44 | 3.49
11.37 | 11.32 | 2.96 | 2.96 | 11.33 | 11.26 | 3.04 | 3.18 | 11.47 | 11.43 | 2.98 | 3.05 | 12.26 | 12.17 | 3.09 | 3.14 | 13.2 | 13.25 | 2.19 | 2.17 | 11.84 | 12.13 | 3.01 | 3.12 | 13.6 | 13.58 | 5.16 | 5.2 | 13.86 | 13.08 | 4.67 | 4.57 | 12.64 | 12.68 | 3.06 | 3.06 | 11.97 | 11.84 | 3.49 | 3.36
11.32 | 11.28 | 2.96 | 2.92 | 11.26 | 11.23 | 3.18 | 3.22 | 11.43 | 11.56 | 3.05 | 3.08 | 12.17 | 12.05 | 3.14 | 3.1 | 13.25 | 13.33 | 2.17 | 2.23 | 12.13 | 11.8 | 3.12 | 3.05 | 13.58 | 13.64 | 5.2 | 5.19 | 13.08 | 13.86 | 4.57 | 4.67 | 12.68 | 12.69 | 3.06 | 3.06 | 11.84 | 11.8 | 3.36 | 3.35
11.28 | 11.25 | 2.92 | 2.94 | 11.23 | 11.22 | 3.22 | 3.08 | 11.56 | 11.4 | 3.08 | 3.05 | 12.05 | 12.07 | 3.1 | 3.01 | 13.33 | 13.3 | 2.23 | 2.24 | 11.8 | 12.13 | 3.05 | 3.15 | 13.64 | 13.65 | 5.19 | 5.21 | 13.86 | 14.42 | 4.67 | 4.59 | 12.69 | 12.71 | 3.06 | 3.06 | 11.8 | 11.61 | 3.35 | 3.37
11.25 | 11.35 | 2.94 | 2.97 | 11.22 | 11.33 | 3.08 | 3.05 | 11.4 | 11.41 | 3.05 | 2.99 | 12.07 | 12.3 | 3.01 | 3.02 | 13.3 | 13.26 | 2.24 | 2.23 | 12.13 | 12.16 | 3.15 | 3.13 | 13.65 | 13.69 | 5.21 | 5.2 | 14.42 | 14.47 | 4.59 | 4.62 | 12.71 | 12.78 | 3.06 | 3.03 | 11.61 | 11.8 | 3.37 | 3.51
11.35 | 11.35 | 2.97 | 2.93 | 11.33 | 11.35 | 3.05 | 3.17 | 11.41 | 11.49 | 2.99 | 3.06 | 12.3 | 12.04 | 3.02 | 3 | 13.26 | 13.28 | 2.23 | 2.18 | 12.16 | 12.4 | 3.13 | 3.05 | 13.69 | 13.72 | 5.2 | 5.2 | 14.47 | 14.66 | 4.62 | 4.75 | 12.78 | 12.75 | 3.03 | 3.1 | 11.8 | 11.86 | 3.51 | 3.43
11.35 | 11.37 | 2.93 | 2.97 | 11.35 | 11.33 | 3.17 | 3.03 | 11.49 | 11.55 | 3.06 | 3.05 | 12.04 | 12.13 | 3 | 3.09 | 13.28 | 13.2 | 2.18 | 2.22 | 12.4 | 12.39 | 3.05 | 3.11 | 13.72 | 13.72 | 5.2 | 5.2 | 14.66 | 14.34 | 4.75 | 4.65 | 12.75 | 12.78 | 3.1 | 3.1 | 11.86 | 11.95 | 3.43 | 3.39
11.37 | 11.38 | 2.97 | 2.94 | 11.33 | 11.29 | 3.03 | 3.14 | 11.55 | 11.51 | 3.05 | 3.06 | 12.13 | 11.94 | 3.09 | 3.16 | 13.2 | 13.14 | 2.22 | 2.24 | 12.39 | 11.88 | 3.11 | 3.07 | 13.72 | 13.71 | 5.2 | 5.17 | 14.34 | 14.35 | 4.65 | 4.7 | 12.78 | 12.74 | 3.1 | 3.12 | 11.95 | 12.38 | 3.39 | 3.34
11.38 | 11.4 | 2.94 | 2.95 | 11.29 | 11.21 | 3.14 | 3.03 | 11.51 | 11.53 | 3.06 | 3.14 | 11.94 | 11.99 | 3.16 | 3.06 | 13.14 | 13.09 | 2.24 | 2.18 | 11.88 | 12.33 | 3.07 | 3.03 | 13.71 | 13.71 | 5.17 | 5.17 | 14.35 | 14.18 | 4.7 | 4.76 | 12.74 | 12.73 | 3.12 | 3.12 | 12.38 | 12.63 | 3.34 | 3.52
11.4 | 11.38 | 2.95 | 2.92 | 11.21 | 11.4 | 3.03 | 3.19 | 11.53 | 11.57 | 3.14 | 3.06 | 11.99 | 12.12 | 3.06 | 3.06 | 13.09 | 13.17 | 2.18 | 2.22 | 12.33 | 12.26 | 3.03 | 3.07 | 13.71 | 13.7 | 5.17 | 5.16 | 14.18 | 14.41 | 4.76 | 4.65 | 12.73 | 12.71 | 3.12 | 3.14 | 12.63 | 12.63 | 3.52 | 3.32
11.38 | 11.37 | 2.92 | 2.96 | 11.4 | 11.35 | 3.19 | 3.03 | 11.57 | 11.51 | 3.06 | 3.02 | 12.12 | 12.13 | 3.06 | 3.06 | 13.17 | 13.22 | 2.22 | 2.23 | 12.26 | 12.22 | 3.07 | 3.02 | 13.7 | 13.69 | 5.16 | 5.17 | 14.41 | 14.23 | 4.65 | 4.64 | 12.71 | 12.69 | 3.14 | 3.11 | 12.63 | 12.54 | 3.32 | 3.33
11.37 | 11.4 | 2.96 | 2.98 | 11.35 | 11.29 | 3.03 | 3.06 | 11.51 | 11.91 | 3.02 | 3.31 | 12.13 | 12.06 | 3.06 | 3.05 | 13.22 | 13.28 | 2.23 | 2.16 | 12.22 | 12.11 | 3.02 | 3.04 | 13.69 | 13.67 | 5.17 | 5.16 | 14.23 | 14.16 | 4.64 | 4.72 | 12.69 | 12.63 | 3.11 | 3.12 | 12.54 | 12.72 | 3.33 | 3.38
11.4 | 11.33 | 2.98 | 2.96 | 11.29 | 11.23 | 3.06 | 3 | 11.91 | 11.37 | 3.31 | 2.99 | 12.06 | 11.98 | 3.05 | 3.13 | 13.28 | 13.27 | 2.16 | 2.16 | 12.11 | 11.86 | 3.04 | 3.03 | 13.67 | 13.63 | 5.16 | 5.17 | 14.16 | 14.03 | 4.72 | 4.64 | 12.63 | 12.68 | 3.12 | 3.09 | 12.72 | 12.79 | 3.38 | 3.48
11.33 | 11.37 | 2.96 | 2.95 | 11.23 | 11.38 | 3 | 3.09 | 11.37 | 11.37 | 2.99 | 2.98 | 11.98 | 12.15 | 3.13 | 3.01 | 13.27 | 13.32 | 2.16 | 2.22 | 11.86 | 12.41 | 3.03 | 3.06 | 13.63 | 13.59 | 5.17 | 5.23 | 14.03 | 14.38 | 4.64 | 4.62 | 12.68 | 12.66 | 3.09 | 2.98 | 12.79 | 12.4 | 3.48 | 3.44
11.37 | 11.33 | 2.95 | 2.96 | 11.38 | 11.33 | 3.09 | 3.32 | 11.37 | 11.42 | 2.98 | 3 | 12.15 | 12.08 | 3.01 | 2.97 | 13.32 | 13.26 | 2.22 | 2.21 | 12.41 | 12.45 | 3.06 | 2.94 | 13.59 | 13.58 | 5.23 | 5.19 | 14.38 | 14.29 | 4.62 | 4.73 | 12.66 | 12.73 | 2.98 | 3.01 | 12.4 | 12.76 | 3.44 | 3.45
11.33 | 11.28 | 2.96 | 2.95 | 11.33 | 11.34 | 3.32 | 3.05 | 11.42 | 11.4 | 3 | 3.04 | 12.08 | 12 | 2.97 | 2.97 | 13.26 | 13.24 | 2.21 | 2.16 | 12.45 | 12.4 | 2.94 | 3.08 | 13.58 | 13.63 | 5.19 | 5.21 | 14.29 | 13.79 | 4.73 | 4.73 | 12.73 | 12.76 | 3.01 | 3.01 | 12.76 | 12.47 | 3.45 | 3.45
11.28 | 11.32 | 2.95 | 2.96 | 11.34 | 11.25 | 3.05 | 3.47 | 11.4 | 11.46 | 3.04 | 3.04 | 12 | 12.07 | 2.97 | 3.04 | 13.24 | 13.18 | 2.16 | 2.17 | 12.4 | 12.3 | 3.08 | 3.1 | 13.63 | 13.68 | 5.21 | 5.18 | 13.79 | 13.83 | 4.73 | 4.61 | 12.76 | 12.78 | 3.01 | 3.03 | 12.47 | 12.71 | 3.45 | 3.48
11.4 | 11.4 | 2.99 | 3.01 | 11.25 | 11.31 | 3.47 | 3.06 | 11.46 | 11.41 | 3.04 | 3 | 12.07 | 12.11 | 3.04 | 3.05 | 13.18 | 13.11 | 2.17 | 2.19 | 12.57 | 12.71 | 1.42 | 1.42 | 13.49 | 13.46 | 5.02 | 5.04 | 13.83 | 12.58 | 4.61 | 4.68 | 12.78 | 12.75 | 3.03 | 3.03 | 12.71 | 12.7 | 3.48 | 3.4
11.4 | 11.45 | 3.01 | 3.01 | 10.78 | 10.92 | 3.16 | 3.02 | 11.55 | 11.55 | 2.94 | 2.93 | 11.25 | 11.46 | 2.98 | 2.95 | 13.31 | 13.35 | 2.13 | 2.1 | 12.71 | 12.58 | 1.42 | 1.42 | 13.46 | 13.5 | 5.04 | 4.97 | 15.54 | 15.8 | 5.27 | 5.33 | 12.68 | 12.68 | 2.89 | 3.09 | 13.23 | 13.29 | 3.1 | 2.95
11.45 | 11.5 | 3.01 | 2.99 | 10.92 | 10.94 | 3.02 | 3.01 | 11.55 | 11.56 | 2.93 | 2.93 | 11.46 | 11.27 | 2.95 | 3.04 | 13.35 | 13.39 | 2.1 | 2.06 | 12.58 | 12.26 | 1.42 | 1.46 | 13.5 | 13.56 | 4.97 | 4.96 | 15.8 | 15.81 | 5.33 | 5.37 | 12.68 | 12.46 | 3.09 | 2.97 | 13.29 | 13.23 | 2.95 | 2.97
11.5 | 11.47 | 2.99 | 3 | 10.94 | 10.79 | 3.01 | 3.17 | 11.56 | 11.45 | 2.93 | 2.94 | 11.27 | 11.15 | 3.04 | 3 | 13.39 | 13.36 | 2.06 | 2.1 | 12.26 | 12.33 | 1.46 | 1.55 | 13.56 | 13.56 | 4.96 | 4.98 | 15.81 | 15.56 | 5.37 | 5.26 | 12.46 | 12.43 | 2.97 | 2.97 | 13.23 | 12.92 | 2.97 | 3.09
11.47 | 11.47 | 3 | 3.02 | 10.79 | 10.66 | 3.17 | 3.1 | 11.45 | 11.42 | 2.94 | 2.93 | 11.15 | 11.1 | 3 | 3.09 | 13.36 | 13.34 | 2.1 | 2.13 | 12.33 | 11.93 | 1.55 | 0.84 | 13.56 | 13.55 | 4.98 | 4.97 | 15.56 | 15.62 | 5.26 | 5.3 | 12.43 | 12.55 | 2.97 | 2.98 | 12.92 | 13.03 | 3.09 | 3.04
11.47 | 11.5 | 3.02 | 3.04 | 10.66 | 10.66 | 3.1 | 2.94 | 11.42 | 11.39 | 2.93 | 2.95 | 11.1 | 11.31 | 3.09 | 3.12 | 13.34 | 13.3 | 2.13 | 2.1 | 11.93 | 12.67 | 0.84 | 1.41 | 13.55 | 13.59 | 4.97 | 5.04 | 15.62 | 15.69 | 5.3 | 5.27 | 12.55 | 12.59 | 2.98 | 2.85 | 13.03 | 12.88 | 3.04 | 2.96
11.5 | 11.52 | 3.04 | 3.02 | 10.66 | 10.85 | 2.94 | 3.15 | 11.39 | 11.41 | 2.95 | 2.94 | 11.31 | 11.32 | 3.12 | 2.92 | 13.3 | 13.16 | 2.1 | 2.08 | 12.67 | 12.87 | 1.41 | 1.45 | 13.59 | 13.6 | 5.04 | 5.02 | 15.69 | 15.66 | 5.27 | 5.3 | 12.59 | 12.64 | 2.85 | 2.81 | 12.88 | 12.86 | 2.96 | 2.94
11.52 | 11.44 | 3.02 | 3.04 | 10.85 | 10.62 | 3.15 | 3 | 11.41 | 11.45 | 2.94 | 2.92 | 11.32 | 10.87 | 2.92 | 2.84 | 13.16 | 13.14 | 2.08 | 2.13 | 12.87 | 12.86 | 1.45 | 1.42 | 13.6 | 13.55 | 5.02 | 4.98 | 15.66 | 15.46 | 5.3 | 5.3 | 12.64 | 12.67 | 2.81 | 2.86 | 12.86 | 13.32 | 2.94 | 2.99
11.44 | 11.39 | 3.04 | 3.01 | 10.62 | 10.74 | 3 | 2.98 | 11.45 | 11.49 | 2.92 | 2.91 | 10.87 | 10.87 | 2.84 | 2.9 | 13.14 | 13.22 | 2.13 | 2.13 | 12.86 | 12.72 | 1.42 | 1.53 | 13.55 | 13.61 | 4.98 | 5.04 | 15.46 | 15.67 | 5.3 | 5.22 | 12.67 | 12.63 | 2.86 | 2.84 | 13.32 | 13.19 | 2.99 | 2.96
11.39 | 11.47 | 3.01 | 2.98 | 10.74 | 10.56 | 2.98 | 3.01 | 11.49 | 11.48 | 2.91 | 2.92 | 10.87 | 11.25 | 2.9 | 2.87 | 13.22 | 13.3 | 2.13 | 2.05 | 12.72 | 12.76 | 1.53 | 1.4 | 13.61 | 13.56 | 5.04 | 5.03 | 15.67 | 15.04 | 5.22 | 5.31 | 12.63 | 12.56 | 2.84 | 2.83 | 13.19 | 12.94 | 2.96 | 2.93
11.47 | 11.36 | 2.98 | 3 | 10.56 | 10.89 | 3.01 | 2.92 | 11.48 | 11.53 | 2.92 | 2.91 | 11.25 | 11.34 | 2.87 | 2.93 | 13.3 | 13.35 | 2.05 | 2.08 | 12.76 | 12.8 | 1.4 | 1.42 | 13.56 | 13.52 | 5.03 | 4.98 | 15.04 | 14.95 | 5.31 | 5.24 | 12.56 | 12.58 | 2.83 | 2.84 | 12.94 | 13.2 | 2.93 | 3
11.36 | 11.38 | 3 | 2.99 | 10.89 | 10.81 | 2.92 | 3.01 | 11.53 | 11.49 | 2.91 | 2.93 | 11.34 | 11.3 | 2.93 | 2.88 | 13.35 | 13.36 | 2.08 | 2.13 | 12.8 | 12.78 | 1.42 | 1.37 | 13.52 | 13.47 | 4.98 | 5.01 | 14.95 | 14.62 | 5.24 | 5.25 | 12.58 | 12.49 | 2.84 | 2.88 | 13.2 | 12.88 | 3 | 2.97
11.38 | 11.4 | 2.99 | 3.02 | 10.81 | 10.61 | 3.01 | 2.98 | 11.49 | 11.53 | 2.93 | 2.93 | 11.3 | 11.35 | 2.88 | 2.95 | 13.36 | 13.33 | 2.13 | 2.1 | 12.78 | 12.7 | 1.37 | 1.42 | 13.47 | 13.5 | 5.01 | 5.03 | 14.62 | 14.73 | 5.25 | 5.32 | 12.49 | 12.48 | 2.88 | 2.88 | 12.88 | 13.18 | 2.97 | 2.94
11.4 | 11.4 | 3.02 | 3.03 | 10.61 | 10.59 | 2.98 | 3.14 | 11.53 | 11.54 | 2.93 | 2.93 | 11.35 | 11.28 | 2.95 | 2.88 | 13.33 | 13.34 | 2.1 | 2.06 | 12.7 | 12.58 | 1.42 | 1.49 | 13.5 | 13.51 | 5.03 | 4.97 | 14.73 | 14.56 | 5.32 | 5.21 | 12.48 | 12.52 | 2.88 | 2.91 | 13.18 | 13.09 | 2.94 | 2.91
11.4 | 11.43 | 3.03 | 3.03 | 10.59 | 10.54 | 3.14 | 2.96 | 11.54 | 11.52 | 2.93 | 2.93 | 11.28 | 11.32 | 2.88 | 2.99 | 13.34 | 13.24 | 2.06 | 2.09 | 12.58 | 12.5 | 1.49 | 1.36 | 13.51 | 13.55 | 4.97 | 4.93 | 14.56 | 14.58 | 5.21 | 5.31 | 12.52 | 12.6 | 2.91 | 2.94 | 13.09 | 13.07 | 2.91 | 3.07
11.43 | 11.47 | 3.03 | 3.06 | 10.54 | 10.49 | 2.96 | 2.96 | 11.52 | 11.47 | 2.93 | 2.91 | 11.32 | 11.33 | 2.99 | 2.75 | 13.24 | 13.19 | 2.09 | 2.11 | 12.5 | 12.5 | 1.36 | 1.44 | 13.55 | 13.57 | 4.93 | 4.94 | 14.58 | 14.6 | 5.31 | 5.23 | 12.6 | 12.62 | 2.94 | 2.96 | 13.07 | 12.79 | 3.07 | 2.91
11.47 | 11.48 | 3.06 | 3.05 | 10.49 | 11.13 | 2.96 | 3 | 11.47 | 11.43 | 2.91 | 2.93 | 11.33 | 11.31 | 2.75 | 2.97 | 13.19 | 13.14 | 2.11 | 2.15 | 12.5 | 12.59 | 1.44 | 1.4 | 13.57 | 13.58 | 4.94 | 4.92 | 14.6 | 14.28 | 5.23 | 5.29 | 12.62 | 12.7 | 2.96 | 2.95 | 12.79 | 12.78 | 2.91 | 2.95
11.48 | 11.49 | 3.05 | 3.01 | 11.13 | 10.93 | 3 | 2.99 | 11.43 | 11.37 | 2.93 | 2.94 | 11.31 | 11.23 | 2.97 | 2.98 | 13.14 | 13.21 | 2.15 | 2.07 | 12.59 | 12.53 | 1.4 | 1.38 | 13.58 | 13.53 | 4.92 | 4.96 | 14.28 | 14.42 | 5.29 | 5.26 | 12.7 | 12.67 | 2.95 | 2.95 | 12.78 | 12.97 | 2.95 | 3.04
11.49 | 11.53 | 3.01 | 3.01 | 10.93 | 10.91 | 2.99 | 3.17 | 11.37 | 11.39 | 2.94 | 2.93 | 11.23 | 11.39 | 2.98 | 2.9 | 13.21 | 13.27 | 2.07 | 2.09 | 12.53 | 12.51 | 1.38 | 1.37 | 13.53 | 13.57 | 4.96 | 4.98 | 14.42 | 14.23 | 5.26 | 5.18 | 12.67 | 12.66 | 2.95 | 2.88 | 12.97 | 13.1 | 3.04 | 2.95
11.53 | 11.5 | 3.01 | 3.01 | 10.91 | 11.29 | 3.17 | 3.12 | 11.39 | 11.43 | 2.93 | 2.92 | 11.39 | 11.42 | 2.9 | 3 | 13.27 | 13.3 | 2.09 | 2.1 | 12.51 | 12.52 | 1.37 | 1.36 | 13.57 | 13.58 | 4.98 | 4.97 | 14.23 | 14.02 | 5.18 | 5.27 | 12.66 | 12.61 | 2.88 | 2.84 | 13.1 | 12.82 | 2.95 | 2.91
11.5 | 11.49 | 3.01 | 3.01 | 11.29 | 11 | 3.12 | 2.99 | 11.43 | 11.48 | 2.92 | 2.93 | 11.42 | 11.41 | 3 | 2.96 | 13.3 | 13.32 | 2.1 | 2.13 | 12.52 | 12.7 | 1.36 | 1.38 | 13.58 | 13.59 | 4.97 | 4.96 | 14.02 | 13.94 | 5.27 | 5.24 | 12.61 | 12.63 | 2.84 | 2.85 | 12.82 | 12.99 | 2.91 | 2.94
11.49 | 11.51 | 3.01 | 3.04 | 11 | 11.54 | 2.99 | 3.01 | 11.48 | 11.47 | 2.93 | 2.93 | 11.41 | 11.22 | 2.96 | 2.31 | 13.32 | 13.4 | 2.13 | 2.07 | 12.7 | 12.61 | 1.38 | 1.44 | 13.59 | 13.5 | 4.96 | 4.98 | 13.94 | 13.97 | 5.24 | 5.24 | 12.63 | 12.6 | 2.85 | 2.83 | 12.99 | 13.02 | 2.94 | 2.98
11.51 | 11.47 | 3.04 | 3.03 | 11.54 | 11.52 | 3.01 | 3.18 | 11.47 | 11.51 | 2.93 | 2.93 | 11.22 | 11.29 | 2.31 | 2.99 | 13.4 | 13.36 | 2.07 | 2.09 | 12.61 | 12.42 | 1.44 | 1.45 | 13.5 | 13.47 | 4.98 | 4.99 | 13.97 | 13.94 | 5.24 | 5.3 | 12.6 | 12.57 | 2.83 | 2.85 | 13.02 | 13.02 | 2.98 | 2.94
11.47 | 11.43 | 3.03 | 2.99 | 11.52 | 11.42 | 3.18 | 2.96 | 11.51 | 11.53 | 2.93 | 2.92 | 11.29 | 11.42 | 2.99 | 3.03 | 13.36 | 13.3 | 2.09 | 2.12 | 12.42 | 12.21 | 1.45 | 1.41 | 13.47 | 13.43 | 4.99 | 5.03 | 13.94 | 13.91 | 5.3 | 5.19 | 12.57 | 12.98 | 2.85 | 3.11 | 13.02 | 12.75 | 2.94 | 2.86
11.43 | 11.39 | 2.99 | 3 | 11.42 | 11.52 | 2.96 | 3.01 | 11.53 | 11.57 | 2.92 | 2.92 | 11.42 | 11.43 | 3.03 | 2.93 | 13.3 | 13.25 | 2.12 | 2.11 | 12.21 | 12.43 | 1.41 | 1.39 | 13.43 | 13.5 | 5.03 | 5.02 | 13.91 | 14.15 | 5.19 | 5.39 | 12.98 | 12.43 | 3.11 | 2.93 | 12.75 | 13.09 | 2.86 | 3.01
11.39 | 11.36 | 3 | 3.02 | 11.52 | 11.49 | 3.01 | 3.11 | 11.57 | 11.55 | 2.92 | 2.92 | 11.43 | 11.31 | 2.93 | 2.97 | 13.25 | 13.12 | 2.11 | 2.04 | 12.43 | 11.95 | 1.39 | 1.38 | 13.5 | 13.55 | 5.02 | 5.01 | 14.15 | 14.06 | 5.39 | 5.27 | 12.43 | 12.44 | 2.93 | 2.92 | 13.09 | 12.31 | 3.01 | 2.95
11.36 | 11.42 | 3.02 | 3.03 | 11.49 | 11.3 | 3.11 | 3 | 11.55 | 11.55 | 2.92 | 2.94 | 11.31 | 11.3 | 2.97 | 3.02 | 13.12 | 13.21 | 2.04 | 2.08 | 11.95 | 12.82 | 1.38 | 1.4 | 13.55 | 13.56 | 5.01 | 4.99 | 14.06 | 13.94 | 5.27 | 5.27 | 12.44 | 12.51 | 2.92 | 2.95 | 12.31 | 12.75 | 2.95 | 2.93
11.42 | 11.44 | 3.03 | 3.04 | 11.3 | 11.39 | 3 | 3.11 | 11.55 | 11.51 | 2.94 | 2.94 | 11.3 | 11.39 | 3.02 | 3.01 | 13.21 | 13.26 | 2.08 | 2.09 | 12.82 | 12.57 | 1.4 | 1.39 | 13.56 | 13.58 | 4.99 | 4.99 | 13.94 | 13.94 | 5.27 | 5.27 | 12.51 | 12.55 | 2.95 | 2.91 | 12.75 | 13.24 | 2.93 | 3.04
11.44 | 11.46 | 3.04 | 3.02 | 11.39 | 11.42 | 3.11 | 3.22 | 11.51 | 11.47 | 2.94 | 2.93 | 11.39 | 11.41 | 3.01 | 2.99 | 13.26 | 13.31 | 2.09 | 2.11 | 12.57 | 12.59 | 1.39 | 1.41 | 13.58 | 13.58 | 4.99 | 5.01 | 13.94 | 13.91 | 5.27 | 5.22 | 12.55 | 12.61 | 2.91 | 2.87 | 13.24 | 13.26 | 3.04 | 2.87
11.46 | 11.49 | 3.02 | 3.03 | 11.42 | 11.44 | 3.22 | 3 | 11.47 | 11.35 | 2.93 | 2.93 | 11.41 | 11.34 | 2.99 | 3.01 | 13.31 | 13.27 | 2.11 | 2.06 | 12.59 | 12.84 | 1.41 | 1.4 | 13.58 | 13.6 | 5.01 | 4.99 | 13.91 | 13.76 | 5.22 | 5.24 | 12.61 | 12.62 | 2.87 | 2.85 | 13.26 | 13.23 | 2.87 | 2.89
11.49 | 11.45 | 3.03 | 3.01 | 11.44 | 11.4 | 3 | 3.06 | 11.35 | 11.41 | 2.93 | 2.9 | 11.34 | 11.26 | 3.01 | 2.98 | 13.27 | 13.41 | 2.06 | 2.06 | 12.84 | 12.68 | 1.4 | 1.37 | 13.6 | 13.62 | 4.99 | 4.98 | 13.76 | 13.92 | 5.24 | 5.23 | 12.62 | 12.65 | 2.85 | 2.86 | 13.23 | 13.23 | 2.89 | 2.94
11.45 | 11.52 | 3.01 | 3.04 | 11.4 | 11.49 | 3.06 | 3.17 | 11.41 | 11.4 | 2.9 | 2.91 | 11.26 | 11.27 | 2.98 | 2.97 | 13.41 | 13.34 | 2.06 | 2.11 | 12.68 | 12.61 | 1.37 | 1.41 | 13.62 | 13.55 | 4.98 | 4.99 | 13.92 | 13.94 | 5.23 | 5.32 | 12.65 | 12.63 | 2.86 | 2.83 | 13.23 | 13.05 | 2.94 | 3
11.52 | 11.52 | 3.04 | 3.05 | 11.49 | 11.39 | 3.17 | 3.02 | 11.4 | 11.42 | 2.91 | 2.92 | 11.27 | 11.4 | 2.97 | 2.96 | 13.34 | 13.33 | 2.11 | 2.12 | 12.61 | 12.64 | 1.41 | 1.34 | 13.55 | 13.51 | 4.99 | 4.95 | 13.94 | 14.04 | 5.32 | 5.17 | 12.63 | 12.62 | 2.83 | 2.86 | 13.05 | 13.01 | 3 | 2.93
11.52 | 11.52 | 3.05 | 3.05 | 11.39 | 11.42 | 3.02 | 3.06 | 11.42 | 11.46 | 2.92 | 2.92 | 11.4 | 11.44 | 2.96 | 1.93 | 13.33 | 13.26 | 2.12 | 2.07 | 12.64 | 12.64 | 1.34 | 1.38 | 13.51 | 13.56 | 4.95 | 4.96 | 14.04 | 13.91 | 5.17 | 5.22 | 12.62 | 12.63 | 2.86 | 2.87 | 13.01 | 12.84 | 2.93 | 2.92
11.52 | 11.53 | 3.05 | 3.03 | 11.42 | 11.5 | 3.06 | 3.13 | 11.46 | 11.48 | 2.92 | 2.92 | 11.44 | 11.38 | 1.93 | 2.95 | 13.26 | 13.13 | 2.07 | 2.11 | 12.64 | 12.5 | 1.38 | 1.34 | 13.56 | 13.45 | 4.96 | 4.96 | 13.91 | 13.81 | 5.22 | 5.14 | 12.63 | 12.53 | 2.87 | 2.87 | 12.84 | 13.26 | 2.92 | 2.94
11.53 | 11.5 | 3.03 | 3.03 | 11.5 | 11.52 | 3.13 | 3.06 | 11.48 | 11.53 | 2.92 | 2.93 | 11.38 | 11.21 | 2.95 | 3 | 13.13 | 13.16 | 2.11 | 2.1 | 12.5 | 12.36 | 1.34 | 1.37 | 13.45 | 13.48 | 4.96 | 4.96 | 13.81 | 13.81 | 5.14 | 5.31 | 12.53 | 12.49 | 2.87 | 2.92 | 13.26 | 12.37 | 2.94 | 2.97
11.5 | 11.47 | 3.03 | 3 | 11.52 | 11.65 | 3.06 | 3.06 | 11.53 | 11.57 | 2.93 | 2.94 | 11.21 | 11.19 | 3 | 2.39 | 13.16 | 13.26 | 2.1 | 2.03 | 13.16 | 13.27 | 1.29 | 1.21 | 13.48 | 13.5 | 4.96 | 4.99 | 13.81 | 13.88 | 5.31 | 5.19 | 12.49 | 12.47 | 2.92 | 2.93 | 12.37 | 13.14 | 2.97 | 2.92
11.47 | 11.42 | 3 | 3.03 | 11.65 | 11.66 | 3.06 | 2.95 | 11.57 | 11.54 | 2.94 | 2.92 | 11.19 | 11.46 | 2.39 | 2.9 | 13.26 | 13.26 | 2.03 | 2.04 | 13.27 | 13.19 | 1.21 | 1.28 | 13.51 | 13.51 | 4.89 | 4.85 | 13.88 | 13.86 | 5.19 | 5.19 | 12.47 | 12.5 | 2.93 | 2.91 | 13.14 | 12.64 | 2.92 | 2.97
11.38 | 11.37 | 2.91 | 2.93 | 11.66 | 11.65 | 3.02 | 3 | 11.37 | 11.32 | 2.64 | 2.64 | 11.96 | 11.84 | 2.9 | 2.91 | 13.26 | 13.34 | 2.04 | 2.09 | 13.19 | 13.25 | 1.28 | 1.21 | 13.51 | 13.49 | 4.85 | 4.84 | 13.86 | 14.19 | 5.19 | 5.24 | 12.52 | 12.48 | 2.71 | 2.73 | 12.64 | 13.08 | 2.97 | 2.98
11.37 | 11.38 | 2.93 | 2.92 | 11.65 | 11.59 | 3 | 3.16 | 11.32 | 11.34 | 2.64 | 2.64 | 11.84 | 11.81 | 2.91 | 3.02 | 13.23 | 13.17 | 1.97 | 1.97 | 13.25 | 13.21 | 1.21 | 1.27 | 13.49 | 13.44 | 4.84 | 4.85 | 13.84 | 14.03 | 5.02 | 4.89 | 12.48 | 12.46 | 2.73 | 2.76 | 13.49 | 13.42 | 2.82 | 2.89
11.38 | 11.37 | 2.92 | 2.9 | 11.59 | 11.55 | 3.16 | 3.1 | 11.34 | 11.39 | 2.64 | 2.64 | 11.81 | 12 | 3.02 | 2.92 | 13.17 | 13.31 | 1.97 | 2 | 13.21 | 13.24 | 1.27 | 1.24 | 13.44 | 13.37 | 4.85 | 4.85 | 14.03 | 13.98 | 4.89 | 4.96 | 12.46 | 12.47 | 2.76 | 2.77 | 13.42 | 13.47 | 2.89 | 2.83
11.37 | 11.36 | 2.9 | 2.9 | 11.55 | 11.66 | 3.1 | 3.08 | 11.39 | 11.28 | 2.64 | 2.62 | 12 | 11.89 | 2.92 | 2.91 | 13.31 | 13.26 | 2 | 2.02 | 13.24 | 13.12 | 1.24 | 1.32 | 13.37 | 13.4 | 4.85 | 4.85 | 13.98 | 14.07 | 4.96 | 4.94 | 12.47 | 12.53 | 2.77 | 2.76 | 13.47 | 13.54 | 2.83 | 2.86
11.36 | 11.39 | 2.9 | 2.9 | 11.66 | 11.54 | 3.08 | 3.11 | 11.28 | 11.32 | 2.62 | 2.67 | 11.89 | 11.89 | 2.91 | 3.01 | 13.26 | 13.25 | 2.02 | 2.03 | 13.12 | 13.05 | 1.32 | 1.22 | 13.4 | 13.44 | 4.85 | 4.88 | 14.07 | 13.87 | 4.94 | 4.93 | 12.53 | 12.53 | 2.76 | 2.78 | 13.54 | 13.35 | 2.86 | 2.92
11.39 | 11.3 | 2.9 | 2.87 | 11.54 | 11.59 | 3.11 | 3.04 | 11.32 | 11.28 | 2.67 | 2.68 | 11.89 | 11.87 | 3.01 | 2.97 | 13.25 | 13.17 | 2.03 | 2.03 | 13.05 | 13.29 | 1.22 | 1.28 | 13.44 | 13.5 | 4.88 | 4.9 | 13.87 | 13.83 | 4.93 | 5.01 | 12.53 | 12.58 | 2.78 | 2.74 | 13.35 | 13.35 | 2.92 | 2.91
11.3 | 11.32 | 2.87 | 2.91 | 11.59 | 11.6 | 3.04 | 3.08 | 11.28 | 11.16 | 2.68 | 2.7 | 11.87 | 11.85 | 2.97 | 2.89 | 13.17 | 13.07 | 2.03 | 1.96 | 13.29 | 13.26 | 1.28 | 1.2 | 13.5 | 13.51 | 4.9 | 4.92 | 13.83 | 13.85 | 5.01 | 5.04 | 12.58 | 12.61 | 2.74 | 2.69 | 13.35 | 13.44 | 2.91 | 2.89
11.32 | 11.28 | 2.91 | 2.93 | 11.6 | 11.59 | 3.08 | 3.06 | 11.16 | 11.12 | 2.7 | 2.68 | 11.85 | 11.84 | 2.89 | 2.93 | 13.07 | 13.03 | 1.96 | 2 | 13.26 | 13.32 | 1.2 | 1.26 | 13.51 | 13.53 | 4.92 | 4.86 | 13.85 | 13.95 | 5.04 | 4.87 | 12.61 | 12.57 | 2.69 | 2.67 | 13.44 | 13.32 | 2.89 | 2.98
11.28 | 11.26 | 2.93 | 2.89 | 11.59 | 11.74 | 3.06 | 3.06 | 11.12 | 11.19 | 2.68 | 2.69 | 11.84 | 11.86 | 2.93 | 2.93 | 13.03 | 13.1 | 2 | 2 | 13.32 | 13.18 | 1.26 | 1.23 | 13.53 | 13.53 | 4.86 | 4.84 | 13.95 | 13.9 | 4.87 | 4.98 | 12.57 | 12.62 | 2.67 | 2.68 | 13.32 | 13.43 | 2.98 | 2.83
11.26 | 11.32 | 2.89 | 2.87 | 11.74 | 11.65 | 3.06 | 3.1 | 11.19 | 11.21 | 2.69 | 2.69 | 11.86 | 11.82 | 2.93 | 2.92 | 13.1 | 13.1 | 2 | 2.02 | 13.18 | 13.39 | 1.23 | 1.28 | 13.53 | 13.52 | 4.84 | 4.84 | 13.9 | 13.99 | 4.98 | 4.94 | 12.62 | 12.61 | 2.68 | 2.7 | 13.43 | 13.45 | 2.83 | 2.83
11.32 | 11.35 | 2.87 | 2.88 | 11.65 | 11.49 | 3.1 | 3.12 | 11.21 | 11.28 | 2.69 | 2.7 | 11.82 | 11.61 | 2.92 | 2.93 | 13.1 | 13.15 | 2.02 | 1.98 | 13.39 | 13.09 | 1.28 | 1.23 | 13.52 | 13.56 | 4.84 | 4.82 | 13.99 | 14.04 | 4.94 | 5.01 | 12.61 | 12.58 | 2.7 | 2.69 | 13.45 | 13.45 | 2.83 | 2.85
11.35 | 11.36 | 2.88 | 2.91 | 11.49 | 11.46 | 3.12 | 3.16 | 11.28 | 11.28 | 2.7 | 2.7 | 11.61 | 11.76 | 2.93 | 2.79 | 13.15 | 13.22 | 1.98 | 2 | 13.09 | 13.14 | 1.23 | 1.3 | 13.56 | 13.45 | 4.82 | 4.84 | 14.04 | 13.92 | 5.01 | 4.96 | 12.58 | 12.48 | 2.69 | 2.69 | 13.45 | 13.52 | 2.85 | 2.93
11.36 | 11.42 | 2.91 | 2.92 | 11.46 | 11.72 | 3.16 | 3.1 | 11.28 | 11.25 | 2.7 | 2.68 | 11.76 | 11.84 | 2.79 | 3.3 | 13.22 | 13.22 | 2 | 2.02 | 13.14 | 13.13 | 1.3 | 1.23 | 13.45 | 13.43 | 4.84 | 4.88 | 13.92 | 14 | 4.96 | 5.05 | 12.48 | 12.5 | 2.69 | 2.73 | 13.52 | 13.4 | 2.93 | 2.83
11.42 | 11.39 | 2.92 | 2.89 | 11.72 | 11.65 | 3.1 | 3.08 | 11.25 | 11.27 | 2.68 | 2.66 | 11.84 | 11.76 | 3.3 | 2.89 | 13.22 | 13.24 | 2.02 | 1.99 | 13.13 | 13.1 | 1.23 | 1.3 | 13.43 | 13.38 | 4.88 | 4.89 | 14 | 9.14 | 5.05 | 3.14 | 12.5 | 12.54 | 2.73 | 2.81 | 13.4 | 13.46 | 2.83 | 2.95
11.39 | 11.36 | 2.89 | 2.88 | 11.65 | 11.66 | 3.08 | 3.02 | 11.27 | 11.35 | 2.66 | 2.66 | 11.76 | 11.64 | 2.89 | 2.9 | 13.24 | 13.2 | 1.99 | 2.02 | 13.1 | 13.08 | 1.3 | 1.24 | 13.38 | 13.38 | 4.89 | 4.91 | 9.14 | 14.02 | 3.14 | 4 | 12.54 | 12.49 | 2.81 | 2.8 | 13.46 | 13.53 | 2.95 | 2.94
11.36 | 11.4 | 2.88 | 2.89 | 11.66 | 11.69 | 3.02 | 3.11 | 11.35 | 11.32 | 2.66 | 2.66 | 11.64 | 11.82 | 2.9 | 2.89 | 13.2 | 13.13 | 2.02 | 1.96 | 13.08 | 13.15 | 1.24 | 1.3 | 13.38 | 13.46 | 4.91 | 4.86 | 14.02 | 14.13 | 4 | 5.06 | 12.49 | 12.47 | 2.8 | 2.77 | 13.53 | 13.41 | 2.94 | 2.86
11.4 | 11.4 | 2.89 | 2.9 | 11.69 | 11.6 | 3.11 | 3.01 | 11.32 | 11.33 | 2.66 | 2.64 | 11.82 | 11.88 | 2.89 | 2.89 | 13.13 | 13.05 | 1.96 | 1.98 | 13.15 | 13.15 | 1.3 | 1.23 | 13.46 | 13.5 | 4.86 | 4.85 | 14.13 | 14.05 | 5.06 | 4.98 | 12.47 | 12.58 | 2.77 | 2.65 | 13.41 | 13.48 | 2.86 | 2.92
11.4 | 11.39 | 2.9 | 2.94 | 11.6 | 11.73 | 3.01 | 3.02 | 11.33 | 11.32 | 2.64 | 2.6 | 11.88 | 11.83 | 2.89 | 2.95 | 13.05 | 13.01 | 1.98 | 2 | 13.15 | 13.19 | 1.23 | 1.29 | 13.5 | 13.51 | 4.85 | 4.85 | 14.05 | 14.08 | 4.98 | 5.12 | 12.58 | 12.56 | 2.65 | 2.65 | 13.48 | 13.65 | 2.92 | 2.87
11.39 | 11.35 | 2.94 | 2.91 | 11.73 | 11.7 | 3.02 | 3.14 | 11.32 | 11.27 | 2.6 | 2.59 | 11.83 | 11.78 | 2.95 | 2.87 | 13.01 | 13.09 | 2 | 2.02 | 13.19 | 13.14 | 1.29 | 1.2 | 13.51 | 13.52 | 4.85 | 4.85 | 14.08 | 13.95 | 5.12 | 4.98 | 12.56 | 12.58 | 2.65 | 2.63 | 13.65 | 13.5 | 2.87 | 2.78
11.35 | 11.34 | 2.91 | 2.87 | 11.7 | 11.57 | 3.14 | 3.14 | 11.27 | 11.28 | 2.59 | 2.6 | 11.78 | 11.62 | 2.87 | 2.98 | 13.09 | 13.12 | 2.02 | 2.04 | 13.14 | 13.14 | 1.2 | 1.28 | 13.52 | 13.47 | 4.85 | 4.81 | 13.95 | 14.04 | 4.98 | 5.06 | 12.58 | 12.58 | 2.63 | 2.67 | 13.5 | 13.49 | 2.78 | 2.84
11.34 | 11.29 | 2.87 | 2.99 | 11.57 | 11.58 | 3.14 | 3.03 | 11.28 | 11.26 | 2.6 | 2.6 | 11.62 | 11.82 | 2.98 | 2.94 | 13.12 | 13.17 | 2.04 | 1.98 | 13.14 | 12.83 | 1.28 | 1.19 | 13.47 | 13.52 | 4.81 | 4.89 | 14.04 | 14.06 | 5.06 | 4.98 | 12.58 | 12.63 | 2.67 | 2.66 | 13.49 | 13.38 | 2.84 | 2.94
11.29 | 11.24 | 2.99 | 2.88 | 11.58 | 11.7 | 3.03 | 3.17 | 11.26 | 11.14 | 2.6 | 2.63 | 11.82 | 11.82 | 2.94 | 2.91 | 13.17 | 13.17 | 1.98 | 1.96 | 12.83 | 13.05 | 1.19 | 1.29 | 13.52 | 13.53 | 4.89 | 4.92 | 14.06 | 14.09 | 4.98 | 5.03 | 12.63 | 12.58 | 2.66 | 2.7 | 13.38 | 13.64 | 2.94 | 2.9
11.24 | 11.26 | 2.88 | 2.91 | 11.7 | 11.65 | 3.17 | 3.03 | 11.14 | 11.17 | 2.63 | 2.62 | 11.82 | 11.96 | 2.91 | 3.02 | 13.17 | 13.24 | 1.96 | 2 | 13.05 | 13.25 | 1.29 | 1.21 | 13.53 | 13.51 | 4.92 | 4.87 | 14.09 | 14.06 | 5.03 | 4.91 | 12.58 | 12.59 | 2.7 | 2.67 | 13.64 | 13.52 | 2.9 | 2.82
11.26 | 11.3 | 2.91 | 2.85 | 11.65 | 11.63 | 3.03 | 3.28 | 11.17 | 11.19 | 2.62 | 2.62 | 11.96 | 11.77 | 3.02 | 2.88 | 13.24 | 13.21 | 2 | 2 | 13.25 | 13.17 | 1.21 | 1.26 | 13.51 | 13.51 | 4.87 | 4.84 | 14.06 | 14.12 | 4.91 | 5.1 | 12.59 | 12.52 | 2.67 | 2.73 | 13.52 | 13.31 | 2.82 | 2.82
11.3 | 11.35 | 2.85 | 2.89 | 11.63 | 11.66 | 3.28 | 3.09 | 11.19 | 11.22 | 2.62 | 2.62 | 11.77 | 11.83 | 2.88 | 2.87 | 13.21 | 13.24 | 2 | 1.95 | 13.17 | 13.21 | 1.26 | 1.22 | 13.51 | 13.44 | 4.84 | 4.85 | 14.12 | 13.93 | 5.1 | 4.93 | 12.52 | 12.5 | 2.73 | 2.73 | 13.31 | 13.37 | 2.82 | 2.83
11.35 | 11.35 | 2.89 | 2.93 | 11.66 | 11.79 | 3.09 | 3.02 | 11.22 | 11.3 | 2.62 | 2.62 | 11.83 | 11.82 | 2.87 | 2.84 | 13.24 | 13.14 | 1.95 | 1.99 | 13.21 | 13.2 | 1.22 | 1.27 | 13.44 | 13.36 | 4.85 | 4.89 | 13.93 | 14.02 | 4.93 | 5.03 | 12.5 | 12.46 | 2.73 | 2.76 | 13.37 | 13.46 | 2.83 | 2.87
11.35 | 11.45 | 2.93 | 2.91 | 11.79 | 11.67 | 3.02 | 3.13 | 11.3 | 11.34 | 2.62 | 2.62 | 11.82 | 11.86 | 2.84 | 1.04 | 13.14 | 13.07 | 1.99 | 2 | 13.2 | 13.11 | 1.27 | 1.18 | 13.36 | 13.4 | 4.89 | 4.89 | 14.02 | 13.84 | 5.03 | 5.04 | 12.46 | 12.45 | 2.76 | 2.76 | 13.46 | 13.48 | 2.87 | 2.99
11.45 | 11.38 | 2.91 | 2.89 | 11.67 | 11.7 | 3.13 | 3.03 | 11.34 | 11.32 | 2.62 | 2.63 | 11.86 | 11.82 | 1.04 | 1.03 | 13.07 | 13.05 | 2 | 2.02 | 13.11 | 13.1 | 1.18 | 1.28 | 13.4 | 13.46 | 4.89 | 4.9 | 13.84 | 13.87 | 5.04 | 5.09 | 12.45 | 12.45 | 2.76 | 2.74 | 13.48 | 13.41 | 2.99 | 2.89
11.38 | 11.37 | 2.89 | 2.86 | 11.7 | 11.66 | 3.03 | 3.12 | 11.32 | 11.33 | 2.63 | 2.62 | 11.82 | 11.95 | 1.03 | 2.89 | 13.05 | 13.01 | 2.02 | 1.95 | 13.1 | 13.11 | 1.28 | 1.21 | 13.46 | 13.49 | 4.9 | 4.91 | 13.87 | 13.93 | 5.09 | 4.96 | 12.45 | 12.54 | 2.74 | 2.63 | 13.41 | 13.35 | 2.89 | 2.81
11.37 | 11.37 | 2.86 | 2.86 | 11.66 | 11.55 | 3.12 | 3.14 | 11.33 | 11.32 | 2.62 | 2.63 | 11.95 | 11.89 | 2.89 | 2.94 | 13.01 | 13.1 | 1.95 | 1.96 | 13.11 | 13.12 | 1.21 | 1.24 | 13.49 | 13.5 | 4.91 | 4.93 | 13.93 | 13.94 | 4.96 | 5 | 12.54 | 12.55 | 2.63 | 2.66 | 13.35 | 13.67 | 2.81 | 2.82
11.37 | 11.39 | 2.86 | 2.91 | 11.55 | 11.7 | 3.14 | 3.03 | 11.32 | 11.34 | 2.63 | 2.63 | 11.89 | 11.84 | 2.94 | 2.85 | 13.1 | 13.19 | 1.96 | 2 | 13.12 | 12.98 | 1.24 | 1.19 | 13.5 | 13.51 | 4.93 | 4.88 | 13.94 | 13.9 | 5 | 4.94 | 12.55 | 12.58 | 2.66 | 2.69 | 13.67 | 13.41 | 2.82 | 2.84
11.39 | 11.36 | 2.91 | 2.9 | 11.7 | 11.69 | 3.03 | 3.11 | 11.34 | 11.28 | 2.63 | 2.66 | 11.84 | 11.68 | 2.85 | 0.96 | 13.19 | 13.23 | 2 | 2 | 12.98 | 13.09 | 1.19 | 1.25 | 13.51 | 13.55 | 4.88 | 4.86 | 13.9 | 13.9 | 4.94 | 5.07 | 12.58 | 12.57 | 2.69 | 2.69 | 13.41 | 13.59 | 2.84 | 2.95
11.36 | 11.3 | 2.9 | 2.84 | 11.69 | 11.65 | 3.11 | 3.15 | 11.28 | 11.29 | 2.66 | 2.67 | 11.68 | 11.82 | 0.96 | 2.88 | 13.23 | 13.2 | 2 | 1.99 | 13.09 | 13.12 | 1.25 | 1.19 | 13.55 | 13.56 | 4.86 | 4.86 | 13.9 | 13.8 | 5.07 | 5.06 | 12.57 | 12.62 | 2.69 | 2.68 | 13.59 | 13.31 | 2.95 | 2.92
11.3 | 11.31 | 2.84 | 2.87 | 11.65 | 11.56 | 3.15 | 3.08 | 11.29 | 11.27 | 2.67 | 2.65 | 11.82 | 11.78 | 2.88 | 2.95 | 13.2 | 13.27 | 1.99 | 1.94 | 13.12 | 13.2 | 1.19 | 1.22 | 13.56 | 13.51 | 4.86 | 4.84 | 13.8 | 13.97 | 5.06 | 5.07 | 12.62 | 12.56 | 2.68 | 2.71 | 13.31 | 13.45 | 2.92 | 2.95
11.31 | 11.24 | 2.87 | 2.88 | 11.56 | 11.62 | 3.08 | 3.13 | 11.27 | 11.16 | 2.65 | 2.66 | 11.78 | 11.87 | 2.95 | 3.01 | 13.27 | 13.18 | 1.94 | 1.94 | 13.2 | 13 | 1.22 | 1.17 | 13.51 | 13.47 | 4.84 | 4.88 | 13.97 | 13.78 | 5.07 | 4.96 | 12.56 | 12.59 | 2.71 | 2.73 | 13.45 | 13.51 | 2.95 | 2.85
11.24 | 11.26 | 2.88 | 2.89 | 11.62 | 11.73 | 3.13 | 3.28 | 11.16 | 11.12 | 2.66 | 2.66 | 11.85 | 11.82 | 2.94 | 2.88 | 13.18 | 13.19 | 1.94 | 1.98 | 13 | 13.07 | 1.17 | 1.23 | 13.47 | 13.54 | 4.88 | 4.91 | 13.78 | 13.8 | 4.96 | 5.02 | 12.59 | 12.57 | 2.73 | 2.73 | 13.51 | 13.53 | 2.85 | 2.8
11.26 | 11.28 | 2.89 | 2.89 | 11.73 | 11.65 | 3.28 | 3.16 | 11.12 | 11.15 | 2.66 | 2.66 | 11.82 | 11.72 | 2.88 | 2.86 | 13.19 | 13.14 | 1.98 | 2.03 | 13.19 | 13.16 | 1.09 | 1.03 | 13.48 | 13.52 | 4.89 | 4.87 | 13.8 | 13.89 | 5.02 | 4.94 | 12.57 | 12.5 | 2.73 | 2.72 | 13.53 | 13.49 | 2.8 | 2.79
11.28 | 11.3 | 2.89 | 2.86 | 11.65 | 11.61 | 3.16 | 3.12 | 11.15 | 11.17 | 2.66 | 2.68 | 11.72 | 11.86 | 2.86 | 2.98 | 13.14 | 13.03 | 2.03 | 2.01 | 13.16 | 13.27 | 1.03 | 1.13 | 13.52 | 13.52 | 4.87 | 4.87 | 13.89 | 12.51 | 4.94 | 5.06 | 12.5 | 12.45 | 2.72 | 2.62 | 13.49 | 13.47 | 2.79 | 2.87
11.07 | 11.11 | 2.71 | 2.68 | 11.29 | 11.36 | 2.94 | 3.08 | 11.18 | 11.16 | 2.58 | 2.6 | 11.86 | 11.81 | 2.98 | 2.95 | 12.93 | 12.87 | 1.81 | 1.85 | 13.27 | 13.23 | 1.13 | 1.07 | 13.52 | 13.52 | 4.87 | 4.89 | 14.02 | 14.16 | 4.93 | 4.64 | 12.52 | 12.45 | 2.55 | 2.57 | 13.47 | 13.54 | 2.87 | 2.89
11.11 | 11.12 | 2.68 | 2.67 | 11.36 | 11.19 | 3.08 | 2.96 | 11.16 | 11.22 | 2.6 | 2.64 | 11.81 | 11.74 | 2.95 | 2.97 | 12.87 | 12.89 | 1.85 | 1.86 | 13.23 | 13.18 | 1.07 | 1.11 | 13.52 | 13.53 | 4.89 | 4.91 | 14.16 | 14.04 | 4.64 | 4.37 | 12.45 | 12.43 | 2.57 | 2.6 | 12.97 | 12.97 | 3.04 | 2.84
11.12 | 11.16 | 2.67 | 2.7 | 11.19 | 11.2 | 2.96 | 2.95 | 11.22 | 11.27 | 2.64 | 2.67 | 11.74 | 11.63 | 2.97 | 1.08 | 12.89 | 12.96 | 1.86 | 1.78 | 13.18 | 13.1 | 1.11 | 1.07 | 13.53 | 13.57 | 4.91 | 4.95 | 14.04 | 13.97 | 4.37 | 4.25 | 12.43 | 12.44 | 2.6 | 2.63 | 12.97 | 12.58 | 2.84 | 3
11.16 | 11.15 | 2.7 | 2.7 | 11.2 | 11.41 | 2.95 | 3.01 | 11.27 | 11.27 | 2.67 | 2.62 | 11.63 | 11.84 | 1.08 | 0.89 | 12.96 | 12.98 | 1.78 | 1.79 | 13.1 | 13.38 | 1.07 | 1.07 | 13.57 | 13.59 | 4.95 | 4.87 | 13.97 | 13.89 | 4.25 | 4.55 | 12.44 | 12.36 | 2.63 | 2.64 | 12.58 | 12 | 3 | 2.9
11.15 | 11.18 | 2.7 | 2.65 | 11.41 | 11.33 | 3.01 | 3.01 | 11.27 | 11.33 | 2.62 | 2.6 | 11.84 | 11.98 | 0.89 | 2.84 | 12.98 | 13.09 | 1.79 | 1.81 | 13.38 | 13.24 | 1.07 | 1.06 | 13.59 | 13.53 | 4.87 | 4.87 | 13.89 | 13.89 | 4.55 | 4.32 | 12.36 | 12.36 | 2.64 | 2.53 | 12 | 12.84 | 2.9 | 2.82
11.18 | 11.19 | 2.65 | 2.69 | 11.33 | 11.34 | 3.01 | 3.29 | 11.33 | 11.33 | 2.6 | 2.6 | 11.98 | 11.87 | 2.84 | 2.88 | 13.09 | 13.09 | 1.81 | 1.82 | 13.24 | 13.14 | 1.06 | 1.09 | 13.53 | 13.43 | 4.87 | 4.88 | 13.89 | 13.96 | 4.32 | 4.43 | 12.36 | 12.31 | 2.53 | 2.54 | 12.84 | 12.79 | 2.82 | 2.87
11.19 | 11.24 | 2.69 | 2.67 | 11.34 | 11.3 | 3.29 | 2.92 | 11.33 | 11.31 | 2.6 | 2.63 | 11.87 | 11.95 | 2.88 | 2.91 | 13.09 | 13.1 | 1.82 | 1.82 | 13.14 | 13.34 | 1.09 | 1.05 | 13.43 | 13.36 | 4.88 | 4.92 | 13.96 | 13.93 | 4.43 | 4.82 | 12.31 | 12.3 | 2.54 | 2.55 | 12.79 | 12.82 | 2.87 | 2.89
11.24 | 11.22 | 2.67 | 2.71 | 11.3 | 11.21 | 2.92 | 3.08 | 11.31 | 11.37 | 2.63 | 2.65 | 11.95 | 12.03 | 2.91 | 2.72 | 13.1 | 13.05 | 1.82 | 1.84 | 13.34 | 13.28 | 1.05 | 1.04 | 13.36 | 13.38 | 4.92 | 4.94 | 13.93 | 13.85 | 4.82 | 4.62 | 12.3 | 12.4 | 2.55 | 2.6 | 12.82 | 12.93 | 2.89 | 2.82
11.22 | 11.23 | 2.71 | 2.71 | 11.21 | 11.43 | 3.08 | 3.09 | 11.37 | 11.39 | 2.65 | 2.67 | 12.03 | 11.86 | 2.72 | 2.91 | 13.05 | 13.02 | 1.84 | 1.77 | 13.28 | 13.45 | 1.04 | 1.04 | 13.38 | 13.47 | 4.94 | 4.87 | 13.85 | 14.01 | 4.62 | 4.38 | 12.4 | 12.44 | 2.6 | 2.61 | 12.93 | 12.88 | 2.82 | 2.9
11.23 | 11.15 | 2.71 | 2.71 | 11.43 | 11.44 | 3.09 | 2.97 | 11.39 | 11.31 | 2.67 | 2.59 | 11.86 | 11.87 | 2.91 | 2.85 | 13.02 | 12.95 | 1.77 | 1.79 | 13.45 | 13.2 | 1.04 | 1.02 | 13.47 | 13.52 | 4.87 | 4.87 | 14.01 | 14 | 4.38 | 4.4 | 12.44 | 12.47 | 2.61 | 2.64 | 12.88 | 12.76 | 2.9 | 2.85
11.15 | 11.14 | 2.71 | 2.68 | 11.44 | 11.37 | 2.97 | 3 | 11.31 | 11.31 | 2.59 | 2.59 | 11.87 | 11.84 | 2.85 | 1 | 12.95 | 12.84 | 1.79 | 1.82 | 13.2 | 13.21 | 1.02 | 1.06 | 13.52 | 13.53 | 4.87 | 4.89 | 14 | 13.9 | 4.4 | 4.61 | 12.47 | 12.46 | 2.64 | 2.58 | 12.76 | 12.86 | 2.85 | 3.03
11.14 | 11.09 | 2.68 | 2.69 | 11.37 | 11.33 | 3 | 2.98 | 11.31 | 11.31 | 2.59 | 2.61 | 11.84 | 12.01 | 1 | 2.84 | 12.84 | 12.88 | 1.82 | 1.82 | 13.21 | 13.33 | 1.06 | 1.04 | 13.53 | 13.55 | 4.89 | 4.91 | 13.9 | 14.05 | 4.61 | 4.61 | 12.46 | 12.52 | 2.58 | 2.54 | 12.86 | 12.92 | 3.03 | 2.92
11.09 | 11.04 | 2.69 | 2.66 | 11.33 | 11.47 | 2.98 | 2.93 | 11.31 | 11.19 | 2.61 | 2.65 | 12.01 | 11.64 | 2.84 | 0.94 | 12.88 | 12.97 | 1.82 | 1.82 | 13.33 | 13.28 | 1.04 | 1.12 | 13.55 | 13.54 | 4.91 | 4.9 | 14.05 | 13.9 | 4.61 | 4.57 | 12.52 | 12.48 | 2.54 | 2.57 | 12.92 | 12.9 | 2.92 | 1.91
11.04 | 11.11 | 2.66 | 2.71 | 11.47 | 11.47 | 2.93 | 3.01 | 11.19 | 11.19 | 2.65 | 2.69 | 11.64 | 11.92 | 0.94 | 2.85 | 12.97 | 13 | 1.82 | 1.77 | 13.28 | 13.25 | 1.12 | 1.05 | 13.54 | 13.6 | 4.9 | 4.87 | 13.9 | 13.82 | 4.57 | 4.6 | 12.48 | 12.45 | 2.57 | 2.62 | 12.9 | 12.73 | 1.91 | 0.59
11.11 | 11.11 | 2.71 | 2.71 | 11.47 | 11.31 | 3.01 | 3.28 | 11.19 | 11.16 | 2.69 | 2.61 | 11.92 | 11.66 | 2.85 | 1.19 | 13 | 13.08 | 1.77 | 1.77 | 13.25 | 13.32 | 1.05 | 1.02 | 13.6 | 13.54 | 4.87 | 4.86 | 13.82 | 13.89 | 4.6 | 4.14 | 12.45 | 12.46 | 2.62 | 2.62 | 12.73 | 12.99 | 0.59 | 2.92
11.11 | 11.15 | 2.71 | 2.67 | 11.31 | 11.35 | 3.28 | 2.98 | 11.16 | 11.17 | 2.61 | 2.59 | 11.66 | 11.87 | 1.19 | 2.95 | 13.08 | 13.07 | 1.77 | 1.79 | 13.32 | 13.08 | 1.02 | 0.2 | 13.54 | 13.5 | 4.86 | 4.89 | 13.89 | 13.96 | 4.14 | 4.68 | 12.46 | 12.42 | 2.62 | 2.52 | 12.99 | 12.94 | 2.92 | 2.98
11.15 | 11.16 | 2.67 | 2.67 | 11.35 | 11.47 | 2.98 | 3.09 | 11.17 | 11.21 | 2.59 | 2.61 | 11.87 | 11.84 | 2.95 | 2.89 | 13.07 | 13.06 | 1.79 | 1.82 | 13.08 | 13.32 | 0.2 | 1.06 | 13.5 | 13.41 | 4.89 | 4.89 | 13.96 | 13.88 | 4.68 | 4.68 | 12.42 | 12.29 | 2.52 | 2.54 | 12.94 | 13.01 | 2.98 | 2.92
11.16 | 11.14 | 2.67 | 2.7 | 11.47 | 11.35 | 3.09 | 3.22 | 11.21 | 11.26 | 2.61 | 2.63 | 11.84 | 11.8 | 2.89 | 3.07 | 13.06 | 13.15 | 1.82 | 1.84 | 13.32 | 13.34 | 1.06 | 1.09 | 13.41 | 13.39 | 4.89 | 4.93 | 13.88 | 13.91 | 4.68 | 4.67 | 12.29 | 12.3 | 2.54 | 2.61 | 13.01 | 12.98 | 2.92 | 3.01
11.14 | 11.2 | 2.7 | 2.72 | 11.35 | 11.4 | 3.22 | 2.96 | 11.26 | 11.29 | 2.63 | 2.65 | 11.8 | 11.85 | 3.07 | 3.02 | 13.15 | 13.01 | 1.84 | 1.81 | 13.34 | 13.24 | 1.09 | 1.12 | 13.39 | 13.44 | 4.93 | 4.91 | 13.91 | 13.94 | 4.67 | 4.46 | 12.3 | 12.39 | 2.61 | 2.57 | 12.98 | 12.96 | 3.01 | 3.07
11.2 | 11.22 | 2.72 | 2.84 | 11.4 | 11.31 | 2.96 | 3.02 | 11.29 | 11.33 | 2.65 | 2.67 | 11.85 | 11.86 | 3.02 | 2.91 | 13.01 | 12.92 | 1.81 | 1.77 | 13.24 | 13.29 | 1.12 | 1.03 | 13.44 | 13.51 | 4.91 | 4.84 | 13.94 | 14.05 | 4.46 | 4.43 | 12.39 | 12.4 | 2.57 | 2.55 | 12.96 | 12.6 | 3.07 | 3
11.22 | 11.28 | 2.84 | 2.72 | 11.31 | 11.34 | 3.02 | 2.93 | 11.33 | 11.32 | 2.67 | 2.62 | 11.86 | 10.49 | 2.91 | 2 | 12.92 | 12.84 | 1.77 | 1.79 | 13.29 | 13.2 | 1.03 | 1.11 | 13.51 | 13.49 | 4.84 | 4.89 | 14.05 | 13.97 | 4.43 | 4.56 | 12.4 | 12.43 | 2.55 | 2.57 | 12.6 | 12.85 | 3 | 2.72
11.28 | 11.19 | 2.72 | 2.71 | 11.34 | 11.38 | 2.93 | 3.1 | 11.32 | 11.34 | 2.62 | 2.58 | 10.49 | 11.83 | 2 | 2.68 | 12.84 | 12.88 | 1.79 | 1.82 | 13.2 | 13.41 | 1.11 | 1.02 | 13.49 | 13.52 | 4.89 | 4.88 | 13.97 | 14.16 | 4.56 | 4.59 | 12.43 | 12.47 | 2.57 | 2.61 | 12.85 | 12.7 | 2.72 | 2.99
11.19 | 11.21 | 2.71 | 2.74 | 11.38 | 11.26 | 3.1 | 2.99 | 11.34 | 11.31 | 2.58 | 2.61 | 11.83 | 11.93 | 2.68 | 3.02 | 12.88 | 12.91 | 1.82 | 1.82 | 13.41 | 13.33 | 1.02 | 1.09 | 13.52 | 13.51 | 4.88 | 4.94 | 14.16 | 13.98 | 4.59 | 4.68 | 12.47 | 12.48 | 2.61 | 2.54 | 12.7 | 12.95 | 2.99 | 2.97
11.21 | 11.19 | 2.74 | 2.71 | 11.26 | 10.96 | 2.99 | 3.04 | 11.31 | 11.36 | 2.61 | 2.65 | 11.93 | 11.86 | 3.02 | 2.89 | 12.91 | 13.05 | 1.82 | 1.83 | 13.33 | 13.33 | 1.09 | 1.05 | 13.51 | 13.55 | 4.94 | 4.94 | 13.98 | 14.28 | 4.68 | 4.52 | 12.48 | 12.49 | 2.54 | 2.59 | 12.95 | 13.18 | 2.97 | 2.88
11.19 | 11.14 | 2.71 | 2.69 | 10.96 | 11.21 | 3.04 | 3.13 | 11.36 | 11.3 | 2.65 | 2.66 | 11.86 | 11.8 | 2.89 | 3.01 | 13.05 | 13.04 | 1.83 | 1.77 | 13.33 | 13.19 | 1.05 | 1.09 | 13.55 | 13.55 | 4.94 | 4.86 | 14.28 | 14.3 | 4.52 | 4.41 | 12.49 | 12.56 | 2.59 | 2.61 | 13.18 | 13.12 | 2.88 | 2.94
11.14 | 11.14 | 2.69 | 2.72 | 11.21 | 11.49 | 3.13 | 2.99 | 11.3 | 11.33 | 2.66 | 2.62 | 11.8 | 11.84 | 3.01 | 2.86 | 13.04 | 13.07 | 1.77 | 1.78 | 13.19 | 13.19 | 1.09 | 1.07 | 13.55 | 13.54 | 4.86 | 4.87 | 14.3 | 14.12 | 4.41 | 4.81 | 12.56 | 12.45 | 2.61 | 2.59 | 13.12 | 13.11 | 2.94 | 2.93
11.14 | 11.11 | 2.72 | 2.68 | 11.49 | 11.29 | 2.99 | 2.97 | 11.33 | 11.23 | 2.62 | 2.58 | 11.84 | 11.79 | 2.86 | 2.88 | 13.07 | 13.1 | 1.78 | 1.8 | 13.19 | 13.06 | 1.07 | 1.13 | 13.54 | 13.48 | 4.87 | 4.92 | 14.12 | 14.07 | 4.81 | 4.56 | 12.45 | 12.47 | 2.59 | 2.57 | 13.11 | 13.21 | 2.93 | 2.8
11.11 | 11.12 | 2.68 | 2.65 | 11.29 | 11.27 | 2.97 | 2.98 | 11.23 | 11.26 | 2.58 | 2.59 | 11.79 | 11.82 | 2.88 | 2.87 | 13.1 | 13.06 | 1.8 | 1.81 | 13.06 | 13.29 | 1.13 | 1.07 | 13.48 | 13.45 | 4.92 | 4.93 | 14.07 | 14.07 | 4.56 | 4.45 | 12.47 | 12.42 | 2.57 | 2.62 | 13.21 | 13.13 | 2.8 | 2.96
11.12 | 11.13 | 2.65 | 2.73 | 11.27 | 11.34 | 2.98 | 2.99 | 11.26 | 11.18 | 2.59 | 2.62 | 11.82 | 11.89 | 2.87 | 2.88 | 13.06 | 13.04 | 1.81 | 1.81 | 13.29 | 13.38 | 1.07 | 1.08 | 13.45 | 13.36 | 4.93 | 4.88 | 14.07 | 14.09 | 4.45 | 4.44 | 12.42 | 12.35 | 2.62 | 2.59 | 13.13 | 12.97 | 2.96 | 2.95
11.13 | 11.17 | 2.73 | 2.7 | 11.34 | 11.35 | 2.99 | 2.94 | 11.18 | 11.13 | 2.62 | 2.64 | 11.89 | 11.69 | 2.88 | 2.84 | 13.04 | 13 | 1.81 | 1.78 | 13.38 | 13.34 | 1.08 | 1.24 | 13.36 | 13.44 | 4.88 | 4.89 | 14.09 | 14.09 | 4.44 | 4.33 | 12.35 | 12.28 | 2.59 | 2.55 | 12.97 | 13.03 | 2.95 | 2.95
11.17 | 11.21 | 2.7 | 2.67 | 11.35 | 11.23 | 2.94 | 2.99 | 11.13 | 11.21 | 2.64 | 2.66 | 11.69 | 11.66 | 2.84 | 2.49 | 13 | 12.81 | 1.78 | 1.78 | 13.34 | 13.31 | 1.24 | 0.98 | 13.44 | 13.47 | 4.89 | 4.9 | 14.09 | 14.04 | 4.33 | 4.67 | 12.28 | 12.4 | 2.55 | 2.6 | 13.03 | 12.92 | 2.95 | 2.96
11.21 | 11.25 | 2.67 | 2.73 | 11.23 | 11.31 | 2.99 | 3.01 | 11.21 | 11.24 | 2.66 | 2.58 | 11.84 | 11.81 | 2.85 | 2.93 | 12.81 | 12.88 | 1.78 | 1.8 | 13.31 | 13.27 | 0.98 | 1.11 | 13.47 | 13.51 | 4.9 | 4.94 | 14.04 | 14.14 | 4.67 | 4.82 | 12.4 | 12.39 | 2.6 | 2.55 | 12.92 | 13.08 | 2.96 | 2.76
11.25 | 11.21 | 2.73 | 2.71 | 11.31 | 11.17 | 3.01 | 2.92 | 11.24 | 11.78 | 2.58 | 2.69 | 11.81 | 11.64 | 2.93 | 2.85 | 12.88 | 12.94 | 1.8 | 1.79 | 13.27 | 13.31 | 1.11 | 1.06 | 13.51 | 13.55 | 4.94 | 4.88 | 14.14 | 14.14 | 4.82 | 4.68 | 12.39 | 12.44 | 2.55 | 2.57 | 13.08 | 12.98 | 2.76 | 2.91
11.21 | 11.19 | 2.71 | 2.72 | 11.17 | 11.39 | 2.92 | 3.05 | 11.78 | 11.08 | 2.69 | 2.62 | 11.64 | 11.78 | 2.85 | 2.94 | 12.94 | 12.97 | 1.79 | 1.83 | 13.31 | 13.19 | 1.06 | 1.13 | 13.55 | 13.55 | 4.88 | 4.83 | 14.14 | 14.09 | 4.68 | 4.57 | 12.44 | 12.51 | 2.57 | 2.6 | 12.98 | 12.93 | 2.91 | 2.9
11.19 | 11.2 | 2.72 | 2.7 | 11.39 | 11.32 | 3.05 | 2.98 | 11.08 | 11.04 | 2.62 | 2.65 | 11.78 | 11.92 | 2.94 | 2.93 | 12.97 | 13.02 | 1.83 | 1.84 | 13.19 | 13.3 | 1.13 | 1.07 | 13.55 | 13.52 | 4.83 | 4.85 | 14.09 | 14.01 | 4.57 | 4.44 | 12.51 | 12.48 | 2.6 | 2.56 | 12.93 | 13.12 | 2.9 | 2.74
11.2 | 11.21 | 2.7 | 2.66 | 11.32 | 11.3 | 2.98 | 2.96 | 11.04 | 11.03 | 2.65 | 2.65 | 11.92 | 11.74 | 2.93 | 2.85 | 13.02 | 13.08 | 1.84 | 1.76 | 12.87 | 12.92 | 1.03 | 0.99 | 13.52 | 13.59 | 4.85 | 4.88 | 14.01 | 14.06 | 4.44 | 4.48 | 12.48 | 12.49 | 2.56 | 2.56 | 13.12 | 13.21 | 2.74 | 2.94
11.21 | 11.08 | 2.66 | 2.71 | 11.3 | 11.36 | 2.96 | 3.05 | 11.03 | 11.09 | 2.65 | 2.68 | 11.74 | 11.75 | 2.85 | 2.89 | 13.08 | 13.05 | 1.76 | 1.78 | 12.92 | 12.88 | 0.99 | 1.01 | 13.54 | 13.54 | 4.77 | 4.82 | 14.06 | 13.91 | 4.48 | 4.53 | 12.49 | 12.51 | 2.56 | 2.59 | 13.21 | 13.13 | 2.94 | 2.99
11.08 | 11.11 | 2.71 | 2.66 | 11.36 | 11.22 | 3.05 | 2.95 | 11.13 | 11.06 | 2.46 | 2.47 | 11.75 | 11.65 | 2.89 | 1.01 | 13.05 | 12.97 | 1.74 | 1.76 | 12.88 | 12.8 | 1.01 | 0.98 | 13.54 | 13.48 | 4.82 | 4.8 | 13.6 | 13.94 | 3.51 | 5.04 | 12.51 | 12.44 | 2.59 | 2.64 | 13.13 | 13.11 | 2.99 | 2.85
11.15 | 11.1 | 2.69 | 2.68 | 9.86 | 9.91 | 2.49 | 2.58 | 11.06 | 11.01 | 2.47 | 2.51 | 11.65 | 11.8 | 1.01 | 2.77 | 12.97 | 12.87 | 1.76 | 1.76 | 12.8 | 12.88 | 0.98 | 1.1 | 13.48 | 13.41 | 4.8 | 4.76 | 13.94 | 14.01 | 5.04 | 4.57 | 12.3 | 12.26 | 2.44 | 2.46 | 13.1 | 12.98 | 2.63 | 2.62
11.1 | 11.07 | 2.68 | 2.68 | 9.91 | 9.97 | 2.58 | 2.82 | 11.01 | 10.95 | 2.51 | 2.54 | 11.8 | 11.77 | 2.77 | 1.66 | 12.87 | 12.88 | 1.76 | 1.78 | 12.88 | 13.08 | 1.1 | 1.02 | 13.41 | 13.36 | 4.76 | 4.75 | 14.01 | 14.15 | 4.57 | 5.35 | 12.26 | 12.29 | 2.46 | 2.51 | 12.98 | 12.97 | 2.62 | 2.61
11.07 | 11.13 | 2.68 | 2.7 | 9.97 | 9.62 | 2.82 | 2.51 | 10.95 | 10.99 | 2.54 | 2.48 | 11.77 | 11.72 | 1.66 | 2.88 | 12.88 | 12.88 | 1.78 | 1.78 | 13.08 | 13.01 | 1.02 | 1 | 13.36 | 13.4 | 4.75 | 4.82 | 14.15 | 14.02 | 5.35 | 4.93 | 12.29 | 12.36 | 2.51 | 2.54 | 12.97 | 9.8 | 2.61 | 2.38
11.13 | 11.15 | 2.7 | 2.62 | 9.62 | 9.53 | 2.51 | 2.62 | 10.99 | 11.11 | 2.48 | 2.45 | 11.72 | 11.56 | 2.88 | 1.44 | 12.88 | 12.96 | 1.78 | 1.8 | 13.01 | 13.08 | 1 | 1.04 | 13.4 | 13.44 | 4.82 | 4.79 | 14.02 | 14.06 | 4.93 | 4.74 | 12.36 | 12.43 | 2.54 | 2.43 | 9.8 | 12.9 | 2.38 | 2.47
11.15 | 11.16 | 2.62 | 2.66 | 9.53 | 9.62 | 2.62 | 2.61 | 11.11 | 11.03 | 2.45 | 2.47 | 11.56 | 11.76 | 1.44 | 2.85 | 12.96 | 13.01 | 1.8 | 1.71 | 13.08 | 12.99 | 1.04 | 1.01 | 13.44 | 13.5 | 4.79 | 4.76 | 14.06 | 14.08 | 4.74 | 5.34 | 12.43 | 12.38 | 2.43 | 2.48 | 12.9 | 12.94 | 2.47 | 2.65
11.16 | 11.19 | 2.66 | 2.7 | 9.62 | 9.52 | 2.61 | 2.71 | 11.03 | 11.09 | 2.47 | 2.52 | 11.76 | 11.72 | 2.85 | 2.75 | 13.01 | 13.04 | 1.71 | 1.73 | 12.99 | 12.83 | 1.01 | 0.96 | 13.5 | 13.51 | 4.76 | 4.81 | 14.08 | 13.94 | 5.34 | 5.09 | 12.38 | 12.44 | 2.48 | 2.51 | 12.94 | 12.93 | 2.65 | 2.66
11.19 | 11.2 | 2.7 | 2.66 | 9.52 | 9.52 | 2.71 | 2.58 | 11.09 | 11.15 | 2.52 | 2.55 | 11.72 | 11.68 | 2.75 | 2.79 | 13.04 | 13.05 | 1.73 | 1.75 | 12.83 | 13.04 | 0.96 | 1.01 | 13.51 | 13.51 | 4.81 | 4.83 | 13.94 | 14.11 | 5.09 | 4.73 | 12.44 | 12.44 | 2.51 | 2.52 | 12.93 | 12.97 | 2.66 | 2.67
11.2 | 11.2 | 2.66 | 2.72 | 9.52 | 9.86 | 2.58 | 2.73 | 11.15 | 11.15 | 2.55 | 2.49 | 11.68 | 11.79 | 2.79 | 2.8 | 13.05 | 13.1 | 1.75 | 1.76 | 13.04 | 12.75 | 1.01 | 0.99 | 13.51 | 13.55 | 4.83 | 4.75 | 14.11 | 14.02 | 4.73 | 5.08 | 12.44 | 12.4 | 2.52 | 2.42 | 12.97 | 13.02 | 2.67 | 2.72
11.2 | 11.24 | 2.72 | 2.66 | 9.86 | 9.85 | 2.73 | 2.68 | 11.15 | 11.21 | 2.49 | 2.46 | 11.79 | 11.71 | 2.8 | 2.86 | 13.1 | 13.04 | 1.76 | 1.78 | 12.75 | 12.87 | 0.99 | 1.05 | 13.55 | 13.47 | 4.75 | 4.74 | 14.02 | 14.1 | 5.08 | 5.38 | 12.4 | 12.36 | 2.42 | 2.43 | 13.02 | 12.95 | 2.72 | 2.71
11.24 | 11.18 | 2.66 | 2.69 | 9.85 | 9.63 | 2.68 | 2.65 | 11.21 | 11.15 | 2.46 | 2.46 | 11.71 | 11.61 | 2.86 | 2.04 | 13.04 | 13.01 | 1.78 | 1.77 | 12.87 | 11.56 | 1.05 | 1.02 | 13.47 | 13.46 | 4.74 | 4.81 | 14.1 | 14.11 | 5.38 | 5.24 | 12.36 | 12.34 | 2.43 | 2.49 | 12.95 | 12.97 | 2.71 | 2.71
11.18 | 11.2 | 2.69 | 2.7 | 9.63 | 9.62 | 2.65 | 2.57 | 11.15 | 11.14 | 2.46 | 2.53 | 11.61 | 11.65 | 2.04 | 2.95 | 13.01 | 12.93 | 1.77 | 1.79 | 11.56 | 12.83 | 1.02 | 0.99 | 13.46 | 13.45 | 4.81 | 4.75 | 14.11 | 14.07 | 5.24 | 5.03 | 12.34 | 12.2 | 2.49 | 2.52 | 12.97 | 12.86 | 2.71 | 2.75
11.2 | 11.2 | 2.7 | 2.65 | 9.62 | 9.5 | 2.57 | 2.69 | 11.14 | 11.14 | 2.53 | 2.5 | 11.65 | 11.61 | 2.95 | 2.87 | 12.93 | 12.81 | 1.79 | 1.78 | 12.83 | 12.28 | 0.99 | 1.01 | 13.45 | 13.32 | 4.75 | 4.76 | 14.07 | 14.13 | 5.03 | 4.78 | 12.2 | 12.25 | 2.52 | 2.44 | 12.86 | 12.87 | 2.75 | 2.72
11.2 | 11.18 | 2.65 | 2.72 | 9.5 | 9.4 | 2.69 | 2.52 | 11.14 | 11.02 | 2.5 | 2.45 | 11.61 | 11.64 | 2.87 | 2.93 | 12.81 | 12.83 | 1.78 | 1.76 | 12.28 | 12.92 | 1.01 | 0.97 | 13.32 | 13.36 | 4.76 | 4.76 | 14.13 | 14.07 | 4.78 | 5.31 | 12.25 | 12.27 | 2.44 | 2.43 | 12.87 | 12.96 | 2.72 | 2.75
11.18 | 11.17 | 2.72 | 2.68 | 9.4 | 9.62 | 2.52 | 2.54 | 11.02 | 10.99 | 2.45 | 2.46 | 11.64 | 11.69 | 2.93 | 2.94 | 12.83 | 12.9 | 1.76 | 1.75 | 12.92 | 13.05 | 0.97 | 1.04 | 13.36 | 13.43 | 4.76 | 4.81 | 14.07 | 14.02 | 5.31 | 5.11 | 12.27 | 12.33 | 2.43 | 2.48 | 12.96 | 13.1 | 2.75 | 2.76
11.17 | 11.07 | 2.68 | 2.71 | 9.62 | 9.83 | 2.54 | 2.62 | 10.99 | 11.03 | 2.46 | 2.48 | 11.69 | 11.63 | 2.94 | 2.86 | 12.9 | 12.96 | 1.75 | 1.77 | 13.05 | 12.93 | 1.04 | 0.99 | 13.43 | 13.43 | 4.81 | 4.75 | 14.02 | 14.18 | 5.11 | 5.21 | 12.33 | 12.38 | 2.48 | 2.49 | 13.1 | 12.83 | 2.76 | 2.72
11.07 | 11.1 | 2.71 | 2.65 | 9.83 | 10.5 | 2.62 | 2.54 | 11.03 | 11.07 | 2.48 | 2.52 | 11.63 | 11.56 | 2.86 | 1.95 | 12.96 | 13 | 1.77 | 1.76 | 12.93 | 13.03 | 0.99 | 0.96 | 13.43 | 13.48 | 4.75 | 4.74 | 14.18 | 14.04 | 5.21 | 5.38 | 12.38 | 12.43 | 2.49 | 2.41 | 12.83 | 13 | 2.72 | 2.64
11.1 | 11.11 | 2.65 | 2.64 | 10.5 | 10.02 | 2.54 | 2.72 | 11.07 | 11.15 | 2.52 | 2.54 | 11.56 | 11.64 | 1.95 | 2.88 | 13 | 13.06 | 1.76 | 1.79 | 13.03 | 13.28 | 0.96 | 1 | 13.48 | 13.52 | 4.74 | 4.77 | 14.04 | 14.16 | 5.38 | 5.23 | 12.43 | 12.38 | 2.41 | 2.43 | 13 | 12.96 | 2.64 | 2.62
11.11 | 11.16 | 2.64 | 2.69 | 10.02 | 9.76 | 2.72 | 2.62 | 11.15 | 11.11 | 2.54 | 2.47 | 11.64 | 11.81 | 2.88 | 2.83 | 13.06 | 13.12 | 1.79 | 1.78 | 13.28 | 13.08 | 1 | 0.96 | 13.52 | 13.49 | 4.77 | 4.79 | 14.16 | 14.04 | 5.23 | 4.81 | 12.38 | 12.39 | 2.43 | 2.44 | 12.96 | 13.05 | 2.62 | 2.61
11.16 | 11.19 | 2.69 | 2.68 | 9.76 | 9.69 | 2.62 | 2.63 | 11.11 | 11.13 | 2.47 | 2.46 | 11.81 | 11.63 | 2.83 | 2.9 | 13.12 | 13.04 | 1.78 | 1.78 | 13.08 | 13.05 | 0.96 | 1.03 | 13.49 | 13.48 | 4.79 | 4.75 | 14.04 | 14.09 | 4.81 | 4.68 | 12.39 | 12.4 | 2.44 | 2.51 | 13.05 | 13.12 | 2.61 | 2.64
11.19 | 11.22 | 2.68 | 2.73 | 9.69 | 9.8 | 2.63 | 2.55 | 11.13 | 11.13 | 2.46 | 2.45 | 11.63 | 11.58 | 2.9 | 2.87 | 13.04 | 13.02 | 1.78 | 1.75 | 13.05 | 13.1 | 1.03 | 1 | 13.48 | 13.51 | 4.75 | 4.75 | 14.09 | 13.89 | 4.68 | 5.44 | 12.4 | 12.36 | 2.51 | 2.42 | 13.12 | 13.15 | 2.64 | 2.63
11.22 | 11.25 | 2.73 | 2.67 | 9.8 | 9.71 | 2.55 | 2.53 | 11.13 | 11.17 | 2.45 | 2.5 | 11.58 | 11.61 | 2.87 | 2.95 | 13.02 | 13.04 | 1.75 | 1.73 | 13.1 | 13.09 | 1 | 1.01 | 13.51 | 13.48 | 4.75 | 4.82 | 13.89 | 14.13 | 5.44 | 5.19 | 12.36 | 12.23 | 2.42 | 2.42 | 13.15 | 13.07 | 2.63 | 2.64
11.25 | 11.22 | 2.67 | 2.7 | 9.71 | 9.86 | 2.53 | 2.66 | 11.17 | 11.12 | 2.5 | 2.54 | 11.61 | 11.74 | 2.95 | 2.86 | 13.04 | 12.88 | 1.73 | 1.75 | 13.09 | 13.01 | 1.01 | 1.07 | 13.48 | 13.36 | 4.82 | 4.81 | 14.13 | 14.01 | 5.19 | 5.45 | 12.23 | 12.2 | 2.42 | 2.5 | 13.07 | 12.95 | 2.64 | 2.63
11.22 | 11.22 | 2.7 | 2.68 | 9.86 | 9.63 | 2.66 | 2.52 | 11.12 | 11.13 | 2.54 | 2.45 | 11.74 | 11.79 | 2.86 | 2.81 | 12.88 | 12.84 | 1.75 | 1.78 | 13.01 | 13.07 | 1.07 | 0.98 | 13.36 | 13.34 | 4.81 | 4.77 | 14.01 | 14.08 | 5.45 | 5.27 | 12.2 | 12.23 | 2.5 | 2.48 | 12.95 | 12.84 | 2.63 | 2.73
11.22 | 11.2 | 2.68 | 2.65 | 9.63 | 9.83 | 2.52 | 2.7 | 11.13 | 11.05 | 2.45 | 2.41 | 11.79 | 11.72 | 2.81 | 2.87 | 12.84 | 12.93 | 1.78 | 1.8 | 13.07 | 13.12 | 0.98 | 1.04 | 13.34 | 13.38 | 4.77 | 4.73 | 14.08 | 13.96 | 5.27 | 4.87 | 12.23 | 12.31 | 2.48 | 2.4 | 12.84 | 12.92 | 2.73 | 2.72
11.2 | 11.23 | 2.65 | 2.72 | 9.83 | 9.71 | 2.7 | 2.64 | 11.05 | 10.93 | 2.41 | 2.47 | 11.72 | 11.77 | 2.87 | 3.07 | 12.93 | 12.95 | 1.8 | 1.8 | 13.12 | 13.16 | 1.04 | 1.01 | 13.38 | 13.48 | 4.73 | 4.82 | 13.96 | 14 | 4.87 | 5.44 | 12.31 | 12.55 | 2.4 | 2.47 | 12.92 | 13.11 | 2.72 | 2.73
11.23 | 11.15 | 2.72 | 2.66 | 9.71 | 9.65 | 2.64 | 2.51 | 10.93 | 11.01 | 2.47 | 2.49 | 11.77 | 11.42 | 3.07 | 1.02 | 12.95 | 12.99 | 1.8 | 1.81 | 13.16 | 13.16 | 1.01 | 0.97 | 13.48 | 13.49 | 4.82 | 4.75 | 14 | 13.83 | 5.44 | 5.25 | 12.55 | 12.05 | 2.47 | 2.49 | 13.11 | 12.89 | 2.73 | 2.7
11.15 | 11.06 | 2.66 | 2.71 | 9.65 | 9.89 | 2.51 | 2.6 | 11.01 | 11.08 | 2.49 | 2.51 | 11.42 | 11.76 | 1.02 | 2.84 | 12.99 | 13.03 | 1.81 | 1.73 | 13.16 | 13.09 | 0.97 | 1.05 | 13.49 | 13.51 | 4.75 | 4.78 | 13.83 | 14.07 | 5.25 | 4.67 | 12.05 | 12.12 | 2.49 | 2.51 | 12.89 | 13.01 | 2.7 | 2.72
11.06 | 11.11 | 2.71 | 2.63 | 9.89 | 9.87 | 2.6 | 2.57 | 11.08 | 11.1 | 2.51 | 2.44 | 11.72 | 11.72 | 2.04 | 2.54 | 13.03 | 13.03 | 1.73 | 1.74 | 13.09 | 12.96 | 1.05 | 1.04 | 13.51 | 13.52 | 4.78 | 4.81 | 14.07 | 14.02 | 4.67 | 4.9 | 12.12 | 12.21 | 2.51 | 2.5 | 13.01 | 12.9 | 2.72 | 2.78
11.11 | 11.1 | 2.63 | 2.68 | 9.87 | 9.84 | 2.57 | 2.64 | 11.1 | 11.14 | 2.44 | 2.44 | 11.72 | 11.84 | 2.54 | 2.62 | 13.03 | 13.08 | 1.74 | 1.75 | 12.96 | 13 | 1.04 | 0.99 | 13.52 | 13.51 | 4.81 | 4.82 | 14.02 | 14 | 4.9 | 5.41 | 12.21 | 12.23 | 2.5 | 2.41 | 12.9 | 12.91 | 2.78 | 2.74
11.1 | 11.14 | 2.68 | 2.63 | 9.84 | 9.76 | 2.64 | 2.48 | 11.14 | 11.11 | 2.44 | 2.48 | 11.84 | 11.82 | 2.62 | 2.61 | 13.08 | 13.05 | 1.75 | 1.8 | 13 | 12.97 | 0.99 | 1.06 | 13.51 | 13.5 | 4.82 | 4.74 | 14 | 14.03 | 5.41 | 5.24 | 12.23 | 12.23 | 2.41 | 2.41 | 12.91 | 12.82 | 2.74 | 2.75
11.14 | 11.14 | 2.63 | 2.65 | 9.76 | 9.85 | 2.48 | 2.55 | 11.11 | 11.15 | 2.48 | 2.52 | 11.82 | 11.9 | 2.61 | 2.63 | 13.05 | 13.01 | 1.8 | 1.81 | 12.97 | 13.04 | 1.06 | 1 | 13.5 | 13.45 | 4.74 | 4.75 | 14.03 | 14.03 | 5.24 | 5.19 | 12.23 | 12.32 | 2.41 | 2.45 | 12.82 | 12.81 | 2.75 | 2.74
11.14 | 11.19 | 2.65 | 2.67 | 9.85 | 10.06 | 2.55 | 2.66 | 11.15 | 11.13 | 2.52 | 2.52 | 11.9 | 11.62 | 2.63 | 2.62 | 13.01 | 12.94 | 1.81 | 1.79 | 13.04 | 13.07 | 1 | 0.99 | 13.45 | 13.43 | 4.75 | 4.8 | 14.03 | 14.01 | 5.19 | 5.44 | 12.32 | 12.32 | 2.45 | 2.42 | 12.81 | 12.84 | 2.74 | 2.78
11.19 | 11.18 | 2.67 | 2.67 | 10.06 | 10.16 | 2.66 | 3.34 | 11.13 | 11.14 | 2.52 | 2.44 | 11.62 | 11.75 | 2.62 | 0.44 | 12.94 | 12.86 | 1.79 | 1.78 | 13.07 | 13 | 0.99 | 1.01 | 13.43 | 13.39 | 4.8 | 4.76 | 14.01 | 14.1 | 5.44 | 5.31 | 12.32 | 12.31 | 2.42 | 2.39 | 12.84 | 12.77 | 2.78 | 2.73
11.18 | 11.23 | 2.67 | 2.71 | 10.16 | 9.93 | 3.34 | 2.74 | 11.14 | 11.08 | 2.44 | 2.46 | 11.75 | 11.77 | 0.44 | 2.65 | 12.86 | 12.85 | 1.78 | 1.75 | 13 | 13.07 | 1.01 | 0.98 | 13.39 | 13.37 | 4.76 | 4.78 | 14.1 | 14.27 | 5.31 | 4.97 | 12.31 | 12.25 | 2.39 | 2.38 | 12.77 | 12.86 | 2.73 | 2.73
11.23 | 11.2 | 2.71 | 2.64 | 9.93 | 9.62 | 2.74 | 2.62 | 11.08 | 11.06 | 2.46 | 2.49 | 11.77 | 11.81 | 2.65 | 2.67 | 12.85 | 12.94 | 1.75 | 1.76 | 13.17 | 13.15 | 0.96 | 0.91 | 13.37 | 13.4 | 4.78 | 4.77 | 14.27 | 13.99 | 4.97 | 5.44 | 12.25 | 12.17 | 2.38 | 2.45 | 12.86 | 12.97 | 2.73 | 2.77
11.2 | 11.16 | 2.64 | 2.65 | 9.62 | 9.82 | 2.62 | 2.64 | 11.06 | 11.02 | 2.49 | 2.52 | 11.81 | 11.72 | 2.67 | 2.66 | 12.94 | 13.01 | 1.76 | 1.78 | 13.15 | 13.29 | 0.91 | 0.99 | 13.34 | 13.32 | 4.67 | 4.65 | 13.99 | 13.96 | 5.44 | 5.22 | 12.17 | 12.14 | 2.45 | 2.48 | 12.97 | 12.67 | 2.77 | 2.74
11.16 | 11.18 | 2.65 | 2.7 | 9.82 | 9.71 | 2.64 | 2.55 | 11.02 | 10.99 | 2.52 | 2.44 | 11.72 | 11.65 | 2.66 | 2.13 | 12.74 | 12.79 | 1.67 | 1.68 | 13.29 | 13.12 | 0.99 | 0.91 | 13.32 | 13.3 | 4.65 | 4.59 | 14.15 | 14.13 | 5.16 | 5 | 12.14 | 12.2 | 2.48 | 2.39 | 12.67 | 12.79 | 2.74 | 2.7
11.18 | 11.12 | 2.7 | 2.66 | 11.61 | 11.3 | 3.01 | 2.87 | 11.05 | 10.95 | 2.33 | 2.34 | 11.65 | 11.72 | 2.13 | 2.61 | 12.79 | 12.7 | 1.68 | 1.67 | 13.12 | 13.2 | 0.91 | 1.02 | 13.3 | 13.17 | 4.59 | 4.59 | 14.13 | 13.86 | 5 | 5.41 | 12.35 | 12.39 | 2.4 | 2.42 | 12.79 | 12.62 | 2.7 | 2.77
11.05 | 11.06 | 2.64 | 2.59 | 11.3 | 11.25 | 2.87 | 2.87 | 10.95 | 11 | 2.34 | 2.38 | 11.72 | 11.8 | 2.61 | 2.52 | 12.7 | 12.75 | 1.67 | 1.67 | 13.2 | 13.17 | 1.02 | 0.91 | 13.17 | 13.23 | 4.59 | 4.65 | 13.86 | 13.95 | 5.41 | 5.23 | 12.39 | 12.36 | 2.42 | 2.32 | 12.67 | 12.25 | 2.48 | 2.53
11.06 | 11.05 | 2.59 | 2.61 | 11.25 | 11.2 | 2.87 | 2.84 | 11 | 10.95 | 2.38 | 2.42 | 11.8 | 11.69 | 2.52 | 2.63 | 12.75 | 12.84 | 1.67 | 1.68 | 13.17 | 13.4 | 0.91 | 0.92 | 13.23 | 13.31 | 4.65 | 4.61 | 13.95 | 13.86 | 5.23 | 5.09 | 12.36 | 12.24 | 2.32 | 2.34 | 12.25 | 12.51 | 2.53 | 2.45
11.05 | 11.06 | 2.61 | 2.63 | 11.2 | 11.41 | 2.84 | 2.74 | 10.95 | 10.94 | 2.42 | 2.36 | 11.69 | 11.56 | 2.63 | 0.73 | 12.84 | 12.89 | 1.68 | 1.66 | 13.4 | 13.4 | 0.92 | 0.95 | 13.31 | 13.32 | 4.61 | 4.61 | 13.86 | 13.96 | 5.09 | 5.45 | 12.24 | 12.23 | 2.34 | 2.35 | 12.51 | 12.55 | 2.45 | 2.53
11.06 | 11.04 | 2.63 | 2.61 | 11.41 | 11.48 | 2.74 | 2.87 | 10.94 | 10.92 | 2.36 | 2.33 | 11.56 | 11.74 | 0.73 | 2.6 | 12.89 | 12.89 | 1.66 | 1.66 | 13.4 | 13.32 | 0.95 | 0.91 | 13.32 | 13.37 | 4.61 | 4.63 | 13.96 | 13.92 | 5.45 | 5.31 | 12.23 | 12.14 | 2.35 | 2.37 | 12.55 | 13.02 | 2.53 | 2.56
11.04 | 10.98 | 2.61 | 2.66 | 11.48 | 11.47 | 2.87 | 2.78 | 10.92 | 10.86 | 2.33 | 2.33 | 11.74 | 11.78 | 2.6 | 0.6 | 12.89 | 12.92 | 1.66 | 1.67 | 13.32 | 13.23 | 0.91 | 1.02 | 13.37 | 13.33 | 4.63 | 4.66 | 13.92 | 13.91 | 5.31 | 5.16 | 12.14 | 12.16 | 2.37 | 2.31 | 13.02 | 13 | 2.56 | 2.55
10.98 | 10.97 | 2.66 | 2.57 | 11.47 | 11.42 | 2.78 | 2.86 | 10.86 | 10.86 | 2.33 | 2.37 | 11.78 | 11.66 | 0.6 | 2.73 | 12.92 | 12.91 | 1.67 | 1.65 | 13.23 | 13.15 | 1.02 | 0.93 | 13.33 | 13.39 | 4.66 | 4.6 | 13.91 | 13.91 | 5.16 | 5.29 | 12.16 | 12.22 | 2.31 | 2.34 | 13 | 12.89 | 2.55 | 2.55
10.97 | 10.95 | 2.57 | 2.62 | 11.42 | 11.41 | 2.86 | 2.91 | 10.86 | 10.9 | 2.37 | 2.37 | 11.66 | 11.68 | 2.73 | 2.6 | 12.91 | 12.9 | 1.65 | 1.64 | 13.15 | 13.12 | 0.93 | 0.89 | 13.39 | 13.3 | 4.6 | 4.58 | 13.91 | 13.92 | 5.29 | 5.33 | 12.22 | 12.3 | 2.34 | 2.38 | 12.89 | 12.92 | 2.55 | 2.55
10.95 | 11.01 | 2.62 | 2.6 | 11.41 | 11.42 | 2.91 | 2.82 | 10.9 | 10.9 | 2.37 | 2.39 | 11.68 | 11.72 | 2.6 | 2.59 | 12.9 | 12.81 | 1.64 | 1.66 | 13.12 | 13.38 | 0.89 | 0.9 | 13.3 | 13.35 | 4.58 | 4.65 | 13.92 | 14 | 5.33 | 5.11 | 12.3 | 12.29 | 2.38 | 2.31 | 12.92 | 12.77 | 2.55 | 2.57
11.01 | 11.03 | 2.6 | 2.6 | 11.42 | 11.38 | 2.82 | 2.78 | 10.9 | 10.96 | 2.39 | 2.32 | 11.72 | 5.31 | 2.59 | 0.34 | 12.81 | 12.69 | 1.66 | 1.67 | 13.38 | 13.26 | 0.9 | 0.89 | 13.35 | 13.25 | 4.65 | 4.62 | 14 | 6.81 | 5.11 | 4.8 | 12.29 | 12.31 | 2.31 | 2.32 | 12.77 | 12.85 | 2.57 | 2.48
11.03 | 11.05 | 2.6 | 2.66 | 11.38 | 11.38 | 2.78 | 2.84 | 10.96 | 10.98 | 2.32 | 2.33 | 5.31 | 11.15 | 0.34 | 1.25 | 12.69 | 12.73 | 1.67 | 1.66 | 13.26 | 13.3 | 0.89 | 0.96 | 13.25 | 13.21 | 4.62 | 4.59 | 6.81 | 13.95 | 4.8 | 3.67 | 12.31 | 12.32 | 2.32 | 2.36 | 12.85 | 12.78 | 2.48 | 2.44
11.05 | 11.09 | 2.66 | 2.56 | 11.38 | 11.41 | 2.84 | 2.78 | 10.98 | 11.02 | 2.33 | 2.32 | 11.15 | 10.19 | 1.25 | 2.66 | 12.73 | 12.83 | 1.66 | 1.68 | 13.3 | 13.19 | 0.96 | 0.87 | 13.21 | 13.25 | 4.59 | 4.59 | 13.95 | 14.1 | 3.67 | 4.52 | 12.32 | 12.3 | 2.36 | 2.32 | 12.78 | 12.79 | 2.44 | 2.52
11.09 | 11.07 | 2.56 | 2.61 | 11.41 | 11.25 | 2.78 | 3.13 | 11.02 | 10.95 | 2.32 | 2.36 | 10.19 | 10.37 | 2.66 | 2.61 | 12.83 | 12.85 | 1.68 | 1.66 | 13.19 | 13.27 | 0.87 | 0.99 | 13.25 | 13.25 | 4.59 | 4.65 | 14.1 | 14.13 | 4.52 | 5.49 | 12.3 | 12.31 | 2.32 | 2.31 | 12.79 | 12.78 | 2.52 | 2.46
11.07 | 11.09 | 2.61 | 2.66 | 11.25 | 11.4 | 3.13 | 2.79 | 10.95 | 10.96 | 2.36 | 2.37 | 10.37 | 11.71 | 2.61 | 1.9 | 12.85 | 12.91 | 1.66 | 1.68 | 13.27 | 11.67 | 0.99 | 0.77 | 13.25 | 13.33 | 4.65 | 4.65 | 14.13 | 13.96 | 5.49 | 5.32 | 12.31 | 12.25 | 2.31 | 2.36 | 12.78 | 13.01 | 2.46 | 2.46
11.09 | 11.07 | 2.66 | 2.61 | 11.4 | 11.36 | 2.79 | 2.89 | 10.96 | 10.93 | 2.37 | 2.38 | 11.71 | 10.56 | 1.9 | 2.61 | 12.91 | 12.93 | 1.68 | 1.69 | 11.67 | 12.94 | 0.77 | 0.98 | 13.33 | 13.39 | 4.65 | 4.6 | 13.96 | 13.98 | 5.32 | 4.73 | 12.25 | 12.18 | 2.36 | 2.37 | 13.01 | 12.86 | 2.46 | 2.47
11.07 | 11.04 | 2.61 | 2.65 | 11.36 | 11.25 | 2.89 | 2.91 | 10.93 | 10.96 | 2.38 | 2.31 | 10.56 | 11.54 | 2.61 | 2.51 | 12.93 | 12.91 | 1.69 | 1.68 | 12.94 | 13.12 | 0.98 | 0.95 | 13.39 | 13.33 | 4.6 | 4.6 | 13.98 | 13.86 | 4.73 | 5.42 | 12.18 | 12.17 | 2.37 | 2.32 | 12.86 | 12.86 | 2.47 | 2.46
11.04 | 11.06 | 2.65 | 2.54 | 11.25 | 11.5 | 2.91 | 2.8 | 10.96 | 10.85 | 2.31 | 2.29 | 11.54 | 11.6 | 2.51 | 2.65 | 12.91 | 12.91 | 1.68 | 1.66 | 13.12 | 13.26 | 0.95 | 0.86 | 13.33 | 13.38 | 4.6 | 4.62 | 13.86 | 13.89 | 5.42 | 5.29 | 12.17 | 12.17 | 2.32 | 2.31 | 12.86 | 12.8 | 2.46 | 2.46
11.06 | 10.96 | 2.54 | 2.62 | 11.5 | 11.47 | 2.8 | 3.17 | 10.85 | 10.83 | 2.29 | 2.31 | 11.6 | 11.63 | 2.65 | 2.6 | 12.91 | 12.8 | 1.66 | 1.64 | 13.26 | 13.26 | 0.86 | 0.94 | 13.38 | 13.35 | 4.62 | 4.61 | 13.89 | 13.99 | 5.29 | 5.44 | 12.17 | 12.25 | 2.31 | 2.38 | 12.8 | 12.84 | 2.46 | 2.51
10.96 | 10.97 | 2.62 | 2.65 | 11.47 | 11.35 | 3.17 | 2.83 | 10.83 | 10.92 | 2.31 | 2.37 | 11.63 | 11.76 | 2.6 | 2.57 | 12.8 | 12.69 | 1.64 | 1.64 | 13.26 | 13.05 | 0.94 | 0.92 | 13.35 | 13.29 | 4.61 | 4.58 | 13.99 | 14.13 | 5.44 | 5.31 | 12.25 | 12.29 | 2.38 | 2.31 | 12.84 | 12.92 | 2.51 | 2.54
10.97 | 10.98 | 2.65 | 2.57 | 11.35 | 11.35 | 2.83 | 2.88 | 10.92 | 10.9 | 2.37 | 2.37 | 11.76 | 11.69 | 2.57 | 1.79 | 12.69 | 12.7 | 1.64 | 1.68 | 13.05 | 13.08 | 0.92 | 0.96 | 13.29 | 13.33 | 4.58 | 4.64 | 14.13 | 14.19 | 5.31 | 5.39 | 12.29 | 12.3 | 2.31 | 2.31 | 12.92 | 12.73 | 2.54 | 2.61
10.98 | 11 | 2.57 | 2.65 | 11.35 | 11.32 | 2.88 | 2.76 | 10.9 | 10.92 | 2.37 | 2.34 | 11.69 | 11.61 | 1.79 | 2.57 | 12.7 | 12.78 | 1.68 | 1.67 | 13.08 | 13.06 | 0.96 | 0.9 | 13.33 | 13.27 | 4.64 | 4.65 | 14.19 | 14.06 | 5.39 | 5.25 | 12.3 | 12.34 | 2.31 | 2.35 | 12.73 | 12.85 | 2.61 | 2.58
11 | 11.04 | 2.65 | 2.56 | 11.32 | 11.25 | 2.76 | 3.24 | 10.92 | 10.99 | 2.34 | 2.3 | 11.61 | 10.18 | 2.57 | 2.53 | 12.78 | 12.86 | 1.67 | 1.64 | 13.06 | 13.03 | 0.9 | 1.08 | 13.27 | 13.19 | 4.65 | 4.58 | 14.06 | 14.01 | 5.25 | 5.31 | 12.34 | 12.33 | 2.35 | 2.39 | 12.85 | 12.88 | 2.58 | 2.53
11.04 | 11.04 | 2.56 | 2.65 | 11.25 | 11.39 | 3.24 | 2.77 | 10.99 | 10.94 | 2.3 | 2.31 | 10.18 | 11.69 | 2.53 | 2.46 | 12.86 | 12.89 | 1.64 | 1.65 | 13.03 | 13.16 | 1.08 | 1.01 | 13.19 | 13.26 | 4.58 | 4.61 | 14.01 | 13.98 | 5.31 | 5.32 | 12.33 | 12.32 | 2.39 | 2.29 | 12.88 | 12.98 | 2.53 | 2.57
11.04 | 11.04 | 2.65 | 2.59 | 11.39 | 11.09 | 2.77 | 2.87 | 10.94 | 10.94 | 2.31 | 2.31 | 11.69 | 10.11 | 2.46 | 0.75 | 12.89 | 12.89 | 1.65 | 1.61 | 13.16 | 13.14 | 1.01 | 0.87 | 13.26 | 13.3 | 4.61 | 4.66 | 13.98 | 13.78 | 5.32 | 5.12 | 12.32 | 12.34 | 2.29 | 2.32 | 12.98 | 12.9 | 2.57 | 2.45
11.04 | 11.05 | 2.59 | 2.59 | 11.09 | 11.24 | 2.87 | 2.79 | 10.94 | 10.97 | 2.31 | 2.35 | 10.11 | 10.7 | 0.75 | 2.47 | 12.89 | 12.9 | 1.61 | 1.62 | 13.14 | 12.98 | 0.87 | 0.97 | 13.3 | 13.38 | 4.66 | 4.59 | 13.78 | 13.91 | 5.12 | 5.37 | 12.34 | 12.27 | 2.32 | 2.36 | 12.9 | 12.92 | 2.45 | 2.43
11.05 | 11.08 | 2.59 | 2.63 | 11.24 | 11.26 | 2.79 | 2.84 | 10.97 | 10.95 | 2.35 | 2.39 | 11.83 | 11.71 | 1.83 | 2.58 | 12.9 | 12.84 | 1.62 | 1.61 | 12.98 | 13.19 | 0.97 | 0.9 | 13.38 | 13.35 | 4.59 | 4.6 | 13.91 | 13.96 | 5.37 | 5.45 | 12.27 | 12.23 | 2.36 | 2.3 | 12.92 | 12.75 | 2.43 | 2.48
11.08 | 11.08 | 2.63 | 2.58 | 11.26 | 11.03 | 2.84 | 2.78 | 10.95 | 10.96 | 2.39 | 2.29 | 11.71 | 11.66 | 2.58 | 2.61 | 12.84 | 12.83 | 1.61 | 1.61 | 13.19 | 13.04 | 0.9 | 0.99 | 13.35 | 13.34 | 4.6 | 4.62 | 13.96 | 14.03 | 5.45 | 5.39 | 12.23 | 12.19 | 2.3 | 2.31 | 12.75 | 12.76 | 2.48 | 2.51
11.08 | 11.06 | 2.58 | 2.64 | 11.03 | 11.11 | 2.78 | 2.96 | 10.96 | 10.94 | 2.29 | 2.3 | 11.66 | 11.69 | 2.61 | 2.55 | 12.83 | 12.74 | 1.61 | 1.61 | 13.04 | 13 | 0.99 | 0.97 | 13.34 | 13.37 | 4.62 | 4.68 | 14.03 | 13.79 | 5.39 | 5.5 | 12.19 | 12.18 | 2.31 | 2.36 | 12.76 | 12.71 | 2.51 | 2.44
11.06 | 11.08 | 2.64 | 2.54 | 11.11 | 11.03 | 2.96 | 2.74 | 10.94 | 10.81 | 2.3 | 2.31 | 11.69 | 11.73 | 2.55 | 2.65 | 12.74 | 12.72 | 1.61 | 1.6 | 13 | 12.93 | 0.97 | 0.89 | 13.37 | 13.35 | 4.68 | 4.63 | 13.79 | 13.98 | 5.5 | 5.47 | 12.18 | 12.21 | 2.36 | 2.36 | 12.71 | 12.69 | 2.44 | 2.54
11.08 | 10.97 | 2.54 | 2.61 | 11.03 | 11.22 | 2.74 | 2.87 | 10.81 | 10.84 | 2.31 | 2.35 | 11.73 | 11.85 | 2.65 | 2.56 | 12.72 | 12.76 | 1.6 | 1.64 | 12.93 | 13.19 | 0.89 | 0.97 | 13.35 | 13.34 | 4.63 | 4.6 | 13.98 | 14 | 5.47 | 5.4 | 12.21 | 12.27 | 2.36 | 2.27 | 12.69 | 12.71 | 2.54 | 2.45
10.97 | 10.97 | 2.61 | 2.57 | 11.22 | 11.44 | 2.87 | 2.78 | 10.84 | 10.89 | 2.35 | 2.38 | 11.85 | 11.7 | 2.56 | 0.63 | 12.76 | 12.86 | 1.64 | 1.62 | 13.19 | 13.34 | 0.97 | 0.9 | 13.34 | 13.27 | 4.6 | 4.6 | 14 | 13.94 | 5.4 | 5.38 | 12.27 | 12.27 | 2.27 | 2.32 | 12.71 | 13 | 2.45 | 2.44
10.97 | 10.93 | 2.57 | 2.57 | 11.44 | 11.16 | 2.78 | 2.96 | 10.89 | 10.91 | 2.38 | 2.33 | 11.7 | 11.68 | 0.63 | 2.63 | 12.86 | 12.85 | 1.62 | 1.61 | 13.34 | 13.25 | 0.9 | 1.01 | 13.27 | 13.2 | 4.6 | 4.62 | 13.94 | 14 | 5.38 | 5.3 | 12.27 | 12.33 | 2.32 | 2.36 | 13 | 12.81 | 2.44 | 2.43
10.93 | 10.96 | 2.57 | 2.65 | 11.16 | 11.21 | 2.96 | 2.76 | 10.91 | 10.94 | 2.33 | 2.29 | 11.68 | 11.79 | 2.63 | 2.58 | 12.85 | 12.9 | 1.61 | 1.63 | 13.25 | 13.18 | 1.01 | 0.92 | 13.2 | 13.22 | 4.62 | 4.64 | 14 | 14.28 | 5.3 | 5.51 | 12.33 | 12.36 | 2.36 | 2.3 | 12.81 | 12.78 | 2.43 | 2.44
10.96 | 11.02 | 2.65 | 2.55 | 11.21 | 11.31 | 2.76 | 2.85 | 10.94 | 10.95 | 2.29 | 2.3 | 11.79 | 11.86 | 2.58 | 2.6 | 12.9 | 12.9 | 1.63 | 1.62 | 13.18 | 13.16 | 0.92 | 0.88 | 13.22 | 13.3 | 4.64 | 4.6 | 14.28 | 14.15 | 5.51 | 5.22 | 12.36 | 12.27 | 2.3 | 2.28 | 12.78 | 12.73 | 2.44 | 2.45
11.02 | 11.09 | 2.55 | 2.62 | 11.31 | 11.31 | 2.85 | 2.84 | 10.95 | 10.92 | 2.3 | 2.33 | 11.86 | 11.86 | 2.6 | 0.94 | 12.9 | 12.86 | 1.62 | 1.62 | 13.16 | 13.03 | 0.88 | 0.94 | 13.3 | 13.32 | 4.6 | 4.59 | 14.15 | 13.97 | 5.22 | 5.36 | 12.27 | 12.32 | 2.28 | 2.34 | 12.73 | 12.73 | 2.45 | 2.59
11.09 | 11.06 | 2.62 | 2.57 | 11.31 | 11.38 | 2.84 | 2.89 | 10.92 | 10.93 | 2.33 | 2.37 | 11.86 | 11.72 | 0.94 | 2.57 | 12.86 | 12.82 | 1.62 | 1.63 | 12.71 | 12.77 | 0.83 | 0.76 | 13.32 | 13.31 | 4.59 | 4.64 | 13.97 | 14.05 | 5.36 | 5.37 | 12.32 | 12.29 | 2.34 | 2.34 | 12.73 | 12.67 | 2.59 | 2.56
11.06 | 11.09 | 2.57 | 2.64 | 11.38 | 11.23 | 2.89 | 2.78 | 10.93 | 10.92 | 2.37 | 2.4 | 11.72 | 11.75 | 2.57 | 2.53 | 12.82 | 12.73 | 1.63 | 1.6 | 12.77 | 11.51 | 0.76 | 0.86 | 13.34 | 13.22 | 4.56 | 4.49 | 14.05 | 14.05 | 5.37 | 5.29 | 12.29 | 12.2 | 2.34 | 2.3 | 12.67 | 12.7 | 2.56 | 2.57
11.09 | 11.1 | 2.64 | 2.57 | 11.23 | 11 | 2.78 | 2.84 | 10.85 | 10.9 | 2.23 | 2.24 | 11.75 | 11.68 | 2.53 | 2.55 | 12.94 | 12.93 | 1.6 | 1.66 | 11.51 | 12.86 | 0.86 | 0.77 | 13.22 | 13.13 | 4.49 | 4.51 | 14.03 | 13.95 | 4.57 | 4.6 | 12.2 | 12.17 | 2.3 | 2.34 | 12.7 | 12.67 | 2.57 | 2.6
10.97 | 11.09 | 2.48 | 2.54 | 12.08 | 12.18 | 3.08 | 2.78 | 10.9 | 10.84 | 2.24 | 2.28 | 11.68 | 11.71 | 2.55 | 0.72 | 12.93 | 12.98 | 1.66 | 1.68 | 12.86 | 12.92 | 0.77 | 0.79 | 13.13 | 13.13 | 4.51 | 4.55 | 13.95 | 14 | 4.6 | 4.69 | 12.13 | 12.02 | 2.17 | 2.2 | 12.67 | 12.84 | 2.6 | 2.5
11.09 | 10.99 | 2.54 | 2.49 | 12.18 | 12.11 | 2.78 | 2.51 | 10.84 | 10.88 | 2.28 | 2.22 | 11.71 | 11.83 | 0.72 | 2.64 | 12.98 | 13.01 | 1.68 | 1.69 | 12.92 | 12.79 | 0.79 | 0.83 | 13.13 | 13.2 | 4.55 | 4.5 | 14 | 13.9 | 4.69 | 4.61 | 12.02 | 12.03 | 2.2 | 2.16 | 12.89 | 13 | 3.04 | 2.91
10.99 | 11.02 | 2.49 | 2.44 | 12.11 | 12.13 | 2.51 | 2.63 | 10.88 | 10.83 | 2.22 | 2.2 | 11.83 | 11.6 | 2.64 | 2.61 | 13.01 | 12.93 | 1.69 | 1.6 | 12.79 | 12.88 | 0.83 | 0.74 | 13.2 | 13.26 | 4.5 | 4.48 | 13.9 | 13.82 | 4.61 | 4.89 | 12.03 | 12.09 | 2.16 | 2.22 | 13 | 12.96 | 2.91 | 2.96
11.02 | 11.11 | 2.44 | 2.7 | 12.13 | 11.99 | 2.63 | 2.58 | 10.83 | 10.8 | 2.2 | 2.26 | 11.6 | 11.66 | 2.61 | 2.53 | 12.93 | 12.87 | 1.6 | 1.54 | 12.88 | 12.86 | 0.74 | 0.82 | 13.26 | 13.28 | 4.48 | 4.51 | 13.82 | 13.94 | 4.89 | 4.38 | 12.09 | 12.18 | 2.22 | 2.17 | 12.96 | 12.91 | 2.96 | 3.13
11.11 | 11.01 | 2.7 | 2.51 | 11.99 | 12.19 | 2.58 | 2.5 | 10.8 | 10.78 | 2.26 | 2.26 | 11.66 | 11.69 | 2.53 | 2.58 | 12.87 | 12.81 | 1.54 | 1.66 | 12.86 | 12.95 | 0.82 | 0.79 | 13.28 | 13.3 | 4.51 | 4.54 | 13.94 | 14.16 | 4.38 | 3.55 | 12.18 | 12.16 | 2.17 | 2.2 | 12.91 | 12.99 | 3.13 | 3.07
11.01 | 10.96 | 2.51 | 2.59 | 12.19 | 12.1 | 2.5 | 2.61 | 10.78 | 10.73 | 2.26 | 2.27 | 11.69 | 11.77 | 2.58 | 2.51 | 12.81 | 12.71 | 1.66 | 1.67 | 12.95 | 12.75 | 0.79 | 0.77 | 13.3 | 13.34 | 4.54 | 4.54 | 14.16 | 13.98 | 3.55 | 4.61 | 12.16 | 12.18 | 2.2 | 2.18 | 12.99 | 13.13 | 3.07 | 2.96
10.96 | 10.98 | 2.59 | 2.52 | 12.1 | 11.99 | 2.61 | 2.61 | 10.73 | 10.7 | 2.27 | 2.2 | 11.77 | 11.79 | 2.51 | 1.62 | 12.71 | 12.81 | 1.67 | 1.68 | 12.75 | 12.1 | 0.77 | 0.79 | 13.34 | 13.27 | 4.54 | 4.54 | 13.98 | 14 | 4.61 | 4.68 | 12.18 | 12.23 | 2.18 | 2.19 | 13.13 | 13.01 | 2.96 | 3.06
10.98 | 11.01 | 2.52 | 2.58 | 11.99 | 12.09 | 2.61 | 2.57 | 10.7 | 10.76 | 2.2 | 2.19 | 11.79 | 11.49 | 1.62 | 2.5 | 12.81 | 12.85 | 1.68 | 1.69 | 12.1 | 12.88 | 0.79 | 0.72 | 13.27 | 13.24 | 4.54 | 4.49 | 14 | 14.02 | 4.68 | 4.67 | 12.23 | 12.19 | 2.19 | 2.18 | 13.01 | 13.07 | 3.06 | 3.11
11.01 | 11.05 | 2.58 | 2.49 | 12.09 | 12.17 | 2.57 | 2.51 | 10.76 | 10.82 | 2.19 | 2.25 | 11.49 | 11.82 | 2.5 | 1.82 | 12.85 | 12.93 | 1.69 | 1.65 | 12.88 | 12.13 | 0.72 | 0.86 | 13.24 | 13.19 | 4.49 | 4.49 | 14.02 | 14 | 4.67 | 4.6 | 12.19 | 12.16 | 2.18 | 2.17 | 13.07 | 12.83 | 3.11 | 3.07
11.05 | 11.03 | 2.49 | 2.52 | 12.17 | 12.14 | 2.51 | 2.56 | 10.82 | 10.86 | 2.25 | 2.28 | 11.82 | 11.83 | 1.82 | 2.49 | 12.93 | 12.99 | 1.65 | 1.7 | 12.13 | 12.62 | 0.86 | 0.77 | 13.19 | 13.12 | 4.49 | 4.48 | 14 | 13.9 | 4.6 | 4.06 | 12.16 | 12.14 | 2.17 | 2.15 | 12.83 | 12.97 | 3.07 | 2.9
11.03 | 11 | 2.52 | 2.43 | 12.14 | 12.18 | 2.56 | 2.6 | 10.86 | 10.86 | 2.28 | 2.23 | 11.83 | 11.83 | 2.49 | 2.64 | 12.99 | 12.99 | 1.7 | 1.68 | 12.62 | 13.02 | 0.77 | 0.79 | 13.12 | 13.17 | 4.48 | 4.49 | 13.9 | 13.89 | 4.06 | 4.71 | 12.14 | 12.02 | 2.15 | 2.2 | 12.97 | 12.88 | 2.9 | 2.92
11 | 11.04 | 2.43 | 2.52 | 12.18 | 12.05 | 2.6 | 2.58 | 10.86 | 10.88 | 2.23 | 2.2 | 11.83 | 11.82 | 2.64 | 2.59 | 12.99 | 12.97 | 1.68 | 1.64 | 13.02 | 13.07 | 0.79 | 0.79 | 13.17 | 13.25 | 4.49 | 4.53 | 13.89 | 14.18 | 4.71 | 5.03 | 12.02 | 12.05 | 2.2 | 2.16 | 12.88 | 12.92 | 2.92 | 3
11.04 | 11.05 | 2.52 | 2.46 | 12.05 | 12.16 | 2.58 | 2.71 | 10.88 | 10.9 | 2.2 | 2.23 | 11.82 | 11.78 | 2.59 | 0.88 | 12.97 | 12.93 | 1.64 | 1.56 | 13.07 | 12.97 | 0.79 | 0.73 | 13.25 | 13.24 | 4.53 | 4.56 | 14.18 | 13.94 | 5.03 | 4.16 | 12.05 | 12.04 | 2.16 | 2.2 | 12.92 | 12.83 | 3 | 3.09
11.05 | 11.02 | 2.46 | 2.51 | 12.16 | 12.08 | 2.71 | 2.55 | 10.9 | 10.88 | 2.23 | 2.25 | 11.78 | 11.84 | 0.88 | 2.59 | 12.93 | 12.82 | 1.56 | 1.62 | 12.97 | 12.98 | 0.73 | 0.85 | 13.24 | 13.29 | 4.56 | 4.49 | 13.94 | 13.93 | 4.16 | 3.65 | 12.04 | 12.11 | 2.2 | 2.23 | 12.83 | 12.92 | 3.09 | 3
11.02 | 11.01 | 2.51 | 2.47 | 12.08 | 12.2 | 2.55 | 2.52 | 10.88 | 10.86 | 2.25 | 2.25 | 11.84 | 11.89 | 2.59 | 2.58 | 12.82 | 12.7 | 1.62 | 1.66 | 12.98 | 12.91 | 0.85 | 0.79 | 13.29 | 13.33 | 4.49 | 4.47 | 13.93 | 13.86 | 3.65 | 4.78 | 12.11 | 12.13 | 2.23 | 2.19 | 12.92 | 13.13 | 3 | 3.1
11.01 | 11.04 | 2.47 | 2.53 | 12.2 | 12.35 | 2.52 | 2.8 | 10.86 | 10.9 | 2.25 | 2.2 | 11.89 | 11.83 | 2.58 | 2.55 | 12.7 | 12.8 | 1.66 | 1.67 | 12.91 | 12.94 | 0.79 | 0.8 | 13.33 | 13.31 | 4.47 | 4.5 | 13.86 | 13.75 | 4.78 | 4.61 | 12.13 | 12.16 | 2.19 | 2.17 | 13.13 | 12.98 | 3.1 | 3
11.04 | 10.95 | 2.53 | 2.46 | 12.35 | 12.23 | 2.8 | 2.62 | 10.9 | 10.83 | 2.2 | 2.21 | 11.83 | 11.64 | 2.55 | 1.55 | 12.8 | 12.89 | 1.67 | 1.68 | 12.94 | 12.71 | 0.8 | 0.81 | 13.31 | 13.28 | 4.5 | 4.56 | 13.75 | 13.91 | 4.61 | 4.62 | 12.16 | 12.26 | 2.17 | 2.38 | 12.98 | 13.11 | 3 | 3.05
10.95 | 10.9 | 2.46 | 2.53 | 12.23 | 12.2 | 2.62 | 2.7 | 10.83 | 10.8 | 2.21 | 2.26 | 11.64 | 11.73 | 1.55 | 2.59 | 12.89 | 12.9 | 1.68 | 1.69 | 12.71 | 12.76 | 0.81 | 0.75 | 13.28 | 13.28 | 4.56 | 4.54 | 13.91 | 13.66 | 4.62 | 4.78 | 12.26 | 12.28 | 2.38 | 2.31 | 13.11 | 13.01 | 3.05 | 3.11
10.9 | 10.93 | 2.53 | 2.43 | 12.2 | 12.1 | 2.7 | 2.6 | 10.8 | 10.77 | 2.26 | 2.22 | 11.73 | 11.87 | 2.59 | 2.47 | 12.9 | 12.94 | 1.69 | 1.6 | 12.76 | 12.67 | 0.75 | 0.82 | 13.28 | 13.14 | 4.54 | 4.53 | 13.66 | 13.86 | 4.78 | 4.71 | 12.28 | 12.23 | 2.31 | 2.19 | 13.01 | 13.08 | 3.11 | 3.06
10.93 | 10.93 | 2.43 | 2.5 | 12.1 | 12.05 | 2.6 | 2.64 | 10.77 | 10.74 | 2.22 | 2.2 | 11.87 | 11.9 | 2.47 | 2.67 | 12.94 | 12.97 | 1.6 | 1.63 | 12.67 | 12.88 | 0.82 | 0.79 | 13.14 | 13.15 | 4.53 | 4.48 | 13.86 | 13.76 | 4.71 | 4.67 | 12.23 | 12.16 | 2.19 | 2.19 | 13.08 | 12.98 | 3.06 | 3.07
10.93 | 10.96 | 2.5 | 2.45 | 12.05 | 11.97 | 2.64 | 2.57 | 10.74 | 10.72 | 2.2 | 2.18 | 11.9 | 11.83 | 2.67 | 1.91 | 12.97 | 12.91 | 1.63 | 1.65 | 12.88 | 12.99 | 0.79 | 0.89 | 13.15 | 13.17 | 4.48 | 4.47 | 13.76 | 13.85 | 4.67 | 4.44 | 12.16 | 12.11 | 2.19 | 2.21 | 12.98 | 12.76 | 3.07 | 2.9
10.96 | 11.02 | 2.45 | 2.51 | 11.97 | 12.16 | 2.57 | 2.64 | 10.72 | 10.8 | 2.18 | 2.24 | 11.83 | 11.84 | 1.91 | 2.52 | 12.91 | 12.92 | 1.65 | 1.64 | 12.99 | 13.01 | 0.89 | 0.82 | 13.17 | 13.19 | 4.47 | 4.48 | 13.85 | 13.97 | 4.44 | 4.22 | 12.11 | 12.01 | 2.21 | 2.15 | 12.76 | 12.96 | 2.9 | 2.91
11.02 | 11.04 | 2.51 | 2.47 | 12.16 | 12.25 | 2.64 | 2.59 | 10.8 | 10.82 | 2.24 | 2.27 | 11.84 | 11.84 | 2.52 | 2.49 | 12.92 | 12.9 | 1.64 | 1.68 | 13.01 | 12.76 | 0.82 | 0.76 | 13.19 | 13.24 | 4.48 | 4.5 | 13.97 | 13.89 | 4.22 | 4.55 | 12.01 | 12.04 | 2.15 | 2.2 | 12.96 | 12.94 | 2.91 | 2.86
11.04 | 11.02 | 2.47 | 2.49 | 12.25 | 12.21 | 2.59 | 2.81 | 10.82 | 10.84 | 2.27 | 2.22 | 11.84 | 11.77 | 2.49 | 2.57 | 12.9 | 12.88 | 1.68 | 1.69 | 12.76 | 12.91 | 0.76 | 0.9 | 13.24 | 13.31 | 4.5 | 4.52 | 13.89 | 13.84 | 4.55 | 3.8 | 12.04 | 12.08 | 2.2 | 2.22 | 12.94 | 13.13 | 2.86 | 2.9
11.02 | 11.03 | 2.49 | 2.42 | 12.21 | 12.19 | 2.81 | 2.71 | 10.84 | 10.85 | 2.22 | 2.21 | 11.77 | 11.85 | 2.57 | 1.38 | 12.88 | 12.77 | 1.69 | 1.67 | 12.91 | 12.96 | 0.9 | 0.78 | 13.31 | 13.28 | 4.52 | 4.53 | 13.84 | 13.8 | 3.8 | 4.62 | 12.08 | 12.13 | 2.22 | 2.15 | 13.13 | 13.07 | 2.9 | 3.07
11.03 | 10.99 | 2.42 | 2.5 | 12.19 | 12.14 | 2.71 | 2.64 | 10.85 | 10.87 | 2.21 | 2.2 |  |  |  |  | 12.77 | 12.77 | 1.67 | 1.59 | 12.96 | 12.82 | 0.78 | 0.76 | 13.28 | 13.21 | 4.53 | 4.44 | 13.8 | 13.81 | 4.62 | 4.92 | 12.13 | 12.14 | 2.15 | 2.2 | 13.07 | 13.07 | 3.07 | 2.93
10.99 | 11.01 | 2.5 | 2.43 | 12.14 | 12.33 | 2.64 | 2.75 | 10.87 | 10.89 | 2.2 | 2.27 |  |  |  |  | 12.77 | 12.84 | 1.59 | 1.61 | 12.82 | 12.75 | 0.76 | 0.81 | 13.21 | 13.28 | 4.44 | 4.44 | 13.81 | 14.01 | 4.92 | 3.87 | 12.14 | 12.15 | 2.2 | 2.21 | 13.07 | 12.97 | 2.93 | 2.88
11.01 | 10.95 | 2.43 | 2.5 | 12.33 | 12.13 | 2.75 | 2.6 | [8 empty] | 12.84 | 12.87 | 1.61 | 1.61 | 12.75 | 12.97 | 0.81 | 0.72 | 13.28 | 13.18 | 4.44 | 4.42 | 14.01 | 14.11 | 3.87 | 3.76 | 12.15 | 12.14 | 2.21 | 2.17 | 12.97 | 13.03 | 2.88 | 3.01
10.95 | 10.93 | 2.5 | 2.48 | 12.13 | 12.14 | 2.6 | 2.69 | [8 empty] | 12.87 | 12.94 | 1.61 | 1.67 | 12.97 | 12.98 | 0.72 | 0.84 | 13.18 | 13.12 | 4.42 | 4.49 | 14.11 | 14.02 | 3.76 | 4.69 | 12.14 | 12.24 | 2.17 | 2.21 | 13.03 | 12.97 | 3.01 | 3.03
10.93 | 10.89 | 2.48 | 2.41 | 12.14 | 12.12 | 2.69 | 2.69 | [8 empty] | 12.94 | 12.96 | 1.67 | 1.68 | 12.98 | 12.74 | 0.84 | 0.78 | 13.12 | 13.08 | 4.49 | 4.5 | 14.02 | 13.86 | 4.69 | 4.64 | 12.24 | 12.18 | 2.21 | 2.16 | 12.97 | 13.08 | 3.03 | 2.95
10.89 | 10.73 | 2.41 | 2.47 | 12.12 | 12.1 | 2.69 | 2.53 | [8 empty] | 12.96 | 13 | 1.68 | 1.66 | 12.74 | 12.81 | 0.78 | 0.82 | 13.08 | 13.14 | 4.5 | 4.48 | 13.86 | 14.07 | 4.64 | 4.5 | 12.18 | 12.15 | 2.16 | 2.16 | 13.08 | 12.85 | 2.95 | 2.9
10.73 | 10.92 | 2.47 | 2.5 | 12.1 | 12.2 | 2.53 | 2.65 | [8 empty] | 13 | 12.9 | 1.66 | 1.57 | 12.81 | 12.68 | 0.82 | 0.76 | 13.14 | 13.15 | 4.48 | 4.45 | 14.07 | 13.96 | 4.5 | 3.99 | 12.15 | 12.17 | 2.16 | 2.22 | 12.85 | 13.02 | 2.9 | 2.92
10.92 | 10.98 | 2.5 | 2.43 | 12.2 | 12.12 | 2.65 | 2.82 | [8 empty] | 12.9 | 12.88 | 1.57 | 1.64 | 12.68 | 12.79 | 0.76 | 0.73 | 13.15 | 13.22 | 4.45 | 4.45 | 13.96 | 13.87 | 3.99 | 4.22 | 12.17 | 12.02 | 2.22 | 2.23 | 13.02 | 12.89 | 2.92 | 2.88
10.98 | 10.92 | 2.43 | 2.57 | 12.12 | 12.11 | 2.82 | 2.59 | [8 empty] | 12.88 | 12.78 | 1.64 | 1.65 | 12.79 | 12.8 | 0.73 | 0.78 | 13.22 | 13.25 | 4.45 | 4.49 | 13.87 | 13.82 | 4.22 | 3.47 | 12.02 | 12.01 | 2.23 | 2.17 | 12.89 | 13.08 | 2.88 | 2.84
10.92 | 11.01 | 2.57 | 2.55 | 12.11 | 12.19 | 2.59 | 2.61 | [8 empty] | 12.78 | 12.8 | 1.65 | 1.68 | 12.8 | 12.79 | 0.78 | 0.75 | 13.25 | 13.23 | 4.49 | 4.48 | 13.82 | 14.02 | 3.47 | 4.76 | 12.01 | 12.08 | 2.17 | 2.2 | 13.08 | 12.95 | 2.84 | 2.97
11.01 | 10.95 | 2.55 | 2.52 | 12.19 | 12.16 | 2.61 | 2.57 | [8 empty] | 12.8 | 12.82 | 1.68 | 1.62 | [8 empty] | 14.02 | 13.95 | 4.76 | 4.58 | 12.08 | 12.13 | 2.2 | 2.22 | 12.95 | 12.99 | 2.97 | 2.93
10.95 | 11 | 2.52 | 2.51 | 12.16 | 12.3 | 2.57 | 2.66 | [8 empty] | 12.82 | 12.92 | 1.62 | 1.71 | [8 empty] | 13.95 | 13.81 | 4.58 | 4.63 | 12.13 | 12.19 | 2.22 | 2.15 | 12.99 | 12.91 | 2.93 | 2.87
[32 empty] | 12.19 | 12.16 | 2.15 | 2.16 | 12.91 | 12.88 | 2.87 | 2.87
[36 empty] | 12.88 | 12.84 | 2.87 | 3

## VILI Data
VILI Scoring (1 to 4)
Rabbit | Weight (Kg) | Wet Lung Weight (g) | Dry Lung Weight (g) | Wet:Dry Weight Ratio | Alveolar Congestion | Hemorrhage | Infiltration of Neutrophils in airspaces or vessel walls | Alveolar wall thickiness/hyaline membrane formation | Overall Average
MV1 | 2.79 | 0.3764 | 0.0739 | 5.0934 | 1 | 4 | 1.6667 | 1 | 1.9167
MV2 | 3.18 | 0.3788 | 0.0641 | 5.9095 | 3 | 3 | 1 | 1 | 2
MV3 | 3.11 | 0.2305 | 0.0427 | 5.3981 | 1.6667 | 4 | 1 | 1.6667 | 2.0833
MV4 | 3.02 | 0.1442 | 0.0219 | 6.5845 | 1 | 0.66667 | 0.33333 | 0 | 0.5
MV5 | 2.98 | 0.1952 | 0.0303 | 6.4422 | 2 | 1 | 2 | 0.66667 | 1.4167
SB1 | 2.65 | 0.2641 | 0.0501 | 5.2715 | 1 | 0.33333 | 0.33333 | 0 | 0.41667
SB2 | 2.65 | 0.2704 | 0.0474 | 5.7046 | 4 | 0.66667 | 0.33333 | 0 | 1.25
SB3 | 2.53 | 0.1554 | 0.0301 | 5.1628 | 4 | 0.66667 | 0.33333 | 0 | 1.25