The overall approach was to build a subject-specific multiphase computational fluid dynamics (CFD) model and to verify model results using a corresponding bench-top model. After verification, the models were used to quantitatively assess tracer removal by Neurapheresis therapy compared to lumbar drain. Fluorescein solution was used as a surrogate tracer and modeled as a bulk fluid phase within the CSF (29-32) to track clearance of blood components, specifically hemoglobin derivatives.

Model geometry

A high-resolution T2-weighted magnetic resonance imaging (MRI) sequence was used to acquire a subject-specific geometry of the CSF system of a 23-year-old healthy female subject (Fig. 1a). This model contained two main regions: 1) the spinal SAS and 2) the intracranial CSF space.  The MRI protocol and image reconstruction method for the spinal SAS were previously described by Sass et al. (33). In brief, the spine model combined the high-resolution MR imaging with anatomic measurements and cadaveric studies in the literature and included 31 pairs of anatomically realistic spinal cord nerve roots, the filum terminale and the thecal sac (Fig. 1b).

A detailed intracranial CSF space geometry obtained with high-resolution MRI was added to the spinal SAS. The CSF space schematic in this region is shown in Fig. 1c. The intracranial CSF was manually segmented using ITK-SNAP (Version 3.4.0, University of Pennsylvania, U.S.A.), exported in .STL (Stereolithography) format, and imported into Blender (Amsterdam, Netherlands, Version 2.79).  The model geometry was simplified within Blender to include the following key intracranial CSF spaces: a) lateral ventricles and foramen of Monro (left and right), b) third ventricle and aqueduct of Sylvius, c) fourth ventricle and foramen of Luschka (lateral apertures), d) cisterna magna, e) pre-pontine and pontine cisterns, f) trigeminal cistern, g) quadrigeminal cistern, h) Sylvian cisterns (left and right), and i) cortical SAS (Fig. 1c1).  The complex geometry of the cortical surface and the ventricular CSF system was smoothed and simplified such that the total cranial CSF volume remained equivalent to that of the original segmentation. The cortical SAS was divided at the midline to represent the falx cerebri and superior sagittal sinus.  The posterior-superior aspect of the cerebellar cistern was terminated by a barrier representing the tentorium cerebelli.  The cortical SAS was considered to cover the entire cortical surface with a uniform thickness of ~2.78 mm.  To represent CSF production from the choroid plexus, CSF production ports with diameters of 5 mm were added to the superior aspect of the lateral ventricles (Fig. 1c1).  A 5.5 mm diameter port was also added at the lumbar spine at the L3-L4 for catheter insertion. Each partitioned cistern was exported as a stereolithography (.STL) file and imported into Autodesk Netfabb (San Rafael, CA, Build 1608) to ascertain volume, surface area, and connectivity. The superior and inferior aspects of the model incorporated 22.5 mm diameter cylindrical passages with a length of 8 mm to accommodate oscillatory CSF pulsation to the system.

The dual-lumen Neurapheresis catheter geometry was identical with our previous publication (21).  In brief, the catheter was inserted at L3-L4 and located within the posterior SAS with the aspiration port at L2 and the return port at T2 vertebral level (Fig. 2a).  Internal and external catheter diameters were 0.5 and 0.7 mm, respectively, for the inner lumen, and 1.7 and 1.5 mm for the outer lumen.  A series of 11 holes were located at the return port and 12 holes at the aspiration port (Fig. 2a1 and 2a2).

Computational fluid dynamic model

Computational mesh

An unstructured tetrahedral computational mesh was generated using ANSYS ICEM 19.2 (Canonsburg, PA). Details on the prism layer and Mesh parameters are included in Table 1 and Fig. 1c2.  The mesh was refined near the catheter aspiration and return ports, since at these regions the geometric dimensions of the holes were relatively small.  The mesh included the internal geometry of the return and aspiration ports to ensure correct flow distribution from each hole.  The final mesh covering the Neurapheresis catheter, cranial, and spinal SAS had 16.8 Million cells.   To allow verification of numerical results, an identical geometry was used for the in vitro model (Fig. 1b).

TABLE 1. Mesh details

| Parameter | Type |
| --- | --- |
| Cell type: | Tetrahedral |
| Face type: | Triangle |
| Prism layer: | 4 |
| Prism Height: | 0.05 (mm) |
| Growth law: | Exponential |
| Growth factor: | 1.2 |
| Mesh cells: | 16.8 M |
| Mesh faces: | 26.4 M |
| Mesh nodes: | 3.7 M |
| Min mesh size: | 0.05 (mm) |
| Max mesh size: | 1.0 (mm) |

Time step: 0.1 (s), Cycle: 5th, Time: 100 (s)

Boundary conditions

Clearance of tracer after SAH with Neurapheresis therapy and lumbar drain was studied under the following boundary conditions (Fig. 2a). The model system was oriented at 30º to horizontal to mimic a typical patient position within a hospital bed. Fluid was aspirated and returned under Neurapheresis therapy at 2.0 and 1.8 mL/min. The 0.2 mL/min difference between return and aspiration matched the drain rate applied under lumbar drain (see details below).  Constant CSF production from the choroid plexus within the left and right ventricles was specified to be 0.1 mL/min (0.2 mL/min total) at the CSF production channel entrance. For the lumbar drain simulation, a drainage rate of 0.2 mL/min was specified, as this is a nominal value typically observed in the literature for that procedure (34, 35).

To represent CSF pulsation around the brain and spinal cord, an oscillatory velocity inlet boundary condition was defined at the model caudal opening using a User Defined Function.  The exact waveform was derived from the C2-C3 CSF flow rate waveform (Fig. 2b) obtained from phase-contrast MR imaging of the healthy 23-year-old female subject (36).  The angular frequency of the waveform was ω = 6.98 s-1. A zero-pressure outlet boundary condition was defined at the cranial opening. No-slip boundary conditions were imposed at the model walls (dural, pial and intraventricular spaces) with the walls modeled as stationary. CSF was modeled as incompressible with a density of 998.3 kg/m3 and viscosity of 0.89 mpa∙s.  Tracer density and viscosity was assumed to be identical to CSF

Flow model

The pulsatile CSF flow field was computed using ANSYS Fluent 19.2 (Canonsburg, PA) by solving the continuity and Navier–Stokes equations

Where  is fluid density,  is fluid viscosity,  is acceleration due to gravity,  is the velocity vector, and  is the pressure field.

Multiphase model

The 3D CFD computations used the ANSYS multi-phase fluid model to track the dispersion of a tracer within the CSF, with tracer volume fraction given by

where  is the bulk fluid phase,  and   are the phase density and volume fraction of phase ,  is the mass-averaged velocity,  is the mixture density, and   is the drift velocity for phase , with phase k=1 being the CSF and k=2 for the tracer.  We assumed the relative velocity (slip velocity) between phase k and the bulk fluid to be zero.  Thus, the drift velocity was considered to be zero.

Based on clinical SAH observations and previous research, we assumed a baseline tracer concentration of 10% (=0.1) throughout the model domain.  Tangen et al. (37) showed that blood debris is evenly dispersed throughout the spinal SAS within the first hour post-SAH.  For the present study, we expected at least 1-hour of time to have passed before a lumbar drain or Neurapheresis therapy system could be applied post-SAH.  The lumbar puncture is a common clinical emergency procedure used to aid in SAH diagnosis (38).  Medical doctors have observed the CSF samples from these SAH patients to be colored with a light pink shade (xanthochromia) (39), indicating that blood has spread throughout the CSF system down to the lumbar spine, justifying the evenly-mixed tracer concentration used in this model as a reasonable approximation.  Steady-streaming velocities were determined based on the average of 10 CSF flow cycles with transient effects eliminated by removal of the first flow cycle.  Similar to Kuttler et al.(40), the fixed velocity field (“frozen flow field”) was used to solve the volume fraction equation for spatial-temporal tracer concentration. The axial distribution of tracer concentration, , for 3 mm thick slices along the z-axis was computed by:

Where  is the cell volume and summations computed for all cells within each 3 mm thick cross-section. Spatial-temporal distribution of tracer concentration was plotted over 24-hours for Neurapheresis therapy and lumbar drain.

To determine concentration profiles over 24-hours after SAH, we determined the solute transport due to the steady-streaming CSF velocity field.  Molecular diffusion of large molecules within the CSF (MW ~ 150 kDa) is much smaller than steady-streaming and oscillatory CSF velocities (40). However, shear and mixing across the cross section has the potential to greatly increase the effective diffusivity in the spinal SAS (41).   Tangen et al.(37) found that molecular diffusion had a negligible impact on tracer spread within an idealized geometry representing the spinal and cranial SAS. Kurtcuoglu et al.(42) also neglected diffusion in their model. To make the computational effort reasonable, molecular diffusion of the tracer was not included in the current study. However, as described in the following paragraphs, the potential impact of neglecting diffusion was estimated.

To help understand the relative importance of diffusive versus advective mass transport, the Sherwood number () was calculated and  provides the ratio of convective mass transport, , to the effective diffusive mass transport,  , where is a characteristic length and  is the effective diffusivity including shear-augmented dispersion.  was computed based on the mean cross-sectional velocity at peak systolic (h=0.26 m/s and 2.4 m/s for cortical and spinal SAS, respectively).  was assumed to be the minimum gap width between the shells in the cortical SAS (~2 mm) and mean of the hydraulic diameter for the spinal SAS (5.87 mm).

To see if the tracer is a good similitude for hemoglobin clearance,  was calculated for both tracer and hemoglobin using an order-of-magnitude model by Sharp, et al. (41):

(5)

(6)

(7)

(8)

Where is the maximum enhancements with optimal mixing, is the characteristic non-dimensional pressure gradient (~152.6 (41)),  is Womersley number and  is oscillatory Peclet numbers. Schmidt number () was described as the ratio of momentum diffusivity of water at room temperature (= 0.89 E-06 m2 /s), to molecular diffusion coefficient (). The molecular diffusivity of the fluorescein tracer is =4.25 E-10 m2/s  (43, 44) and hemoglobin is  =10.2 E-11 m2/s (45).

Solver settings

Simulations were carried out using the PISO Scheme (pressure-implicit with splitting of operators) to solve the flow equations with second-order upwind for momentum discretization, PRESTO! (PREssure STaggering Options) for pressure discretization, and first-order upwind for volume fraction discretization. Default values were used for under relaxation factors. The implicit formulation was used for volume fraction parameters and a dispersed model was used for phase interface modeling. The convergence criteria for velocity, continuity, momentum, and phase volume fraction were set to 1E-06 with time step = 0.1 s and maximum of 100 iterations per time-step. Total simulation time was 14 days to compute 24-hours of real-world time in parallel mode with 126 GB RAM and 38 processors.  Time required to solve the fixed flow field (11 flow cycles) was 36 hours.

Numerical sensitivity studies

Axial distribution of average tracer concentration at one hour was verified by numerical sensitivity studies for time-step size and mesh resolution.  Results were computed for a “coarse”, “medium”, and “fine” mesh with wall prism layers.   For the medium mesh, time-step sensitivity was then checked with time-step resolution given by fractions of the cardiac cycle, T = 0.9 (s), for T/18, T/9, T/5. Based on these results, a “Medium” mesh with time step = T/9 was carried forward for computation of final results.

In vitro experimental model

To help verify the numerical model results, an in vitro CSF model was constructed with a fluid domain geometry identical to the numerical model.  For the in vitro model, the fluid domain was encased in a 2 mm thick transparent shell.  The shell was divided into cranial (Fig. 1b), upper thoracic (Fig. 1b1), and lower spine (Fig. 1b2) pieces not exceeding the maximum build size of the 3D printer.  The final digital geometry was submitted for STL 3D printing with a layer and in-plane resolution of 127 μm and 500 μm, respectively, for the cranial volume and 127 μm and 250 μm in the spinal column. The caudal end of the model was connected to a custom-built oscillatory flow pump imparting an identical CSF waveform utilized for the numerical simulation (Fig. 2b).  The dual-lumen Neurapheresis catheter was inserted through an access port such that aspiration and return flow occurred at L2 and T2, respectively (Fig. 2a). These locations were congruent with the numerical model.

To mimic SAH blood distribution within the CSF, an aqueous solution of fluorescein sodium (25 mg / 1000 mL) was used to fill the entire model representing a uniform distribution of blood. A uniform distribution was assumed because previous research from Tangen et al (37) showed that blood debris is evenly dispersed throughout the spinal SAS within the first hour post-SAH.  Usage of fluorescein as a surrogate tracer for blood is in agreement with previous studies that sought to track blood components in SAH, specifically hemoglobin derivatives (29-32). Experiments were conducted with deionized water, since the transport of SAH products is thought to be passive, i.e., the effects of chemical reactions are negligible (37). Fluorescein also does not mimic hydrodynamic effects on finite-size particles, such as the lift exerted on red cells near a boundary.

A continuous flow syringe pump (KD Scientific Legato 270, Holliston, MA) was used to aspirate fluid via the Neurapheresis catheter at 2.0 mL min-1.  This same pump was also configured in parallel to return fresh deionized water at a rate of 2.0 mL/min.  This clean return flow was split and 0.2 mL/min was diverted by a second continuous flow syringe pump for CSF production originating from the ventricles.  Thus a net flow of 1.8 mL/min was delivered to the return port of the Neurapheresis catheter. Another syringe pump was used to generate the subject-specific CSF flow wave form (Fig. 2c).The caudal end of the model was chosen as the input location for SAS CSF oscillations.  This location allowed a uniform CSF flow rate at any cross-section throughout the rigid subarachnoid space.  Also, this location allowed CSF outflow to occur at the top of the brain, a region approximately located near the arachnoid granulations, where CSF outflow is believed to occur into the superior sagittal sinus in vivo. Pump flow rates were verified by a graduated cylinder and stopwatch before each experiment.  Identical to the numerical model, all experiments were conducted with the model at 30 degrees (head up) from horizontal .

The following digital image subtraction technique was applied to quantify spatial-temporal tracer concentration.  The entire system was enclosed to block ambient light sources. The model was surrounded within the enclosure by blue wavelength (450 (nm)) light emitting diodes triggered to a camera flash.  For model calibration, a baseline image was obtained without any tracer introduced to the model.  To establish baseline decay of the fluorescein tracer, time-lapse imaging was obtained at 5-minute intervals over a total of 24-hours with the tracer solution at a known concentration 15 μM.  To eliminate aliasing artifacts due to CSF oscillations, images were collected with a trigger set to the cardiac systole.  An identical series of images was then collected for the same experiment under Neurapheresis therapy and also with lumbar drain.

Spatial-temporal distribution of tracer concentration for each case was computed as , where  is the signal intensity for each case,  is the background image intensity at zero concentration and  is the baseline decay. To account for lens distortion, the distance and orientation of the camera to the model was measured and used to correct axial pixel position for comparison to CFD results.  Similar to the CFD simulations, spatial distribution of tracer concentration was averaged for 3 mm thick slices along the model (z-direction).

12-bit images were collected with 12 MP resolution using a digital camera (Sony alpha a7s ii) with a 50 mm focal length prime lens with f/10 aperture (Sony E 50 mm f/1.8 OSS).  Imaging settings were 2000 ISO, white balance 5500 Kelvin, and 1/100 s shutter speed. A 525 nm bandpass filter (Midwest optical systems BP525-49) was used in front of the lens to improve image signal-to-noise ratio.

Geometric and Hydrodynamic Quantification

Based on the 3D reconstruction and meshing, the following geometric and hydrodynamic parameters were calculated along the spine at 1 mm intervals using our previously described methods (46).  Reynolds number based on hydraulic diameter was calculated as , where  is the absolute value of the peak flow rate from the flow rate waveform at each cross section, is hydraulic diameter,  is the cross-sectional area and  is kinematic viscosity. Womersley number was quantified as    where  is angular velocity and T is the cycle period. Mean cross sectional velocity at peak systolic and diastolic flow, was computed as , with  defined as the maximum flow rate at peak systole and diastole at each slice.

Quantification of steady-streaming CSF flow

To quantify steady-streaming, the cyclic mean velocity in the z-direction, , was computed for each node in the computational mesh as a summation of z-velocity on each node during one cycle similar to our previous publication (25).  was visualized at a mid-sagittal slice for Neurapheresis therapy and lumbar drain.  A positive value for indicates steady-streaming in the rostral direction.  was used in  number calculations as a convective mass transfer term, , since the dominant velocity field in the simulation was assumed to be equal to . The axial distribution of steady-streaming, , was estimated by computing the cross-sectional average of  magnitude:

(9)

where the summations are over all cells in the cross section at a given z location.  was calculated for z-slices at 1 mm intervals along the spine. To further quantify the magnitude of steady-streaming flow, a non-dimensional fraction of the specified flow rate amplitude was defined as:

Verification of numerical results

A detailed comparison of numerical and in vitro results was performed by the following correlation analysis similar to that previously conducted by our group (47). In vitro experiments were conducted using fluorescein as a tracer (as described above) under Neurapheresis therapy and lumbar drain.  In vitro and CFD spatial-temporal tracer concentration results were compared by linear regression and Bland-Altman plot analysis with 95% confidence interval calculations.