## Individual adaptation of mechanical ventilation in ARDS with Electrical Impedance Tomography

Tobias H. Becher, Valerie Buchholz, Daniel Hassel, Timo A. Meinel, Dirk Schädler, Inéz Frerichs and Norbert J. Weiler

## Electronic Supplementary Material

## Algorithm for adjustment of tidal volume and PEEP with EIT

After initial assessment of ventilation delay, stress and strain, ventilator settings were optimized according to the EIT-based algorithm (cf. main manuscript figure 1). Initially, an arterial blood gas (ABG) sample was collected for assessment of arterial partial pressures of oxygen (PaO2), carbon dioxide (PaCO2) and pH. Respiratory rate (RR) was increased if pH was < 7.30, provided no auto-peep was present. Auto-PEEP was assessed by visual analysis of the flow curve and by performing an end-expiratory occlusion maneuver for direct measurement of air trapping. If auto-PEEP was present or pH >7.40 or arterial partial pressure of carbon dioxide (PaCO2) < 35 mmHg, respiratory rate was decreased.

If RR could not be increased due to the presence of auto-PEEP and pH was below 7.20, VT was increased by 0.5 to1 ml per kg predicted body weight (PBW).

Recruitablity was assessed using a sustained-inflation recruitment maneuver (RM) with airway pressure (Paw) of 40 mbar followed by a PEEP increase of 3 mbarRegional Crs was assessed by dividing the EIT image in four horizontal regions of interest (ROIs) and by multiplying global Crs with the relative tidal impedance change in each of the ROIs. An increase in regional Crs of more than 5% (normalized to global Crs) was interpreted as recruitment. If recruitable lung tissue was identified following the RM, the higher PEEP level was kept. Tidal recruitment and overdistension were analyzed with a short reduction in VT that was achieved by halving the inspiratory pressure difference (∆P) during pressure-controlled-ventilation. Changes in regional Crs during this maneuver were assessed as follows:  Regional increases and decreases in Crs were summed up separately and normalized to global Crs with previously set VT. Any regional increase in Crs with reduced VT was interpreted as indicative of overdistension with the previously applied VT, whereas any regional decrease in Crs with reduced VT was interpreted as indicator for tidal recruitment. If tidal recruitment was identified, PEEP was elevated by 3 mbar. If overdistension was identified, VT was decreased by 1 ml / kg PBW provided this did not lead to severe acidosis (pH <7.20). PEEP was reduced by 2 mbar if no recruitability and no tidal recruitment had been identified during the last two hours. After any PEEP-reduction, an assessment of derecruiment was performed and the PEEP-reduction was reversed if relevant derecruitment was detected.

After four hours of prospective optimization of mechanical ventilation with the EIT-based algorithm, another assessment of ventilation delay, stress and strain was performed for analyzation of primary and secondary outcome measures.

Protocol rules and explanations (cf. main manuscript figure 1)

1) Criteria for VT reduction:

VT can be reduced by 0.5 to 1 ml / kg PBW, if

- pH ≥ 7,30 AND

- PaCO2 ≤ 75 mmHg AND

- current VT > 4 ml / kg PBW AND

- no elevated intracranial pressure

2) Criteria for VT increase:

VT must be increased by 0.5 to 1 ml / kg PBW, if

- pH <7,20 AND

- RR cannot be increased 4

OR

- if current VT < 8 ml / kg PBW AND

- no overdistension AND

no VT – reduction for overdistension during last hour

3) Criteria for recruitment maneuver (RM)

RM shall be performed, if

- PaO2/FiO2 < 300 mmHg AND

- no unsuccessful RM during last 2 hours AND

- no successful RM during last 30 minutes

4) Criteria for adaptation of respiratory rate (RR):

increase RR, if

- pH <7,30 AND

- no Auto-PEEP (observe flow curve, perform PEEPi –measurement maneuver if necessary: Auto-PEEP is present if PEEPi > PEEP + 1 mbar)

decrease RR, if

- Auto-PEEP is present (see above) OR

- pH ≥ 7,40 OR

- PaCO2 < 35 mmHg

5) Recruitment

RM shall be classified as successful, if after 1 min

- regional Crs in any ROI increased by ≥ 3% of global Crs (measured before RM) OR

- if global Crs increased by ≥ 10%

in all other cased RM is classified as „unsuccessful“

6) Derecruitment

Derecruitment is present if 5 minutes after a PEEP decrease

- regional Crs in any ROI decreased by > 3% of global Crs (measured before PEEP decrease) OR

- global Crs decreased by > 10%

7) Alveolar Cycling

Alveolar Cycling is present, if during a reduction in driving pressure (∆P) of about 50% for a few breaths,

- Crs at lower ∆P decreases by more than 3% in any ROI (in comparison to global Crs at normal ∆P)

8) Overdistension

Overdistension is present, if during a reduction in ∆P of about 50% for a few breaths

- Crs at lower ∆P increases by more than 3% in any ROI (in comparison to global Crs at normal ∆P)

For diagnosing alveolar cycling and overdistension, a reduction in ∆P by 50% for at least 5 breaths must always be performed.

9) PEEP Changes

Driving pressure (∆P) must always be kept constant during any changes in PEEP by changing inspiratory pressure by the same value.

„Last PEEP-increase“ refers to

- the last successful recruitment maneuver (that is always accompanied by PEEP increase) OR

- the last PEEP increase because of „alveolar cycling“.

Unsuccessful recruitment maneuvers are not referred to as „Last PEEP increase“.