Supporting Information

Supporting Information is available from the Wiley Online Library or from the author.

Title Pulse Physiology Engine to simulate transient oxygen changes and closed-loop, autonomous pump control during venovenous extracorporeal membrane oxygenation

Rei Ukita1, 2, *, Sriya Bapatla2, Jeffrey Webb3, Aaron Bray3, Yerin Woo1, Andrew Kumpfbeck1, Steve P Keller4, Rachel Clipp3

- Department of Cardiac Surgery, Vanderbilt University Medical Center,  Nashville, USA

- Department of Biomedical Engineering, Vanderbilt University, Nashville, USA

- Kitware Inc, Carrboro, 27510, USA

- Department of Medicine, Johns Hopkins University, Division of Pulmonary and Critical Care Medicine, Baltimore, 21287, USA

*Corresponding Author: rei.ukita@vumc.org

Table S1: Circuit parameters represented in the Pulse digital twin model. Labels correspond to Figure 1A in the manuscript.

| Label | Parameter | Component |
| --- | --- | --- |
| 1 | Inlet Location | Cannulas |
| 2 | Outlet Location |  |
| 3 | Drainage Cannula Volume |  |
| 4 | Return Cannula Volume |  |
| 5 | Drainage Cannula Resistance |  |
| 6 | Return Cannula Resistance |  |
| 7 | Pump Flow | Pump |
| 8 | Pump Volume |  |
| 9 | Sweep Gas Flow | Oxygenator |
| 10 | Sweep Gas Pressure |  |
| 11 | Fraction of Sweep Gas Oxygen |  |
| 12a | Oxygen Membrane Diffusing Capacity |  |
| 12b | Carbon Dioxide Membrane Diffusing Capacity |  |
| 13 | Oxygenator Liquid Volume |  |
| 14 | Oxygenator Gas Volume |  |
| 15 | Oxygenator Resistance |  |
| 16 | Bag Compound | Fluid Administration |
| 16 | Bag Volume |  |
| 17 | Infuser Rate |  |
| 18 | Infuser Volume |  |
| 19 | Drainage Limb Volume | Peripherals |
| 20 | Return Limb Volume |  |
| 21 | Inlet Probe Volume |  |
| 22 | Outlet Probe Volume |  |
| 23 | Drainage Limb Resistance |  |
| 24 | Return Limb Resistance |  |

Table S2: Summary of extracorporeal circuit parameters used in the Pulse simulations

| Subsystem | Parameter | Value | Specification Basis |
| --- | --- | --- | --- |
| Cannulation | Inlet Location | Inferior Vena Cava | Drain cannulation |
| Cannulation | Outlet Location | Right Atrium | Return cannulation |
| Cannulation | Drainage Cannula Volume | 7 – 18 mL for 20-31 French | The cross-sectional areas of Avalon drainage and return lumens have an approximate ratio of 2:1. Drainage and return lumen volumes are calculated as 2/3 or 1/3 of the total volume, respectively. Total volume is approximated as a cylinder with a diameter equal to French size and a length of 31 cm |
| Cannulation | Return Cannula Volume | 3 – 9 mL for 20-31 Fr | See above |
| Cannulation | Drainage Cannula Resistance | Varies | Based on Avalon dual-lumen cannulas’ pressure-flow curves shown in prior publications19 |
| Cannulation | Return Cannula Resistance | Varies | See above |
| Pump | Pump Flow | Varies | Calculated based on patient’s circuit and pump characteristics taken from prior publications17 |
| Pump | Pump Volume | 32 mL | Rotaflow pump |
| Pump | Pump Minimum Speed | 500 RPM | Rotaflow |
| Pump | Pump Maximum Speed | 4000 RPM | Rotaflow |
| Sweep Gas | Sweep Gas Flow | Varies | Scenario specific |
| Sweep Gas | Fraction of Sweep Gas Oxygen | 21-100% | Scenario specific |
| Oxygenator | Oxygen Membrane Diffusing Capacity | 0.53 mL/min/mmHg | Quadrox PLS |
| Oxygenator | Carbon Dioxide Membrane Diffusing Capacity | 14 mL/min/mmHg | Quadrox PLS |
| Oxygenator | Oxygenator Liquid Volume | 250 mL | Quadrox PLS |
| Oxygenator | Oxygenator Gas Volume | 215 mL | Typical |
| Oxygenator | Oxygenator Resistance | 5.5 mmHg·min/L | Prior publication20 |
| Circuit | Drainage Limb Volume | 71 mL | 3/8" ID tubing |
| Circuit | Return Limb Volume | 71 mL | 3/8" ID tubing |
| Circuit | Drainage Limb Resistance | 1.5 mmHg·min/L | 3/8" ID tubing |
| Circuit | Return Limb Resistance | 1.5 mmHg·min/L | 3/8" ID tubing |
| Circuit | Inlet Probe Volume | 5 mL | Typical |
| Circuit | Outlet Probe Volume | 5 mL | Typical |

Table S3: Summary of proportional integral performance based on cannula sizes (Fr) and gain coefficients (Kp and Ti)

| Fr | Kp | Ti | Rise Time (seconds) | % Flow Overshoot | Error (SvO2 – Setpoint) | Time to Steady State (seconds) |
| --- | --- | --- | --- | --- | --- | --- |
| 20 | 1000 | 10 | 221 | 2.86E-12 | 0.6918 | 440 |
| 20 | 1000 | 36 | 899 | 1.7853 | 2.1785* | Does not reach steady state |
| 20 | 4000 | 10 | 27 | 2.43E-12 | 0.6886 | 345 |
| 20 | 4000 | 36 | 192 | 3.02E-12 | 0.694 | 412 |
| 27 | 1000 | 10 | 82 | 27.0328 | 0.0201 | 195 |
| 27 | 1000 | 36 | 698 | 0.6484 | 0.468 | 780 |
| 27 | 4000 | 10 | 8 | 144.5452 | 0.0029 | 260 |
| 27 | 4000 | 36 | 46 | 20.3093 | 0.0076 | 178 |
| 31 | 1000 | 10 | 50 | 47.8647 | -7.43E-04 | 400 |
| 31 | 1000 | 36 | 483 | 0.1866 | 0.106 | 509 |
| 31 | 4000 | 10 | 3 | 194.5251 | 3.06E-04 | 303 |
| 31 | 4000 | 36 | 12 | 37.3319 | 0.0042 | 138 |

Fr: French, SvO2: venous oxyhemoglobin saturation,

* Though this simulation scenario did not reach a steady state, error was estimated based on the last 100 seconds of the simulated duration