Supplementary Material

to

Comparison of image quality and spatial resolution between 18F, 68Ga and 64Cu phantom measurements using a digital Biograph Vision PET/CT

Authors:

Anja Braune; Department of Nuclear Medicine, University Hospital Carl Gustav Carus at the Technische Universität Dresden, Institution under public law of the free state of Saxony, Dresden, Germany; anja.braune@ukdd.de (corresponding author)

Liane Oehme; Department of Nuclear Medicine, University Hospital Carl Gustav Carus at the Technische Universität Dresden, Institution under public law of the free state of Saxony, Dresden, Germany; liane.oehme@ukdd.de

Robert Freudenberg; Department of Nuclear Medicine, University Hospital Carl Gustav Carus at the Technische Universität Dresden, Institution under public law of the free state of Saxony, Dresden, Germany; robert.freudenberg@ukdd.de

Frank Hofheinz; PET Center, Institute of Radiopharmaceutical Cancer Research, Helmholtz-Zentrum Dresden-Rossendorf, Dresden, Germany; hofheinz@hzdr.de

Jörg van den Hoff; PET Center, Institute of Radiopharmaceutical Cancer Research, Helmholtz-Zentrum Dresden-Rossendorf, Dresden, Germany; j.van_den_hoff@hzdr.de

Jörg Kotzerke; Department of Nuclear Medicine, University Hospital Carl Gustav Carus at the Technische Universität Dresden, Institution under public law of the free state of Saxony, Dresden, Germany; Carl Gustav Carus Faculty of Medicine at the Technische Universität Dresden; joerg.kotzerke@ukdd.de

Sebastian Hoberück; 
Department of Nuclear Medicine, University Hospital Carl Gustav Carus at the Technische Universität Dresden, Institution under public law of the free state of Saxony, Dresden, Germany; 
Department of Nuclear Medicine, Helios Klinikum Erfurt, Erfurt, Germany. 
Sebastian.Hoberueck@uniklinikum-dresden.de

Keywords: PET/CT, 18F-FDG, 68Ga, 64Cu, NEMA PET Body phantom, Jaszczak phantom, spatial resolution, image quality

Journal: EJNMMI Physics

Attachment – Table 1A: Mean and standard deviation of quantitative image quality parameters determined for all NEMA PET Body phantom measurements.

|  |  | Contrast: 4:1 normal weight setup |  | Contrast: 4:1 obese setup |  |  | Contrast: 8:1 normal weight setup |  | Contrast: 8:1 obese setup |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| RCmean | 18F-FDG | 1.06 ± | 0.01 |  | 1.08 ± | 0.02 | 1.00 ± | 0.04 |  | 1.02 ± | 0.05 |  |
|  | 68Ga-HCl | 0.95 ± | 0.10 |  | 0.99 ± | 0.08 | 0.94 ± | 0.06 |  | 0.97 ± | 0.05 |  |
|  | 64Cu-HCl | 1.04 ± | 0.05 |  | 1.07 ± | 0.02 | 1.10 ± | 0.03 |  | 1.12 ± | 0.03 |  |
| RCmax | 18F-FDG | 1.38 ± | 0.05 |  | 1.45 ± | 0.04 | 1.33 ± | 0.09 |  | 1.37 ± | 0.11 |  |
|  | 68Ga-HCl | 1.27 ± | 0.13 |  | 1.36 ± | 0.15 | 1.22 ± | 0.04 |  | 1.30 ± | 0.07 |  |
|  | 64Cu-HCl | 1.38 ± | 0.06 |  | 1.43 ± | 0.05 | 1.42 ± | 0.09 |  | 1.51 ± | 0.06 |  |
| RCpeak | 18F-FDG | 1.12 ± | 0.23 |  | 1.10 ± | 0.25 | 1.00 ± | 0.33 |  | 1.03 ± | 0.34 |  |
|  | 68Ga-HCl | 0.85 ± | 0.43 |  | 1.04 ± | 0.16 | 1.00 ± | 0.16 |  | 1.00 ± | 0.21 |  |
|  | 64Cu-HCl | 1.10 ± | 0.22 |  | 1.10 ± | 0.27 | 1.09 ± | 0.33 |  | 1.13 ± | 0.34 |  |
| RCpeak (d ≥ 17mm) | 18F-FDG | 0.98 ± | 0.10 |  | 0.93 ± | 0.04 | 0.80 ± | 0.08 |  | 0.82 ± | 0.07 |  |
|  | 68Ga-HCl | 0.93 ± | 0.06 |  | 0.95 ± | 0.04 | 0.90 ± | 0.07 |  | 0.88 ± | 0.10 |  |
|  | 64Cu-HCl | 0.97 ± | 0.08 |  | 0.93 ± | 0.05 | 0.89 ± | 0.08 |  | 0.92 ± | 0.11 |  |
| CRC (%) | 18F-FDG | 100.99 ± | 1.72 |  | 101.57 ± | 2.99 | 96.46 ± | 4.07 |  | 97.86 ± | 5.24 |  |
|  | 68Ga-HCl | 84.11 ± | 11.79 |  | 84.53 ± | 9.70 | 83.30 ± | 5.75 |  | 83.26 ± | 4.89 |  |
|  | 64Cu-HCl | 97.32 ± | 5.23 |  | 98.79 ± | 2.18 | 103.90 ± | 3.32 |  | 106.09 ± | 3.09 |  |
| CNR | 18F-FDG | 34.26 ± | 0.58 |  | 24.60 ± | 0.72 | 86.55 ± | 3.65 |  | 64.19 ± | 3.44 |  |
|  | 68Ga-HCl | 28.56 ± | 4.00 |  | 21.27 ± | 2.44 | 72.51 ± | 5.00 |  | 56.54 ± | 3.32 |  |
|  | 64Cu-HCl | 46.08 ± | 2.48 |  | 33.24 ± | 0.73 | 76.40 ± | 2.44 |  | 57.87 ± | 1.69 |  |
| CoVBG (%) | 18F-FDG | 8.96 ± | 0.33 |  | 12.55 ± | 0.15 | 9.72 ± | 0.21 |  | 13.29 ± | 0.77 |  |
|  | 68Ga-HCl | 9.60 ± | 0.43 |  | 12.95 ± | 0.47 | 9.67 ± | 0.45 |  | 12.39 ± | 0.69 |  |
|  | 64Cu-HCl | 9.30 ± | 0.64 |  | 13.09 ± | 0.63 | 9.68 ± | 0.69 |  | 13.04 ± | 0.77 |  |

Legend Attachment – Table 1A: Mean and standard deviation of mean, maximal and peak Recovery Coefficient (RCmean, RCmax and RCpeak), percent contrast between each sphere and the background (Contrast Recovery Coefficient (CRC)), Contrast-Noise-Ratio (CNR) and percent background variability (CoVBG). Each parameter is presented as mean ± standard deviation of the six spheres for PET/CT measurements of the NEMA PET Body phantom filled with 18F-FDG, 68Ga-HCl and 64Cu-HCl, respectively, sphere-to-background activity concentration ratios of 4:1 and 8:1, respectively, and with (obese setup) and without (normal weight setup) the simulation of an obese patient by applying cooling packs around the phantom.

References: