Supporting Information

Table 1 Description,Time list and in process control of [11C]CO2 trapping and release.

| Time |  | Device |  | Value | Comment |
| --- | --- | --- | --- | --- | --- |
| 0 | Set | Power | = | on | V.210416 |
| 2 | Set | Flow rate Set Point | = | 5 |  |
| 4 | Set | V28 | = | Open | Cooling on |
| 6 | Set | V24 | = | a (down) |  |
| 8 | Set | V25 | = | Open |  |
| 10 | Set | V27 | = | Open |  |
| 12 | Set | Process Control | = | Show Message and wait | System ready to receive activity |
| t1+2 | Set | V25 | = | Close |  |
| t1+3 | Set | V26 | = | Open | C11 activity release |
| t1+5 | Set | V27 | = | Open |  |
| t1+10 | Set | Select counter | = | 3B on, 3a Off |  |
| t1+11 | Get | Activity Product |  |  |  |
| t1+4’0 | Set | V25 | = | Open |  |
| t1+4’0 | Set | V26 | = | Close | C11 Activity stop |
| t1+4’1 | Set | V28 | = | Close | Cooling off |
| t1+4’2 | Set | V27 | = | Open |  |
| t1+4’3 | Set | Temp. Set Point CH4 | = | 300 |  |
| t1+4’3 | Wait | Temp. Reg. Status CH4 | = | Temp OK |  |
| t2+0 | Set | VX | = | b (up) |  |
| t2+2 | Set | Temp. Set Point CH4 | = | 380 |  |
| t2+3 | Wait | Temp. Reg. Status CH4 | = | Temp. OK |  |
| t3+1 | Set | Flow Rate Set point | = | 10 |  |
| t3+45 | Set | Flow Rate Set Point | = | 15 |  |
| t3+1’30 | Set | Flow Rate Set Point | = | 25 |  |
| t3+2’15 | Set | Flow Rate Set Point | = | 40 |  |
| t3+3’0 | Set | V27 | = | Close |  |
| t3+3’0 | Set | V28 | = | Open |  |
| t3+3’1 | Set | V25 | = | Reset |  |
| t3+3’1 | Set | V25 | = | Close |  |

Table 2 Description, Time list and in process control of [11C]CH3I Production and release.

| Time |  | Device |  | Value | Comment |
| --- | --- | --- | --- | --- | --- |
| 0 | Set | Power | = | on | MeI for Choline Modified |
| 0.4 | Set | Temp. Reg. Switch MeI | = | on |  |
| 0.6 | Set | Flow Rate Set Point | = | 20 |  |
| 1 | Set | V25 | = | open |  |
| 5 | Set | V28 | = | open |  |
| 6 | Wait | Temp. CH4 | <= | 45 |  |
| t1+0 | Set | Process Control | = | Show message and wait | Ready for CH4 trap cooling |
| t2+0 | Set | Temp. Set Point CH4 Trap | = | -75 |  |
| t2+1 | Wait | Temp. Reg. Status CH4 Trap | = | Temp. OK |  |
| t3+0 | Set | Process Control | = | Show message and wait | System ready to receive activity |
| t4+0 | Set | Temp. Set Point I2 | = | 100 |  |
| t4+1 | Get | Flow Rate Set Point | = | 0 |  |
| t4+2 | Set | V25 | = | Close |  |
| t4+2.5 | Set | V26 | = | Open | C11 activity release |
| t4+3 | Set | V27 | = | Open |  |
| t4+4’9 | Set | V26 | = | Close | C11activity stop |
| t4+4’10 | Set | V24 | = | B (up) | H2 flow |
| t4+4’11 | Wait | Flow Rate Set Point | = | 200 |  |
| t4+4’12 | Set | V25 | = | Open |  |
| t4+4’13 | Set | Flow Rate | >= | 45 |  |
| t5+5 | Set | V27 | = | Close |  |
| t5+10 | Set | V28 | = | Close |  |
| t5+15 | Set | V25 | = | Close |  |
| t5+16 | Set | V24 | = | A (down) | Switch to He |
| t5+17 | Set | Flow Rate Set Point | = | 50 |  |
| t5+18 | Set | Temp. Set Point CH4 | = | 360 |  |
| t5+19 | Wait | Temp. Reg. Status CH4 | = | Temp. OK |  |
| t6+40 | Set | V10 Exhaust | = | Open |  |
| t6+41 | Set | V15 CH4 Trap Input | = | B (left) |  |
| t6+42 | Set | V29 | = | Open |  |
| t6+43 | Set | V25 | = | Open | Purge and transport with He |
| t6+1’40 | Set | V09 | = | B(right) |  |
| t6+1’41 | Set | V16 Exhaust | = | B(up) |  |
| t6+1’5 | Set | V25 | = | Close |  |
| t6+1’51 | Set | V29 | = | Close |  |
| t6+1’53 | Set | V15 CH4 Trap Input | = | A (right) |  |
| t6+1’54 | Set | V16 Exhaust | = | A(down) |  |
| t6+1’55 | Set | Gas Pump | = | On |  |
| t6+1’56 | Set | V16 Exhaust | = | B (up) |  |
| t6+1’56 | Set | Temp. Set Point CH4 Trap | = | 80 |  |
| t6+1’56 | Set | Temp. Reg. Switch CH4 | = | Off |  |
| t6+1'58 | Set | Select Heater | = | 3B on, 3 A off |  |
| t6+1'58 | Set | V16 Exhaust | = | A (down) |  |
| t6+4'20 | Set | Process Control | = | Show message and wait | Press o.k. when activity on MeI-Trap constant |
| t7+0 | Set | Temp. Set Point MeI Trap | = | 190 |  |
| t7+30 | Set | Gas Pump | = | Off | Circulation |
| t7+31.9 | Set | V15 CH4 Trap Input | = | B (left) |  |
| t7+32.9 | Set | V29 | = | Open |  |
| t7+33.9 | Set | V25 | = | Open |  |
| t7+34.9 | Set | V16 Exhaust | = | B (up) |  |
| t7+36 | Set | Flow Rate Set Point | = | 5 |  |
| t7+40 | Set | V17 | = | B (down) |  |
| t7+41 | Set | V16 Exhaust | = | A (down) |  |
| t7+42 | Set | V08 | = | Open |  |
| t7+43 | Wait | Temp. MeI Trap | >= | 185 |  |
| t8+1 | Set | Flow Rate Set Point | = | 15 |  |
| t8+1'10 | Set | Process Control | = | Show message and wait | Press Okay when MeI has been transferred |