Online supplement 2: MRI protocol

| Sequence | SPGR inspiration/ expiration | T2 weighted FS PROPELLER | ZTE | UTE | Fourier Decomposition |
| --- | --- | --- | --- | --- | --- |
| Acquisition plane | Sagittal | Axial | Axial | Axial | Coronal |
| TR/TE (ms) Flip angle (°) RF | 1.5/0.6 2 Selective | */73 90/120 Selective | 1.1/0 2 Non-selective | 5.2/0.032 3 Selective | 2.3/0.8 5 Selective |
| In-plane matrix | 120x120 | 340x340 | 200x200 | 228x228 | 128x128 |
| k-space trajectory | Cartesian | Blades | Radial | Cones | Cartesian |
| In-plane Field-of-view (FOV) RecFOV Actual voxel resolution (mm3) Slices Slice thickness | 36 0.75 3.0x3.0x3.0 130 3.0 | 34 - 1.0x1.0x3.5 50-70 3.5 | 30 - 1.5x1.5x1.5 200 1.5 | 34 - 1.5x1.5x1.5 230 1.5 | 50 - 3.9x3.9x15.0 7-10 15.0 |
| Receiver bandwidth (KHz) | 100/90/90 | 83.33 | 62.5 | 125 | 83.33 |
| Parallel imaging (ARC) | 1.5x1.2 | 3.0 | None | None | 2.0 |
| Number of averages Number of phases No. of spokes per segment | 1 1 - | 2.0 1 14 | 2.5 1 900 | 1 1 - | 1 200/186 ms per phase - |
| Physiological triggering | BH | Prospective pencil beam navigator | Prospective projection navigator | Prospective pneumobelt | Free breathing |
| Scan time (sec) RR=20 | +/- 6sec | +/- 5 min | +/- 6 min | +/- 6 min | 7-10x 38 sec |

Online supplement 2: MRI protocol. FS: fat suppression, FOV; Field Of View, NEX; Number of Excitations, PROPELLER; Periodically Overlapping ParalLEL Lines with Enhanced Reconstruction, RF; radio frequency, SPGR; Spoiled Gradient Echo Sequence, TE; Echo Time, TR; Repetition Time.*TR chosen varied on the respiratory frequency of the patient.