Table 1. Comparison of Common Emergency Use Authorized Diagnostic Methods for Evaluating COVID-19 Tests

A. Molecular Assays

Manufacturer / Platform			Method		RNA Targets	LoD (NDU/mL)		PPA(%)	NPA(%)

Abbott Molecular / Alinity m			RT-PCR	N1/N2		2700			>99.0		>99.0

Becton Dickenson / BDMax			RT-PCR	N1/N2		1800			>99.0		>99.0

BioFire Defense / FilmArray			RT-PCR	ORF1ab/ORF8	5400			>99.0		>99.0

Bio-Rad / QX200				ddPCR		N1/N2		600			>99.0		>99.0

CDC / ABI 7500					RT-PCR	N1/N2		18000			>99.0		>99.0

Cepheid / GenXpert				RT-PCR	N2/E		5400			>99.0		>99.0

Hologic / Panther				RT-TMA	ORF1ab	600			>99.0		>99.0

Perkin Elmer / ABI QuantStudio 5		RT-PCR	ORF1ab/N	180			>99.0		>99.0

Roche Molecular Systems / cobas 6800		RT-PCR	ORF1ab/E 	1800			>99.0		>99.0

Roche Molecular systems / cobas Liat		RT-PCR	ORF1ab/N	5400			>99.0		>99.0

Table 1. Comparison of Common Emergency Use Authorized Diagnostic Methods for Evaluating COVID-19 Tests (Continued)

B. Antigen Assays

Manufacturer / Platform			Method		Ag Targets	LoD (TCID50/mL)	PPA(%)	NPA(%)

Abbott Diagnostics / BinaxNow COVID-19 Ag	ICMA		N		140.6			84.6		98.5

AccessBio / CareStart COVID-19 Ag		ICMA		N		6.4 x 103		83.3		100.0

Becton Dickenson / Veritor			IMCA		N		1.4 x 102		84.0		100.0

Ellume Limited / Elumme COVID-19 Home	LFIA		N		1.0 x 103.8		91.0		96.0

Lumira Dx / LumeriaDx SARS-CoV-2 Ag test	LFIA		N 		2.8 x 105		97.6		96.6

Quidel / Sofia-2					LFIA		N		3.4 x 105		96.7		100.0

Abbreviations: ABI, Applied Biosciences; Ag, antigen; CDC, Centers for Disease Control and Prevention; ddPCR, digital droplet PCR; E, envelope protein gene; IMCA, immunochromographic membrane assay; LFIA, lateral flow immunofluorescent assay; LoD, limit of detection; N, nucleoprotein gene; NDU, nucleic acid test detectable units; NPA, negative percent agreement; ORF, open reading frame; PCR, polymerase chain reaction; PPA, positive percent agreement; RNA, ribonucleic acid; RT, reverse transcription; TCID50, median tissue culture infective dose; TMA, transcription mediated amplification.