Additional Table 1 (A-E): ProNGF ELISA Validation Experiments – Serum and Rinse

1A: Effect of heterophilic blocking antibodies (Ab) (BL-003-1000, Biosensis, Australia) on rate of positivity of serum proNGF levels.  No changes were observed in performance of the standard curve with presence or absence of blocking antibody (data not shown).

| Disease category | Serum samples assayed (n) | % of samples positive |  | p-value |
| --- | --- | --- | --- | --- |
|  |  | No Heterophilic Blocking Ab | With Heterophilic Blocking Ab |  |
| Thyroid cancer | 20 | 45% | 15% | 0.04 |
| Benign thyroid disease | 20 | 30% | 15% | 0.25 |
| No thyroid disease | 18 | 44% | 6% | 0.006 |

1B: Spike and Recovery Experiments. 4 serum samples were assayed before and after a 2.5ng/mL spike (50ng/mL following adjustment for dilution) of recombinant proNGF (Biosensis, Australia) at 1:20 dilution, in the presence of Heterophilic Blocking Antibodies (BL-003-1000).

| Sample | ProNGF (ng/mL) |  | % recovery of spike |
| --- | --- | --- | --- |
|  | Pre-spike | Post-spike |  |
| Sample A | 2.7 | 66.7 | 128% |
| Sample B | 0.0 | 47.8 | 96% |
| Sample C | 0.1 | 41.0 | 82% |
| Sample D | 0.0 | 40.1 | 80% |
| Mean |  |  | 96% |

1C: Inter-plate Quality Control Samples A quality-control solution (QC) was prepared by spiking 1:20 diluted proNGF-negative serum with 1.61ng/mL (32.15ng/mL after correction for dilution) recombinant proNGF with heterophilic blocking antibodies.  Aliquots of the QC sample were included on each ELISA plate.  Plate numbers reflect two different batches of ELISA plates.

| Plate | 1.1 | 1.2 | 1.3 | 1.4 | 1.5 | 1.6 | 2.1 | 2.2 | 2.3 | 2.4 | 2.5 | Mean (SD) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| proNGF (ng/mL) | 39.9 | 31.9 | 34.5 | 41.4 | 41.1 | 29.3 | 20.1 | 29.0 | 29.5 | 18.6 | 32.2 | 31.6 ± 7.3 |
| Percent Recovery | 124% | 99% | 107% | 129% | 128% | 91% | 63% | 90% | 92% | 58% | 100% | 98 ± 22% |

For the biopsy rinse, a similar QC solution was prepared by spiking 1.6ng/mL recombinant proNGF into phosphate buffered saline with the addition of protease inhibitors and heterophilic blocking antibodies (38µg/mL).  Aliquots of the QC sample were included on each ELISA plate.  Plate number represent different batches.

| Plate | 1.5 | 1.6 | 1.8 | 3.1 | 3.2 | 4.1 | 4.2 |  |  |  |  | Mean (SD) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| proNGF (ng/mL) | 2.03 | 1.53 | 1.77 | 2.27 | 2.41 | 1.56 | 1.55 |  |  |  |  | 1.87 ± 0.34 |
| Percent Recovery | 127% | 96% | 111% | 142% | 151% | 98% | 97% |  |  |  |  | 117 ± 20% |

1D: Linearity of Dilution 3 serum samples positive for endogenous proNGF were assayed in serial dilution using Assay Diluent A (Biosensis, Australia) and Heterophilic Blocking Antibodies (BL-003-1000, Biosensis, Australia).  The measured concentration at 1:20 dilution was defined as 100%. Reasonable linearity of dilution was observed.  The zero values observed correspond to expected values below the limit of detection.

| Sample | Endogenous ProNGF (ng/mL), corrected for dilution |  |  |  |
| --- | --- | --- | --- | --- |
|  | 1:10 | 1:20 | 1:40 | 1:80 |
| Sample A | 18.9 (109%) | 17.3 (100%) | 10.4 (60%) | 14.6 (84%) |
| Sample B | 2.6 (106%) | 2.4 (100%) | 0 | 0 |
| Sample C | 13.5 (124%) | 10.9 (100%) | 6.7 (62%) | 0 |

1E: Linearity of Dilution of Biopsy Rinse Diluent

In addition to the spike and recovery experiments performed above,  a 1 in 2 dilution series from 5ng/mL was prepared as the standard curve for biopsy rinse specimens using the same phosphate-buffered saline with the addition of protease inhibitors and heterophilic antibody blockers that was used to prepare the biopsy rinse specimens.  4 paramater logisitic regression curves are shown below for the first three ELISA plates, showing excellent linearity of dilution.