Comparison of the Optimal Design of Spinal Hybrid Elastic Rod for Dynamic Stabilization: A Finite Element Analysis.

Supplementary Figure 1.

The von-Mises stress distribution of the implanted L3–L4 disc in flexion. The disc stress in all models decreased when compared with INT.

INT, intact lumbar spine model.

Supplementary Figure 2.

The von-Mises stress on the inner NS markedly increased on the lower portion as the diameter decreased  in flexion, extension and lateral bending.

NS, nitinol stick.

Supplementary Table 1.

The biomechanical analysis of the five finite element models in flexion.

|  | INT |  | Nt45 |  | Nt35 |  | Nt25 |  | Nt15 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Preload (N) | 150 |  |  |  |  |  |  |  |  |  |
| Moment (Nm) | 14.4 |  | 18.7 |  | 18.7 |  | 18.7 |  | 18.4 |  |
| Intervertebral range of motion (degree) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 5.39 |  | 6.43 |  | 6.43 |  | 6.43 |  | 6.35 |  |
| L2-L3 | 5.71 |  | 6.81 |  | 6.81 |  | 6.81 |  | 6.72 |  |
| L3-L4 | 5.71 |  | 2.28 |  | 2.31 |  | 2.38 |  | 2.47 |  |
| L4-L5 | 7.28 |  | 8.60 |  | 8.60 |  | 8.60 |  | 8.51 |  |
| Total | 24.09 |  | 24.12 |  | 24.15 |  | 24.22 |  | 24.05 |  |
| Stress of intervertebral disc (Kpa) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 1080 |  | 1220 |  | 1220 |  | 1230 |  | 1270 |  |
| L2-L3 | 893 |  | 1130 |  | 1130 |  | 1110 |  | 1110 |  |
| L3-L4 | 810 |  | 543 |  | 547 |  | 548 |  | 559 |  |
| L4-L5 | 959 |  | 1200 |  | 1200 |  | 1180 |  | 1180 |  |
| Facet contact forces (N) |  |  |  |  |  |  |  |  |  |  |
|  | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right |
| L1-L2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| L2-L3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| L3-L4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| L4-L5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
| Stress of screws (MPa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 139 |  | 140 |  | 134 |  | 130 |  |
| Stress of PCU shell (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 4.7 |  | 3 |  | 3.5 |  | 4.9 |  |
| Stress of Nitinol stick (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 44 |  | 54.6 |  | 72.4 |  | 118 |  |

Supplementary Table 2.

The biomechanical analysis of the five finite element models in extension.

|  | INT |  | Nt45 |  | Nt35 |  | Nt25 |  | Nt15 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Preload (N) | 150 |  |  |  |  |  |  |  |  |  |
| Moment (Nm) | 15.1 |  | 18.4 |  | 18.4 |  | 18.4 |  | 18 |  |
| Intervertebral range of motion (degree) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 3.33 |  | 3.86 |  | 3.86 |  | 3.86 |  | 3.80 |  |
| L2-L3 | 3.17 |  | 3.65 |  | 3.65 |  | 3.65 |  | 3.60 |  |
| L3-L4 | 2.85 |  | 1.09 |  | 1.12 |  | 1.18 |  | 1.24 |  |
| L4-L5 | 3.27 |  | 3.97 |  | 3.95 |  | 3.97 |  | 3.88 |  |
| Total | 12.62 |  | 12.57 |  | 12.58 |  | 12.66 |  | 12.52 |  |
| Stress of intervertebral disc (Kpa) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 475 |  | 542 |  | 542 |  | 542 |  | 533 |  |
| L2-L3 | 488 |  | 556 |  | 556 |  | 556 |  | 548 |  |
| L3-L4 | 414 |  | 236 |  | 236 |  | 239 |  | 240 |  |
| L4-L5 | 312 |  | 439 |  | 436 |  | 439 |  | 430 |  |
| Facet contact forces (N) |  |  |  |  |  |  |  |  |  |  |
|  | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right |
| L1-L2 | 76 | 76 | 95 | 95 | 95 | 95 | 95 | 95 | 93 | 93 |
| L2-L3 | 94 | 94 | 117 | 117 | 118 | 118 | 118 | 118 | 115 | 115 |
| L3-L4 | 105 | 105 | 6 | 5 | 7 | 7 | 11 | 11 | 14 | 15 |
| L4-L5 | 102 | 102 | 129 | 128 | 128 | 128 | 128 | 128 | 125 | 125 |
| Stress of screws (MPa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 105 |  | 110 |  | 112 |  | 110 |  |
| Stress of PCU shell (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 5.1 |  | 2.7 |  | 3.1 |  | 4.9 |  |
| Stress of Nitinol stick (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 39.5 |  | 46 |  | 58.4 |  | 101 |  |

Supplementary Table 3.

The biomechanical analysis of the five finite element models in torsion.

|  | INT |  | Nt45 |  | Nt35 |  | Nt25 |  | Nt15 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Preload (N) | 150 |  |  |  |  |  |  |  |  |  |
| Moment (Nm) | 29.0 |  | 31.9 |  | 31.3 |  | 30.2 |  | 29.6 |  |
| Intervertebral range of motion (degree) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 3.85 |  | 4.13 |  | 4.07 |  | 3.96 |  | 3.91 |  |
| L2-L3 | 4.08 |  | 4.28 |  | 4.24 |  | 4.16 |  | 4.12 |  |
| L3-L4 | 4.60 |  | 3.79 |  | 4.01 |  | 4.22 |  | 4.40 |  |
| L4-L5 | 6.27 |  | 6.62 |  | 6.55 |  | 6.41 |  | 6.33 |  |
| Total | 18.80 |  | 18.82 |  | 18.87 |  | 18.75 |  | 18.76 |  |
| Stress of intervertebral disc (Kpa) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 683 |  | 740 |  | 728 |  | 707 |  | 694 |  |
| L2-L3 | 678 |  | 722 |  | 722 |  | 687 |  | 675 |  |
| L3-L4 | 751 |  | 611 |  | 611 |  | 696 |  | 731 |  |
| L4-L5 | 927 |  | 1020 |  | 1020 |  | 968 |  | 950 |  |
| Facet contact forces (N) |  |  |  |  |  |  |  |  |  |  |
|  | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right |
| L1-L2 | 0 | 336 | 0 | 371 | 0 | 363 | 0 | 349 | 0 | 342 |
| L2-L3 | 0 | 336 | 0 | 374 | 0 | 366 | 0 | 351 | 0 | 343 |
| L3-L4 | 0 | 336 | 0 | 213 | 0 | 245 | 0 | 285 | 0 | 323 |
| L4-L5 | 0 | 344 | 0 | 389 | 0 | 380 | 0 | 363 | 0 | 354 |
| Stress of screws (MPa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 211 |  | 162 |  | 107 |  | 98.9 |  |
| Stress of PCU shell (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 17.4 |  | 7.6 |  | 6 |  | 4.6 |  |
| Stress of Nitinol stick (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 219 |  | 324 |  | 383 |  | 381 |  |

Supplementary Table 4.

The biomechanical analysis of the five finite element models in lateral bending.

|  | INT |  | Nt45 |  | Nt35 |  | Nt25 |  | Nt15 |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Preload (N) | 150 |  |  |  |  |  |  |  |  |  |
| Moment (Nm) | 19.1 |  | 23.2 |  | 23.2 |  | 23.2 |  | 22.6 |  |
| Intervertebral range of motion (degree) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 6.12 |  | 7.25 |  | 7.25 |  | 7.24 |  | 7.10 |  |
| L2-L3 | 6.12 |  | 7.15 |  | 7.15 |  | 7.14 |  | 7.01 |  |
| L3-L4 | 6.17 |  | 2.63 |  | 2.77 |  | 2.96 |  | 3.16 |  |
| L4-L5 | 6.45 |  | 7.79 |  | 7.78 |  | 7.76 |  | 7.62 |  |
| Total | 24.86 |  | 24.82 |  | 24.95 |  | 25.10 |  | 24.89 |  |
| Stress of intervertebral disc (Kpa) |  |  |  |  |  |  |  |  |  |  |
| L1-L2 | 1230 |  | 1510 |  | 1500 |  | 1500 |  | 1470 |  |
| L2-L3 | 1160 |  | 1440 |  | 1430 |  | 1430 |  | 1400 |  |
| L3-L4 | 1130 |  | 512 |  | 544 |  | 585 |  | 621 |  |
| L4-L5 | 1040 |  | 1320 |  | 1320 |  | 1320 |  | 1290 |  |
| Facet contact forces (N) |  |  |  |  |  |  |  |  |  |  |
|  | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right |
| L1-L2 | 21 | 5 | 30 | 10 | 30 | 10 | 30 | 9 | 28 | 9 |
| L2-L3 | 53 | 31 | 82 | 38 | 82 | 37 | 80 | 35 | 74 | 33 |
| L3-L4 | 41 | 9 | 0 | 0 | 0 | 0 | 1 | 0 | 9 | 0 |
| L4-L5 | 55 | 0 | 61 | 0 | 61 | 0 | 61 | 0 | 59 | 0 |
| Stress of screws (MPa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 158 |  | 178 |  | 181 |  | 175 |  |
| Stress of PCU shell (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 6.8 |  | 5.1 |  | 5 |  | 7.2 |  |
| Stress of Nitinol stick (Mpa) |  |  |  |  |  |  |  |  |  |  |
| Maximum | - |  | 127 |  | 145 |  | 187 |  | 275 |  |