METHODS

Inclusion Criteria

- Ability to provide written informed consent

- Age ≥18 years

- Documented diagnosis of COPD according to guidelines

- Current use of non-invasive ventilation for ≥3 months

Exclusion Criteria

- Patients with unstable hemodynamic and respiratory conditions (e.g., recent acute heart failure, respiratory exacerbation or infection with worsening symptoms within the last month, or the need for immediate intubation

- Not compliant on NIV (i.e. usage <4 h/night)

- Pre-existing conditions: severe bullous lung disease, pneumothorax or pneumomediastinum, pulmonary embolism, Malignancy)

- Sleep disordered breathing (overlap syndrome, obesity hypoventilation)

- Pregnancy

- Tracheostomy

- PAP therapy contraindication

- Surgery of the upper airway, nose, sinus, or middle ear within the previous 90 days

- Inability to provide an informed consent

-Spirometry (VMax 6150, SensorMedics, Yorba Linda, CA) at the time of the study were performed. Lung function parameters (FEV1, FVC, FEV1/FVC) were obtained from all patients. In addition, predicted lung volumes and voluntary respiratory muscle strength values) MIP and MEP) were recorded as recommended

Arterial blood gas  (ABG) measurements

ABG (Model 850, Chiron Diagnostics, Medfield, MA) (pH, PaCO2, PaO2, PaO2/FiO2) were performed on room air for all patients. The sample was taken during the day (9am-5 pm). The blood gas was done during the initial assessment (after study consent) to ensure that patients were clinically stable prior to the study. The test was performed while the patient in the sitting position.

During an in-hospital stay of several days, eligible patients were commenced for treatment with Fixed EPAP (Group I) and auto-titrating EPAP (iVAPS -AE) (Group II) for 5 consecutive days according to hospital-internal clinical routine processes. Randomization was performed (1:1) by a laboratory scientist not involved with the study using the technique of shuffled sealed envelopes containing equal numbers of each treatment arm (Figure1).

Ventilatory settings: Astral 150, ResMed Corp. iVPAPs device algorithms

Intelligent volume assured pressure support ventilation (iVAPS) is an algorithm for the automatic adjustment of pressure support (PS), within pre-set boundaries, in order to achieve a target alveolar ventilation (Va).

Target Va (l/min) is the non-anatomical-deadspace ventilation. During therapy, actual Va passed through a low pass filter is compared to the target Va. PS is then adjusted to produce ventilatory changes equivocal to the magnitude and direction of the difference over the following 2-3 breaths. A maximum and minimum PS (IPAP minus EPAP) is set in accordance with patient tolerance (AutoVPAP max and min pressures 4 and 20cmH2O respectively).

Backup RR is variable to maintain optimal patient effort, with the least interference with triggering. The instantaneous backup rate is 2/3 of the preset target rate during spontaneous ventilation (i.e., it remains ‘in the background’ to give the patient the maximum opportunity to trigger spontaneously).

iVAPS settings were automatically determined following a 20 minute ‘learn’ period during awake spontaneous breathing, on the ventilator with 4 cmH2O EPAP. The ventilator automatically determined the treatment target alveolar ventilation (Va) and backup RR based on the patient’s spontaneous RR and flow rates, together with the estimated anatomical dead space: Vd = 100*(H/1.7)3 (Vd = dead space ventilation and H = height (or arm span of 2 outstretched arms) in meters). The EPAP was initially set to 4 cmH2O and IPAP minimum and maximum were set to allow the ventilator the greatest degree of freedom to achieve the target ventilation, according to patient tolerance. Mean target Va, was automatically determined by the intelligent algorithm. In both modes, the target tidal volume (VT) was set as 6–8 ml/kg (ideal body weight),

Nasal or oronasal ResMed masks were used. Patients were observed during a trial wake treatment and any issues with ventilation were addressed including pressure titration to comfort. Further titration of settings during the trial was permitted but record was kept of adjustments. Major adjustments were considered those that would directly affect the minute ventilation (ie. target Va, PS, EPAP, RR, inspiratory:expiratory ratio). Minor adjustments were those related to patient comfort (ie. rise time, fall time).

Ventilator data was downloaded onto ‘smartcards’ and returned after approximately 5 days of therapy on the initial ventilator.

Table I:. FixedEPAP iVAPS and AutoEPAP iVAPS setting guidelines for device parameters

| Parameter | Fixed-EPAP iVAPS | Auto-EPAP iVAPS |
| --- | --- | --- |
| EPAP (cmH2O) | As per optimum EPAP setting seen in ST mode | N/A |
| Min EPAP (cmH2O) | N/A | 5.0 cmH2O1 |
| Max EPAP (cmH2O) | N/A | Recommended: Default setting (15 cmH2O) |
| Min PS (cmH2O) | Recommended: Default setting (2 cmH2O) | Recommended: Default setting (2 cmH2O) |
| Max PS (cmH2O) | Recommended: Default setting (25 cmH2O) (Alternative: Clinician’s discretion*) | Recommended: Default setting (25 cmH2O) (Alternative: Clinician’s discretion*) |
| Patient height/ arm span (cm) | As measured by clinician | As measured by clinician |
| Ti min | Recommended: Default setting (0.3 sec) (Alternative: Clinician’s discretion*) | Recommended: Default setting (0.3 sec) (Alternative: Clinician’s discretion*) |
| Ti max | Recommended: Default setting (2.0 sec) (Alternative: Clinician’s discretion*) | Recommended: Default setting (2.0 sec) (Alternative: Clinician’s discretion*) |
| Rise Time | Recommended: Default setting (300 msec) (Alternative: Clinician’s discretion*) | Recommended: Default setting (300 msec) (Alternative: Clinician’s discretion*) |
| Trigger | Recommended: Default setting (Med) (Alternative: Clinician’s discretion*) | Recommended: Default setting (Med) (Alternative: Clinician’s discretion*) |
| Cycle | Recommended: Default setting (Med) (Alternative: Clinician’s discretion*) | Recommended: Default setting (Med) (Alternative: Clinician’s discretion*) |
| Ramp | Participant’s normal ramp | Participant’s normal ramp |
| Oxygen concentration | ≤4 L/min** | ≤4 L/min** |
| Humidification setting | At patient/clinician’s discretion | At patient/clinician’s discretion |
| Circuit settings | ‘AB filter’: ‘yes’ Mask type: Clinician’s discretion* | ‘AB filter’: ‘yes’ Mask type: Clinician’s discretion* |

Auto-EPAP= automatically adjusting EPAP; EPAP, expiratory positive airway pressure; iVAPS= intelligent volume-assured PS; COPD= chronic obstructive pulmonary disease; N/A= not applicable; PS=pressure support; Ti= inspiratory time; Va= alveolar ventilation.

*Clinician’s discretion = the investigator may change the setting if it is clinically appropriate and the default setting does not meet the clinical needs of the participant.

1Based on mean EPAP value from: Kohnlein T, et al. Non-invasive positive pressure ventilation for the treatment of severe stable chronic obstructive pulmonary disease: a prospective, multicenter, randomized, controlled clinical trial. Lancet Respir Med. 2014;2:698-705.