# S3: Analytical performance of the multiplex PCR (M-PCR) with the micro-organisms identified in culture (irrespective of threshold)

|  | Organism | True positive (Culture = M-PCR) | False positive (M-PCR + / Culture -) | False negative (Culture + / M-PCR -) | Se (%) [95% CI] | Sp (%) [95% CI] | PPV (%) [95% CI] | NPV (%) [95% CI] |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| # Gram-positive bacteria | Staphylococcus aureus | 8 | 0 | 4 | 67 | 100 | 100 | 95 |
|  | Streptococcus pneumoniae | 0 | 0 | 2 | 0 | 100 | - | 98 |
| # Enterobacteriaceae | Citrobacter freundii | 0 | 0 | 0 | - | 100 | - | 100 |
|  | Escherichia coli | 14 | 1 | 1 | 93 | 99 | 93 | 99 |
|  | Enterobacter cloacae complex | 4 | 2 | 4 | 50 | 98 | 67 | 96 |
|  | Enterobacter aerogenes | 1 | 0 | 0 | 100 | 100 | 100 | 100 |
|  | Proteus spp. | 7 | 0 | 4 | 64 | 100 | 100 | 95 |
|  | Klebsiella pneumoniae | 9 | 0 | 5 | 64 | 100 | 100 | 94 |
|  | Klebsiella oxytoca | 2 | 2 | 0 | 100 | 98 | 50 | 100 |
|  | Klebsiella variicola | 1 | 0 | 0 | 100 | 100 | 100 | 100 |
|  | Serratia marcescens | 5 | 0 | 0 | 100 | 100 | 100 | 100 |
|  | Morganella morganii | 2 | 0 | 2 | 50 | 100 | 100 | 98 |
| # Non-fermenting bacteria | Moraxella catarrhalis | 1 | 1 | 0 | 100 | 99 | 50 | 100 |
|  | Pseudomonas aeruginosa | 32 | 1 | 0 | 100 | 98 | 97 | 100 |
|  | Actinetobacter baumanii complex | 3 | 0 | 0 | 100 | 100 | 100 | 100 |
|  | Strenotrophomonas maltophilia | 1 | 2 | 0 | 100 | 98 | 33 | 100 |
|  | Legionella pneumophila | 2 | 0 | 0 | 100 | 100 | 100 | 100 |
| # Others | Pneumocystis jirovecii | 0 | 0 | 0 | - | 100 | - | 100 |
|  | Haemophilus influenzae | 3 | 0 | 1 | 75 | 100 | 100 | 99 |
|  | Mycoplasma pneumoniae | 0 | 0 | 0 | - | 100 | - | 100 |
|  | Chlamydophila pneumoniae | 0 | 0 | 0 | - | 100 | - | 100 |
| # Total |  | 95 | 9 | 23 | 81 [72-87] | 100 [99-100] | 91 [85-95] | 98 [98-99] |