Supplemental Table 1 Analytical Results for CTC enumeration assays on MiSelect R System.

Overall recovery efficiency across the entire spiked cell range.

| Number of spike-in cell | Test number | Mean number of cells recovered | Average recovery (%) | Standard deviation (%) | Coefficient of variation (%) |
| --- | --- | --- | --- | --- | --- |
| 1 | 50 | 0.98 | 98% | 14% | 14.3% |
| 8 | 6 | 7.5 | 93.8% | 6.3% | 6.7% |
| 32 | 66 | 27.9 | 87.2% | 7.6% | 8.7% |
| 128 | 6 | 112.5 | 87.9% | 6.5% | 7.4% |
| 512 | 6 | 479.5 | 93.7% | 4.4% | 4.7% |
| 1024 | 60 | 920 | 89.9% | 7.0% | 7.8% |

Limit of Detection and Limit of Background

| Number of spike-in cell | Test number | % positive for recovered cells |
| --- | --- | --- |
| 0 | 28 | 0% |
| 1 | 44 | 97.7% |
| 2 | 9 | 100% |
| 3 | 21 | 100% |

Precision

| Number of spike-in cell | Test number | Average recovery (%) | Coefficient of variation (%) |
| --- | --- | --- | --- |
| 1024 | 60 | 89.9% | 7.8% |
| 32 | 60 | 87.1% | 9.2% |

Supplemental Figure 1 SkBr3 recovery rate across the entire spiked cell range.