Supplementary information

Analytical performance studies

Precision was evaluated by following the CLSI EP05-A3 [1]. In brief, healthy human nasopharyngeal specimens were spiked with commercialized artificial SARS-CoV-2 particles (NanJing Leading Medical Co., Ltd), including wildtype, variant alpha, variant beta, variant delta, variant gamma and variant omicron. Two different concentrations for each spiked-in were prepared. Samples were measured in duplicate and were run 2 times per day for a total of 20 consecutive days. Three batches of the testing kits were used and the precision evaluation was performed in three different laboratories in parallel.

The limit of quantitation (LoQ) was determined as described by the Clinical and Laboratory Standards Institute (CLSI) EP17-A2 [2]. To establish the LoQ, five samples with concentrations of 0.318 pg/mL, 0.637 pg/mL, 0.955 pg/mL, 1.274 pg/mL and 1.592 pg/mL were prepared. The study was performed on one MAGLUMI X8 analyzer using one calibration curve and one lot of reagent (lot No. 43322031001). Each sample was run in 3 days, with 5 replicates per run each day. A calibration curve was plotted to determine the LoQ, which was the concentration corresponding to the 20% coefficient of variation (CV).

Linearity was assessed by using a series of mixed 6 sample pools. As specified in the CLSI EP06 [3], a high-level of SARS-CoV-2 antigen concentration from one patient’s nasopharyngeal sample was serially diluted with a low nasopharyngeal sample. Diluted sample pools were measured in triplicate. Linear regression analysis was employed to assess the performance of recovery.

In order to assess the analytical specificity of the system, forty-eight commercially available virus forms were studied respectively. For detailed information of the 48 virus forms, please see Supplemental Table 2. Healthy human nasopharyngeal specimens were spiked with one of the virus to the target concentration, and tested with the MAGLUMI SARS-CoV-2 Ag assay. In addition, 10 respiratory specimens from clinical patients with upper respiratory virus infection were measured, 2 of whom were diagnosed with seasonal coronavirus infection, 2 were adenovirus, 2 influenza A, 2 influenza B and 2 were respiratory syncytial infection. None of these samples were SARS-CoV-2 positive. Specimens used in analytical validation were not used for subsequent studies.

To comprehensively evaluate the detection performance on different virus variants, 5 different variants of SARS-CoV-2 were added to a pool of nasopharyngeal samples obtained from healthy donors respectively. The variants were from ZeptoMetrix, LLC., with the initial concentration unit of TCID50/mL. Serial dilutions with 5 concentrations for each variants were prepared for SARS-CoV-2 Antigen testing. The mean of results obtained from triplicate tests were plotted. And a fitted linear regression model was plotted against the corresponding theoretical variant concentration.

Analytical performance of MAGLUMI SARS-CoV-2 Ag (CLIA) tests

The default unit for the analyte detected by the MAGLUMI SARS-CoV-2 antigen tests was manufacturer’s arbitrary unit (AU)/mL. To ensure our final results are shown in an SI derived unit, we carried out a unit conversion experiment prior to the studies. A linear relationship between AU/mL and pg/mL was found (correlation coefficient factor R2=0.9976) in the range between 0.763 pg/mL to 6.104 pg/ml (Figure 1A), and the conversion factor to obtain pg/mL is 1.592 (i.e. 1 AU/mL = 1.592 pg/mL).

At the concentration of 2.2pg/mL and 8.5pg/mL, the average percentage of coefficient of variation (CV%) for within-run and between-run were both shown to be less than 5%, and the total imprecision was less than 8% (Supplemental Table 1).

Analytical validation established limit of quantitation (LoQ) at 0.399 pg/mL and linear range from 0.40 to 3184.00 pg/mL (Figure 1B,1C). A maximum of 20-fold dilution can be performed to achieve the clinical reportable range of 0.40 to 63680.00 pg/mL.

Analytical specificity was assessed with forty-eight commercially available virus forms in spiked-in human nasopharyngeal samples, respectively. Results showed negligible reaction to the SARS-CoV-2 antigen assay when compared to control specimens from healthy donors (Figure 1D). Similarly, compared with the control specimens, no significant change in relative light unit (RLU) was found in the tested respiratory samples of 10 SARS-CoV-2 negative respiratory specimens, including 2 Seasonal coronavirus, 2 adenovirus, 2 influenza A, 2 influenza B and 2 respiratory syncytial virus positive specimens (Figure 1D). These results indicated a well-performed analytical specificity of the assay.

Five serial diluents of five SARS-CoV-2 variants of concerns (VOCs) were detected respectively with MAGLUMI SARS-CoV-2 Ag (CLIA) assay. The fitted linear regression model plotted against the corresponding theoretical variant concentrations showed a remarkable linear correlation between the antigen concentration versus the infectious dose for different VOCs (Figure 1E).

Supplemental Table 1. Precision study on the MAGLUMI SARS-CoV-2 antigen assay. The mean ± standard deviation (SD) was shown in the unit of pg/mL.

| Sample IDa | Mean (n=720)b | Within-Run |  | Between-Run |  | Between-Day |  | Total |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | SD | %CV | SD | %CV | SD | %CV | SD | %CV |
| 1 | 2.207 | 0.086 | 3.90 | 0.049 | 2.22 | 0.046 | 2.08 | 0.134 | 6.07 |
| 2 | 8.468 | 0.358 | 4.23 | 0.116 | 1.37 | 0.239 | 2.82 | 0.489 | 5.77 |
| 3 | 2.058 | 0.073 | 3.55 | 0.059 | 2.87 | 0.056 | 2.72 | 0.124 | 6.03 |
| 4 | 8.764 | 0.349 | 3.98 | 0.193 | 2.20 | 0.217 | 2.48 | 0.514 | 5.86 |
| 5 | 2.192 | 0.094 | 4.29 | 0.051 | 2.33 | 0.062 | 2.83 | 0.134 | 6.11 |
| 6 | 8.543 | 0.342 | 4.00 | 0.180 | 2.11 | 0.209 | 2.45 | 0.513 | 6.00 |
| 7 | 2.211 | 0.091 | 4.12 | 0.056 | 2.53 | 0.038 | 1.72 | 0.124 | 5.61 |
| 8 | 8.554 | 0.328 | 3.83 | 0.204 | 2.38 | 0.188 | 2.20 | 0.514 | 6.01 |
| 9 | 2.082 | 0.086 | 4.13 | 0.041 | 1.97 | 0.059 | 2.83 | 0.127 | 6.10 |
| 10 | 8.797 | 0.339 | 3.85 | 0.220 | 2.50 | 0.220 | 2.50 | 0.570 | 6.48 |
| 11 | 2.211 | 0.089 | 4.03 | 0.045 | 2.04 | 0.057 | 2.58 | 0.124 | 5.61 |
| 12 | 8.466 | 0.360 | 4.25 | 0.209 | 2.47 | 0.174 | 2.06 | 0.661 | 7.81 |

a. Healthy human nasopharyngeal specimens spiked with commercialized artificial SARS-CoV-2 particles, including wildtype, variant alpha, variant beta, variant delta, variant gamma and variant omicron. Two different concentrations for each spiked-in were prepared.

b. Samples were measured in duplicate and were run 2 times per day for a total of 20 consecutive days. Three batches of the testing kits were used and the precision evaluation was performed in 3 different laboratories.

Supplemental Table 2. Overview of virus forms examined in the cross-reaction study.

| Virus type | Manufacturer | Catalogue number | Final concentration for the assay |
| --- | --- | --- | --- |
| HKU1 | Invent Diagnostica GmbH | DSPA 4.1.9.14.1 | 1.5×105 TCID50/mL |
| OC43 | ZeptoMetrix, LLC | 0810024CF | 1.5×105 TCID50/mL |
| NL63 | ZeptoMetrix, LLC | 0810228CF | 1.5×105 TCID50/mL |
| 229E | ZeptoMetrix, LLC | 0810229CF | 1.5×105 TCID50/mL |
| H1N1 (seasonal) | ZeptoMetrix, LLC | 0810244CF | 1.5×105 TCID50/mL |
| H3N2 | ZeptoMetrix, LLC | 0810238CF | 1.5×105 TCID50/mL |
| H5N1 | BEI Resources | NR-31131 | 1.5×105 TCID50/mL |
| H7N9 | BEI Resources | NR-44078 | 1.5×105 TCID50/mL |
| H1N1 (2009) | American Type Culture Collection | VR-1894™ | 5.2×105 TCID50/mL |
| Influenza B (Yamagata) | ZeptoMetrix, LLC | 0810518CF | 3.8×105 TCID50/mL |
| Influenza B (Victoria) | ZeptoMetrix, LLC | 0810571CF | 1.5×105 TCID50/mL |
| HPIV I | ZeptoMetrix, LLC | 0810014CF | 1.26×105 TCID50/mL |
| HPIV II | ZeptoMetrix, LLC | 0810015CF | 1.5×105 TCID50/mL |
| HPIV III | ZeptoMetrix, LLC | 0810016CF | 8.51×105 TCID50/mL |
| RSV type A | ZeptoMetrix, LLC | 0810040ACF | 1.5×105 TCID50/mL |
| RSV type B | ZeptoMetrix, LLC | 0810040CF | 1.5×105 TCID50/mL |
| Rhinovirus A | ZeptoMetrix, LLC | 0810012CFN | 1.5×105 TCID50/mL |
| Rhinovirus B | ZeptoMetrix, LLC | 0810038CF | 1.5×105 TCID50/mL |
| Rhinovirus C | ZeptoMetrix, LLC | 0810053CF | 1.9×105 TCID50/mL |
| Adenovirus 1 | ZeptoMetrix, LLC | 0810050CF | 3.09×105 TCID50/mL |
| Adenovirus 2 | ZeptoMetrix, LLC | 0810110CF | 1.5×105 TCID50/mL |
| Adenovirus 3 | ZeptoMetrix, LLC | 0810062CF | 1.5×105 TCID50/mL |
| Adenovirus 4 | ZeptoMetrix, LLC | 0810070CF | 1.5×105 TCID50/mL |
| Adenovirus 5 | ZeptoMetrix, LLC | 0810020CF | 2.7×105 TCID50/mL |
| Adenovirus 7 | ZeptoMetrix, LLC | 0810021CF | 1.5×105 TCID50/mL |
| Adenovirus 55 | ZeptoMetrix, LLC | 0810022CF | 3.6×105 TCID50/mL |
| Enterovirus A | ZeptoMetrix, LLC | 0810236CF | 1.5×105 TCID50/mL |
| Enterovirus B | ZeptoMetrix, LLC | 0810543CF | 1.5×105 TCID50/mL |
| Enterovirus C | ZeptoMetrix, LLC | 0810343CF | 5.5×105 TCID50/mL |
| Enterovirus D | ZeptoMetrix, LLC | 0810300CF | 2.2×105 TCID50/mL |
| EBV | ZeptoMetrix, LLC | 0810008CF | 1.5×105 TCID50/mL |
| Measles virus | ZeptoMetrix, LLC | 0810025CF | 1.5×105 TCID50/mL |
| HCMV | ZeptoMetrix, LLC | 0810003CF | 1.5×105 TCID50/mL |
| Rotavirus | ZeptoMetrix, LLC | 0810041CF | 2.4×105 TCID50/mL |
| Norwalk virus | BIOIVT & ELEVATING SCIENCE | HMN48675 | 1.5×105 TCID50/mL |
| Mumps virus | ZeptoMetrix, LLC | 0810079CF | 1.5×105 TCID50/mL |
| VZV | BIOIVT & ELEVATING SCIENCE | HMN48824 | 3.2×105 TCID50/mL |
| hMPV | ZeptoMetrix, LLC | 0810159CF | 1.5×105 TCID50/mL |
| M. pneumonia | ZeptoMetrix, LLC | 0801579 | 2.70×106 CCU/mL |
| C. pneumoniae | ZeptoMetrix, LLC | 0804392 | 1.70×106 IFU/mL |
| Haemophilus influenzae | ZeptoMetrix, LLC | 0801679 | 6.97×106 CFU/mL |
| S. aureus | ZeptoMetrix, LLC | 0804125 | 2.51×106 CFU/mL |
| S. pneumoniae | ZeptoMetrix, LLC | 0801439 | 1.34×106 CFU/mL |
| K. pneumoniae | ZeptoMetrix, LLC | 0801506 | 1.14×106 CFU/mL |
| M. tuberculosis | ZeptoMetrix, LLC | 0801660 | 1.49×106 CFU/mL |
| Canidia Albicans | ZeptoMetrix, LLC | 0801504 | 4.76×106 CFU/mL |
| MERS-CoV | ZeptoMetrix, LLC | 0810575UV | 1.5×105 TCID50/mL |
| SARS-CoV | BEI Resources | NR-9323 | 1.5×105 TCID50/mL |

References

- McEnroe RJ, Durham AP, Goldford MD, Kondratovich MV, Lababidi S, Magari R, et al. Evaluation of Precision of Quantitative Measurement Procedure; Approved Guideline. Clinical Laboratory Standards Institute 2014.

- Pierson-Perry JF, Vaks JE, Vore TEK, Durham AP, Fischer C, Gutenbrunner C, et al. Evaluation of detection capability for clinical laboratory measurement procedures; approved guideline. Clinical Laboratory Standards Institute 2012.

- McEnroe RJ, Durham AP, Kondratovich MV, Johansen JV, Meyers PG, Souers RJ, Vaks JE. Evaluation of linearity of quantitative measurement procedures. EP06. Clinical and Laboratory Standards Institute 2020.