Supplementary

Rapid Detection of SARS-CoV-2 Variants by Molecular Clamping Technology Based RT-qPCR

Shuo Shen1#, Andrew Y. Fu1#, Maidar Jamba1# Jonathan Li1, Zhen Cui1, Larry Pastor1, Daniel Cataldi1, Qing Sun1, Joseph A. Pathakamuri2, Daniel Kuebler2, Michael Rohall2, Madison Krohn2, Daniel Kissinger2, Jocelyn Neves2, Isaac Archibeque 2, Mike J. Powell1, Aiguo Zhang1, Chuanyi M. Lu3, and Michael Y. Sha1*

1DiaCarta Inc., 4385 Hopyard Rd., Pleasanton, CA 94588. 2 Franciscan University of Steubenville, Steubenville, OH. 43952. 3 University of California San Francisco and San Francisco VA Healthcare System, San Francisco, CA 94121.

# Equal contribution

*Correspondence: msha@diacarta.com

S. Table 1. Interpretation of the qPCR Data

| PCR Tube | Target | Variants |  |  |  |  |  |  |  | Wildtype |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Alpha B.1.1.7 | Beta B.1.351 | Gamma P.1 | Delta B.1.617.2 | Delta plus | Omicron B.1.1.529 | Epsilon B.1.427/B.1.429 | Kappa B.1.617.1 | Positive | Negative | Invalid |
| A | D614G | √ | √ | √ | √ | √ | √ | √ | √ |  |  |  |
|  | L452R |  |  |  | √ | √ |  | √ | √ |  |  |  |
| B | K417T |  |  | √ |  |  |  |  |  |  |  |  |
|  | N501Y | √ | √ | √ |  |  | √ |  |  |  |  |  |
|  | RP | √ | √ | √ | √ | √ | √ | √ | √ | √ | √ | NEG |
| C | T478K |  |  |  | √ | √ | √ |  |  |  |  |  |
|  | K417N |  | √ |  |  | √ | √ |  |  |  |  |  |
|  | ORF1ab | √ | √ | √ | √ | √ | √ | √ | √ | √ |  |  |
| Status |  | Valid | Valid | Valid | Valid | Valid | Valid | Valid | Valid | Valid | Valid | Invalid |
| Result |  | Alpha | Beta | Gamma | Delta | Delta Plus | Omicron | Epsilon | Kappa | POS | NEG |  |
| Action |  | Report result |  |  |  |  |  |  |  | Report result as wildtype | Report result as negative | Repeat test one more time. If the repeat result retains invalid, consider collection new specimen |

S. Table 2a. XNA Synthesis (Part I) Matrix-assisted laser desorption/ionization mass spectrometry (MALDI-MS) data confirmed the structure and molecular weights (MW) of D614 / N501 XNAs synthesized by the SPPS method.

| XNA Code | XNA Length | Theoretical MW (Dalton) | Measured MW (Dalton) |
| --- | --- | --- | --- |
| D614 XNA001 | 25-mer | 7245 | 7244 |
| D614 XNA002 | 22-mer | 6453 | 6451 |
| D614 XNA003 | 20-mer | 5914 | 5910 |
| D614 XNA004 | 18-mer | 5372 | 5376 |
| N501 XNA001 | 17-mer | 5097 | 5105 |
| N501 XNA002 | 19-mer | 5599 | 5599 |
| N501 XNA003 | 20-mer | 5936 | 5933 |
| Bovine Cytochrome C (MW Standard) | N/A | 12232 | 12228 |

S. Table 2b. XNA Synthesis (Part II) Matrix-assisted laser desorption/ionization mass spectrometry (MALDI-MS) data confirmed the structure and molecular weights (MW) of T478 / L452 / K417 XNAs synthesized by the SPPS method.

| XNA Code | XNA Length | Theoretical MW (Dalton) | Measured MW (Dalton) |
| --- | --- | --- | --- |
| T478 XNA001 | 19-mer | 5624 | 5609 |
| T478 XNA002 | 17-mer | 5098 | 5082 |
| L452 XNA001 | 23-mer | 6742 | 6721 |
| L452 XNA002 | 25-mer | 7283 | 7264 |
| L452 XNA003 | 24-mer | 7033 | 7009 |
| K417 XNA001 | 23-mer | 6673 | 6657 |
| K417 XNA002 | 21-mer | 6147 | 6133 |
| Bovine Cytochrome C (MW Standard) | N/A | 12232 | 12224 |

S. Table 3 Summary of SARS-CoV-2 D614G XNA Evaluation and Selection

| XNA Concentration (µM) | XNA001 |  |  |  |  | XNA002 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Wild Type |  | Mutant |  |  | Wild Type |  | Mutant |  |  |
|  | Ct | Ct mean | Ct | Ct mean | Δ Ct | Ct | Ct mean | Ct | Ct mean | Δ Ct |
| 8 µM | 50.0 | 50.0 | 33.7 | 34.1 | 15.9 | 31.6 | 31.3 | 30.9 | 30.7 | 0.6 |
|  | 50.0 |  | 34.5 |  |  | 30.9 |  | 30.7 |  |  |
|  | 50.0 |  | 34.0 |  |  | 31.3 |  | 30.5 |  |  |
| 4 µM | 50.0 | 41.2 | 30.9 | 30.9 | 10.3 | 31.1 | 30.8 | 30.3 | 30.3 | 0.5 |
|  | 36.5 |  | 30.9 |  |  | 30.7 |  | 30.3 |  |  |
|  | 37.2 |  | 30.8 |  |  | 30.5 |  | 30.3 |  |  |
| 2 µM | 36.6 | 36.5 | 30.3 | 30.8 | 5.6 | 30.8 | 30.5 | 30.1 | 32.5 | –2.0 |
|  | 36.2 |  | 31.9 |  |  | 30.4 |  | 37.3 |  |  |
|  | 36.6 |  | 30.3 |  |  | 30.3 |  | 29.9 |  |  |
| 1 µM | 36.5 | 36.2 | 30.1 | 30.0 | 6.3 | 30.5 | 30.7 | 29.9 | 30.0 | 0.7 |
|  | 36.5 |  | 30.0 |  |  | 30.7 |  | 30.2 |  |  |
|  | 35.7 |  | 29.8 |  |  | 30.9 |  | 29.9 |  |  |
| 0 µM | 29.4 | 29.3 | 30.0 | 30.2 | –0.8 | 30.2 | 30.6 | 30.6 | 31.3 | –0.7 |
|  | 29.0 |  | 30.0 |  |  | 30.3 |  | 31.6 |  |  |
|  | 29.5 |  | 30.5 |  |  | 31.2 |  | 31.6 |  |  |

S. Table 4. Summary of SARS-CoV-2 N501Y XNA Evaluation and Selection

| XNA Concentration (mM) | XNA 001 |  |  |  |  | XNA 002 |  |  |  |  | XNA-003 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Wild Type |  | Mutant |  |  | Wild Type |  | Mutant |  |  | Wild Type |  | Mutant |  |  |
|  | Ct | Ct mean | Ct | Ct mean | Δ Ct | Ct | Ct mean | Ct | Ct mean | Δ Ct | Ct | Ct mean | Ct | Ct mean | Δ Ct |
| 0.5mM | 50.0 | 50 | 17.7 | 15.4 | 34.6 | 50.0 | 50.0 | 17.4 | 15.9 | 34.1 | 50.0 | 50.0 | 17.8 | 17.9 | 32.1 |
|  | 50.0 |  | 17.3 |  |  | 50.0 |  | 17.0 |  |  | 50.0 |  | 17.9 |  |  |
|  | 50.0 |  | 11.3 |  |  | 50.0 |  | 13.3 |  |  | 50.0 |  | 18.1 |  |  |
| 0.25mM | 21.4 | 21.8 | 17.8 | 17.6 | 3.9 | 50.0 | 50.0 | 17.6 | 17.7 | 32.3 | 50.0 | 50.0 | 16.8 | 17.3 | 32.7 |
|  | 22.7 |  | 16.8 |  |  | 50.0 |  | 17.3 |  |  | 50.0 |  | 17.5 |  |  |
|  | 21.1 |  | 18.1 |  |  | 50.0 |  | 18.2 |  |  | 50.0 |  | 17.6 |  |  |
| 0.00mM | 17.7 | 17.4 | 17.0 | 17.0 | 0.8 | 17.7 | 17.3 | 17.3 | 16.6 | 1.0 | 16.9 | 17.5 | 16.5 | 16.5 | 0.5 |
|  | 17.4 |  | 16.9 |  |  | 17.3 |  | 16.7 |  |  | 17.8 |  | 16.6 |  |  |
|  | 16.9 |  | 17.0 |  |  | 17.0 |  | 15.8 |  |  | 17.8 |  | 16.4 |  |  |

S. Table 5. Summary of Cross-Reactivity Between SARS-CoV-2 Variant Detection Kit and Organisms Tested

| Organisms | Tube A |  | Tube B |  |  | Tube C |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | D614G | L452R (Cy5) | K417T (FAM) | RP | N501Y (Cy5) | Orf | K417N (HEX) | T478K |
|  | (TEX) |  |  | (HEX) |  | (FAM) |  | (Cy5) |
| Adenovirus Type 3 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| B. pertussis | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| C. pneumoniae | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Coronavirus 229E | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Coronavirus HKU-1 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Coronavirus NL63 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Coronavirus OC43 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Influenza A 2009 H1N1pdm | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Influenza A H1N1 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Influenza A H3N2 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Influenza B | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| M. pneumoniae | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Metapneumovirus | 45.0 | 45.0 | 45.0 | 36.5 | 45.0 | 45.0 | 45.0 | 45.0 |
| Parainfluenza Type 1 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Parainfluenza Type 2 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Parainfluenza Type 3 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Parainfluenza Type 4 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Rhinovirus 1A | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| RSV A | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| MERS-coronavirus | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| SARS-coronavirus | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| Negative Control | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |
| PC | 29.4 | 27.0 | 22.1 | 45.0 | 22.4 | 21.3 | 26.5 | 23.7 |
| EC | 45.0 | 45.0 | 45.0 | 36.5 | 45.0 | 45.0 | 45.0 | 45.0 |
| NTC | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 | 45.0 |

*PC, positive control; EC, extraction control; NTC, No target control.

S. Table 6. Summary of SARS-CoV-2 Alpha Variant Detection in Early 2021*

| Sample Collection Time | RT-qPCR Test | N501Y | D614G | ORF1ab ** | RNase P gene |
| --- | --- | --- | --- | --- | --- |
| Jan 2021 | Positive | 0 | 58 | 132 | 139 |  |
|  | Total (N) | 139 | 139 | 139 | 139 |  |
|  | Alpha variant Detection Rate (%) | 0% | 41.7% | 95.0% | 100.0% |  |
| Feb 2021 | Positive | 7 | 78 | 139 | 139 |  |
|  | Total | 139 | 139 | 139 | 139 |  |
|  | Alpha variant Detection Rate (%) | 5.0% | 56.1% | 100.0% | 100.0% |  |
| Mar-2021 | Positive | 10 | 41 | 96 | 96 |  |
|  | Total | 96 | 96 | 96 | 96 |  |
|  | Alpha variant Detection Rate (%) | 10.4% | 42.7% | 100.0% | 100.0% |  |

*374 clinical samples collected in middle January, late February and March 2021 in California bay area

**ORF1ab gene were targeted for wild type SARS-CoV-2.

S. Table 7. Summary of Delta Variant Detection for Breakthrough Patients by SARS-CoV-2 Delta Plus Detection Kit

| Sample ID | Quantivirus TM SARS-CoV-2 delta + Detection |  |  |  | Result |
| --- | --- | --- | --- | --- | --- |
|  | ORF (FAM) | K417N(HEX) | T478K(Cy5) | RP(TEX) |  |
| DC001 | 18.8 | 50 | 22.3 | 25.6 | Delta |
| DC002 | 50.0 | 50 | 50 | 25.3 | Negative |
| DC003 | 19.7 | 50 | 23.2 | 27.2 | Delta |
| DC004 | 18.1 | 50 | 21.6 | 27.2 | Delta |
| DC005 | 19.2 | 50 | 23 | 25.5 | Delta |
| DC006 | 17.5 | 50 | 21.1 | 25 | Delta |
| DC007 | 19.4 | 50 | 23.6 | 30.1 | Delta |
| DC008 | 14.8 | 50 | 18.8 | 24.6 | Delta |
| DC009 | 50.0 | 50 | 48.5 | 26.2 | Negative |
| DC010 | 16.6 | 50 | 20.9 | 24.3 | Delta |
| DC011 | 24.5 | 50 | 29.2 | 23.8 | Delta |
| DC012 | 20.7 | 50 | 24.5 | 24.5 | Delta |
| DC013 | 18.3 | 50 | 22 | 26.8 | Delta |
| DC014 | 18.1 | 50 | 21.8 | 22 | Delta |
| DC015 | 19.9 | 50 | 23.5 | 21 | Delta |
| DC016 | 24.9 | 50 | 27.8 | 22.9 | Delta |
| DC017 | 17.1 | 50 | 20.6 | 23.6 | Delta |
| DC018 | 17.4 | 50 | 21.2 | 26.3 | Delta |
| DC019 | 19.3 | 50 | 23 | 27.4 | Delta |
| DC020 | 18.1 | 50 | 21.7 | 25.7 | Delta |
| DC021 | 31.1 | 50 | 36.2 | 30.2 | Delta |
| DC022 | 23.1 | 50 | 27.2 | 28.2 | Delta |
| DC023 | 17.2 | 50 | 20.8 | 22.4 | Delta |
| DC024 | 22.0 | 50 | 26.1 | 30.7 | Delta |
| DC025 | 22.8 | 50 | 26.7 | 22.8 | Delta |
| DC026 | 20.9 | 50 | 25.3 | 24.6 | Delta |
| DC027 | 21.1 | 50 | 25.3 | 24.8 | Delta |
| DC028 | 16.0 | 20.1 | 17.12 | 23.2 | Delta Plus |
| DC029 | 17.0 | 50 | 20.5 | 23 | Delta |
| DC030 | 16.7 | 50 | 20.2 | 26.5 | Delta |
| DC031 | 19.1 | 50 | 22.9 | 21.1 | Delta |
| DC032 | 50.0 | 50 | 50 | 21 | Negative |
| DC033 | 50.0 | 50 | 50 | 22.7 | Negative |
| DC034 | 17.3 | 50 | 20.9 | 26.4 | Delta |
| DC035 | 50.0 | 50 | 50 | 23.6 | Negative |
| DC036 | 17.7 | 22.38 | 19.42 | 24.07 | Delta Plus |
| DC037 | 18.6 | 50 | 22.4 | 22.5 | Delta |
| DC038 | 21.9 | 50 | 26.1 | 25.2 | Delta |
| DC039 | 19.6 | 50 | 23.6 | 24 | Delta |
| DC040 | 22.7 | 50 | 26.9 | 23.7 | Delta |
| DC041 | 50.0 | 50 | 50 | 23.2 | Negative |
| DC042 | 24.4 | 50 | 28.8 | 25.1 | Delta |
| DC043 | 26.7 | 50 | 31.9 | 25.3 | Delta |
| DC044 | 14.9 | 19.33 | 16.88 | 22.43 | Delta Plus |
| DC045 | 19.0 | 50 | 23 | 26.4 | Delta |

* All data were confirmed by Sanger Sequence

S. Table 8. Summary of Omicron Variant Detection for Breakthrough Patients by SARS-CoV-2 Variant Detection Kit

| Sample # | Patient ID | QuantiVirus SARS-CoV-2 Variant Detection |  |  |  |  |  |  |  | Variant Identify | QuantiVirus SARS-CoV-2 Test |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | K417T | ORF | RP | K417N | D614G | L452R | N501Y | T478K |  |  |
| 1 | UCSF 2129242 | x | 27.7 | 30.1 | 33.5 | 45.0 | x | 34.5 | 30.2 | Omicron | positive |
| 2 | UCSF 2129224 | x | 32.6 | 28.1 | 39.5 | x | x | 37.9 | 34.8 | Omicron | positive |
| 3 | UCSF 2129226 | x | x | 28.5 | x | x | x | x | x | Negative | Negative |
| 4 | UCSF 2128946 | x | 27.9 | 29.2 | 33.7 | 44.3 | x | 33.8 | 31.8 | Omicron | positive |
| 5 | UCSF 2129234 | x | 29.2 | 29.4 | 35.8 | 39.2 | x | 35.6 | 30.1 | Omicron | positive |
| 6 | UCSF 2129277 | x | 24.3 | 27.0 | 29.2 | 36.7 | x | 29.2 | 25.7 | Omicron | positive |
| 7 | UCSF 2128799 | x | 23.9 | 28.5 | 28.5 | 36.0 | x | 30.5 | 26.3 | Omicron | positive |
| 8 | UCSF 2128823 | x | 25.1 | 29.6 | x | 38.9 | x | 37.9 | 26.0 | Omicron? | positive |
| 9 | UCSF 2120265 | x | 29.2 | 29.4 | 35.8 | 36.1 | x | 37.4 | 17.1 | Omicron | positive |
| 10 | UCSF 2129243 | x | 17.9 | 27.6 | 21.5 | 29.1 | x | 24.1 | 19.4 | Omicron | positive |
| 11 | UCSF 2129249 | x | 27.6 | 30.5 | 33.4 | 44.8 | x | 33.5 | 31.1 | Omicron | positive |
| 12 | UCSF 2129233 | x | x | 29.7 | x | x | x | x | X | Negative | Negative |
| 13 | UCSF 2128806 | x | 22.1 | 27.3 | 26.3 | 35.5 | x | 28.8 | 24.6 | Omicron | positive |
| 14 | UCSF 2129036 | x | 18.7 | 30.1 | 23.0 | 30.5 | x | 23.6 | 20.4 | Omicron | positive |
| 15 | UCSF 2128815 | x | 23.6 | 32.3 | 28.6 | 36.5 | x | 27.4 | 27.5 | Omicron | positive |
| 16 | DC 2201040613 | x | 27.4 | 23.8 | 32.9 | 45.6 | x | 35.2 | 31.2 | Omicron | positive |
| 17 | DC2201040623 | x | 24.9 | 25.3 | 30.1 | 37.4 | x | 31.7 | 26.7 | Omicron | positive |
| 18 | DC2201040058 | x | 27.6 | 22.3 | 33.2 | x | x | 36.1 | 32.7 | Omicron | positive |
| 19 | DC2201040126 | x | 24.1 | 26.0 | 28.7 | 37.0 | x | 29.7 | 27.5 | Omicron | positive |
| 20 | DC2201050186 | x | 18.5 | 24.1 | 22.6 | 30.1 | x | 25.4 | 21.4 | Omicron | positive |
| 21 | DC2201050213 | x | 28.8 | 25.0 | 36.9 | 45.1 | x | 48.7 | 32.1 | Omicron | positive |
| 22 | DC2201070132 | x | 15.6 | 22.1 | 19.5 | 28.4 | x | 25.1 | 18.3 | Omicron | positive |
| 23 | DC2201070137 | x | x | 25.1 | x | x | x | x | x | Negative | Negative |
| 24 | FSU0001 | x | 21.2 | 24.4 | 25.1 | 29.7 | x |  | 26.5 | OMICRON | positive |
| 25 | FSU0002 | x | 21.4 | 23.8 | 36.7 | 33.5 | x |  | 39.5 | OMICRON | positive |
| 26 | FSU0003 | x | 16.2 | 22.6 | 19.1 | 23.4 | x |  | 21.6 | OMICRON | positive |
| 27 | FSU0004 | x | 20.1 | 25.4 | 23.1 | 28.5 | x |  | 25.4 | OMICRON | positive |
| 28 | FSU0005 | x | 19.2 | 18.3 | 23.8 | 27.9 | x |  | 25.2 | OMICRON | positive |
| 29 | FSU0006 | x | 19.4 | 25.8 | 22.2 | 25.4 | x |  | 23.9 | OMICRON | positive |
| 30 | FSU0007 | x | 21.6 | 24.0 | 24.9 | 31.6 | x |  | 27.0 | OMICRON | positive |
| 31 | FSU0008 | x | 20.0 | 26.9 | 22.8 | 29.0 | x |  | 25.1 | OMICRON | positive |
| 32 | FSU0009 | x | 16.3 | 25.4 | 21.0 | 23.8 | x |  | 22.0 | OMICRON | positive |
| 33 | FSU0010 | x | 17.2 | 24.2 | 21.3 | 20.1 | x |  | 23.5 | OMICRON | positive |
| 34 | FSU0011 | x | 20.1 | 25.4 | 23.7 | 28.3 | x |  | 24.7 | OMICRON | positive |
| 35 | FSU0012 | x | 24.1 | 27.8 | 28.1 | 35.0 | x |  | 30.6 | OMICRON | positive |
| 36 | FSU0013 | x | 22.0 | 26.6 | 26.1 | 33.3 | x |  | 28.0 | OMICRON | positive |
| 37 | FSU0014 | x | 25.7 | 31.8 | 30.2 | 34.0 | x |  | 31.3 | OMICRON | positive |

S. Table 9. Comparison of SARS-CoV-2 Variant Detection RT-qPCR, Sanger Sequencing and NGS

| Patient ID | QuantiVirus SARS-CoV-2 Variant Detection test |  |  |  |  |  |  |  | RT-qPCR /Sanger Confirmed | NGS |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | D614G | N501Y | T478K | L452R | K417N | K417T | ORF | RP |  |  |
| 2129242 | 45.0 | 34.5 | 30.2 | x | 33.5 | x | 27.7 | 30.1 | Omicron | Omicron |
| 2128799 | 36.0 | 30.5 | 26.3 | x | 28.5 | x | 23.9 | 28.5 | Omicron | Omicron |
| 2128806 | 35.5 | 28.8 | 24.6 | x | 26.3 | x | 22.1 | 27.4 | Omicron | Omicron |
| 2128815 | 36.5 | 27.4 | 27.5 | x | 28.6 | x | 23.6 | 32.3 | Omicron | Omicron |
| 2128823 | 38.9 | x | 26.1 | x | x | x | 25.1 | 29.6 | Delta | Delta |
| 2128946 | 44.3 | 33.8 | 31.8 | x | 33.7 | x | 27.9 | 29.2 | Omicron | Omicron |
| 2129036 | 30.5 | 23.6 | 20.4 | x | 23.0 | x | 18.7 | 30.1 | Omicron | Omicron |
| 2129224 | x | 37.9 | 34.9 | x | 39.5 | x | 32.6 | 28.1 | Omicron | Omicron |
| 2129233 | x | x | x | x | x | x | x | 29.7 | Negative | NA |
| 2129234 | 39.2 | 35.6 | 30.1 | x | 35.8 | x | 29.2 | 29.4 | Omicron | Omicron |
| 2129243 | 29.1 | 24.1 | 19.4 | x | 21.5 | x | 17.9 | 27.6 | Omicron | Omicron |
| 2129249 | 44.8 | 33.5 | 31.1 | x | 33.4 | x | 27.6 | 30.5 | Omicron | Omicron |
| 2129265 | 36.1 | 37.4 | 17.1 | x | 35.8 | x | 29.2 | 29.4 | Omicron | Omicron |
| 2129266 | x | x | x | x | x | x | x | 28.5 | Negative | NA |
| 2129277 | 36.7 | 29.2 | 25.7 | x | 29.2 | x | 24.3 | 27.0 | Omicron | Omicron |
| DC1 | 24.3 | 25.7 | x | x | x | x | 24.7 | 28.9 | Alpha | Alpha |
| DC2 | 33.4 | 33.7 | x | x | x | x | 34.0 | 32.6 | Alpha | Alpha |
| DC3 | 20.5 | 18.6 | x | x | x | x | 20.0 | 26.5 | Alpha | Alpha |
| DC4 | 21.5 | 22.7 | x | x | x | x | 22.3 | 25.4 | Alpha | Alpha |
| DC5 | 26.9 | 23.6 | x | x | x | x | 25.9 | 27.0 | Alpha | Alpha |
| DC6 | 34.8 | 32.0 | x | x | x | x | 35.4 | 23.6 | Alpha | Alpha |
| DC7 | 29.3 | 20.7 | x | x | x | x | 29.4 | 27.7 | Alpha | Alpha |
| DC8 | 35.0 | 24.4 | x | x | x | x | 30.7 | 25.7 | Alpha | Alpha |
| DC9 | x | x | x | x | x | x | x | 26.5 | Negative | Negative |
| DC10 | x | x | x | x | x | x | x | 24.6 | Negative | Negative |