Supplementary Information

Assessment of different experimental setups to determine viral filtration efficiency of face masks

Arijana Filipić1, Katja Fric1, Maja Ravnikar1, Polona Kogovšek1

1Department of Biotechnology and Systems Biology, National Institute of Biology, Ljubljana, Slovenia

Supplementary tables S1-S6.  Filtration efficiency results for all mask samples in experimental setups I-VI. CFU, colony forming units; PFU, plaque forming units; PC, positive control; NC, negative control; - this measurement was not included in the experiment. When the BFE and VFE values were between 99% and 100%, more decimal places were included to show the exact filtration efficiency of the mask.

Supplementary Table S1. Results for all mask samples in experimental setup I

| Mask sample | Average concentration of two PCs (CFU) | BFE in subsamples (%) |  |  |  |  | NC (CFU) | Average mean particle size (µm) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | 2.70E+03 | 99.93 | 99.9 | 99.93 | - | - | 0 | 3.1 |
|  | 2.51E+03 | 99.4 | 99.7 | 99.7 | - | - | 0 | 3.1 |
|  | 2.42E+03 | 99.9 | 99.8 | 99.7 | - | - | - | 3.0 |
| B | 2.80E+03 | 97 | 98 | 97 | 98 | 97 | 0 | 2.9 |
|  | 3.06E+03 | 97 | 97 | 97 | 97 | 96 | 0 | 3.0 |
|  | 2.92E+03 | 95 | 94 | 96 | 97 | 96 | 0 | 3.0 |
|  | 1.28E+03 | 97 | 95 | 97 | - | - | 1 | 3.0 |
| C | 1.96E+03 | 92 | 89 | 93 | - | - | 0 | 3.1 |
| D | 2.38E+03 | 73 | 76 | 79 | 80 | 86 | 0 | 2.9 |
| E | 3.07E+03 | 99.93 | 99.97 | 99.8 | 100 | 99.97 | 1 | 3.0 |
| F | 2.41+E03 | 88 | 93 | 91 | 88 | 92 | 0 | 3.0 |

Supplementary Table S2. Results for all mask samples in experimental setup II

| Mask sample | Average concentration of two PCs (CFU/mL) | BFE in subsamples (%) |  |  |  |  | NC (CFU/mL) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | 4.46E+03 | 99.8 | 99.8 | 99.9 | - | - | 0 |
|  | 1.38E+04 | 99.96 | 99.9 | 99.96 | 99.96 | 99.96 | 0 |
|  | 8.66E+03 | 99.9 | 100.0 | 100.00 | 99.94 | 99.94 | 0 |
| B | 1.52E+04 | 99 | 99 | 99,0 | 99 | 99,4 | 0 |
|  | 7.30E+03 | 96 | 98 | 97 | 98 | 97 | 0 |
|  | 1.74E+04 | 98 | 97 | 98 | 98 | 99 | 0 |

Supplementary Table S3. Results for all mask samples in experimental setup III

| Mask sample | Average concentration of two PCs (PFU) | VFE in subsamples (%) |  |  |  |  | NC (PFU) | Average mean particle size (µm) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | 2.86E+03 | 99.6 | 99.4 | 99.8 | - | - | 0 | 3.4 |
|  | 2.19E+03 | 99.2 | 99 | 99.4 | - | - | 0 | 3.1 |
| B | 2.36E+03 | 98 | 98 | 90a | - | - | 0 | 3.1 |
|  | 2.93E+03 | 99 | 98 | 97 | - | - | 0 | 3.2 |

a This result was not considered in the calculation of the final VFE because there was likely a problem in the preparation of this mask that influenced the lower VFE. Namely, mask type B is a three-layer material that we had to assemble before each experiment. This VFE proved to be a significant outlier (P < 0.01) when compared to other VFE values from the same experimental setup (III) as well as VFE values from all experimental setups (https://www.graphpad.com/quickcalcs/grubbs2/).

Supplementary Table S4. Results for all mask samples in experimental setup IV

| Mask sample | Average concentration of two PCs (PFU/mL) | VFE in subsamples (%) |  |  |  |  | NC (PFU/mL) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | 1.17E+05 | 99.98 | 100 | 99.4 | - | - | 0 |
|  | 8.57E+04 | 99.99 | 99.98 | 99.98 | 99.98 | 99.98 | 0 |
|  | 1.21E+05 | 99.98 | 99.98 | 99.98 | 99.96 | 99.97 | 0 |
| B | 6.52E+04 | 99.3 | 99.3 | 99.2 | 99.5 | 99.3 | 0 |
|  | 1.24E+05 | 99.2 | 99.3 | 99.4 | 99.3 | 99.3 | 0 |
| C | 1.28E+05 | 98 | 96 | 96 | - | - | 0 |

Supplementary Table S5. Results for all mask samples in experimental setup V

| Mask sample | Average concentration of two PCs (PFU/mL) | VFE in subsamples (%) |  |  |  |  | NC (PFU/mL) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | 1.35E+06 | 100 | 100 | 99.99 | - | - | 0 |
|  | 1.16E+06 | 99.8 | 99.92 | 99.9 | - | - | 0 |
|  | 5.41E+05 | 99.5 | 99.7 | 99.7 | - | - | 0 |
| B | 6.58E+05 | 98 | 98 | 98 | 98 | - | 10 |
|  | 1.79E+06 | 99.93 | 99.4 | 99.3 | - | - | 0 |
|  | 2.56E+06a | 99.7 | 99.1 | 99.2 | - | - | 0 |
| D | 6.11E+05 | 83 | 86 | 88 | 91 | 88 | 0 |
| E | 6.96E+05 | 99.999 | 100 | 99.996 | 100 | 99.998 | 0 |
| F | 1.36E+06 | 91 | 92 | 89 | 94 | 91 | 0 |

aThis value is only from the first positive control as the second one was not performed during the experiment.

Supplementary Table S6. Results for all mask samples in experimental setup VI

| Mask sample | Average concentration of two PCs (PFU/mL) | VFE in subsamples (%) |  |  |  |  | NC (PFU/mL) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | 1.29E+06 | 99.9 | 99.8 | 99.9 | - | - | 0 |
|  | 9.36E+05 | 99.9 | 99.9 | 99.9 | - | - | 0 |
|  | 3.98E+05 | 99.8 | 99.7 | 99.4 | - | - | 0 |

Supplementary Figure S1. Plates with Staphylococcus aureus colonies (left) or MS2 plaques (right) after testing bacterial filtration efficiency (BFE) or viral filtration efficiency (VFE) in the Andersen sampler (experimental setups I and III). The observed bacterial/virus pattern follows the pattern of the holes on each stage. The top two rows of plates (top 6 plates) show the positive control plates (from one of the two positive controls), i.e., the plates without a mask, whereas the bottom two rows (bottom 6 plates) show the plates where the bacterial or viral droplets passed through the mask. The mask on the left had a BFE of 97%, while the mask on the right had a VFE of 99.2%.