Supplementary Files

Appendix 1

HeraBEAT system specification and safety claims.

| Characteristic | Measure | Specifications |
| --- | --- | --- |
| Safety | Complies with | IEC/EN 60601-1, 60601-1-2, 60601-1-11, 60601-2-37 |
| Classification | Antielectric shock type | Class II electrical device when AC/DC adapter connected. Otherwise, internally powered equipment |
|  | Antielectric shock degree | Type BF equipment |
|  | Degree of protection against harmful ingress of water | IP22; Protection against falling drops of water when unit is tilted 15o |
| Physical characteristics | Device size | 88 x 37 mm; 3.5 x 1.5 inches (Diameter × Height, ± 0.08 inches |
|  | Device weight | Approximately 4.58 ounces |
| Operating environment | Temperature | From 41°F up to 104°F |
|  | Humidity | From 5% up to 90% RH (noncondensing) |
| Storage/transport environment | Temperature | From -4°F up to 140°F |
|  | Humidity | From 5% up to 95% (noncondensing) |
|  | Light intensity | No direct sunlight |
| FHR performance | Pregnancy week | 12 to 42 |
|  | FHR measuring range; accuracy; resolution | 50 to 240 bpm; ± 2 bpm; 1 bpm |
|  | MHR measuring range; accuracy; resolution | 45 to 240 BPM; ± 2% or 1 bpm, whichever is greater; 1 bpm |
| Auto acquisition stop | NA | 5 minutes of successful measurement |
| Recommended ultrasound transmission gel | NA | Aquasonic 100 Ultrasound Transmission Gel (Parker Laboratories, Fairfield, NJ) |
| Power consumption | NA | <2 W |
| Rechargeable lithion-ion battery | Nominal capacity | 3.7 V DC, 1250 mAH |
|  | Continuous work time | 4 hours (with a new battery) |
|  | Power input | 5 V DC, 0.3 A |
|  | Charge time | 4 hours |
| Ultrasound (NEMA/FDA) | Nominal frequency | 2 MHz ± 10% |
|  | Ultrasonic output power (P) | 70 mW |
|  | Peak rarefactional pressure (pr) | 0.03 MPa |
|  | Ultrasonic output intensity (Isata) | ≤20 mW/cm2 |
|  | Mechanical index (MI) | 0.02 |
|  | Thermal index (TIS; TIB) | 0.26; 0.7 |
|  | Measurement mode | Continuous wave ultrasound doppler |
|  | Effective radiating area of transducer | 4.9 ± 0.5 cm2 |
| BLE specification | Frequency band of transmission | 2.4–2.5 GHz Channels (2 MHz spacing) 3 advertising channels @ 2402, 2426, 2480 Mh 36 data channels |
|  | Frequency characteristics of the modulation | DSSS: GFSK (modulation index=0.5) |
|  | Maximum RF input | -10 dBm |
|  | Typical receive sensitivity | -94 dBm |
|  | Maximum RF Tx output power | +4 dBm |

HeraBEAT safety claims:

- HeraBEAT works at low voltage (5 V), which is supplied from an internal rechargeable battery (tested per IEC 60601-1).

- HeraBEAT device material is isolated and made of electric nonconducting material.  In addition, the device does not operate while charging.

- HeraBEAT transmits ultrasonic energy at a maximum intensity of 20 mW/c “2, according to IEC 60601-2-37 “Medical electrical equipment – Part 2-37: Particular requirements for the safety of ultrasonic medical diag “ostic and monitoring equipment.”

- The device turns off if not connected to the mobile app for several seconds.

- All materials are biocompatible and approved for use on the skin surface.

- HeraBEAT controls the temperature level inside the device to assure that the device temperature remains below the safe temperature limit.  In addition, a built-in test (BIT) is implemented to verify the correct functioning of the temperature sensor.

- The device conforms to risk management best practices according to ISO 14971:2007 – Medical Devices – Application of Risk Management to Medical Devices.

Appendix 2

Conversion table for System Usability Scale raw scores into percentile and grades.

| SUS Score | Percentile | Grade |
| --- | --- | --- |
| 84.1–100 | 96–100 | A+ |
| 80.8–84.0 | 90–95 | A |
| 78.9–80.7 | 85–89 | A- |
| 77.2–78.8 | 80–84 | B+ |
| 74.1–77.1 | 70–79 | B |
| 72.6–74.0 | 65–69 | B- |
| 71.1–72.5 | 60–64 | C+ |
| 65.0–71.0 | 41–59 | C |
| 62.7–64.9 | 35–40 | C- |
| 51.7–62.6 | 15–34 | D |
| <51.7 | 0–14 | F |

Appendix 3

All positive version of the System Usability Scale and the adjectival enhancement question used in the study.

| Please mark the box that reflects your immediate response to each statement. Don’t think too long about each statement. Please make sure you respond to every statement. If you don’t know how to respond, just mark box ‘3.’ |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Strongly disagree |  |  |  | Strongly agree |
| 1. I think I would like to use this system frequently. | 1 | 2 | 3 | 4 | 5 |
| 2. I found this system to be simple. | 1 | 2 | 3 | 4 | 5 |
| 3. I thought this system was easy to use. | 1 | 2 | 3 | 4 | 5 |
| 4. I think I could use this system without the support of a technical person. | 1 | 2 | 3 | 4 | 5 |
| 5. I found the various functions of this system were well integrated. | 1 | 2 | 3 | 4 | 5 |
| 6. I thought there was a lot of consistency in this system. | 1 | 2 | 3 | 4 | 5 |
| 7. I imagine most people would learn to use it very quickly. | 1 | 2 | 3 | 4 | 5 |
| 8. I found it very intuitive. | 1 | 2 | 3 | 4 | 5 |
| 9. I felt very confident using this system | 1 | 2 | 3 | 4 | 5 |
| 10. I could use this system without having to learn anything new. | 1 | 2 | 3 | 4 | 5 |

| Adjectival assessment: Overall, I would rate the user-friendliness of HBM as…. | Worst imaginable | Awful | Poor | Okay | Good | Excellent | Best imaginable |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 |

## Authors’ Data Sharing Statement

- Will individual participant data be available (including data dictionaries)? Individual, de-identified participant data may be shared by the corresponding author, providing appropriate reciprocal ethical approval has been sort for use of the data.

- What data in particular will be shared? Available data will include individual FHR metrics, SUS scores and accuracy measurements.

- What other documents will be available? The study protocol is available.

- When will data be available (start and end dates)? Data will be available after manuscript publication for a period of 24 months.

- By what access criteria will data be shared (including with whom, for what types of analyses, and by what mechanism)? Data will be shared for reanalysis to confirm results.  All use of data must comply with HREC requirements and Good Clinical Practice.

Appendix 4

Factors related to clinical outcomes (Phase 1)

| User |  | Clinician Administered |  |
| --- | --- | --- | --- |
| Site of recording |  | Antenatal Clinic (Phase 1) (n = 63) |  |
|  |  | Median (IQR) | P |
| Time to 1st detection of FHR (Seconds) |  |  |  |
| BMI at enrolment (kg/m2) | <23.5 | 29 (14 – 44) | 0.78 |
|  | 23.5 to < 30 | 14 (14 – 25) |  |
|  | 30 to <35 | 14 (14 – 25.6) |  |
|  | 35 to <45 | 14 (15 – 18) |  |
| BMI at enrolment ≤ 35 (kg/m2) | < 35 | 14 (14 – 15) | 0.86 |
|  | ≥ 35 | 14 (14 – 18) |  |
| Anterior Placenta Location | Yes | 14 (14 - 14) | 0.09 |
|  | No | 14 (14 -1 6) |  |
| Continuous FHR Trace Duration (Minutes) |  |  |  |
| BMI at enrolment (kg/m2) | <23.5 | 12.7 (3.1 – 22.2) | 0.44 |
|  | 23.5 to < 30 | 8.8 (6.8 – 15.5) |  |
|  | 30 to <35 | 7.1 (4.3 – 11.2) |  |
|  | 35 to <45 | 10.1 (6.9 – 13.2) |  |
| BMI at enrolment ≤ 35 (kg/m2) | < 35 | 7.9 (5.1 – 13.4) | 0.48 |
|  | ≥ 35 | 10.1 (6.9 – 13.2) |  |
| Anterior Placenta Location | Yes | 7.3 (4.7 – 12.1) | 0.27 |
|  | No | 9.2 (6.5 – 16.2) |  |
| Signal Loss (%) |  |  |  |
| BMI at enrolment (kg/m2) | <23.5 | 10.0 (0.4 – 19.6) | 0.39 |
|  | 23.5 to < 30 | 5.8 (3.3 – 10.5) |  |
|  | 30 to <35 | 11.3 (4.2 – 16.9) |  |
|  | 35 to <45 | 12.4 (4.8 – 20.2) |  |
| BMI at enrolment ≤ 35 (kg/m2) | < 35 | 8.2 (3.8 – 16.1) | 0.28 |
|  | ≥ 35 | 12.4 (4.8 – 30.9) |  |
| Anterior Placenta Location | Yes | 9.3 (4.5 – 13.4) | 0.93 |
|  | No | 6.6 (2.7 – 21.3) |  |

- *Gestational age was not included as all participants are in Trimester 3.

Appendix 5

Factors related to clinical outcomes (Phase 2)

| User |  | Clinician Administered |  |
| --- | --- | --- | --- |
| Site of recording |  | Participant’s Home (Phase 2) n = 34 |  |
|  |  | Median (IQR) | P |
| Time to 1st detection of FHR (Seconds) |  |  |  |
| BMI at enrolment (kg/m2) | 23.5 to < 30 | 30 (15 – 60) | 0.053 |
|  | 30 to <35 | 15 (15 – 15) |  |
|  | 35 to <45 | 15 (15 – 15) |  |
| BMI at enrolment ≤ 35 (kg/m2) | < 35 | 15 (15 – 30) | 0.45 |
|  | ≥ 35 | 15 (15 – 75) |  |
| Anterior Placenta Location | Yes | 15 (15 – 30) | 0.92 |
|  | No | 15 (15 – 75) |  |
| Gestation Age | 2nd Trimester (Week 14-26) | 382.5 (15 – 750) | 0.33 |
|  | 3rd Trimester (Week 27+) | 15 (15 – 30) |  |
| Continuous FHR Trace Duration |  |  |  |
| BMI at enrolment (kg/m2) | 23.5 to < 30 | 8.8 (4.7 – 13.9) | 0.59 |
|  | 30 to <35 | 6.4 (3.1 – 8.74) |  |
|  | 35 to <45 | 7.7 (5.3 – 14.2) |  |
| BMI at enrolment ≤ 35 (kg/m2) | < 35 | 6.4 (4.6 – 9.8) | 0.40 |
|  | ≥ 35 | 7.7 (6.1 – 12.8) |  |
| Anterior Placenta Location | Yes | 7.6 (4.4 – 13.6) | 0.76 |
|  | No | 6.5 (5.0 – 9.4) |  |
| Gestation Age | 2nd Trimester (Week 14-26) | 14.8 (8.6 – 20.9) | 0.12 |
|  | 3rd Trimester (Week 27+) | 6.5 (4.6 -11.3) |  |
| Signal Loss (%) |  |  |  |
| BMI at enrolment (kg/m2) | 23.5 to < 30 | 2.6 (1.3 – 6.6) | 0.17 |
|  | 30 to <35 | 6.0 (2.7 – 9.9) |  |
|  | 35 to <45 | 6.7 (1.5 – 8.4) |  |
| BMI at enrolment ≤ 35 (kg/m2) | < 35 | 4.3 (2.1 – 9.7) | 0.99 |
|  | ≥ 35 | 6.7 (1.5 – 8.4) |  |
| Anterior Placenta Location | Yes | 2.9 (1.0 – 9.5) | 0.31 |
|  | No | 6.6 (3.3 – 9.7) |  |
| Gestation Age | 2nd Trimester (Week 14-26) | 1.4 (1.0 – 1.7) | 0.09 |
|  | 3rd Trimester (Week 27+) | 5.9 (2.4 – 9.6) |  |