Supplementary Information

Multi-focus light-field microscopy for high-speed large-volume imaging

# Yi Zhang1,2*, Yuling Wang1,2*, Mingrui Wang2,3, Yuduo Guo2,3, Xinyang Li1,2,3, Yifan Chen1,2, Zhi Lu1,2, Jiamin Wu1,2,4,5,6†, Xiangyang Ji1,2,7† & Qionghai Dai1,2,4,5,8†

# 1Department of Automation, Tsinghua University, Beijing, 100084, China

2Institute for Brain and Cognitive Sciences, Tsinghua University, Beijing 100084, China

3Tsinghua Shenzhen International Graduate School, Tsinghua University, Shenzhen, 518055, China.

4Beijing Key Laboratory of Multi-dimension & Multi-scale Computational Photography (MMCP), Tsinghua University, Beijing 100084, China.

5IDG/McGovern Institute for Brain Research, Tsinghua University, Beijing, China.

6Hangzhou Zhuoxi Institute of Brain and Intelligence, Hangzhou, 311100, China.

7Beijing National Research Center for Information Science and Technology, Tsinghua University, 100084 Beijing, China.

8School of Information and Technology, Fudan University, Shanghai 200433, China.

*These authors contributed equally to this work

†Correspondence: wujiamin@tsinghua.edu.cn (J.W.), xyji@tsinghua.edu.cn (X.J.), qhdai@tsinghua.edu.cn (Q.D.)

| Supplementary Figure 1 | DOF characterization with different levels of spherical aberrations |
| --- | --- |
| Supplementary Figure 2 | Sub-aperture PSFs elongation under different objectives |
| Supplementary Figure 3 | DOF characterization with different numbers of sup-aperture components |
| Supplementary Figure 4 | Analysis of the angular redundancy |
| Supplementary Figure 5 | Experimental comparisons on the USAF-1951 resolution target between sLFM and SAsLFM |
| Supplementary Figure 6 | Numerical simulation for comparisons on DOFs between sLFM, cubic-phase-assisted sLFM and SAsLFM |
| Supplementary Figure 7 | Illustration of phase-space deconvolution algorithm with the circular trajectory |
| Supplementary Figure 8 | Neural network architecture |
| Supplementary Note 1 | Phase-space PSF engineering |
| Supplementary Table 1 | Imaging conditions for all experiments |
| Supplementary Table 2 | Rendering parameters and models for all experiments |
| Supplementary Video 1 | In vivo calcium imaging of mouse brain |
| Supplementary Video 2 | Volumetric imaging of freely-behaving jellyfish |

Supplementary Figure 1 | DOF characterization with different levels of spherical aberrations.  Relationship between the Zernike coefficient of spherical aberration and DOF. The phase-space PSFs of a 6/0.5NA SAsLFM system with different levels of spherical aberrations appended at the pupil plane were simulated based on the wave optical theory1. The DOF was calculated from the union of the high-resolution ranges of all sub-aperture components. The curve shows that the DOF increases linearly with the absolute value of the Zernike coefficient of spherical aberration.

Supplementary Figure 2 | Sub-aperture PSFs elongation under different objectives. We simulated the beam propagation of SAsLFM systems using 20/0.5NA, 40/1.0NA, and 60/1.4NA objectives respectively with the same level of spherical aberration (Zernike coefficient = 10) added on the pupil plane. Each microlens covers 1313 pixels, so the component with angular coordinate (7, 7) is the central component, of which the focal position along the z-axis is termed as the native focal plane. The xz-MIPs of typical sub-aperture PSFs show that the focal positions of the different angular components exhibit different degrees of offsets from the native focal plane, resulting in unilateral extension of the DOFs. The corresponding normalized intensities of sub-aperture PSFs along the z-axis are shown on the bottom, proving the effectiveness of SAsLFM to extend the DOFs from ~50 μm, ~20 μm, and ~10 μm to ~185 μm, ~52 μm, and ~23 μm, respectively. Scale bar: 50 μm

Supplementary Figure 3 | DOF characterization with different numbers of sup-aperture components. a We simulated the beam propagation of SAsLFM (Zernike coefficient = 5) under a 22/0.5NA objective, with each microlens covering NN pixels on the camera and the size of the sensor pixel is fixed at 6.5 μm. Because the effective NA of the system is reduced by a factor of N, the DOF of SAsLFM rises monotonically as the number of sub-aperture components increases from 9 to 19. b When the pitch size of each microlens is fixed at 100 μm, a larger N means a smaller pixel size. When the sampling rate is not sufficient, the DOF increases with the increase of N. When the N is set to 9, the lateral size of the area corresponding to one pixel is ~0.51 μm, which reaches the diffraction limit of the system (~0.53 μm), resulting in the convergence of the expansion.

Supplementary Figure 4 | Analysis of the angular redundancy. a We captured high-resolution confocal image stacks of the blood vessels in zebrafish larvae, and generated the corresponding sLFM data using the wave optics theory1. The curves show the structural similarity index (SSIM) between the reconstructed volume and the confocal ground truth when different numbers of sub-aperture (angular) components are used for deconvolution. The curves indicate that about 6 images from different viewing directions contain sufficient information to reconstruct a 3D volume with high fidelity. b MIPs of a ground truth image stack captured by confocal microscopy and phase-space reconstruction results using different numbers of angular components are shown for comparison. The SSIM values between the reconstruction results and the confocal ground truth are marked at the bottom of each sub-panel. Volumes reconstructed with 6 and 20 components exhibit equally good performance, demonstrating the redundancy of sub-aperture components.

Supplementary Figure 5 | Experimental comparisons on the USAF-1951 resolution target between sLFM and SAsLFM.  a Reconstructed slices of sLFM with a 10/0.28NA objective. The resolution suffers from significant degradation due to the increasingly dispersed sub-aperture PSFs at the layers far from the native objective plane. b Reconstruction results of the UASF-1951 resolution target of SAsLFM. During the imaging procedure, the resolution chart was immersed in 2-cm-thick water. Due to the extended DOF, all the elements in the 6th group can still be discerned even at a depth of -600 μm. c Magnified views of boxed regions in a-b. SAsLFM shows eliminated ringing artifacts at the native focal plane. d Estimated phases on the pupil planes for sLFM and SAsLFM, respectively. Conspicuous spherical aberration exists in the latter one.

Supplementary Figure 6 | Numerical simulation for comparisons on DOFs between sLFM, cubic-phase-assisted sLFM and SAsLFM. a Simulated 14 μm fluorescence beads were randomly located at 150 μm away from the native focal plane. We performed forward projection with various types of PSFs. b The light-field data was generated by the 3D convolution of the simulated ground truth and the sLFM PSFs without any additional phases. The reconstructed slice shows that the two beads cannot be distinguished due to the resolution limitation.  c We convolved the same ground truth with the cubic-phase-assisted sLFM PSFs to generate the scanning light-field data and then performed phase-space deconvolution with the same PSFs. The SSIM between the reconstructed results and the simulated ground truth is marked on the bottom, illustrating that the reconstruction performance is on par with that of traditional sLFM. d We performed the same imaging projection and deconvolution on simulated data with PSFs of SAsLFM. The reconstructed slice reveals the spherical structure with higher fidelity, demonstrating the effectiveness of our method.

Supplementary Figure 7 | Illustration of phase-space deconvolution algorithm with the circular trajectory a The sub-aperture components are divided into 7 groups according to their position relative to the center of the pupil plane (SAsLFM with each microlens covering 1515 pixels for demonstration). Each group is marked with a specific color. Components in the same group contain high-resolution information from the same depth of range. b The deconvolution algorithm consists of inner and outer loops. During the inner iteration, the output of the last outer iteration is used as the initial estimation of the 3D volume. Starting from the margin of the pupil plane, all the components in group 7 are used to update the volume as a whole. The updated volume becomes the initial value of the next inner iteration and the angular components in group 6 are used to continuously update the volume. When going through all the angular components, the high-resolution information at different axial positions is fused and the merged large-scale volume is used as the initial estimation for the next outer iteration. The aforementioned steps are repeated until convergence, which usually takes ~3-5 outer iterations.

Supplementary Figure 8 | Neural network architecture. The network takes sub-aperture components with a patch size of 51351321 pixels as the input, and the output is the corresponding 3D reconstructed volume with a size of 513513201 pixels. We used a simplified 2D U-net as our infrastructure and converted the angular dimension of the input data to the channel dimension. The network consists of an encoding path and a decoding path, and each convolution block consists of multiple operations including 2D convolution, batch normalization, and activation with ReLU functions. The training procedure is carried out using the weighted summation of L1 loss, L2 loss and SSIM loss as the loss function, and a small validation dataset was adopted to prevent overfitting.

Supplementary note 1 | Phase-space PSF engineering

Based on the scanning path, the pixels sharing the same relative position  to the center of each microlens can be realigned together to form an intensity projection  of the volume information from a specific viewing angle, which is referred to as a sub-aperture component in the phase-space, see Fig.1(a). Following the Debye theory1, the original PSF  of an arbitrary point at the depth  with a lateral coordinate  is formulated as follows:

(1)

Where  is the 2D Fourier transform function,  is the lateral coordinates on the sensor plane,  represents the refractive index of the sample,  is the focal length of the micorlens, and  is the wavelength of emission fluorescence. represents the wavefront at the native image plane with phase modulation induced by the microlens array:

(2)

Where  is the wavefront at the native image plane:

(3)

is the magnification of the microscopy,  is the focal length of the objective lens,  is the half-angle of the NA, and  is the zeroth-order Bessel function of the first kind.  is the phase modulation caused by the microlens array:

(4)

is the pitch size of a microlens,  denotes the convolution operation and  is a 2D comb function. In addition, a point on the sensor plane with coordinate  has the relative displacement  to the corresponding microlens with center position located at .  can be written as:

(5)

A pixel with coordinates  at the sensor plane can be identified by the four-dimensional vector . The final light intensity accumulated on the pixel can be represented as:

(6)

Where  is the size of a pixel on the sensor plane. Evidently, , demonstrating the 2D shift-invariance property. In essence, this property would not be affected by external phase modification on the pupil plane and it opens up the opportunity of extending the DOF of sLFM with sub-aperture PSF engineering.

Reference

1 	Broxton M, Grosenick L, Yang S, Cohen N, Andalman A, Deisseroth K et al. Wave optics theory and 3-D deconvolution for the light field microscope. Optics Express 2013; 21: 25418–25439.

Supplementary table 1 | Imaging conditions for all experiments

| Figure # | Objective | Sub-aperture | Scan trajectory |
| --- | --- | --- | --- |
| 3a-b, | 10/0.28NA | 1515 | 1515 |
| 4a, 4d-f, 5b | 10/0.5NA | 1515 | 33 |
| 5d-e | 10/0.5NA | 1515 | 11 |
| 5g-h | 20/0.5NA | 1313 | 55 |
| 6c-d | 10/0.5NA | 1515 | 55 |
| S2 | 20/0.5NA | 1313 | 1313 |
| S2 | 40/1.0NA | 1313 | 1313 |
| S2 | 60/1.4NA | 1313 | 1313 |
| S5 | 10/0.28NA | 1515 | 1515 |

Supplementary table 2 | Rendering parameters and models for all experiments

| Figure # | software | colormap | Display mode |
| --- | --- | --- | --- |
| 2a-b | Matlab | Green | Slices & MIPs |
| 3a-b | ImageJ | Grays | Slices |
| 4a, f | ImageJ | Green | Slices |
| 4d, e | Amira | Green | Voltex (MIP) |
| 5b | Matlab | Green | MIPs |
| 5d | ImageJ | Cyan | Slices |
| 5e | ImageJ | Cyan | MIPs |
| 5g-h | ImageJ | Red, Green & Yellow | MIPs |
| 6c | Amira | Yellow | Voltex (MIP) |
| 6d | Matlab | Yellow | MIPs |
| S2 | Matlab | Fire | MIPs |
| S4b | Matlab | Green | MIPs |