Supplementary Materials for

Development and Validation of a duplex real-time PCR for the rapid detection and quantitation of HTLV-1

Huimin Ji 12, Le Chang 12, Ying Yan12, Lunan Wang123 ＃

1. National Center for Clinical Laboratories, Beijing Hospital, National Center of Gerontology; Institute of Geriatric Medicine, Chinese Academy of Medical Sciences, Beijing, P. R. China

2. Beijing Engineering Research Center of Laboratory Medicine, Beijing Hospital, Beijing, P. R. China

3. Graduate School, Peking Union Medical College, Chinese Academy of Medical Sciences, Beijing, P. R. China

＃Corresponding Author:

# Lunan Wang, Email: lunan99@163.com

Tel.: + 86 10 85133609

Table S2. Details of results by LIA, WB, qPCR, Murex-ELISA and Roche-ECLIA

| NO. |  | Murex |  | Roche |  |  | LIA |  | WB |  | qPCR |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | S/CO |  |  | COI |  | Results |  | Results |  | copies/100 cells |
| 1 |  | 11.06 |  |  | 450.2 |  | HTLV-1 |  | HTLV-1 |  | 3.333 |
| 2 |  | 11.98 |  |  | 427.9 |  | HTLV-1 |  | HTLV-1 |  | 0.339 |
| 3 |  | 12.25 |  |  | 365.2 |  | HTLV-1 |  | HTLV-1 |  | 1.345 |
| 4 |  | 13.39 |  |  | 364.6 |  | HTLV-1 |  | HTLV-1 |  | 2.278 |
| 5 |  | 12.61 |  |  | 70.62 |  | HTLV-1 |  | HTLV-1 |  | 0.273 |
| 6 |  | 13.94 |  |  | 277.4 |  | HTLV-1 |  | HTLV-1 |  | 3.001 |
| 7 |  | 12.26 |  |  | 317.3 |  | HTLV-1 |  | HTLV-1 |  | 5.288 |
| 8 |  | 12.52 |  |  | 287.7 |  | HTLV-1 |  | HTLV-1 |  | 1.11 |
| 9 |  | 12.32 |  |  | 329.1 |  | HTLV-1 |  | HTLV-1 |  | 0.67 |
| 10 |  | 15.35 |  |  | 415.8 |  | HTLV-1 |  | HTLV-1 |  | 0.795 |
| 11 |  | 12.72 |  |  | 303.6 |  | HTLV-1 |  | HTLV-1 |  | 0.12 |
| 12 |  | 14.15 |  |  | 465 |  | HTLV-1 |  | HTLV-1 |  | 0.026 |
| 13 |  | 13.31 |  |  | 214.7 |  | HTLV-1 |  | HTLV-1 |  | 0.071 |
| 14 |  | 13.37 |  |  | 213.1 |  | HTLV-1 |  | HTLV-1 |  | 0.678 |
| 15 |  | 13.12 |  |  | 136.5 |  | HTLV-1 |  | HTLV-1 |  | 1.23 |
| 16 |  | 12.23 |  |  | 235.5 |  | HTLV-1 |  | HTLV-1 |  | 0.001 |
| 17 |  | 12.55 |  |  | 458.2 |  | HTLV-1 |  | HTLV-1 |  | 0.013 |
| 18 |  | 12.48 |  |  | 235.4 |  | HTLV-1 |  | HTLV-1 |  | 0.033 |
| 19 |  | 12.24 |  |  | 371.7 |  | HTLV-1 |  | HTLV-1 |  | 1.956 |
| 20 |  | 9.80 |  |  | 398.4 |  | HTLV-1 |  | HTLV-1 |  | 0.913 |
| 21 |  | 13.89 |  |  | 361.8 |  | HTLV-1 |  | HTLV-1 |  | 2.815 |
| 22 |  | 11.62 |  |  | 125.7 |  | HTLV-1 |  | HTLV-1 |  | 0.01 |
| 23 |  | 10.84 |  |  | 136.6 |  | HTLV-1 |  | HTLV-1 |  | 0.132 |
| 24 |  | 11.57 |  |  | 124.2 |  | HTLV-1 |  | HTLV-1 |  | 0.695 |
| 25 |  | 12.51 |  |  | 176.3 |  | HTLV-1 |  | HTLV-1 |  | 6.886 |
| 26 |  | 20.38 |  |  | 150.5 |  | HTLV-1 |  | HTLV-1 |  | 1.562 |
| 27 |  | 11.50 |  |  | 132.3 |  | HTLV-1 |  | HTLV-1 |  | Neg |
| 28 |  | 11.69 |  |  | 77.8 |  | HTLV-1 |  | IND |  | 11.397 |
| 29 |  | 0.34 |  |  | 3.26 |  | HTLV-1 |  | IND |  | Neg |
| 30 |  | 11.37 |  |  | 2.27 |  | HTLV-1 |  | Neg |  | Neg |
| 31 |  | 13.03 |  |  | 214 |  | HTLV |  | HTLV-1 |  | 1.11 |
| 32 |  | 12.38 |  |  | 178.2 |  | HTLV |  | IND |  | 0.009 |
| 33 |  | 7.15 |  |  | 8.32 |  | IND |  | Neg |  | 0.017 |
| 34 |  | 8.59 |  |  | 28.18 |  | IND |  | IND |  | Neg |
| 35 |  | 10.44 |  |  | 21.07 |  | IND |  | IND |  | Neg |
| 36 |  | 10.34 |  |  | 8.17 |  | IND |  | IND |  | Neg |
| 37 |  | 1.51 |  |  | 171.5 |  | IND |  | Neg |  | Neg |
| 38 |  | 0.19 |  |  | 197.1 |  | IND |  | Neg |  | Neg |
| 39 |  | 0.56 |  |  | 27.3 |  | IND |  | Neg |  | Neg |
| 40 |  | 0.13 |  |  | 8.05 |  | IND |  | Neg |  | Neg |
| 41 |  | 4.21 |  |  | 1.03 |  | IND |  | Neg |  | Neg |
| 42 |  | 2.66 |  |  | 1.37 |  | IND |  | Neg |  | Neg |
| 43 |  | 0.89 |  |  | 2.3 |  | IND |  | Neg |  | Neg |
| 44 |  | 0.35 |  |  | 2.19 |  | IND |  | Neg |  | Neg |
| 45 |  | 1.12 |  |  | 1.75 |  | IND |  | Neg |  | Neg |
| 46 |  | 1.84 |  |  | 0.415 |  | IND |  | Neg |  | Neg |
| 47 |  | 0.34 |  |  | 3.77 |  | IND |  | Neg |  | Neg |
| 48 |  | 8.72 |  |  | 3.62 |  | IND |  | Neg |  | Neg |
| 49 |  | 6.45 |  |  | 12.71 |  | IND |  | Neg |  | Neg |
| 50 |  | 1.95 |  |  | 0.485 |  | IND |  | Neg |  | Neg |
| 51 |  | 1.30 |  |  | 0.648 |  | IND |  | Neg |  | Neg |
| 52 |  | 1.39 |  |  | 0.289 |  | IND |  | Neg |  | Neg |
| 53 |  | 0.33 |  |  | 5.25 |  | IND |  | Neg |  | Neg |
| 54 |  | 0.44 |  |  | 4.77 |  | IND |  | Neg |  | Neg |
| 55 |  | 0.29 |  |  | 1.31 |  | IND |  | Neg |  | Neg |
| 56 |  | 9.28 |  |  | 1.43 |  | IND |  | Neg |  | Neg |
| 57 |  | 0.31 |  |  | 1.29 |  | IND |  | Neg |  | Neg |
| 58 |  | 0.70 |  |  | 32.28 |  | IND |  | Neg |  | Neg |
| 59 |  | 0.30 |  |  | 5.96 |  | IND |  | Neg |  | Neg |
| 60 |  | 7.75 |  |  | 0.362 |  | IND |  | Neg |  | Neg |
| 61 |  | 0.88 |  |  | 2.09 |  | IND |  | Neg |  | Neg |
| 62 |  | 0.08 |  |  | 3.04 |  | IND |  | Neg |  | Neg |
| 63 |  | 0.30 |  |  | 14.47 |  | IND |  | Neg |  | Neg |
| 64 |  | 1.52 |  |  | 0.979 |  | IND |  | Neg |  | Neg |
| 65 |  | 0.77 |  |  | 17.1 |  | IND |  | Neg |  | Neg |
| 66 |  | 0.37 |  |  | 4.72 |  | IND |  | Neg |  | Neg |
| 67 |  | 0.32 |  |  | 3.03 |  | IND |  | Neg |  | Neg |
| 68 |  | 0.25 |  |  | 14.61 |  | IND |  | Neg |  | Neg |
| 69 |  | 1.11 |  |  | 0.574 |  | IND |  | Neg |  | Neg |
| 70 |  | 2.92 |  |  | 0.512 |  | IND |  | Neg |  | Neg |
| 71 |  | 5.14 |  |  | 3.9 |  | IND |  | Neg |  | Neg |
| 72 |  | 4.87 |  |  | 0.786 |  | IND |  | Neg |  | Neg |
| 73 |  | 0.30 |  |  | 2.08 |  | IND |  | Neg |  | Neg |
| 74 |  | 0.42 |  |  | 2.39 |  | IND |  | Neg |  | Neg |
| 75 |  | 0.37 |  |  | 71.91 |  | IND |  | Neg |  | Neg |
| 76 |  | 0.59 |  |  | 13.93 |  | IND |  | Neg |  | Neg |
| 77 |  | 0.87 |  |  | 0.565 |  | IND |  | Neg |  | Neg |
| 78 |  | 9.92 |  |  | 47.55 |  | IND |  | Neg |  | Neg |
| 79 |  | 0.41 |  |  | 2.74 |  | IND |  | Neg |  | Neg |
| 80 |  | 0.89 |  |  | 25.6 |  | IND |  | Neg |  | Neg |
| 81 |  | 1.00 |  |  | 14.37 |  | IND |  | Neg |  | Neg |

Note: WB, Western blot; LIA, line immunoassay; Ind, indeterminate, Pos, Positive; Neg, Negative;