## Computed Dynamic Method
Model results using Maquet 1 liter Test Lung 190 and a Maquet Servo S ventilator | [10 empty] | Transition epochs
Two tidal volumes - 300 and 400 mL | [9 empty] | RVI = Respiratory rate variability %
Three PEEP levels - 0, 5 and 10 cmH2O | [9 empty] | C_rs = Respiratory system compliance cmH2O·mL-1
Analyzed breaths = 880 | [9 empty] | R_rs = Respiratory system resistance cmH2O·s· L-1
Input | [6 empty] | Results
Epoch | Mode | Breaths | Peak_P | RR | TV | PEEP | RVI | C_rs | R_rs | Plateau
1 | VC | 32 | 17.12 | 15.1 | 298.5 | 0.1 | 8.5 | 21.551 | 7.753 | 13.9509
2 | VC | 30 | 17.1 | 14.6 | 302 | 0.1 | 9 | 21.555 | 7.77 | 14.1107
3 | VC | 31 | 17.12 | 14.6 | 301.97 | 0.1 | 9.5 | 21.473 | 7.723 | 14.1628
4 | VC | 30 | 17.15 | 14.6 | 301.74 | 0.1 | 9 | 21.4 | 7.63 | 14.2
5 | VC | 30 | 17.15 | 14.6 | 301.94 | 0.1 | 9 | 21.503 | 7.75 | 14.1418
6 | VC | 30 | 17.16 | 14.6 | 301.21 | 0.1 | 9 | 21.413 | 7.56 | 14.1667
7 | VC | 31 | 19.8 | 15.1 | 300.81 | 0.1 | 9 | 22.726 | 8.158 | 13.3364
8 | VC | 30 | 20.94 | 14.6 | 300.31 | 4.9 | 9.5 | 23.114 | 8.16 | 17.8926
9 | VC | 31 | 20.92 | 15.1 | 300.48 | 4.9 | 9.5 | 23.134 | 8.187 | 17.8887
10 | VC | 29 | 20.9 | 14.6 | 301.93 | 4.9 | 9 | 23.232 | 8.328 | 17.8963
11 | VC | 30 | 20.9 | 15.1 | 300.74 | 4.9 | 9.5 | 23.275 | 8.32 | 17.8212
12 | VC | 30 | 20.86 | 14.6 | 301.9 | 4.9 | 9 | 23.285 | 8.31 | 17.8654
13 | VC | 28 | 23.23 | 14.6 | 300.85 | 4.9 | 9.5 | 25.566 | 8.25 | 16.6676
14 | VC | 31 | 24.28 | 14.6 | 301.88 | 9.8 | 9 | 26.429 | 8.303 | 21.2223
15 | VC | 30 | 24.3 | 14.6 | 301.18 | 9.8 | 9 | 26.385 | 8.13 | 21.2148
16 | VC | 26 | 24.27 | 15.1 | 301.04 | 9.9 | 9 | 26.4 | 8.123 | 21.303
17 | VC | 11 | 24.32 | 14.6 | 302.18 | 10 | 8 | 26.555 | 8.264 | 21.3794
18 | VC | 15 | 24.29 | 15.1 | 303.87 | 9.9 | 9 | 26.294 | 8.06 | 21.4566
19 | VC | 0 | 22.96 | 14.6 | 367.83 | 10 | -26 | 999 | 999 | 10.3682
20 | VC | 31 | 22.36 | 14.6 | 401.07 | 0.1 | 10 | 23.082 | 10.065 | 17.4759
21 | VC | 31 | 22.43 | 14.6 | 401.24 | 0.1 | 9.5 | 22.998 | 10.026 | 17.5467
22 | VC | 31 | 22.47 | 14.6 | 402 | 0.1 | 10 | 23.082 | 10.335 | 17.5162
23 | VC | 31 | 22.53 | 14.6 | 402.1 | 0.1 | 9 | 22.979 | 10.094 | 17.5986
24 | VC | 31 | 22.49 | 14.6 | 399.8 | 0.1 | 10 | 22.846 | 10.035 | 17.5998
25 | VC | 31 | 25.01 | 14.6 | 402.28 | 0.1 | 9 | 24.111 | 10.703 | 16.7845
26 | VC | 29 | 26.13 | 14.6 | 400.72 | 4.9 | 9.5 | 24.621 | 11.038 | 21.1755
27 | VC | 32 | 26.12 | 15.1 | 395.66 | 4.9 | 9 | 24.774 | 10.987 | 20.8708
28 | VC | 30 | 26.05 | 14.6 | 400.12 | 4.9 | 9.5 | 24.599 | 10.9 | 21.1657
29 | VC | 32 | 26.1 | 15.1 | 399.72 | 4.8 | 9 | 24.77 | 10.997 | 20.9373
30 | VC | 31 | 26.12 | 14.6 | 400.6 | 4.8 | 9.5 | 24.708 | 11.148 | 21.0134
31 | VC | 32 | 28.29 | 15.1 | 400.51 | 4.9 | 9.5 | 26.773 | 10.819 | 19.8595
32 | VC | 31 | 29.34 | 14.6 | 399.32 | 9.9 | 9.5 | 27.706 | 10.829 | 24.3128
33 | VC | 32 | 29.31 | 15.1 | 401.39 | 9.9 | 9.5 | 28.098 | 10.987 | 24.1854
34 | VC | 31 | 29.33 | 14.6 | 400.28 | 9.8 | 9.5 | 27.95 | 10.935 | 24.1213
35 | VC | 32 | 29.29 | 15.1 | 401.43 | 9.9 | 9.5 | 28.206 | 10.969 | 24.1321
36 | VC | 31 | 29.28 | 14.6 | 400.91 | 9.9 | 9.5 | 28.106 | 11.042 | 24.1642

## Measured Flow
Airway Flow L/min
Epoch No. | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36
Time (s)
0 | 21.45 | -0.29 | 21.5 | -3.93 | -0.06 | -5.68 | -0.06 | -28.48 | -0.06 | 0 | 0 | 21.33 | 0.06 | 21.39 | 0.18 | 20.98 | -11.78 | 0.06 | -30.7 | -0.06 | -47.76 | -0.06 | -67.97 | -0.18 | 3.75 | 0.06 | 28.42 | 0.35 | 28.59 | -17.58 | 0 | -21.15 | -0.18 | -22.62 | 0 | -24.08
0.032 | 21.45 | -0.23 | 21.5 | -2.29 | 4.8 | -4.34 | -0.06 | -25.9 | -0.06 | 0 | 0.06 | 17.64 | 0.12 | 21.45 | 0.29 | 21.45 | -10.2 | 0.06 | -28.48 | -0.18 | -46.53 | -0.06 | -54.08 | -0.06 | 0.12 | 0.06 | 28.65 | 0.29 | 28.65 | -15.7 | 0 | -19.4 | -0.06 | -20.98 | -0.18 | -22.68
0.064 | 21.45 | -0.23 | 21.45 | -1.29 | 10.9 | -3.16 | -0.06 | -23.91 | -0.06 | 0 | 0.12 | 0.12 | 0.06 | 21.39 | 0.23 | 21.62 | -8.32 | -0.06 | -27.42 | -0.18 | -44.53 | -0.06 | -45.53 | 0 | 0 | 0.06 | 28.65 | 0.53 | 28.54 | -13.77 | 0 | -18.16 | -0.06 | -19.45 | 0 | -20.86
0.096 | 21.45 | -0.18 | 21.45 | -0.64 | 16.23 | -1.64 | 4.57 | -21.45 | -0.06 | -42.31 | 0 | 0 | 0.18 | 21.39 | 0.18 | 21.39 | -5.98 | 0.06 | -25.55 | -0.18 | -40.9 | -0.06 | -44.94 | 0 | 0.06 | 0 | 28.48 | 0.35 | 28.13 | -11.89 | 0.06 | -16.23 | -0.18 | -17.64 | 0 | -19.28
0.128 | 21.39 | -0.23 | 21.39 | -0.18 | 21.97 | -0.94 | 10.72 | -18.99 | -0.06 | -68.56 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.21 | -4.69 | -0.06 | -23.61 | -0.06 | -40.55 | 0 | -42.31 | -0.18 | 0 | 0.06 | 28.24 | 0.18 | 28.48 | -9.67 | 5.68 | -14.24 | -0.06 | -15.59 | -0.06 | -17.93
0.16 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -0.47 | 15.94 | -17.29 | -0.18 | -47.7 | -0.06 | 0 | 0.06 | 0.12 | 0.12 | 21.39 | -3.28 | 0 | -21.56 | -0.18 | -37.44 | -0.18 | -42.01 | 0 | 0 | -0.06 | 28.65 | 0.35 | 28.59 | -7.38 | 13.71 | -12.77 | 0 | -14.53 | 5.74 | -15.65
0.192 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.5 | -15.35 | -0.06 | -39.14 | -0.06 | 0 | 0 | 0 | 0.18 | 21.21 | -2.23 | 0.06 | -20.51 | -0.06 | -35.57 | -0.06 | -38.67 | 0 | 0 | 0.06 | 28.65 | 0.35 | 28.42 | -5.33 | 23.97 | -11.13 | 5.1 | -12.83 | 13.65 | -14.06
0.224 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | -13.71 | -0.06 | -33.87 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -1.46 | 0 | -19.75 | -0.18 | -34.75 | -0.06 | -37.09 | -0.18 | 0 | 0.06 | 28.42 | 0.29 | 28.42 | -3.4 | 28.71 | -9.43 | 13.18 | -11.07 | 21.39 | -12.72
0.256 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | -10.66 | -0.06 | -31.7 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | -0.94 | 0.12 | -16.88 | 0 | -32.99 | 0 | -35.39 | 0 | 0 | -0.06 | 28.54 | 0.18 | 28.65 | -1.99 | 28.59 | -7.27 | 20.68 | -9.2 | 29.53 | -10.2
0.288 | 21.5 | -0.23 | 21.45 | -0.41 | 21.45 | -0.35 | 21.5 | -8.14 | -0.06 | -32.93 | 0.06 | 0 | 0 | 0 | 0.12 | 21.45 | -0.41 | 0.06 | -14.77 | 0 | -30.47 | 0 | -32.93 | -0.06 | 0 | -0.06 | 28.54 | 0.18 | 28.36 | -1.46 | 28.36 | -5.51 | 28.89 | -7.32 | 28.48 | -8.85
0.32 | 21.39 | -0.18 | 21.39 | -0.47 | 21.39 | -0.29 | 21.5 | -6.68 | -0.06 | -32.46 | -0.18 | 0 | 0 | 0 | 0.12 | 21.15 | -0.29 | 0.12 | -13.24 | 0 | -28.36 | 0 | -31.58 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.65 | -0.53 | 28.83 | -3.87 | 28.54 | -5.8 | 28.48 | -7.27
0.352 | 21.45 | -0.18 | 21.39 | -0.18 | 21.45 | -0.47 | 21.45 | -4.28 | 0 | -30.41 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.33 | -0.06 | 0 | -12.19 | -0.06 | -25.67 | 0 | -28.89 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.59 | -0.18 | 28.42 | -2.52 | 28.48 | -4.04 | 28.42 | -5.68
0.384 | 0.18 | -0.18 | 21.39 | -0.29 | 21.45 | -0.35 | 21.5 | -2.46 | -0.18 | -28.36 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.33 | -0.06 | 0 | -10.02 | -0.06 | -23.79 | 0 | -26.6 | 0 | 0 | 0 | 28.48 | 0.18 | 28.48 | 0 | 28.89 | -1.82 | 28.42 | -2.75 | 28.42 | -3.98
0.416 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.39 | -1.46 | -0.18 | -26.31 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | 0.06 | 0.06 | -7.79 | -0.06 | -21.86 | 0 | -25.25 | -0.06 | -8.32 | -0.06 | 0.18 | 0.06 | 28.48 | 0 | 28.48 | -1 | 28.65 | -1.41 | 28.65 | -2.23
0.448 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.53 | -0.06 | -24.26 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.15 | 0.12 | 3.69 | -6.15 | -0.06 | -20.16 | 0 | -22.85 | 0 | -77.05 | -0.18 | 0 | 0 | 28.48 | 0.18 | 28.54 | -0.7 | 28.48 | -0.94 | 28.42 | -1.64
0.48 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | -0.18 | -22.09 | 0 | -37.85 | -0.18 | 0 | 0.12 | 21.33 | 0.12 | 11.48 | -4.34 | -0.06 | -17.75 | 0 | -21.21 | -0.06 | -56.96 | -0.18 | 0 | 0 | 28.83 | 0.29 | 28.71 | -0.53 | 28.54 | -0.82 | 28.42 | -1.05
0.512 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | -0.18 | -19.86 | -0.18 | -61.76 | -0.18 | 0 | 0 | 21.39 | 0.18 | 17.11 | -2.99 | 0 | -15.47 | -0.06 | -18.4 | 0 | -45.18 | -0.06 | 0 | 0 | 28.54 | 0.29 | 28.59 | -0.06 | 28.48 | -0.29 | 28.48 | -0.53
0.544 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.23 | 21.45 | 0.23 | -0.18 | -17.87 | -0.06 | -43.71 | 0 | 0 | 0.12 | 21.39 | 0.29 | 21.62 | -1.99 | 0 | -14.65 | 0 | -16.93 | -0.18 | -45.29 | -0.18 | 0 | 0 | 28.48 | 0.29 | 28.59 | -0.18 | 28.48 | -0.47 | 28.59 | -0.53
0.576 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | 0.35 | -0.18 | -15.59 | -0.06 | -39.61 | -0.18 | -50.16 | 0.06 | 21.39 | 0.12 | 21.45 | -1.11 | -0.06 | -12.89 | -0.06 | -15.35 | 0 | -42.42 | 0 | 0 | -0.06 | 28.42 | 0.35 | 28.48 | -0.06 | 28.83 | 0 | 28.59 | -0.06
0.608 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | 0.35 | 0 | -13.95 | -0.06 | -33.99 | -0.06 | -62.99 | 0 | 21.33 | 0.29 | 21.45 | -0.59 | -0.06 | -11.54 | 0 | -13.54 | 0 | -41.37 | -0.06 | 0 | -0.06 | 28.65 | 0.23 | 28.48 | 0 | 28.59 | -0.06 | 28.48 | -0.06
0.64 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.29 | 21.45 | 0.29 | -0.18 | -11.89 | -0.18 | -32.23 | -0.18 | -39.73 | -0.06 | 21.39 | 0.23 | 21.33 | -0.59 | -0.06 | -9.84 | 6.27 | -12.42 | -0.06 | -38.85 | 0 | 0 | 0 | 28.65 | 0.35 | 28.36 | 0.18 | 28.42 | 0.06 | 28.48 | 0
0.672 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.29 | -0.18 | -9.08 | -0.18 | -32.76 | 0 | -27.01 | 0 | 21.39 | 0.29 | 21.45 | -0.18 | 0 | -8.56 | 14.06 | -10.37 | 0 | -37.15 | -0.06 | 0 | 0 | 28.13 | 0.23 | 28.59 | 0.06 | 28.65 | 0.12 | 28.48 | 0
0.704 | 0 | -0.06 | 21.45 | -0.23 | 21.39 | -0.23 | 21.45 | 0.47 | -0.06 | -7.03 | -0.18 | -31.76 | -0.06 | -31.76 | -0.06 | 21.39 | 0.23 | 21.39 | -0.06 | 7.62 | -7.09 | 22.03 | -8.91 | 0.06 | -35.63 | -0.18 | 0 | 0 | 0.18 | 0.18 | 28.59 | 0.23 | 28.48 | 0.06 | 28.48 | 0.06
0.736 | 0 | -0.18 | 21.45 | -0.06 | 21.45 | -0.23 | 21.39 | 0.53 | 5.68 | -4.75 | -0.06 | -30.18 | -0.06 | -35.98 | 0 | 21.39 | 0.18 | 21.39 | 0 | 15 | -6.09 | 29.24 | -7.97 | 0 | -33.34 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.54 | 0.12 | 28.71 | -0.06 | 28.48 | 0.18
0.768 | 0 | -0.18 | 0.23 | -0.18 | 21.45 | -0.23 | 21.39 | 0.53 | 11.66 | -2.81 | -0.06 | -28.07 | -0.06 | -33.17 | 0 | 21.39 | 0.23 | 21.39 | 0.06 | 22.74 | -4.92 | 28.89 | -6.68 | -0.06 | -30.94 | -0.06 | 0 | 0 | 0 | 0.23 | 28.59 | 0.18 | 28.65 | 0.06 | 28.42 | 0.18
0.8 | -31.93 | -0.06 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 19.04 | -1.82 | -0.06 | -26.43 | -0.06 | -30 | 0 | 0.23 | 0.23 | 21.27 | 0.12 | 28.95 | -3.87 | 28.65 | -5.74 | 5.22 | -29.24 | 0 | 0.06 | 0.06 | 0 | 0.12 | 28.59 | -0.06 | 28.59 | 0.18 | 28.36 | 0.18
0.832 | -59.42 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | 21.45 | -0.76 | -0.06 | -23.67 | -0.06 | -28.24 | 0 | 0.06 | 0.23 | 21.56 | 0.18 | 28.65 | -3.16 | 28.54 | -4.45 | 12.95 | -27.31 | -0.06 | -43.48 | 0 | 0 | 0.06 | 28.36 | 0.12 | 28.48 | 0.18 | 28.48 | 0.12
0.864 | -43.65 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 21.39 | -0.7 | 0.06 | -21.45 | -0.06 | -27.07 | 0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.54 | -2.64 | 28.89 | -3.4 | 21.09 | -25.37 | -0.06 | -80.98 | -0.06 | 0 | 0 | 28.36 | 0.06 | 28.48 | 0.12 | 28.48 | 0.23
0.896 | -40.14 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.33 | 0.18 | -0.06 | -20.27 | -0.18 | -25.67 | -0.06 | 0 | 0.23 | 21.33 | 0.18 | 28.36 | -2.05 | 28.65 | -2.75 | 28.89 | -22.44 | -0.06 | -48.52 | -0.06 | 0 | 0.18 | 28.36 | -0.06 | 28.65 | 0.18 | 28.48 | 0.12
0.928 | -37.97 | -0.06 | 0 | -0.18 | 0.23 | -0.06 | 21.39 | 0.29 | 21.33 | 0.06 | -0.06 | -17.99 | -0.06 | -23.61 | 0 | 0 | 0.12 | 21.33 | 0.18 | 28.54 | -1.52 | 28.71 | -2.23 | 28.71 | -21.09 | -0.18 | -47.46 | 0 | 0 | 0.06 | 28.54 | -0.06 | 28.36 | 0.12 | 28.54 | 0.12
0.96 | -35.22 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.45 | 0.18 | 21.33 | 0.18 | 0 | -16 | 0 | -21.74 | 0.12 | 0 | 0.12 | 21.39 | 0.29 | 28.48 | -1.17 | 28.71 | -1.64 | 28.71 | -18.69 | -0.06 | -40.9 | 0 | 0 | 0 | 28.54 | 0 | 28.3 | 0.12 | 28.48 | 0.06
0.992 | -33.63 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.23 | 0.18 | 21.33 | 0.18 | -0.06 | -13.24 | -0.18 | -19.98 | -0.18 | 0 | 0.18 | 21.39 | 0.23 | 28.42 | -0.82 | 28.42 | -1.11 | 28.48 | -16.99 | -0.06 | -36.45 | 0.06 | 0 | -0.06 | 28.48 | -0.06 | 28.59 | 0.12 | 28.42 | 0.18
1.024 | -31.06 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.06 | 0.29 | 21.39 | 0.35 | -0.06 | -11.13 | -0.18 | -19.1 | -0.06 | 0 | 0.29 | 21.39 | 0.23 | 28.65 | -0.7 | 28.48 | -0.76 | 28.54 | -15.76 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0.18 | -0.06 | 28.42 | 0.06 | 28.42 | 0.12
1.056 | -29.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.41 | -0.06 | -9.08 | -0.06 | -16.58 | 0 | 0 | 0.18 | 21.15 | 0.12 | 28.83 | -0.53 | 28.42 | -0.47 | 28.77 | -13.65 | -0.06 | -37.97 | 0 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 0.23 | 0.23
1.088 | -27.77 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 0 | -6.97 | -0.06 | -14.41 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 28.65 | -0.47 | 28.71 | -0.29 | 28.59 | -12.54 | 0 | -37.09 | 0 | 0 | -0.06 | 0 | 0 | 0.29 | 0 | 0.12 | 0.18
1.12 | -24.9 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.41 | 5.86 | -4.86 | -0.06 | -12.77 | -0.06 | 0 | 0.18 | 21.45 | 0.18 | 28.3 | -0.41 | 28.71 | -0.23 | 28.54 | -10.37 | 0 | -34.75 | 0 | -44.77 | -0.06 | 0.06 | 0 | 0.12 | 0.06 | 0 | 0.12
1.152 | -23.09 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.27 | 0.41 | 11.72 | -2.23 | -0.06 | -11.25 | -0.06 | 0 | 0.12 | 21.33 | 0.18 | 28.89 | -0.29 | 28.54 | -0.18 | 28.77 | -9.14 | -0.06 | -33.05 | 0 | -84.55 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0.18
1.184 | -21.86 | -0.06 | -0.06 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | 0.29 | 17.23 | -1.93 | 0.12 | -9.61 | -0.06 | 0 | 0.18 | 21.39 | 0.29 | 28.3 | -0.29 | 28.59 | -0.23 | 28.48 | -8.14 | 0 | -31.17 | 0 | -50.28 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0
1.216 | -17.87 | 0.06 | -63.69 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.91 | -0.53 | -0.06 | -7.44 | 0 | -0.06 | 0.06 | 21.39 | 0.18 | 28.3 | -0.29 | 28.3 | -0.18 | 28.42 | -6.5 | 10.2 | -28.54 | 0 | -46.64 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0
1.248 | -16.47 | 0 | -48.05 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | -0.35 | 0 | -5.68 | 0 | -58.01 | 0.12 | 21.39 | 0.18 | 28.3 | -0.29 | 28.77 | -0.23 | 28.42 | -5.39 | 17.46 | -27.6 | 0 | -41.43 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0
1.28 | -14.94 | -0.06 | -40.02 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.15 | 0.35 | 21.39 | -0.06 | 1.99 | -3.98 | -0.06 | -53.03 | 0.06 | 21.39 | 0.06 | 28.3 | -0.23 | 28.59 | -0.18 | 28.42 | -4.45 | 24.96 | -25.55 | 0 | -36.51 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0.06
1.312 | -13.71 | -0.06 | -39.14 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.23 | 6.62 | -2.46 | -0.06 | -31.47 | 0.06 | 21.39 | 0.23 | 28.65 | -0.23 | 28.54 | -0.18 | 28.42 | -3.69 | 28.59 | -24.14 | 0.06 | -37.21 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0
1.344 | -11.78 | 0 | -36.97 | -0.06 | -24.14 | -0.18 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.29 | 14.18 | -1.7 | -0.06 | -28.13 | 0.06 | 0.18 | 0.06 | 28.3 | -0.23 | 28.36 | -0.23 | 28.71 | -2.87 | 28.65 | -22.03 | -0.06 | -38.44 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0
1.376 | -10.25 | 0 | -34.45 | -0.06 | -62.11 | -0.06 | 0 | -0.06 | 21.21 | 0.18 | 21.39 | 0.47 | 19.57 | -0.88 | -0.06 | -34.1 | 0.06 | 0 | 0.18 | 28.42 | -0.18 | 28.42 | -0.18 | 28.24 | -1.88 | 28.59 | -18.87 | 0 | -37.15 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
1.408 | -8.73 | 0 | -32.4 | -0.06 | -45.88 | -0.06 | -20.51 | -0.06 | 21.33 | 0.18 | 21.33 | 0.47 | 21.39 | -0.59 | 0.12 | -35.1 | 0.06 | 0 | 0.18 | 28.54 | -0.18 | 28.42 | -0.18 | 28.48 | -1.46 | 28.54 | -17.23 | -0.06 | -35.16 | -0.06 | 0.06 | -0.06 | 0 | -0.18 | 0 | 0.06
1.44 | -7.15 | 0 | -31.11 | -0.06 | -40.2 | -0.06 | -62.17 | -0.06 | 21.39 | 0.18 | 21.45 | 0.53 | 21.39 | -0.18 | 0.06 | -31.47 | 0 | 0 | 0.12 | 28.42 | -0.18 | 28.48 | -0.06 | 28.65 | -0.94 | 28.48 | -15.59 | -0.06 | -33.69 | -0.06 | -38.97 | -0.18 | 0 | -0.18 | 0 | -0.06
1.472 | -5.8 | 0 | -28.65 | -0.06 | -38.32 | -0.06 | -46.47 | -0.06 | 21.39 | 0.18 | 21.39 | 0.53 | 21.33 | -0.23 | -0.06 | -29.24 | 0 | 0 | 0.06 | 28.65 | -0.18 | 28.54 | -0.18 | 28.48 | -0.94 | 28.95 | -13.71 | 0 | -31.23 | -0.06 | -73.6 | 0 | 0 | -0.06 | -29.94 | 0
1.504 | -4.1 | 0 | -26.08 | 0 | -35.51 | -0.06 | -40.37 | -0.06 | 21.39 | 0 | 21.45 | 0.47 | 21.39 | -0.06 | -0.18 | -27.48 | -0.06 | 0 | 0.12 | 28.42 | -0.18 | 28.59 | -0.18 | 28.42 | -0.59 | 28.77 | -11.66 | 0 | -29.12 | -0.18 | -51.39 | -0.18 | -25.31 | 0 | -84.91 | -0.06
1.536 | -2.81 | 0 | -23.85 | -0.06 | -34.1 | -0.06 | -38.85 | -0.06 | 21.33 | 0.06 | 21.45 | 0.47 | 21.39 | 0.06 | 0 | -27.19 | -0.06 | 0 | 0.18 | 28.42 | -0.06 | 0.29 | -0.18 | 28.3 | -0.35 | 28.48 | -8.79 | 0 | -27.19 | -0.06 | -47.23 | -0.12 | -88.13 | 0 | -56.31 | -0.06
1.568 | -1.64 | 4.51 | -21.62 | -0.06 | -31.76 | -0.18 | -35.33 | -0.06 | 21.15 | 0 | 21.39 | 0.41 | 21.33 | 0.23 | 7.03 | -25.55 | 0 | 0 | 0.18 | 28.13 | -0.18 | 0.06 | -0.18 | 28.48 | -0.23 | 28.36 | -6.68 | -0.06 | -25.55 | -0.18 | -41.54 | -0.18 | -58.01 | -0.06 | -33.17 | -0.06
1.6 | -0.94 | 10.55 | -18.99 | -0.06 | -29.42 | -0.06 | -34.86 | -0.06 | 17.52 | 0.06 | 21.45 | 0.41 | 21.33 | 0.23 | 12.77 | -22.09 | 0 | 0 | 0.18 | 0.23 | -0.18 | 0 | -0.06 | 28.42 | -0.29 | 28.42 | -4.51 | -0.18 | -23.85 | -0.18 | -36.45 | -0.06 | -33.46 | 0 | -34.45 | -0.06
1.632 | -0.53 | 15.88 | -17.99 | -0.06 | -27.77 | -0.06 | -32.23 | 0.06 | 0.12 | 0.12 | 21.39 | 0.35 | 21.39 | 0.29 | 18.22 | -20.98 | 0 | 0 | 0.18 | 0.06 | -0.18 | 0 | -0.06 | 28.65 | -0.29 | 28.42 | -2.87 | 0 | -22.15 | 0 | -37.56 | -0.06 | -34.63 | 0 | -42.25 | 0
1.664 | -0.23 | 21.68 | -15.7 | -0.18 | -25.55 | -0.06 | -30.29 | 0 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.29 | 21.45 | -19.69 | 0 | 0 | 0.18 | 0 | -0.18 | 0 | -0.06 | 28.42 | -0.35 | 28.65 | -1.82 | 3.4 | -19.1 | 0 | -38.38 | -0.06 | -42.25 | 0 | -40.78 | -0.18
1.696 | -0.18 | 21.45 | -13.77 | -0.06 | -23.61 | -0.06 | -27.48 | 0 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0.18 | 21.33 | -18.05 | -0.06 | 0 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0.23 | -0.29 | 28.36 | -0.82 | 11.84 | -17.17 | -0.18 | -37.03 | -0.06 | -40.84 | 0 | -36.92 | -0.18
1.728 | -0.18 | 21.45 | -12.07 | -0.06 | -21.5 | -0.06 | -25.9 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.21 | 0.23 | 21.39 | -16.35 | -0.06 | 0 | 0.18 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.29 | 28.71 | -0.29 | 19.4 | -15.76 | 0 | -35.57 | -0.06 | -37.74 | -0.06 | -32.93 | -0.18
1.76 | -0.35 | 21.45 | -11.31 | 0 | -18.46 | -0.06 | -23.38 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 21.5 | -14.06 | 0.06 | -35.98 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.23 | 28.42 | -0.06 | 28.83 | -13.54 | -0.06 | -33.22 | -0.18 | -33.93 | 0 | -32.99 | -0.06
1.792 | -0.41 | 21.45 | -9.61 | -0.06 | -17.17 | -0.06 | -21.86 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | 0.23 | 21.39 | -12.54 | 0 | -67.33 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0.06 | -0.23 | 28.48 | 0.06 | 28.48 | -11.6 | -0.06 | -31.11 | -0.06 | -33.17 | 0 | -32.93 | -0.18
1.824 | -0.18 | 21.5 | -8.03 | -0.06 | -15.94 | -0.18 | -19.81 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | 0.23 | 21.33 | -10.49 | 0 | -45 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.12 | 28.48 | -9.2 | -0.06 | -29.47 | -0.06 | -31.99 | 0 | -30.7 | 0
1.856 | -0.41 | 21.45 | -6.45 | -0.06 | -13.59 | -0.06 | -17.05 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.27 | 0.23 | 21.39 | -9.26 | -0.06 | -29.71 | 0.12 | 0 | 0 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.35 | 28.71 | -6.86 | -0.06 | -27.48 | -0.18 | -30.59 | 0 | -28.65 | -0.18
1.888 | -0.35 | 21.45 | -4.86 | -0.06 | -12.42 | -0.06 | -15.53 | 0 | 0 | -0.06 | 21.33 | 0 | 21.39 | 0.18 | 21.33 | -7.56 | -0.06 | -29.06 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.29 | 28.48 | 0.35 | 28.48 | -5.16 | -0.06 | -25.43 | -0.06 | -29.24 | -0.06 | -27.01 | 0
1.92 | -0.23 | 21.45 | -3.63 | 0.06 | -10.08 | -0.06 | -13.95 | 0 | 0 | 0 | 21.33 | 0.18 | 21.27 | 0.23 | 21.39 | -5.33 | -0.06 | -35.63 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.48 | 0.29 | 28.48 | -2.87 | 0 | -23.67 | 0 | -27.07 | 0 | -25.14 | 0
1.952 | -0.18 | 21.5 | -2.46 | 2.29 | -8.67 | -0.06 | -12.48 | 0 | 0 | -0.06 | 21.45 | -0.06 | 21.33 | 0.18 | 21.39 | -3.34 | -0.06 | -34.22 | 0.12 | 0 | -0.06 | -33.22 | -0.18 | 0 | -0.23 | 28.65 | 0.29 | 28.65 | -1.88 | 0 | -21.09 | 0 | -25.84 | -0.06 | -24.08 | -0.06
1.984 | -0.23 | 21.45 | -1.52 | 7.79 | -7.62 | 0 | -11.13 | 0 | 0 | -0.06 | 21.15 | -0.06 | 21.33 | 0.06 | 21.39 | -2.23 | -0.06 | -31.29 | 0.12 | 0.06 | 0 | -75.06 | -0.06 | 0 | -0.18 | 28.54 | 0.41 | 28.48 | -0.64 | 0 | -19.45 | -0.06 | -24.43 | 0 | -22.33 | -0.06
2.016 | -0.29 | 21.45 | -0.7 | 13.42 | -6.21 | 6.8 | -9.2 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.27 | 0 | 21.15 | -1.82 | 0.06 | -28.48 | 0.12 | -38.91 | -0.06 | -51.45 | -0.06 | 0 | -0.18 | 28.54 | 0.35 | 28.42 | -0.59 | -0.06 | -17.64 | -0.18 | -22.74 | 0.06 | -20.92 | -0.18
2.048 | -0.23 | 21.45 | -0.35 | 18.99 | -4.63 | 12.42 | -7.68 | -0.18 | -38.61 | 0.06 | 0 | 0.06 | 21.09 | 0.06 | 21.39 | -1.05 | 0 | -27.48 | 0 | -69.2 | -0.06 | -46.23 | -0.06 | 0 | -0.18 | 28.42 | 0.35 | 28.42 | 0 | -0.06 | -15.7 | -0.06 | -20.98 | 0 | -19.4 | 0
2.08 | -0.29 | 21.45 | -0.18 | 21.56 | -2.99 | 17.99 | -6.68 | -0.06 | -61.47 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | -0.53 | -0.06 | -26.13 | 0.12 | -49.57 | 0 | -43.77 | -0.18 | 0 | -0.18 | 23.79 | 0.35 | 28.3 | -0.18 | 0 | -13.83 | -0.06 | -19.34 | 0 | -17.29 | 0
2.112 | -0.29 | 21.45 | -0.18 | 21.45 | -2.05 | 21.74 | -4.63 | -0.06 | -43.24 | -0.06 | 0 | 0 | 21.15 | 0.06 | 21.15 | -0.23 | 2.99 | -24.26 | 0.06 | -44.88 | -0.06 | -42.01 | -0.06 | -25.84 | -0.18 | 0.18 | 0.18 | 28.48 | 0.18 | 3.87 | -11.31 | -0.06 | -17.29 | 0 | -15.94 | -0.06
2.144 | -0.35 | 21.45 | -0.29 | 21.45 | -1.05 | 21.45 | -3.57 | -0.06 | -39.38 | 0 | 0 | -0.06 | 21.33 | 0 | 21.39 | -0.18 | 9.38 | -21.86 | 0.06 | -43.71 | -0.06 | -40.43 | -0.18 | -78.81 | -0.18 | 0 | 0.23 | 28.65 | 0.12 | 12.31 | -9.38 | -0.12 | -15.7 | 0 | -14.06 | 0
2.176 | -0.23 | 21.45 | -0.35 | 21.45 | -0.53 | 21.45 | -1.99 | -0.06 | -33.99 | 0 | 0 | -0.18 | 11.07 | -0.06 | 21.15 | -0.06 | 16.41 | -20.39 | 0.18 | -40.9 | 0 | -38.32 | -0.06 | -48.46 | -0.06 | 0 | 0.29 | 28.65 | 0.41 | 19.92 | -7.09 | 11.72 | -14.18 | 4.22 | -12.48 | -0.06
2.208 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -1.17 | -0.06 | -32.58 | 0 | 0 | -0.06 | 0.12 | 0 | 21.33 | 0.12 | 21.5 | -19.28 | 0.12 | -39.85 | 0.06 | -36.45 | -0.06 | -46.47 | -0.18 | 0 | 0.23 | 28.36 | 0.35 | 27.36 | -5.33 | 19.28 | -12.77 | 12.36 | -10.9 | 4.39
2.24 | -0.23 | 21.45 | -0.41 | 21.45 | -0.06 | 21.45 | -0.76 | 0 | -32.46 | 0 | 0 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.45 | -17.23 | 0 | -37.85 | 0 | -34.28 | 0.06 | -44.42 | 0 | 0 | 0.23 | 28.48 | 0.35 | 28.36 | -3.34 | 26.95 | -10.9 | 20.16 | -8.79 | 12.72
2.272 | -0.18 | 21.45 | -0.35 | 21.39 | -0.18 | 21.45 | -0.29 | -0.06 | -31.88 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -15.35 | 0.06 | -35.22 | 0 | -32.99 | 0 | -41.43 | -0.06 | 0 | 0.18 | 28.48 | 0.18 | 28.54 | -1.7 | 28.65 | -8.61 | 27.66 | -7.27 | 20.22
2.304 | -0.06 | 21.39 | -0.23 | 21.45 | -0.18 | 21.39 | -0.06 | -0.18 | -30.59 | 0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -13.54 | 0.12 | -33.93 | 0 | -30 | 0 | -40.9 | -0.18 | 0 | 0 | 28.48 | 0.23 | 28.48 | -1.11 | 28.54 | -6.97 | 28.48 | -5.63 | 28.13
2.336 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | -0.18 | 0 | -28.24 | -0.06 | 0 | 0 | 0 | -0.06 | 21.27 | 0.12 | 21.33 | -12.31 | 0.06 | -31.52 | 0.06 | -28.36 | 0.06 | -38.2 | 0 | 0 | 0.12 | 28.48 | 0.18 | 28.48 | -0.47 | 28.59 | -5.39 | 28.36 | -3.93 | 28.65
2.368 | -0.18 | 21.39 | -0.23 | 21.45 | -0.41 | 21.45 | -0.35 | 0 | -26.31 | 0.06 | 0 | 0 | 0 | 0 | 21.27 | 0.23 | 21.27 | -10.37 | -0.06 | -29.83 | 0.06 | -26.54 | -0.18 | -36.1 | 0 | 0 | 0 | 28.48 | 0.18 | 28.42 | -0.29 | 28.3 | -4.04 | 28.48 | -2.75 | 28.42
2.4 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | 21.45 | -0.35 | 6.8 | -23.56 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.39 | -8.14 | 0.06 | -28.07 | -0.06 | -23.91 | 0 | -34.75 | 0 | 0 | 0.06 | 28.42 | 0.18 | 28.42 | -0.18 | 28.36 | -2.75 | 28.42 | -1.76 | 28.65
2.432 | -0.18 | 21.39 | -0.29 | 21.5 | -0.29 | 21.45 | -0.35 | 12.42 | -21.68 | -0.06 | -41.9 | -0.06 | 0 | 0 | 21.21 | 0.23 | 21.33 | -5.98 | 0 | -26.43 | 0 | -22.21 | 0 | -32.76 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.48 | 0.06 | 28.36 | -1.7 | 28.59 | -1.11 | 28.36
2.464 | -0.06 | 0.23 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 18.05 | -19.4 | 0 | -69.49 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.27 | -4.63 | 4.63 | -23.85 | 0 | -19.98 | 0.06 | -30.59 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.65 | 0.12 | 28.48 | -1.05 | 28.65 | -0.7 | 28.48
2.496 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.56 | -17.75 | 0.06 | -47.52 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -3.4 | 12.48 | -21.86 | -0.06 | -18.63 | -0.06 | -29.01 | 0 | 0 | 0 | 28.48 | 0 | 28.42 | 0.23 | 28.48 | -0.82 | 28.3 | -0.35 | 28.48
2.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -15.94 | -0.06 | -42.54 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.33 | -2.29 | 18.05 | -19.63 | 0.06 | -16.47 | 0 | -25.9 | -0.06 | -79.05 | -0.06 | 28.42 | 0.06 | 28.54 | 0.29 | 28.42 | -0.23 | 28.65 | -0.23 | 28.36
2.56 | -0.06 | 0.06 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | -13.36 | -0.06 | -33.81 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.27 | -1.35 | 21.56 | -18.69 | 0 | -14.47 | -0.06 | -23.97 | -0.06 | -69.79 | 0.06 | 0.23 | 0.12 | 28.48 | 0.35 | 28.59 | -0.35 | 28.48 | -0.18 | 28.18
2.592 | -0.06 | 0 | -0.29 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -11.07 | 0 | -31.64 | 0.06 | 0 | -0.06 | 0 | 0.29 | 21.39 | -0.76 | 21.39 | -15.76 | 0 | -13.59 | 0 | -22.21 | -0.06 | -48.63 | 0 | 0.06 | 0 | 28.48 | 0.35 | 28.65 | 0 | 28.65 | 0.06 | 28.42
2.624 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -9.08 | -0.06 | -32.93 | 0.06 | -73.54 | -0.06 | 0 | 0.29 | 21.45 | -0.41 | 21.39 | -14.41 | 5.04 | -11.72 | 0 | -19.98 | -0.06 | -46.23 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.54 | -0.18 | 28.42 | 0.12 | 28.42
2.656 | 0 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.23 | 21.33 | -6.68 | 0.12 | -31.99 | -0.06 | -58.89 | -0.06 | 0 | 0.12 | 21.39 | -0.18 | 21.45 | -13.01 | 13.13 | -10.2 | 0.41 | -17.99 | 0 | -39.44 | -0.06 | 0.06 | 0 | 28.89 | 0.23 | 28.24 | 0.18 | 28.59 | 0.06 | 28.83
2.688 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | -0.23 | 21.39 | -4.69 | -0.18 | -31.06 | 0 | -43.3 | 0 | 0 | 0.12 | 21.39 | 0 | 21.39 | -11.6 | 20.98 | -9.02 | 10.2 | -16.29 | 0 | -36.92 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.48 | 0.12 | 28.42 | 0.06 | 28.13
2.72 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | -2.4 | -0.06 | -28.07 | -0.06 | -38.44 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 21.39 | -10.08 | 28.89 | -7.27 | 17.46 | -14.65 | 0 | -37.44 | -0.18 | 0 | -0.06 | 28.54 | 0.06 | 28.07 | 0.12 | 28.83 | 0.12 | 28.48
2.752 | -0.06 | 0 | -0.18 | 21.45 | -0.12 | 21.45 | -0.18 | 21.39 | -1.64 | 0 | -26.43 | -0.06 | -33.28 | -0.06 | 0 | 0.23 | 21.39 | 0.12 | 21.39 | -8.67 | 28.83 | -6.09 | 27.83 | -13.36 | -0.06 | -38.15 | 0 | 0.06 | -0.06 | 28.59 | 0.18 | 28.54 | 0.06 | 28.59 | 0.06 | 28.42
2.784 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -0.94 | 5.22 | -24.26 | -0.18 | -32.7 | -0.06 | 0 | 0.12 | 21.33 | 0.06 | 21.15 | -6.86 | 28.77 | -5.1 | 28.59 | -12.07 | 0 | -36.86 | 0 | 0 | 0 | 28.59 | 0.12 | 28.42 | 0.18 | 28.77 | 0.23 | 28.36
2.816 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | -0.59 | 11.19 | -22.15 | -0.06 | -32.7 | 0 | 0 | 0.18 | 21.39 | 0.06 | 21.39 | -5.98 | 28.48 | -4.22 | 28.77 | -10.02 | 0 | -34.1 | 0.06 | 0 | -0.06 | 28.65 | 0.18 | 28.48 | 0.12 | 28.48 | 0.06 | 28.48
2.848 | -0.06 | 0 | -0.06 | 18.05 | -0.18 | 21.39 | -0.18 | 21.39 | -0.06 | 16.52 | -20.39 | -0.06 | -31.76 | 0 | 0.06 | 0.23 | 21.39 | 0.23 | 21.39 | -4.92 | 28.71 | -3.34 | 28.65 | -8.85 | 0 | -32.34 | 0 | 0 | -0.06 | 28.42 | 0.06 | 28.48 | 0.12 | 28.48 | 0.12 | 28.07
2.88 | -0.18 | -22.27 | -0.18 | 0.12 | -0.06 | 21.39 | -0.18 | 21.45 | 0.12 | 21.86 | -17.64 | -0.06 | -30.47 | 0 | -49.1 | 0.23 | 21.33 | 0.18 | 21.45 | -4.16 | 28.42 | -2.52 | 28.3 | -7.85 | 0 | -30.82 | -0.06 | 0 | -0.06 | 28.54 | 0 | 28.42 | 0.12 | 28.07 | 0.18 | 28.42
2.912 | -0.06 | -64.28 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.33 | 0.18 | 21.45 | -16.06 | -0.06 | -27.77 | 0 | -64.46 | 0.18 | 21.39 | 0.18 | 21.39 | -3.28 | 29.01 | -1.93 | 28.65 | -5.98 | -0.06 | -29.06 | -0.18 | 0 | -0.06 | 28.36 | 0.06 | 28.3 | 0.06 | 28.48 | 0.18 | 28.36
2.944 | -0.06 | -46.35 | -0.06 | 0 | -0.18 | 0 | -0.12 | 21.33 | 0.35 | 21.39 | -13.83 | -0.06 | -25.25 | -0.06 | -40.43 | 0.18 | 21.27 | 0.23 | 21.39 | -2.75 | 28.42 | -1.64 | 28.65 | -5.39 | 0 | -27.31 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.36 | 0.06 | 28.48 | 0.18 | 28.42
2.976 | -0.06 | -40.14 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -11.72 | 0.06 | -23.32 | -0.06 | -27.54 | 0.18 | 21.39 | 0.12 | 21.39 | -1.93 | 28.42 | -1.11 | 28.71 | -3.93 | 0 | -24.61 | -0.06 | -9.79 | 0 | 28.42 | 0 | 28.42 | 0.06 | 28.54 | 0.06 | 28.48
3.008 | -0.06 | -38.5 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.45 | 0.47 | 21.39 | -10.02 | -0.18 | -21.09 | 0 | -32.64 | 0.12 | 0.12 | 0.23 | 21.21 | -1.82 | 28.83 | -0.82 | 28.83 | -3.16 | 0 | -23.32 | -0.06 | -92.58 | 0 | 0.7 | 0 | 28.48 | -0.06 | 28.48 | 0.12 | 28.42
3.04 | 0 | -36.04 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.27 | 0.53 | 21.21 | -6.91 | -0.06 | -19.34 | -0.06 | -35.68 | 0.18 | 0.06 | 0.23 | 21.21 | -1.11 | 28.65 | -0.59 | 28.59 | -2.52 | 0.06 | -21.09 | -0.06 | -54.96 | -0.18 | 0.06 | -0.06 | 3.28 | -0.06 | 28.42 | 0.12 | 28.48
3.072 | -0.06 | -34.22 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.45 | 0.41 | 21.39 | -4.63 | -0.06 | -17.23 | -0.06 | -34.1 | 0.12 | 0 | 0.23 | 21.39 | -0.94 | 28.48 | -0.35 | 28.65 | -1.82 | -0.06 | -18.75 | -0.06 | -49.81 | 0 | 0 | 0.06 | 0.12 | 0 | 0.29 | 0.12 | 28.48
3.104 | -0.06 | -32.05 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -2.75 | -0.06 | -15.06 | -0.06 | -30.29 | 0.06 | 0 | 0.18 | 21.39 | -0.76 | 28.42 | -0.29 | 28.48 | -1.29 | -0.18 | -17.23 | -0.06 | -44.3 | -0.06 | 0.06 | 0 | 0 | -0.06 | 0.12 | 0.12 | 0.29
3.136 | -0.18 | -29.53 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -1.7 | -0.06 | -13.07 | 0 | -27.83 | 0.12 | 0 | 0.18 | 21.39 | -0.59 | 28.54 | -0.29 | 28.54 | -0.76 | 10.9 | -15 | -0.06 | -37.91 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 0 | 0.06 | 0.06
3.168 | -0.06 | -27.77 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.47 | 21.39 | -1.11 | 4.98 | -11.02 | 0 | -26.72 | 0.06 | 0 | 0.23 | 21.39 | -0.53 | 28.13 | -0.29 | 28.42 | -0.7 | 18.22 | -12.72 | -0.06 | -36.68 | -0.18 | 0 | 0 | 0 | 0.06 | 0.06 | 0 | 0
3.2 | -0.06 | -25.31 | -0.06 | 0 | 0 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -0.41 | 10.9 | -8.61 | -0.06 | -25.02 | 0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.65 | -0.29 | 28.77 | -0.53 | 26.13 | -10.61 | -0.06 | -37.97 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 0.06
3.232 | -0.06 | -23.79 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -0.06 | 18.28 | -6.68 | -0.06 | -23.15 | 0.12 | 0 | 0.18 | 21.39 | -0.29 | 28.42 | -0.23 | 28.48 | -0.47 | 28.59 | -8.67 | 0 | -37.33 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
3.264 | 0 | -21.5 | -0.06 | -0.12 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.12 | 21.62 | -3.52 | 4.22 | -22.09 | 0.06 | 0 | 0.18 | 21.33 | -0.29 | 28.77 | -0.29 | 28.83 | -0.35 | 28.59 | -6.68 | 0 | -35.92 | -0.06 | 0 | -0.06 | 0 | 0 | 0.06 | 0 | 0
3.296 | -0.06 | -18.87 | -0.06 | -45.59 | -0.06 | -0.06 | -0.06 | 0.06 | 0.18 | 21.39 | 0.29 | 21.39 | -2.23 | 12.13 | -19.69 | -0.06 | 0 | 0.23 | 21.39 | -0.41 | 28.59 | -0.29 | 28.42 | -0.18 | 28.65 | -4.16 | 0 | -34.22 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 0
3.328 | 0 | -17.05 | 0 | -54.55 | 0 | -36.15 | -0.18 | 0 | 0.18 | 21.5 | 0.35 | 21.45 | -1.17 | 17.46 | -18.34 | 0 | 0 | 0.18 | 21.33 | -0.35 | 28.54 | -0.29 | 28.36 | -0.23 | 29.12 | -2.64 | -0.06 | -31.23 | -0.18 | 0 | 0.06 | 0 | 0 | 0 | 0 | 0
3.36 | -0.18 | -15.29 | -0.06 | -42.13 | 0 | -58.01 | -0.18 | 0 | 0.23 | 21.39 | 0.41 | 21.39 | -0.76 | 21.74 | -16.52 | 0.06 | 0 | 0.18 | 0.12 | -0.35 | 28.77 | -0.23 | 28.07 | -0.23 | 28.48 | -1.46 | 0 | -29.59 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
3.392 | -0.06 | -13.95 | 0 | -39.44 | -0.06 | -42.78 | -0.18 | 0 | 0 | 21.45 | 0.53 | 21.39 | -0.23 | 21.39 | -14.71 | 0 | 0 | 0.23 | 0 | -0.41 | 28.59 | -0.18 | 28.59 | -0.18 | 28.42 | -1.11 | 0 | -28.3 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 0
3.424 | 0 | -12.48 | 0 | -37.56 | -0.06 | -39.85 | -0.06 | 0 | 0.18 | 21.39 | 0.53 | 21.39 | -0.06 | 21.33 | -13.07 | 0.06 | -5.33 | 0.18 | 0 | -0.29 | 28.95 | -0.18 | 28.48 | -0.18 | 28.65 | -0.53 | -0.12 | -26.49 | 0 | -14.53 | -0.18 | 0 | -0.06 | 0 | 0 | 0
3.456 | -0.06 | -10.25 | -0.06 | -34.51 | -0.06 | -37.74 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.21 | 0.18 | 21.39 | -10.72 | -0.06 | -70.08 | 0.23 | 0 | -0.23 | 28.59 | -0.23 | 28.59 | -0.29 | 28.42 | -0.18 | -0.18 | -24.61 | -0.06 | -93.58 | 0 | -4.16 | -0.18 | 0 | -0.06 | 0
3.488 | 5.04 | -9.2 | -0.06 | -33.81 | 0 | -35.68 | -0.06 | 0 | 0.06 | 21.45 | 0.53 | 21.39 | 0.18 | 21.39 | -9.32 | -0.18 | -48.4 | 0.18 | 0 | -0.18 | 28.83 | -0.18 | 28.59 | -0.18 | 28.42 | 0.12 | -0.18 | -22.62 | -0.06 | -60.12 | -0.06 | -89.59 | -0.18 | -18.99 | 0 | 0
3.52 | 11.07 | -7.85 | -0.06 | -31.23 | 0 | -33.34 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.33 | 0.29 | 21.39 | -7.21 | -0.06 | -29.24 | 0.18 | 0 | -0.18 | 0.18 | -0.06 | 28.65 | -0.18 | 28.36 | 0.06 | 5.74 | -19.92 | -0.06 | -48.28 | 0 | -64.46 | -0.06 | -89.89 | -0.06 | -24.67
3.552 | 16.17 | -6.33 | -0.06 | -28.54 | -0.06 | -31.41 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | 0.41 | 21.39 | -5.8 | 0 | -28.65 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.35 | -0.23 | 28.42 | 0.18 | 13.71 | -17.93 | -0.06 | -43.89 | 0 | -36.1 | 0 | -62.17 | -0.06 | -89.53
3.584 | 21.74 | -4.86 | -0.06 | -26.78 | -0.06 | -29.53 | -0.06 | 0 | -0.06 | 21.5 | 0.35 | 21.39 | 0.47 | 21.39 | -4.22 | -0.06 | -35.16 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0.12 | -0.23 | 28.36 | 0.18 | 21.33 | -16.29 | 0 | -37.62 | 0 | -29.06 | -0.06 | -34.81 | -0.18 | -58.89
3.616 | 21.5 | -2.93 | -0.06 | -24.38 | -0.06 | -26.78 | -0.06 | 0 | -0.06 | 21.21 | 0.35 | 21.39 | 0.53 | 21.39 | -2.23 | -0.06 | -34.81 | 0.23 | 0 | -0.23 | 0.06 | -0.06 | 0 | -0.18 | 28.36 | 0.18 | 28.83 | -14.06 | -0.18 | -36.97 | -0.18 | -41.54 | -0.06 | -33.17 | -0.18 | -30.94
3.648 | 21.62 | -2.05 | -0.06 | -22.33 | 0 | -24.38 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.21 | 0.41 | 21.39 | -2.05 | -0.06 | -31.88 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.47 | 28.48 | -12.25 | 0 | -38.32 | -0.06 | -41.66 | -0.18 | -41.84 | 0 | -33.05
3.68 | 21.39 | -1.17 | 0 | -20.74 | 0 | -22.91 | 0 | 0 | -0.18 | 0.23 | 0.18 | 21.39 | 0.41 | 21.21 | -0.94 | -0.06 | -28.89 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.36 | 0.41 | 28.48 | -10.14 | -0.18 | -37.09 | -0.06 | -37.85 | -0.18 | -41.37 | 0.06 | -41.84
3.712 | 21.45 | -0.76 | 0 | -17.99 | -0.06 | -20.92 | -0.06 | -51.33 | -0.06 | 0 | 0.18 | 21.27 | 0.29 | 21.39 | -0.76 | 0.06 | -27.31 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.35 | 28.48 | -7.79 | -0.06 | -36.15 | -0.06 | -34.04 | -0.06 | -37.44 | -0.18 | -41.84
3.744 | 21.45 | -0.35 | -0.06 | -16.06 | 0.06 | -18.93 | -0.06 | -66.21 | -0.06 | 0 | 0.23 | 21.21 | 0.35 | 21.21 | -0.23 | 0 | -26.43 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.3 | -5.63 | 0 | -33.63 | 0 | -32.93 | -0.18 | -34.28 | -0.06 | -37.21
3.776 | 21.45 | -0.18 | 0 | -14.82 | -0.06 | -16.17 | -0.06 | -43.3 | -0.06 | 0.06 | 0.18 | 21.27 | 0.35 | 21.39 | -0.29 | -0.06 | -24.32 | 0.12 | -45.76 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.35 | 28.48 | -3.69 | 0 | -32.23 | -0.06 | -32.58 | -0.18 | -33.17 | 0 | -33.63
3.808 | 21.45 | -0.06 | -0.06 | -13.36 | 0 | -14.82 | 0.06 | -38.44 | -0.06 | 0 | 0.12 | 21.21 | 0.29 | 21.33 | 0.18 | -0.06 | -22.62 | 0.06 | -65.1 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.35 | 28.42 | -2.17 | 5.1 | -30.18 | -0.06 | -31.82 | -0.06 | -32.81 | 0 | -33.05
3.84 | 21.45 | -0.23 | -0.06 | -12.01 | 0 | -12.83 | -0.18 | -33.69 | 0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.12 | -0.06 | -20.92 | 0.06 | -41.6 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.54 | 0.35 | 28.42 | -1.17 | 13.3 | -27.6 | 0.12 | -29.59 | -0.06 | -30.94 | -0.06 | -32.93
3.872 | 21.45 | -0.23 | 0.23 | -9.84 | 0 | -11.89 | 0 | -32.29 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | 0.18 | 0.06 | -19.63 | 0.06 | -29.06 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.42 | -0.7 | 20.92 | -26.31 | -0.06 | -28.07 | -0.06 | -29.01 | -0.06 | -30.94
3.904 | 21.45 | -0.29 | 7.5 | -8.56 | -0.06 | -10.61 | 0 | -33.05 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | 0.06 | 1.99 | -17.87 | 0.18 | -30.53 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | -0.41 | 28.95 | -24.2 | 0 | -26.31 | -0.06 | -27.72 | -0.06 | -28.24
3.936 | 21.5 | -0.35 | 15.06 | -6.97 | 0 | -8.85 | 0 | -31.64 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 21.39 | 0.18 | 7.09 | -15.23 | 0.12 | -35.51 | -0.18 | -25.14 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.59 | -0.18 | 28.77 | -22.62 | -0.06 | -24.55 | 0 | -26.02 | 0 | -26.9
3.968 | 21.39 | -0.29 | 20.39 | -5.22 | -0.06 | -7.15 | -0.06 | -30.35 | 0 | 0 | 0 | 21.33 | 0.18 | 21.39 | 0.23 | 12.89 | -13.65 | 0.12 | -34.16 | -0.18 | -78.81 | -0.06 | -0.06 | -0.06 | 28.54 | 0 | 28.48 | 0 | 28.89 | -20.63 | -0.06 | -22.62 | 0 | -24.2 | 0 | -25.9
4 | 21.39 | -0.41 | 21.5 | -4.16 | 0 | -5.68 | -0.06 | -28.24 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.23 | 20.39 | -12.01 | 0.06 | -31.29 | -0.18 | -48.69 | -0.06 | -61.17 | -0.18 | 18.16 | 0 | 28.59 | 0.29 | 28.54 | -17.75 | -0.18 | -21.33 | -0.06 | -22.74 | -0.18 | -24.43
4.032 | 21.39 | -0.23 | 21.45 | -2.75 | 4.69 | -4.16 | 0.06 | -26.08 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.45 | 0.12 | 21.5 | -10.78 | 0.12 | -28.01 | -0.06 | -45.59 | -0.06 | -56.37 | -0.06 | 0.18 | 0.06 | 28.65 | 0.18 | 28.48 | -16.29 | 0 | -19.45 | 0 | -21.33 | 0 | -23.03
4.064 | 21.45 | -0.23 | 21.5 | -1.46 | 10.9 | -3.22 | -0.06 | -24.55 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.33 | 0.18 | 21.5 | -8.85 | 0.06 | -27.01 | -0.06 | -45 | -0.06 | -44.94 | -0.06 | 0 | 0.12 | 28.59 | 0.35 | 28.54 | -14.06 | 0 | -18.34 | 0 | -19.75 | 0 | -20.86
4.096 | 21.45 | -0.35 | 21.45 | -0.94 | 16.06 | -1.82 | 3.98 | -21.56 | -0.06 | -34.45 | 0 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | -6.39 | 0 | -25.55 | -0.18 | -41.31 | -0.06 | -45.06 | 0 | 0 | -0.06 | 28.89 | 0.35 | 28.65 | -12.48 | 0 | -16.23 | 0 | -17.81 | 0 | -19.16
4.128 | 21.45 | -0.18 | 21.45 | -0.53 | 21.91 | -0.94 | 10.31 | -19.22 | 0 | -71.31 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -5.1 | 0.12 | -23.79 | -0.18 | -40.31 | -0.06 | -43.48 | -0.18 | 0 | 0.06 | 28.59 | 0.41 | 28.65 | -10.37 | 4.8 | -14.24 | 0 | -16.47 | 0 | -17.75
4.16 | 21.45 | -0.29 | 21.45 | -0.06 | 21.5 | -0.41 | 15.41 | -17.58 | 0 | -48.75 | -0.06 | 0 | 0 | 0.12 | 0.12 | 21.45 | -3.57 | 0.06 | -21.91 | -0.06 | -37.91 | -0.18 | -41.43 | 0 | 0 | 0 | 28.48 | 0.35 | 28.65 | -7.79 | 13.07 | -12.89 | 0 | -13.95 | 4.39 | -15.7
4.192 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -0.18 | 21.27 | -15.47 | -0.18 | -39.9 | -0.06 | 0 | 0.12 | 0 | 0.12 | 21.39 | -2.46 | 0 | -20.51 | -0.18 | -36.1 | 0 | -38.61 | -0.18 | 0 | -0.06 | 28.59 | 0.29 | 28.59 | -5.51 | 23.26 | -11.13 | 4.22 | -12.77 | 12.66 | -14.12
4.224 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | -13.89 | -0.06 | -34.98 | -0.06 | 0 | 0 | 0 | 0.18 | 21.33 | -1.7 | 0.06 | -19.1 | 0 | -34.4 | -0.06 | -36.92 | -0.18 | 0 | -0.06 | 28.48 | 0.29 | 28.71 | -3.52 | 29.01 | -9.61 | 12.48 | -10.78 | 20.45 | -12.48
4.256 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -10.84 | 0 | -31.76 | -0.06 | 0 | 0 | 0.06 | 0.12 | 21.39 | -0.94 | -0.06 | -16.88 | -0.06 | -32.76 | -0.06 | -35.22 | -0.18 | 0 | 0 | 28.59 | 0.12 | 28.65 | -1.99 | 28.65 | -7.56 | 20.22 | -9.26 | 28.13 | -10.61
4.288 | 21.45 | -0.18 | 21.39 | -0.41 | 21.45 | -0.41 | 21.5 | -8.79 | 0.06 | -32.76 | 0 | 0 | 0.06 | 0 | 0.12 | 21.33 | -0.41 | 0 | -14.82 | -0.06 | -30.12 | -0.06 | -34.04 | 0 | 0 | 0 | 28.48 | 0.23 | 28.83 | -1.41 | 28.48 | -5.92 | 27.95 | -7.68 | 28.65 | -9.32
4.32 | 21.45 | -0.06 | 21.5 | -0.47 | 21.45 | -0.35 | 21.45 | -6.8 | 0 | -32.46 | 0.06 | 0 | 0 | 0 | 0.06 | 21.39 | -0.41 | -0.06 | -13.54 | -0.06 | -27.95 | -0.06 | -31.76 | -0.06 | 0 | -0.18 | 28.48 | 0.12 | 28.71 | -0.64 | 28.48 | -4.22 | 28.48 | -5.57 | 28.59 | -7.32
4.352 | 21.45 | -0.18 | 21.45 | -0.35 | 21.39 | -0.47 | 21.45 | -4.45 | -0.06 | -30.12 | 0 | 0 | 0 | 0 | 0.06 | 21.39 | -0.18 | -0.18 | -11.89 | 0 | -25.9 | 0 | -28.77 | -0.06 | 0 | -0.06 | 28.36 | 0.12 | 28.48 | -0.35 | 28.48 | -2.75 | 28.48 | -4.22 | 28.65 | -5.74
4.384 | 0.12 | -0.06 | 21.45 | -0.18 | 21.39 | -0.29 | 21.39 | -2.64 | -0.06 | -28.59 | 0.06 | 0 | 0.06 | 0 | 0.06 | 21.33 | 0.06 | -0.06 | -10.2 | -0.06 | -24.61 | -0.06 | -27.01 | -0.18 | 0 | -0.06 | 28.54 | 0.06 | 28.48 | 0 | 28.48 | -1.93 | 28.48 | -2.75 | 28.65 | -4.04
4.416 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | -0.35 | 21.45 | -1.52 | -0.06 | -26.49 | 0 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 0.12 | -7.85 | 0 | -21.97 | 0 | -25.37 | -0.06 | -0.7 | 0 | 0.23 | 0.12 | 28.54 | 0 | 28.59 | -0.94 | 28.83 | -1.76 | 28.42 | -2.7
4.448 | 0 | -0.18 | 21.5 | -0.23 | 21.5 | -0.23 | 21.45 | -0.76 | -0.06 | -23.91 | 0 | 0 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 2.99 | -5.92 | 0 | -20.45 | 0 | -23.2 | -0.06 | -68.38 | 0 | 0 | 0.12 | 28.54 | 0.12 | 28.48 | -0.7 | 28.42 | -1 | 28.65 | -1.82
4.48 | 0 | 0 | 21.39 | -0.23 | 21.45 | -0.23 | 21.45 | -0.53 | -0.18 | -22.09 | -0.06 | -29.36 | 0.06 | 0 | 0.06 | 21.39 | 0 | 10.9 | -4.8 | 0 | -17.99 | 0.06 | -21.27 | -0.06 | -61.58 | -0.06 | 0 | 0.06 | 28.65 | 0.35 | 28.65 | -0.35 | 28.59 | -0.64 | 28.42 | -0.94
4.512 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.29 | 21.39 | 0.12 | -0.18 | -19.63 | 0 | -65.1 | 0 | 0 | 0.06 | 21.39 | 0.18 | 16.41 | -3.52 | 0 | -15.94 | -0.06 | -18.4 | -0.18 | -45.18 | 0 | 0 | 0.06 | 28.54 | 0.35 | 28.65 | -0.35 | 28.59 | -0.47 | 28.65 | -0.7
4.544 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | -0.23 | 21.5 | 0.06 | 0 | -17.58 | 0 | -45.24 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 21.5 | -2.11 | 0 | -14.77 | -0.06 | -17.29 | 0 | -45.76 | 0 | 0 | 0 | 28.3 | 0.35 | 28.65 | -0.18 | 28.59 | -0.29 | 28.65 | -0.53
4.576 | 0 | -0.06 | 21.45 | -0.29 | 21.39 | -0.18 | 21.39 | 0.23 | 0 | -16.17 | -0.06 | -40.61 | -0.06 | -39.85 | 0.06 | 21.33 | 0.23 | 21.39 | -1.11 | -0.06 | -13.59 | -0.06 | -15.18 | 0 | -42.31 | -0.06 | 0 | 0 | 28.36 | 0.35 | 28.83 | -0.06 | 28.54 | -0.18 | 28.48 | -0.23
4.608 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.29 | 21.45 | 0.29 | 0 | -14.24 | 0 | -34.69 | -0.06 | -66.21 | 0.06 | 21.45 | 0.23 | 21.33 | -0.7 | -0.06 | -11.72 | -0.06 | -14.06 | 0 | -41.08 | -0.06 | 0 | 0 | 28.59 | 0.18 | 28.48 | 0 | 28.48 | -0.06 | 28.48 | -0.18
4.64 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | 0.47 | 0.06 | -12.07 | -0.06 | -31.82 | -0.06 | -42.83 | -0.06 | 21.39 | 0.29 | 21.45 | -0.47 | 0 | -10.08 | 5.04 | -12.36 | -0.06 | -39.55 | -0.06 | 0 | -0.06 | 28.59 | 0.29 | 28.59 | 0.12 | 28.59 | 0.06 | 28.65 | 0
4.672 | 0 | 0 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | 0.47 | 0 | -9.32 | 0 | -32.4 | -0.06 | -27.6 | -0.18 | 21.33 | 0.18 | 21.27 | -0.23 | -0.06 | -8.56 | 13.18 | -10.66 | -0.06 | -37.38 | -0.06 | 0 | -0.06 | 28.59 | 0.29 | 28.36 | 0.12 | 28.42 | 0.06 | 28.48 | 0
4.704 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.29 | 21.39 | 0.47 | 0 | -7.15 | -0.18 | -32.17 | -0.06 | -30.29 | -0.06 | 21.45 | 0.12 | 21.33 | -0.06 | 6.74 | -7.32 | 20.98 | -9.14 | -0.06 | -36.04 | -0.06 | 0 | 0.06 | 0.23 | 0.12 | 28.71 | 0.06 | 28.48 | 0.06 | 28.42 | 0
4.736 | 0 | -0.06 | 21.39 | -0.18 | 21.39 | -0.18 | 21.39 | 0.47 | 4.92 | -5.1 | 0.06 | -30.7 | -0.06 | -36.21 | -0.06 | 21.27 | 0.12 | 21.33 | -0.06 | 14.41 | -6.56 | 28.77 | -7.79 | -0.06 | -33.63 | -0.06 | 0 | 0.06 | 0.06 | 0.29 | 28.54 | 0.12 | 28.48 | 0.12 | 28.65 | 0.12
4.768 | 0 | -0.06 | 0.59 | -0.18 | 21.39 | -0.18 | 21.39 | 0.47 | 11.07 | -2.99 | 0.06 | -28.71 | -0.06 | -34.22 | -0.06 | 21.39 | 0.18 | 21.33 | 0.12 | 22.21 | -5.39 | 28.48 | -6.8 | -0.06 | -31.47 | -0.06 | 0 | 0 | 0 | 0.18 | 28.36 | 0.12 | 28.59 | 0.18 | 28.48 | 0.06
4.8 | -25.55 | -0.06 | 0.06 | -0.18 | 21.39 | -0.18 | 21.5 | 0.41 | 18.28 | -1.7 | 0 | -26.6 | -0.06 | -29.77 | -0.06 | 2.99 | 0.23 | 21.39 | 0.06 | 28.89 | -4.28 | 28.48 | -5.74 | 3.81 | -29.36 | 0 | 0 | -0.06 | 0.06 | 0.12 | 28.3 | 0.12 | 28.54 | 0 | 28.18 | 0.18
4.832 | -61.53 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 21.56 | -0.76 | 0.06 | -23.91 | 0 | -27.83 | -0.06 | 0.06 | 0.12 | 21.27 | 0.18 | 28.54 | -3.22 | 28.54 | -4.39 | 12.25 | -27.6 | -0.06 | -30.94 | 0 | 0 | 0.06 | 28.3 | 0.06 | 28.48 | 0.12 | 28.48 | 0.06
4.864 | -46.23 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | 0.35 | 21.39 | -0.47 | -0.06 | -21.91 | -0.06 | -27.13 | -0.06 | 0 | 0.23 | 21.15 | 0.23 | 28.54 | -2.46 | 28.89 | -3.34 | 19.98 | -25.55 | -0.06 | -85.96 | 0 | 0 | 0.12 | 28.42 | 0.06 | 28.59 | 0.06 | 28.54 | 0.12
4.896 | -39.79 | 0 | 0 | -0.06 | 21.39 | -0.18 | 21.39 | 0.35 | 21.39 | 0.06 | -0.06 | -20.22 | -0.06 | -25.9 | -0.06 | 0 | 0.18 | 21.21 | 0.23 | 28.42 | -1.93 | 28.65 | -2.75 | 29.01 | -23.09 | 0 | -50.39 | 0.06 | 0 | 0.06 | 28.42 | 0 | 28.42 | 0.18 | 28.3 | 0.18
4.928 | -38.44 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 21.45 | 0.29 | 21.39 | 0.12 | 0 | -18.16 | -0.06 | -23.67 | -0.06 | 0 | 0.23 | 21.33 | 0.23 | 28.54 | -1.7 | 28.54 | -2.17 | 28.59 | -21.68 | -0.06 | -47.35 | 0 | 0 | 0 | 28.54 | 0 | 28.83 | 0.06 | 28.48 | 0.23
4.96 | -35.92 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.45 | 0.23 | 0 | -16.17 | 0.06 | -22.38 | -0.18 | 0 | 0.12 | 21.27 | 0.18 | 28.42 | -1.17 | 28.77 | -1.82 | 28.42 | -18.93 | -0.06 | -41.9 | -0.06 | 0 | 0 | 28.59 | 0 | 28.36 | 0.12 | 28.42 | 0.06
4.992 | -34.1 | 0 | 0 | -0.06 | 0 | -0.18 | 0.23 | 0.23 | 21.39 | 0.35 | -0.06 | -13.54 | 0 | -20.51 | -0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.54 | -0.82 | 28.65 | -1.17 | 28.65 | -17.46 | -0.06 | -36.62 | 0 | 0 | 0 | 28.48 | 0.06 | 28.48 | 0.06 | 28.71 | 0.18
5.024 | -32.17 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0.06 | 0.18 | 21.39 | 0.23 | -0.06 | -11.25 | 0 | -18.63 | -0.06 | 0 | 0.18 | 21.33 | 0.23 | 29.18 | -0.76 | 28.83 | -0.76 | 28.3 | -15.94 | -0.06 | -37.15 | 0 | 0 | -0.06 | 0.23 | -0.06 | 28.48 | 0.12 | 28.48 | 0.18
5.056 | -29.24 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | -0.06 | -9.43 | 0 | -16.41 | -0.18 | 0 | 0.18 | 21.33 | 0.29 | 28.48 | -0.59 | 28.77 | -0.53 | 28.83 | -13.77 | 0 | -38.2 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 28.48 | 0.18 | 0.29 | 0.12
5.088 | -27.6 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | -0.06 | -7.09 | -0.06 | -14.82 | -0.06 | 0 | 0.18 | 21.27 | 0.29 | 28.54 | -0.53 | 28.59 | -0.35 | 28.42 | -12.36 | 0 | -36.92 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.29 | 0.06 | 0.12 | 0.06
5.12 | -25.08 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | 4.92 | -5.39 | -0.06 | -12.89 | 0.06 | 0 | 0.18 | 21.39 | 0.23 | 28.65 | -0.41 | 28.42 | -0.29 | 28.48 | -11.02 | 0.06 | -36.15 | -0.06 | -36.33 | 0 | 0.06 | -0.18 | 0.12 | 0.12 | 0 | 0.06
5.152 | -23.15 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.47 | 10.84 | -2.7 | -0.06 | -11.48 | 0 | 0 | 0.06 | 21.39 | 0.23 | 28.36 | -0.41 | 28.48 | -0.29 | 28.42 | -9.61 | 0 | -33.63 | -0.06 | -91.06 | -0.06 | 0 | -0.06 | 0 | 0 | 0.06 | 0
5.184 | -21.5 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.47 | 16.35 | -1.41 | 0.06 | -9.84 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 28.65 | -0.35 | 28.59 | -0.35 | 28.48 | -8.32 | 0 | -31.17 | -0.06 | -51.15 | -0.18 | 0 | -0.18 | 0.06 | 0.12 | 0 | 0.18
5.216 | -18.75 | -0.06 | -60.76 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.68 | -1.29 | -0.06 | -7.68 | 0.06 | 0 | 0.06 | 21.39 | 0.12 | 28.48 | -0.29 | 28.54 | -0.29 | 28.65 | -6.56 | 9.96 | -29.24 | -0.06 | -47.05 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.12
5.248 | -16.99 | -0.06 | -48.28 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.33 | -0.23 | 0 | -5.68 | 0 | -48.69 | 0.18 | 21.33 | 0.23 | 28.54 | -0.35 | 28.65 | -0.29 | 28.42 | -5.86 | 17.17 | -27.66 | -0.06 | -42.25 | 0 | 0 | -0.18 | 0 | 0.18 | 0 | 0.06
5.28 | -15.29 | -0.06 | -39.61 | -0.06 | 0 | -0.18 | 0 | 0 | 21.27 | 0.29 | 21.39 | -0.18 | -0.06 | -4.22 | 0.06 | -55.67 | 0.12 | 21.33 | 0.23 | 28.48 | -0.18 | 28.3 | -0.35 | 28.42 | -4.51 | 24.96 | -26.13 | 0 | -36.86 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06
5.312 | -13.71 | -0.06 | -39.55 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | 0.18 | 6.27 | -2.7 | -0.06 | -33.17 | 0.18 | 21.15 | 0.18 | 28.54 | -0.29 | 28.65 | -0.23 | 28.42 | -3.63 | 28.54 | -23.97 | -0.06 | -36.68 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06
5.344 | -12.31 | -0.06 | -37.56 | -0.06 | -20.04 | -0.06 | 0 | -0.06 | 21.45 | 0.29 | 21.39 | 0.23 | 13.89 | -1.93 | -0.06 | -27.54 | 0 | 0.12 | 0.06 | 28.48 | -0.23 | 28.95 | -0.18 | 28.13 | -2.99 | 28.71 | -22.44 | -0.06 | -38.44 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06
5.376 | -10.02 | 0 | -34.98 | -0.18 | -62.81 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.35 | 19.69 | -1.17 | -0.06 | -33.87 | 0 | 0.06 | 0.18 | 28.59 | -0.23 | 28.65 | -0.18 | 28.42 | -1.93 | 28.65 | -19.63 | 0 | -37.38 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0
5.408 | -8.61 | -0.06 | -32.64 | -0.18 | -46.64 | -0.18 | -16 | -0.18 | 21.39 | 0.12 | 21.45 | 0.29 | 21.45 | -0.59 | -0.06 | -35.98 | 0.12 | 0 | 0.18 | 28.48 | -0.18 | 28.48 | -0.29 | 28.65 | -1.64 | 28.42 | -17.81 | 0 | -35.16 | 0 | 0 | -0.23 | 0 | -0.06 | 0 | 0
5.44 | -7.79 | -0.06 | -30.94 | -0.06 | -39.38 | 0 | -63.34 | -0.06 | 21.39 | 0.12 | 21.39 | 0.35 | 21.39 | -0.53 | 0 | -32.52 | 0 | 0 | 0.12 | 28.48 | -0.23 | 28.89 | -0.18 | 28.65 | -1.05 | 28.48 | -15.82 | -0.06 | -33.28 | 0 | -30.24 | -0.18 | 0 | -0.18 | 0 | 0
5.472 | -5.63 | 0 | -28.36 | -0.18 | -38.15 | -0.06 | -47.17 | -0.06 | 21.39 | 0.18 | 21.39 | 0.47 | 21.39 | -0.06 | -0.06 | -29.06 | -0.06 | 0 | 0.23 | 28.65 | -0.23 | 28.59 | -0.18 | 28.65 | -0.7 | 28.59 | -14.18 | 0 | -31.58 | -0.06 | -76.76 | -0.18 | 0 | -0.18 | -14.12 | 0
5.504 | -4.39 | -0.06 | -26.72 | -0.06 | -35.68 | -0.06 | -40.72 | 0 | 21.27 | 0.12 | 21.39 | 0.29 | 21.33 | -0.06 | 0 | -28.07 | 0.06 | 0 | 0.29 | 28.77 | -0.18 | 28.42 | -0.06 | 28.42 | -0.64 | 28.71 | -11.84 | -0.06 | -29.12 | 0 | -52.91 | -0.18 | -17.64 | -0.18 | -89.18 | 0.06
5.536 | -2.81 | 0 | -23.85 | -0.06 | -34.22 | -0.06 | -38.56 | -0.06 | 21.45 | 0.12 | 21.39 | 0.41 | 21.21 | 0.18 | 1.99 | -27.25 | 0.06 | 0.06 | 0.18 | 28.65 | -0.29 | 0.29 | -0.18 | 28.42 | -0.35 | 28.42 | -9.26 | 0 | -27.31 | -0.06 | -47.93 | -0.06 | -89.53 | -0.18 | -60.76 | 0
5.568 | -1.64 | 4.34 | -21.33 | -0.06 | -32.46 | -0.18 | -35.39 | 0 | 21.39 | 0.12 | 21.39 | 0.41 | 21.39 | 0.23 | 6.45 | -25.61 | 0 | 0 | 0.12 | 28.65 | -0.29 | 0.06 | -0.18 | 28.3 | -0.41 | 28.83 | -7.56 | 0 | -25.78 | 0 | -42.66 | -0.18 | -60.24 | -0.06 | -34.98 | -0.06
5.6 | -1.05 | 10.37 | -19.63 | -0.06 | -29.71 | -0.06 | -34.75 | -0.06 | 21.5 | 0.06 | 21.39 | 0.29 | 21.39 | 0.18 | 12.25 | -23.09 | -0.06 | 0 | 0.18 | 0.29 | -0.06 | 0 | -0.06 | 28.54 | -0.18 | 28.65 | -4.8 | -0.06 | -24.14 | -0.18 | -37.15 | -0.06 | -34.22 | -0.18 | -33.69 | 0
5.632 | -0.59 | 15.76 | -17.7 | -0.06 | -27.6 | -0.06 | -31.93 | -0.06 | 0.12 | 0.12 | 21.39 | 0.35 | 21.33 | 0.06 | 17.7 | -20.92 | 0.06 | 0 | 0.23 | 0.06 | -0.18 | 0 | -0.18 | 28.36 | -0.35 | 28.42 | -3.22 | -0.06 | -22.15 | -0.06 | -37.38 | -0.18 | -33.69 | -0.06 | -41.66 | 0.06
5.664 | -0.29 | 21.5 | -16.58 | -0.06 | -25.37 | -0.06 | -30.18 | -0.06 | 0 | 0.06 | 21.45 | 0.18 | 21.39 | 0.29 | 21.39 | -19.63 | 0 | 0 | 0.18 | 0 | -0.18 | 0 | 0 | 28.36 | -0.23 | 28.59 | -2.11 | 1.29 | -19.4 | -0.06 | -38.15 | -0.18 | -41.54 | -0.18 | -41.66 | -0.18
5.696 | -0.06 | 21.45 | -13.95 | -0.18 | -23.15 | 0 | -27.6 | 0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.09 | 0.23 | 21.39 | -18.16 | -0.06 | 0 | 0.29 | 0 | -0.06 | 0 | -0.18 | 0.23 | -0.29 | 28.54 | -0.7 | 11.13 | -17.87 | -0.06 | -37.21 | -0.06 | -41.66 | 0 | -37.56 | -0.18
5.728 | -0.18 | 21.45 | -12.72 | -0.06 | -22.15 | -0.18 | -26.31 | 0 | 0.06 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 21.39 | -16.76 | 0 | 0 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0.06 | -0.23 | 28.18 | -0.59 | 18.57 | -15.88 | -0.06 | -35.74 | -0.06 | -37.21 | -0.06 | -33.87 | 0
5.76 | -0.35 | 21.45 | -11.13 | -0.06 | -18.63 | -0.06 | -23.32 | 0 | 0 | 0 | 21.45 | 0.29 | 21.39 | 0.18 | 21.39 | -14.12 | -0.06 | -33.93 | 0.12 | 0 | -0.06 | 0 | -0.18 | 0 | -0.23 | 28.54 | -0.18 | 28.89 | -13.89 | -0.06 | -33.46 | -0.18 | -33.99 | -0.06 | -33.05 | -0.18
5.792 | -0.35 | 21.45 | -9.61 | 0.06 | -17.05 | -0.06 | -22.09 | -0.06 | 0 | -0.06 | 21.45 | 0.18 | 21.39 | 0.12 | 21.33 | -12.6 | -0.06 | -69.85 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 28.48 | -0.06 | 28.59 | -12.07 | -0.06 | -30.47 | -0.06 | -33.34 | -0.06 | -32.81 | -0.18
5.824 | -0.35 | 21.45 | -8.38 | -0.18 | -15.7 | -0.06 | -19.28 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.18 | 21.39 | -10.96 | 0 | -47.87 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.23 | 28.48 | 0.06 | 28.65 | -9.26 | -0.06 | -28.89 | -0.18 | -32.76 | -0.06 | -31.11 | 0
5.856 | -0.35 | 21.45 | -6.5 | -0.06 | -13.54 | -0.06 | -17.58 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.33 | 0.12 | 21.39 | -9.08 | 0 | -30.7 | 0.23 | 0 | -0.18 | 0 | 0 | 0 | -0.18 | 28.48 | 0.23 | 28.59 | -7.32 | -0.06 | -27.19 | 0 | -31 | -0.06 | -28.59 | -0.06
5.888 | -0.23 | 21.5 | -4.86 | -0.06 | -12.25 | 0 | -15.65 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | 0.23 | 21.39 | -7.44 | -0.06 | -29.3 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.59 | 0.35 | 28.42 | -4.86 | 0 | -26.13 | -0.06 | -28.95 | -0.06 | -27.95 | -0.18
5.92 | -0.18 | 21.45 | -3.63 | -0.06 | -10.37 | -0.06 | -14.24 | 0.06 | 0 | -0.06 | 20.98 | 0.06 | 21.45 | 0.18 | 21.33 | -4.92 | -0.18 | -35.51 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.23 | 28.48 | -3.16 | -0.06 | -23.97 | -0.18 | -27.54 | 0 | -25.9 | -0.06
5.952 | -0.23 | 21.39 | -2.64 | 0.12 | -8.85 | -0.06 | -12.6 | -0.06 | 0 | -0.06 | 21.39 | -0.06 | 21.27 | 0.06 | 21.39 | -3.69 | 0.06 | -34.51 | 0.18 | 0 | -0.06 | -19.63 | -0.18 | 0 | -0.23 | 28.83 | 0.29 | 28.65 | -2.05 | 0 | -21.45 | -0.06 | -25.61 | 0 | -24.2 | -0.06
5.984 | -0.35 | 21.45 | -1.52 | 7.44 | -7.62 | -0.06 | -11.31 | -0.06 | 0 | -0.06 | 21.39 | -0.06 | 21.39 | 0.06 | 21.39 | -2.7 | 0 | -31.47 | 0.12 | -0.06 | -0.06 | -78.81 | -0.06 | 0 | -0.06 | 28.54 | 0.35 | 28.42 | -1.29 | 0 | -19.63 | -0.18 | -24.43 | -0.06 | -23.15 | -0.06
6.016 | -0.23 | 21.45 | -0.76 | 13.18 | -6.21 | 5.74 | -9.2 | -0.06 | 0 | 0 | 0.12 | 0 | 21.33 | 0.12 | 21.45 | -1.41 | -0.06 | -28.54 | 0.06 | -34.69 | -0.06 | -53.97 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.42 | -0.53 | -0.06 | -17.64 | -0.18 | -23.03 | -0.06 | -21.5 | -0.06
6.048 | -0.18 | 21.39 | -0.35 | 18.69 | -4.45 | 11.66 | -7.56 | -0.06 | -30.82 | -0.06 | 0.06 | -0.06 | 21.33 | 0.06 | 21.39 | -1.29 | 0.06 | -27.54 | 0.12 | -71.72 | 0 | -45 | -0.06 | 0 | -0.18 | 28.59 | 0.29 | 28.54 | -0.18 | 0 | -15.94 | -0.06 | -20.98 | 0 | -19.63 | -0.06
6.08 | -0.23 | 21.39 | -0.06 | 21.74 | -2.99 | 17.17 | -6.39 | -0.06 | -64.75 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -0.7 | -0.06 | -26.6 | 0 | -50.51 | -0.06 | -44.77 | -0.06 | 0 | -0.06 | 25.78 | 0.41 | 28.65 | 0 | -0.06 | -14 | -0.18 | -19.63 | 0 | -17.52 | 0
6.112 | -0.23 | 21.45 | 0 | 21.45 | -1.82 | 21.74 | -4.92 | -0.18 | -44.59 | 0.06 | 0.06 | -0.06 | 21.45 | 0.06 | 21.39 | -0.29 | 2.64 | -24.14 | 0.06 | -44.88 | -0.06 | -41.08 | -0.18 | -14 | -0.06 | 0.18 | 0.35 | 28.65 | 0.06 | 2.99 | -11.48 | -0.18 | -17.29 | -0.06 | -15.94 | -0.06
6.144 | -0.23 | 21.39 | -0.23 | 21.45 | -1.05 | 21.5 | -3.75 | -0.06 | -40.2 | -0.06 | 0 | -0.06 | 21.33 | -0.06 | 21.33 | -0.18 | 8.85 | -22.33 | 0.12 | -44.24 | -0.06 | -40.9 | -0.06 | -78.87 | -0.23 | 0 | 0.29 | 28.59 | 0.29 | 11.13 | -9.43 | 1.99 | -15.76 | 0 | -14.59 | -0.06
6.176 | -0.23 | 21.45 | -0.23 | 21.45 | -0.53 | 21.5 | -2.23 | -0.06 | -34.98 | -0.06 | 0 | -0.18 | 13.65 | 0 | 21.33 | -0.18 | 15.94 | -20.74 | 0.06 | -41.43 | 0 | -37.97 | 0 | -49.69 | -0.06 | 0 | 0.35 | 28.54 | 0.23 | 18.81 | -7.32 | 10.72 | -14.3 | 3.87 | -12.77 | -0.18
6.208 | -0.23 | 21.39 | -0.41 | 21.5 | -0.35 | 21.5 | -1.17 | -0.18 | -31.99 | -0.18 | 0 | -0.06 | 0.12 | -0.06 | 21.27 | 0.23 | 21.68 | -19.63 | 0.06 | -39.55 | 0 | -36.74 | -0.06 | -46.23 | -0.06 | 0 | 0.23 | 28.59 | 0.29 | 26.43 | -5.57 | 18.16 | -12.72 | 11.95 | -11.07 | 3.69
6.24 | -0.18 | 21.45 | -0.47 | 21.5 | -0.18 | 21.45 | -0.59 | -0.18 | -32.76 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.27 | -0.06 | 21.45 | -17.05 | 0.06 | -37.74 | 0 | -34.75 | -0.06 | -45.06 | 0 | 0 | 0.06 | 28.48 | 0.29 | 28.48 | -3.34 | 25.9 | -11.13 | 19.86 | -8.85 | 11.89
6.272 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -0.23 | -0.18 | -32.23 | 0 | 0 | 0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | -15.53 | 0.06 | -35.92 | 0 | -32.58 | 0 | -41.9 | -0.18 | 0 | 0.29 | 28.48 | 0.35 | 28.42 | -1.88 | 28.48 | -8.91 | 27.66 | -7.56 | 19.57
6.304 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | 0 | -0.06 | -30.82 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.39 | -13.59 | 0 | -33.87 | -0.06 | -30.29 | 0.06 | -41.13 | -0.06 | 0 | 0.18 | 28.48 | 0.35 | 28.65 | -1 | 28.65 | -7.44 | 28.65 | -5.98 | 27.19
6.336 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | -0.06 | -28.48 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | -12.13 | 0 | -32.23 | 0 | -29.12 | 0 | -37.97 | -0.18 | 0 | 0.06 | 28.48 | 0.23 | 28.24 | -0.59 | 28.3 | -5.8 | 28.54 | -4.45 | 28.65
6.368 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.29 | -0.06 | -26.49 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 21.21 | -10.14 | 0.06 | -30.7 | 0.06 | -27.42 | 0 | -36.97 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.48 | -0.29 | 28.36 | -4.1 | 28.54 | -2.99 | 28.48
6.4 | -0.06 | 21.45 | -0.23 | 21.39 | -0.18 | 21.45 | -0.35 | 5.86 | -23.67 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -8.14 | 0.12 | -28.48 | -0.06 | -24.08 | 0 | -35.45 | 0 | 0 | 0 | 28.71 | 0.12 | 28.42 | -0.06 | 28.36 | -3.11 | 28.42 | -1.88 | 28.59
6.432 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.35 | 11.6 | -21.68 | 0 | -33.28 | -0.06 | 0 | 0 | 21.33 | 0.35 | 21.39 | -6.56 | 0 | -26.31 | 0 | -22.33 | 0.06 | -31.99 | 0 | 0 | 0 | 28.3 | 0.12 | 28.83 | 0.18 | 28.65 | -1.88 | 28.54 | -1.17 | 28.83
6.464 | -0.06 | 0.23 | -0.23 | 21.45 | -0.29 | 21.45 | -0.35 | 17.11 | -20.22 | -0.18 | -74.48 | -0.06 | 0 | -0.18 | 0.18 | 0.35 | 21.33 | -4.8 | 3.69 | -23.03 | 0.06 | -20.45 | -0.06 | -31.88 | -0.18 | 0 | 0.06 | 28.3 | 0.06 | 28.48 | 0.18 | 28.95 | -1.17 | 28.54 | -0.7 | 28.54
6.496 | -0.06 | 0 | -0.23 | 21.39 | -0.23 | 21.45 | -0.23 | 21.45 | -18.22 | 0 | -49.57 | 0 | 0 | 0 | 0 | 0.35 | 21.39 | -3.28 | 11.48 | -21.97 | 0.12 | -18.63 | 0 | -28.48 | -0.06 | 0.06 | 0.06 | 28.54 | 0.18 | 28.42 | 0.29 | 28.48 | -0.76 | 28.42 | -0.53 | 28.42
6.528 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | -16.29 | -0.06 | -42.78 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.33 | -2.17 | 17.17 | -19.63 | 0 | -16.82 | -0.06 | -26.25 | 0 | -73.01 | 0 | 28.59 | 0 | 28.48 | 0.35 | 28.48 | -0.41 | 28.65 | -0.23 | 28.3
6.56 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -13.48 | -0.06 | -34.81 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -1.23 | 21.56 | -17.93 | -0.06 | -14.82 | -0.06 | -24.38 | 0 | -70.78 | 0 | 0.29 | 0 | 28.65 | 0.35 | 28.71 | -0.35 | 28.65 | -0.18 | 28.71
6.592 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.39 | -11.25 | 0 | -31.99 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.39 | -0.94 | 21.39 | -16.47 | -0.06 | -13.36 | 0.06 | -22.38 | -0.06 | -50.16 | 0 | 0.12 | 0.06 | 28.59 | 0.23 | 28.65 | 0 | 28.24 | -0.06 | 28.71
6.624 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -9.61 | -0.06 | -32.64 | -0.06 | -71.08 | -0.06 | 0 | 0.23 | 21.39 | -0.53 | 21.56 | -14.36 | 3.93 | -12.07 | 0.06 | -20.63 | -0.06 | -47.11 | 0 | 0 | 0 | 28.48 | 0.29 | 28.59 | -0.06 | 28.59 | 0.12 | 28.42
6.656 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.29 | 21.39 | -7.27 | -0.06 | -32.05 | 0.06 | -59.77 | 0 | 0 | 0.29 | 21.15 | -0.23 | 21.39 | -12.72 | 12.13 | -10.37 | 2.29 | -17.93 | -0.06 | -40.31 | -0.06 | 0.06 | 0 | 28.89 | 0.29 | 28.42 | 0.18 | 28.59 | 0.12 | 28.54
6.688 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.33 | -5.1 | 0 | -31.29 | -0.06 | -43.71 | 0 | 0 | 0.29 | 21.39 | -0.29 | 21.27 | -11.95 | 19.92 | -9.26 | 9.96 | -16.41 | 0 | -36.68 | 0 | 0 | 0 | 28.42 | 0.35 | 28.48 | 0.12 | 28.65 | 0.29 | 28.42
6.72 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.35 | 21.21 | -2.46 | 0.06 | -28.3 | 0 | -39.08 | 0 | 0 | 0.23 | 21.39 | 0.06 | 21.45 | -10.25 | 29.18 | -7.56 | 17.11 | -15 | 0 | -37.5 | -0.06 | 0 | 0 | 28.65 | 0.23 | 28.54 | 0.18 | 28.42 | 0.18 | 28.54
6.752 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | -0.29 | 21.39 | -1.82 | -0.06 | -26.6 | -0.06 | -33.63 | -0.18 | 0 | 0.23 | 21.21 | 0.06 | 21.45 | -8.91 | 28.59 | -5.92 | 27.36 | -13.59 | -0.06 | -37.97 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.48 | 0.23 | 28.36 | 0.23 | 28.95
6.784 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.45 | -0.29 | 21.33 | -0.64 | 4.92 | -24.61 | -0.06 | -32.76 | 0 | 0 | 0.23 | 21.39 | 0.12 | 21.39 | -7.44 | 28.42 | -5.33 | 28.71 | -12.13 | 0 | -36.86 | 0 | 0 | -0.06 | 28.48 | 0.29 | 28.65 | 0.18 | 29.06 | 0.18 | 28.42
6.816 | -0.06 | 0 | -0.18 | 21.39 | -0.06 | 21.39 | -0.23 | 21.39 | -0.53 | 11.07 | -22.74 | -0.06 | -32.76 | -0.18 | 0 | 0.23 | 21.15 | 0.18 | 21.33 | -6.15 | 28.65 | -4.16 | 28.59 | -10.31 | -0.06 | -34.16 | 0 | 0 | 0 | 28.59 | 0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.18
6.848 | -0.06 | 0 | -0.18 | 19.51 | -0.18 | 21.45 | -0.23 | 21.39 | -0.06 | 16.41 | -21.09 | -0.06 | -31.35 | 0 | 0 | 0.18 | 21.33 | 0.18 | 21.33 | -5.1 | 28.77 | -3.57 | 28.48 | -9.02 | 0 | -32.76 | 0 | 0 | 0 | 28.48 | 0 | 28.59 | 0.12 | 28.59 | 0.18 | 28.42
6.88 | -0.06 | -19.04 | -0.06 | 0.12 | -0.06 | 21.39 | -0.18 | 21.33 | 0.06 | 21.68 | -18.16 | 0.06 | -30.24 | -0.06 | -45.12 | 0.12 | 21.39 | 0 | 21.39 | -4.34 | 28.48 | -2.46 | 28.48 | -7.32 | 0 | -30.18 | 0 | 0 | 0.06 | 28.59 | 0.06 | 28.48 | 0.12 | 28.24 | 0.12 | 28.48
6.912 | -0.18 | -64.22 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.39 | 0.23 | 21.39 | -16.06 | 0.06 | -27.95 | -0.06 | -65.74 | 0.18 | 21.39 | 0.18 | 21.33 | -3.69 | 28.77 | -2.17 | 28.83 | -6.68 | 0 | -28.48 | -0.06 | 0 | -0.06 | 28.3 | 0.06 | 28.54 | 0.18 | 28.65 | 0.12 | 28.54
6.944 | -0.06 | -47.76 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | -14.24 | -0.18 | -25.61 | -0.06 | -41.6 | 0.12 | 21.39 | 0.23 | 21.33 | -2.46 | 28.77 | -1.46 | 28.36 | -5.22 | -0.06 | -27.31 | -0.18 | 0 | -0.06 | 28.42 | 0 | 28.3 | 0.12 | 28.59 | 0.23 | 28.42
6.976 | -0.06 | -39.96 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.45 | -12.13 | 0 | -23.32 | 0 | -27.31 | 0.12 | 21.21 | 0.23 | 21.27 | -2.05 | 28.59 | -1.23 | 28.65 | -3.98 | -0.06 | -24.61 | -0.06 | -3.81 | -0.06 | 28.54 | 0.06 | 28.42 | 0.06 | 28.48 | 0.29 | 28.59
7.008 | -0.06 | -38.79 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.47 | 21.39 | -10.14 | -0.06 | -21.21 | -0.06 | -32.34 | 0.12 | 0.23 | 0.23 | 21.33 | -1.41 | 28.54 | -1.05 | 28.54 | -3.16 | 0 | -23.09 | -0.06 | -89.36 | 0 | 14 | -0.06 | 28.48 | 0 | 28.48 | 0.23 | 28.48
7.04 | -0.06 | -35.68 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.41 | 21.39 | -7.21 | 0 | -19.34 | -0.18 | -35.74 | 0.18 | 0.06 | 0.18 | 21.39 | -1.29 | 28.24 | -0.59 | 28.42 | -2.46 | -0.06 | -21.09 | 0.06 | -55.9 | -0.06 | 0.12 | -0.06 | 17.93 | 0 | 28.42 | 0.18 | 28.48
7.072 | -0.06 | -34.22 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -5.22 | 0 | -17.29 | -0.18 | -33.52 | 0.06 | 0 | 0.23 | 21.33 | -0.94 | 28.65 | -0.53 | 28.42 | -1.88 | 0 | -19.45 | -0.06 | -48.75 | 0 | 0 | -0.06 | 0.18 | -0.06 | 0.64 | 0.12 | 28.48
7.104 | -0.06 | -31.99 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -3.28 | 0 | -14.77 | -0.06 | -30.7 | 0 | 0 | 0.23 | 21.39 | -0.76 | 29.01 | -0.47 | 28.24 | -1.46 | 1.99 | -17.4 | 0 | -43.54 | 0 | 0 | -0.06 | 0 | 0.06 | 0.12 | 0.06 | 1.41
7.136 | -0.18 | -29.65 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -1.88 | 0.06 | -12.95 | 0 | -27.83 | 0.06 | 0 | 0.12 | 21.21 | -0.64 | 28.36 | -0.41 | 28.54 | -0.94 | 10.2 | -14.94 | -0.06 | -38.38 | 0 | 0 | -0.18 | 0.06 | 0.12 | 0 | 0.12 | 0.06
7.168 | -0.06 | -27.54 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.23 | 4.22 | -10.72 | -0.06 | -27.48 | 0 | 0 | 0.23 | 21.15 | -0.53 | 28.71 | -0.35 | 28.89 | -0.59 | 17.46 | -13.3 | -0.06 | -36.62 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0.06 | 0 | 0
7.2 | -0.06 | -25.67 | 0 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.35 | 21.39 | -0.59 | 10.25 | -8.91 | -0.06 | -25.61 | 0 | 0 | 0.18 | 21.39 | -0.53 | 28.42 | -0.23 | 28.48 | -0.53 | 25.31 | -10.9 | -0.06 | -38.2 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0.06
7.232 | 0 | -23.56 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.15 | 0.29 | 21.39 | -0.23 | 17.52 | -6.62 | 0.12 | -23.67 | 0 | 0 | 0.06 | 21.39 | -0.53 | 28.18 | -0.35 | 28.42 | -0.41 | 28.65 | -9.08 | 0 | -37.38 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
7.264 | -0.06 | -22.21 | -0.06 | -0.06 | 0 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.18 | 21.74 | -3.98 | 3.63 | -21.91 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.48 | -0.35 | 28.65 | -0.35 | 28.77 | -6.68 | -0.06 | -35.92 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0
7.296 | -0.06 | -18.99 | 0 | -42.01 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 21.45 | 0.06 | 21.39 | -2.17 | 11.48 | -19.51 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.59 | -0.23 | 28.89 | -0.23 | 28.42 | -4.1 | -0.06 | -33.4 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 0.06
7.328 | -0.06 | -17.23 | -0.06 | -54.49 | -0.06 | -31.41 | -0.18 | 0 | 0.12 | 21.39 | 0.18 | 21.39 | -1.46 | 17.11 | -18.16 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.65 | -0.23 | 28.54 | -0.29 | 28.65 | -2.52 | 0 | -32.17 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0
7.36 | -0.18 | -15.35 | 0 | -41.72 | -0.18 | -61.41 | -0.18 | 0 | 0.23 | 21.39 | 0.41 | 21.39 | -0.82 | 21.68 | -16.17 | 0 | 0 | 0.06 | 0.18 | -0.35 | 28.59 | -0.29 | 28.54 | -0.23 | 28.71 | -1.46 | 0 | -30 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | -0.18 | 0
7.392 | -0.06 | -14 | -0.06 | -39.79 | -0.06 | -43.95 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.33 | -0.06 | 21.45 | -14.82 | 0.06 | 0 | 0.23 | 0 | -0.35 | 28.36 | -0.29 | 28.42 | -0.29 | 28.36 | -1 | -0.06 | -28.54 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0
7.424 | -0.06 | -12.42 | -0.06 | -37.38 | 0 | -40.08 | -0.06 | 0 | -0.06 | 21.45 | 0.29 | 21.39 | 0.18 | 21.27 | -13.18 | -0.06 | -2.93 | 0.12 | 0 | -0.29 | 28.54 | -0.23 | 28.48 | -0.23 | 28.59 | -0.53 | -0.06 | -26.08 | 0 | -3.11 | 0 | 0 | 0 | 0 | -0.06 | 0
7.456 | 0 | -10.08 | -0.06 | -34.81 | -0.06 | -38.32 | -0.18 | 0 | 0 | 21.27 | 0.41 | 21.39 | 0.23 | 21.33 | -10.84 | -0.06 | -71.84 | 0.18 | 0 | -0.29 | 28.48 | -0.23 | 28.42 | -0.41 | 28.54 | 0.06 | -0.06 | -24.55 | 0 | -88.42 | 0 | -0.41 | 0 | 0 | 0 | 0
7.488 | 4.16 | -8.91 | 0 | -32.81 | -0.06 | -35.51 | -0.06 | 0 | 0 | 21.5 | 0.47 | 21.39 | 0.35 | 21.39 | -9.49 | 0 | -49.75 | 0.23 | 0 | -0.23 | 28.48 | -0.18 | 28.65 | -0.18 | 28.59 | -0.06 | 0 | -23.56 | -0.18 | -63.81 | -0.06 | -86.08 | -0.18 | -10.61 | -0.18 | 0
7.52 | 10.08 | -7.44 | -0.06 | -31.17 | 0 | -33.69 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | 0.35 | 21.33 | -7.56 | 0 | -30.41 | 0.23 | 0 | -0.18 | 0.29 | -0.18 | 28.65 | -0.06 | 28.54 | 0.29 | 4.34 | -20.33 | 0 | -48.17 | -0.18 | -69.32 | -0.18 | -89.89 | 0 | -15.23
7.552 | 15.65 | -6.21 | 0 | -28.71 | -0.06 | -31.47 | -0.06 | 0 | 0.12 | 21.5 | 0.41 | 21.39 | 0.41 | 21.39 | -6.15 | 0 | -28.77 | 0.18 | 0.06 | -0.18 | 0.06 | -0.18 | 0.64 | -0.18 | 28.59 | 0.29 | 12.72 | -18.46 | 0 | -44.24 | -0.06 | -38.85 | -0.06 | -62.81 | 0 | -88.71
7.584 | 21.33 | -4.75 | 0 | -26.6 | -0.06 | -30.18 | -0.06 | 0 | 0.06 | 21.56 | 0.35 | 21.39 | 0.47 | 21.39 | -4.1 | 0 | -34.75 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.12 | -0.23 | 28.65 | 0.41 | 20.22 | -16.47 | 0 | -37.79 | -0.06 | -27.83 | 0 | -35.68 | -0.06 | -61.47
7.616 | 21.45 | -3.16 | 0 | -24.55 | -0.06 | -27.13 | -0.18 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.53 | 21.39 | -2.46 | 0 | -34.81 | 0.12 | 0 | -0.23 | 0.06 | -0.18 | 0 | -0.23 | 28.83 | 0.29 | 28.95 | -14.53 | 0 | -36.45 | -0.18 | -41.08 | -0.06 | -32.93 | -0.06 | -31.7
7.648 | 21.45 | -2.05 | -0.06 | -22.21 | -0.06 | -25.37 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.39 | 0.53 | 21.27 | -1.7 | 0 | -31.29 | 0.06 | 0 | -0.23 | 0 | -0.18 | 0.06 | -0.18 | 28.59 | 0.29 | 28.54 | -12.77 | -0.06 | -38.15 | -0.18 | -42.6 | -0.06 | -41.84 | -0.06 | -32.17
7.68 | 21.45 | -1.05 | -0.06 | -20.57 | -0.06 | -23.5 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | 0.53 | 21.39 | -0.94 | -0.06 | -28.65 | 0.18 | 0 | -0.23 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.23 | 28.48 | -10.14 | -0.18 | -37.68 | -0.06 | -37.91 | 0 | -41.78 | 0 | -41.84
7.712 | 21.5 | -0.53 | -0.18 | -17.81 | 0 | -20.98 | -0.06 | -40.84 | 0 | 0 | 0.12 | 21.33 | 0.53 | 21.27 | -0.7 | -0.06 | -27.6 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.54 | -8.03 | -0.18 | -36.39 | -0.18 | -34.45 | -0.18 | -37.62 | 0 | -41.54
7.744 | 21.5 | -0.35 | 0 | -16.29 | 0 | -18.99 | 0 | -69.91 | 0 | 0 | 0.12 | 21.39 | 0.47 | 21.39 | -0.18 | 0.06 | -26.31 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.42 | -5.63 | -0.06 | -34.04 | -0.18 | -32.76 | 0 | -34.4 | 0 | -37.33
7.776 | 21.45 | -0.23 | -0.06 | -14.41 | -0.06 | -16.93 | 0 | -44.88 | 0.06 | 0.06 | 0.23 | 21.39 | 0.23 | 21.39 | -0.06 | 0.06 | -23.91 | 0.12 | -39.38 | -0.18 | 0 | -0.06 | 0 | -0.23 | 28.42 | 0.35 | 28.65 | -3.87 | -0.18 | -31.99 | 0 | -33.34 | -0.18 | -32.7 | -0.18 | -32.99
7.808 | 21.5 | -0.18 | 0 | -13.01 | -0.06 | -15.23 | 0 | -39.67 | -0.18 | 0 | 0.18 | 21.04 | 0.23 | 21.39 | -0.06 | 0 | -22.33 | 0.18 | -67.97 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.71 | -2.4 | 3.81 | -30.35 | 0 | -32.23 | -0.06 | -32.46 | 0 | -32.76
7.84 | 21.45 | -0.18 | -0.06 | -11.66 | 0 | -13.59 | -0.18 | -34.4 | -0.06 | 0 | 0.12 | 21.15 | 0.29 | 21.33 | 0.06 | 0 | -20.98 | 0.06 | -44.06 | -0.18 | 0 | -0.06 | 0 | -0.23 | 28.48 | 0.23 | 28.59 | -1 | 12.13 | -27.95 | 0 | -29.59 | -0.06 | -31.11 | -0.06 | -32.64
7.872 | 21.45 | -0.53 | 1.99 | -9.79 | 0 | -11.84 | -0.06 | -31.88 | -0.06 | 0 | 0.12 | 21.27 | 0.18 | 21.27 | 0.18 | 0.06 | -19.51 | 0.18 | -29.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.18 | 28.42 | -0.7 | 19.92 | -26.13 | 0 | -27.48 | 0 | -29.59 | -0.06 | -30.94
7.904 | 21.45 | -0.35 | 6.62 | -8.38 | 0 | -10.37 | -0.06 | -32.58 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | 0.23 | 1.99 | -17.93 | 0.06 | -30.12 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.18 | 28.42 | -0.47 | 29.01 | -24.2 | -0.06 | -26.43 | 0 | -26.9 | 0 | -28.48
7.936 | 21.39 | -0.35 | 14 | -7.03 | -0.06 | -8.79 | -0.18 | -31.82 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.33 | 0.29 | 6.27 | -15.59 | 0.12 | -36.1 | -0.18 | -13.48 | 0 | 0 | -0.12 | 28.65 | 0.12 | 28.42 | -0.06 | 28.65 | -23.03 | 0 | -24.79 | 0 | -25.78 | -0.06 | -27.19
7.968 | 21.45 | -0.35 | 19.98 | -5.33 | -0.06 | -7.21 | -0.18 | -30.59 | -0.18 | 0 | 0 | 21.33 | 0.12 | 21.39 | 0.29 | 11.78 | -14 | 0.12 | -34.51 | -0.06 | -77.7 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.54 | -0.06 | 28.42 | -21.09 | 0 | -23.32 | 0 | -24.43 | -0.06 | -25.61
8 | 21.45 | -0.35 | 21.5 | -3.98 | -0.06 | -5.98 | -0.06 | -28.59 | -0.18 | 0 | 0 | 21.33 | 0.06 | 21.39 | 0.29 | 19.4 | -11.84 | 0.12 | -30.94 | -0.18 | -49.75 | -0.06 | -56.84 | -0.06 | 23.79 | 0.12 | 28.48 | 0.12 | 28.65 | -18.63 | -0.06 | -21.74 | 0 | -23.15 | 0 | -24.26
8.032 | 21.45 | -0.23 | 21.45 | -2.7 | 3.75 | -4.45 | -0.18 | -26.02 | -0.06 | 0 | -0.06 | 21.27 | 0.12 | 21.39 | 0.23 | 21.56 | -10.49 | 0.06 | -28.36 | -0.18 | -45.53 | -0.18 | -57.37 | -0.06 | 0.18 | -0.06 | 28.65 | 0.18 | 28.65 | -16.76 | -0.06 | -19.92 | -0.06 | -21.09 | 0 | -23.15
8.064 | 21.45 | -0.23 | 21.5 | -1.64 | 9.96 | -3.28 | -0.18 | -24.14 | -0.06 | 0 | -0.06 | 0.12 | 0.06 | 21.21 | 0.29 | 21.27 | -8.56 | 0.12 | -27.42 | -0.18 | -44.06 | -0.06 | -45.35 | -0.06 | 0 | 0.12 | 28.54 | 0.18 | 28.54 | -14.36 | -0.06 | -18.4 | -0.18 | -19.69 | 0 | -21.39
8.096 | 21.45 | -0.18 | 21.5 | -0.94 | 15.23 | -1.7 | 3.75 | -21.62 | -0.06 | -35.22 | 0 | 0 | 0 | 21.21 | 0.29 | 21.39 | -6.45 | 0.12 | -25.72 | -0.06 | -41.19 | -0.18 | -45.47 | -0.06 | 0 | 0.06 | 28.65 | 0.23 | 28.59 | -12.95 | -0.06 | -16.29 | 0 | -17.58 | 0 | -19.45
8.128 | 21.45 | -0.18 | 21.45 | -0.59 | 20.92 | -0.94 | 10.08 | -19.51 | -0.18 | -72.48 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | -4.51 | 0.12 | -23.5 | -0.18 | -40.67 | 0 | -43.48 | -0.06 | 0 | 0 | 28.36 | 0.35 | 28.54 | -10.55 | 4.45 | -15.23 | -0.06 | -15.94 | 0 | -17.81
8.16 | 21.45 | -0.23 | 21.5 | -0.29 | 21.5 | -0.41 | 15.18 | -17.7 | 0 | -50.51 | -0.06 | 0.06 | -0.06 | 0.18 | 0.23 | 21.15 | -3.46 | 0 | -21.68 | -0.06 | -38.09 | -0.06 | -41.02 | -0.06 | 0 | 0.06 | 28.24 | 0.18 | 28.65 | -7.68 | 12.83 | -12.89 | 0 | -14.24 | 3.34 | -15.94
8.192 | 21.39 | -0.23 | 21.45 | -0.06 | 21.45 | -0.18 | 21.09 | -15.47 | 0 | -40.9 | -0.18 | 0 | 0 | 0.06 | 0.18 | 21.45 | -2.4 | 0.06 | -20.92 | -0.06 | -36.27 | -0.06 | -39.44 | -0.06 | 0 | 0 | 28.36 | 0.29 | 28.48 | -5.98 | 23.03 | -11.48 | 3.75 | -13.01 | 11.6 | -14.24
8.224 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -0.18 | 21.5 | -13.89 | 0.06 | -34.81 | 0 | 0 | 0 | 0 | 0.12 | 21.45 | -1.52 | 0 | -19.63 | -0.18 | -34.98 | -0.06 | -37.03 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.48 | -3.93 | 28.65 | -9.55 | 11.89 | -11.07 | 19.22 | -12.77
8.256 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.45 | -11.25 | -0.06 | -31.99 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.76 | -0.06 | -16.82 | -0.06 | -32.4 | -0.06 | -35.57 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.83 | -2.23 | 28.59 | -7.91 | 19.69 | -9.08 | 26.95 | -10.66
8.288 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | -0.23 | 21.45 | -8.79 | -0.06 | -32.58 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.45 | -0.64 | 0 | -15.06 | -0.06 | -30.94 | 0 | -33.69 | -0.06 | 0 | -0.18 | 28.59 | 0.18 | 28.54 | -1.7 | 28.42 | -5.8 | 27.31 | -7.68 | 28.65 | -9.26
8.32 | 21.39 | -0.18 | 21.39 | -0.35 | 21.39 | -0.41 | 21.5 | -6.97 | -0.06 | -32.46 | -0.06 | 0 | 0 | 0 | 0.18 | 21.33 | -0.29 | -0.06 | -13.48 | -0.06 | -28.54 | -0.06 | -31.99 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.65 | -0.47 | 28.48 | -4.45 | 28.71 | -5.8 | 28.48 | -7.56
8.352 | 21.39 | -0.18 | 21.45 | -0.29 | 21.45 | -0.47 | 21.45 | -4.63 | 0 | -30.94 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.06 | 0 | -12.07 | 0 | -26.84 | 0 | -29.06 | -0.18 | 0 | 0 | 28.65 | 0.12 | 28.54 | -0.35 | 28.65 | -3.4 | 28.48 | -4.16 | 28.48 | -5.8
8.384 | 0.23 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | 21.45 | -2.99 | -0.06 | -29.01 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | -0.06 | 0.06 | -10.25 | -0.06 | -24.38 | 0 | -26.84 | -0.06 | 0 | -0.18 | 28.48 | 0.06 | 28.54 | -0.18 | 28.36 | -2.29 | 28.54 | -2.99 | 28.42 | -4.16
8.416 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.39 | -1.7 | 0 | -26.49 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0 | -0.06 | -8.09 | -0.06 | -22.15 | 0 | -25.37 | -0.18 | -0.12 | 0 | 0.23 | 0.06 | 28.54 | 0.12 | 28.65 | -1.46 | 28.48 | -1.88 | 28.48 | -2.58
8.448 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.76 | -0.06 | -24.26 | 0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0 | 2.46 | -6.39 | -0.06 | -20.68 | -0.06 | -23.15 | 0 | -64.69 | 0 | 0.06 | 0.06 | 28.65 | 0.12 | 28.65 | -0.29 | 28.48 | -1.29 | 28.65 | -1.64
8.48 | 0.06 | -0.18 | 21.5 | -0.23 | 21.45 | -0.18 | 21.39 | -0.53 | -0.18 | -22.38 | -0.06 | -17.46 | -0.06 | 0 | 0.12 | 21.33 | 0.12 | 10.14 | -4.92 | -0.06 | -17.93 | 0 | -21.39 | 0 | -63.11 | 0 | 0 | -0.18 | 28.48 | 0.23 | 28.83 | -0.59 | 28.48 | -0.53 | 28.48 | -1
8.512 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.06 | -0.06 | -19.69 | 0.06 | -68.03 | -0.06 | 0 | 0 | 21.39 | 0.23 | 15.65 | -3.34 | -0.12 | -16.06 | -0.06 | -19.04 | 0 | -46.23 | 0 | 0.06 | -0.06 | 28.83 | 0.35 | 28.24 | -0.06 | 28.42 | -0.41 | 28.48 | -0.94
8.544 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.18 | 21.5 | 0 | -0.06 | -17.87 | -0.06 | -46.76 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.21 | -2.23 | -0.06 | -14.53 | -0.06 | -16.93 | -0.06 | -45.53 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.65 | -0.06 | 28.42 | -0.35 | 28.65 | -0.41
8.576 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | -0.29 | 21.45 | 0.18 | 0 | -15.94 | -0.06 | -41.19 | 0.06 | -35.63 | 0 | 21.39 | 0.23 | 21.27 | -1.46 | -0.06 | -13.48 | -0.06 | -15.76 | 0 | -42.95 | 0 | 0 | -0.06 | 28.89 | 0.35 | 28.59 | 0 | 28.59 | -0.18 | 28.59 | -0.23
8.608 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.29 | -0.06 | -14 | 0 | -35.16 | -0.06 | -67.15 | 0 | 21.33 | 0.35 | 21.33 | -0.7 | 0 | -12.07 | -0.06 | -14.36 | -0.06 | -41.43 | -0.06 | 0 | -0.06 | 28.36 | 0.41 | 28.42 | 0.06 | 28.65 | 0.12 | 28.42 | -0.23
8.64 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.41 | 0 | -11.6 | -0.06 | -32.17 | -0.06 | -44.83 | -0.06 | 21.33 | 0.12 | 21.39 | -0.35 | -0.06 | -9.9 | 4.28 | -12.36 | -0.06 | -39.44 | 0 | 0 | -0.18 | 28.42 | 0.41 | 28.36 | 0.18 | 28.42 | 0.18 | 28.65 | -0.06
8.672 | 0 | 0 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.29 | -0.06 | -9.2 | 0 | -32.76 | -0.06 | -28.89 | 0.06 | 21.39 | 0.18 | 21.33 | -0.35 | 0 | -8.91 | 12.48 | -10.55 | -0.06 | -37.79 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.54 | 0.23 | 28.24 | 0.23 | 28.65 | -0.06
8.704 | 0 | -0.06 | 21.5 | -0.23 | 21.39 | -0.29 | 21.45 | 0.47 | -0.06 | -7.44 | -0.06 | -32.52 | -0.06 | -29.77 | 0 | 21.33 | 0.18 | 21.39 | 0 | 6.27 | -7.79 | 20.39 | -9.32 | -0.06 | -36.15 | -0.06 | 0 | -0.06 | 0.18 | 0.12 | 28.42 | 0.23 | 28.71 | 0.06 | 28.59 | 0
8.736 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.35 | 4.45 | -5.22 | -0.06 | -30.76 | -0.06 | -35.74 | -0.06 | 21.39 | 0.29 | 21.39 | 0.06 | 14.18 | -6.39 | 27.89 | -8.56 | 0 | -33.17 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 28.54 | 0.23 | 28.54 | 0.18 | 28.65 | 0.06
8.768 | 0 | -0.06 | 10.78 | -0.18 | 21.39 | -0.18 | 21.45 | 0.47 | 10.55 | -3.22 | -0.06 | -28.89 | -0.06 | -34.45 | 0 | 21.15 | 0.29 | 21.27 | 0.06 | 22.03 | -5.63 | 28.48 | -6.86 | -0.06 | -31.64 | 0 | 0 | -0.18 | 0 | 0.18 | 28.42 | 0.12 | 28.42 | 0.12 | 28.42 | 0.18
8.8 | -12.77 | -0.06 | 0.12 | -0.18 | 21.45 | -0.18 | 21.45 | 0.41 | 17.75 | -1.82 | -0.06 | -26.31 | 0.06 | -30.7 | -0.06 | 13.89 | 0.23 | 21.45 | 0.06 | 28.83 | -4.22 | 28.65 | -5.45 | 3.34 | -29.77 | 0 | 0 | -0.18 | 0.06 | 0.06 | 28.54 | 0.12 | 28.48 | 0.23 | 28.65 | 0.06
8.832 | -64.16 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.29 | 21.68 | -0.76 | 0 | -24.32 | 0 | -27.95 | 0 | 0.06 | 0.23 | 21.27 | 0.18 | 28.54 | -3.4 | 28.65 | -4.22 | 11.48 | -27.6 | 0 | -12.13 | -0.06 | 0 | 0.06 | 28.54 | 0.12 | 28.13 | 0.12 | 28.48 | 0.12
8.864 | -48.28 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.45 | -0.53 | 0 | -22.09 | -0.06 | -27.25 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 28.59 | -2.75 | 28.3 | -3.75 | 19.22 | -25.78 | 0 | -92.82 | 0 | 0 | 0 | 28.48 | 0.06 | 28.3 | 0.18 | 28.48 | 0.18
8.896 | -39.61 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 21.39 | -0.18 | -0.06 | -20.74 | -0.06 | -26.13 | 0 | 0 | 0.18 | 21.39 | 0.29 | 28.48 | -1.99 | 28.48 | -2.93 | 29.06 | -23.32 | 0 | -52.56 | -0.18 | 0 | 0.06 | 28.48 | 0 | 28.42 | 0.18 | 28.48 | 0
8.928 | -38.2 | -0.06 | 0 | -0.06 | 0.23 | -0.06 | 21.5 | 0.29 | 21.39 | 0.23 | 0 | -18.22 | 0 | -23.67 | -0.06 | 0 | 0.23 | 21.33 | 0.18 | 28.18 | -1.64 | 28.65 | -2.23 | 28.59 | -21.68 | -0.06 | -49.4 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.42 | 0.18 | 28.48 | 0.18
8.96 | -35.68 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.5 | 0.29 | 21.39 | 0.23 | -0.06 | -16.41 | -0.06 | -22.09 | 0 | 0 | 0.12 | 21.45 | 0.29 | 28.42 | -1.29 | 28.71 | -1.76 | 28.36 | -19.45 | 0 | -43.13 | -0.18 | 0 | 0.12 | 28.77 | 0 | 27.83 | 0.18 | 28.36 | 0.06
8.992 | -34.22 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.23 | 0.18 | 21.39 | 0.29 | -0.06 | -13.59 | 0.06 | -21.09 | -0.18 | 0 | 0.18 | 21.39 | 0.12 | 28.48 | -0.94 | 28.3 | -1.05 | 28.42 | -17.29 | 0 | -38.15 | -0.06 | 0 | -0.06 | 28.48 | -0.06 | 28.48 | 0.06 | 28.48 | 0.12
9.024 | -32.52 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.06 | 0.12 | 21.39 | 0.41 | 0 | -11.48 | 0 | -19.22 | -0.18 | 0 | 0.18 | 21.33 | 0.12 | 28.3 | -0.59 | 28.71 | -0.94 | 28.59 | -16 | 0 | -37.33 | 0 | 0 | -0.06 | 0.23 | 0 | 28.48 | 0.18 | 28.48 | 0.06
9.056 | -29.77 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | 0 | -9.67 | 0 | -17.17 | -0.18 | 0 | 0.18 | 21.39 | 0.18 | 28.24 | -0.47 | 28.42 | -0.7 | 28.54 | -13.95 | 0 | -37.56 | -0.06 | 0 | -0.06 | 0.06 | -0.18 | 28.48 | 0.12 | 4.51 | 0.06
9.088 | -27.89 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.12 | 21.39 | 0.47 | -0.06 | -7.68 | 0 | -15.06 | -0.18 | 0 | 0.18 | 21.33 | 0.12 | 28.42 | -0.47 | 28.59 | -0.47 | 28.65 | -12.6 | -0.06 | -37.21 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.94 | 0.12 | 0.12 | 0.06
9.12 | -25.31 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.45 | 0.47 | 4.16 | -5.74 | 0 | -13.24 | -0.06 | 0 | 0.23 | 21.27 | 0.18 | 28.42 | -0.35 | 28.42 | -0.35 | 28.48 | -11.48 | -0.06 | -35.33 | -0.06 | -16.41 | 0 | 0.06 | -0.18 | 0.12 | 0.18 | 0 | 0.06
9.152 | -23.61 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.45 | 0.41 | 10.31 | -2.64 | -0.06 | -11.89 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 28.65 | -0.29 | 28.48 | -0.35 | 28.48 | -9.9 | -0.06 | -33.87 | -0.06 | -92.7 | 0 | 0 | -0.06 | 0 | 0.18 | 0.06 | 0
9.184 | -21.68 | -0.06 | -0.06 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.53 | 15.53 | -1.93 | -0.06 | -9.79 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.42 | -0.23 | 28.48 | -0.23 | 28.65 | -8.38 | -0.06 | -31.47 | -0.06 | -53.32 | 0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0.18
9.216 | -19.04 | -0.06 | -54.67 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.45 | -0.94 | 0 | -7.62 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 28.59 | -0.29 | 28.48 | -0.23 | 28.59 | -6.97 | 9.14 | -29.36 | 0 | -48.52 | 0 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06
9.248 | -17.17 | -0.06 | -51.27 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.27 | 0.35 | 21.45 | -0.53 | -0.06 | -5.74 | -0.06 | -40.37 | 0.12 | 21.39 | 0.23 | 28.3 | -0.29 | 28.3 | -0.29 | 28.48 | -5.86 | 16.11 | -27.77 | 0 | -43.65 | 0 | 0 | 0 | 0 | 0 | 0 | 0.12
9.28 | -15.65 | 0.06 | -40.37 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -0.23 | 0.06 | -4.39 | -0.06 | -59.24 | 0.12 | 21.33 | 0.23 | 28.36 | -0.35 | 28.48 | -0.29 | 28.42 | -4.75 | 23.85 | -25.72 | 0 | -37.62 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0
9.312 | -13.77 | 0 | -39.55 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.18 | 21.45 | 0.35 | 5.39 | -2.7 | -0.18 | -35.22 | -0.06 | 21.21 | 0.29 | 28.54 | -0.23 | 28.54 | -0.18 | 28.65 | -3.98 | 28.65 | -23.91 | -0.06 | -36.8 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0
9.344 | -12.83 | -0.06 | -37.68 | -0.06 | -8.09 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.27 | 0.23 | 13.24 | -2.05 | -0.06 | -27.77 | 0.18 | 0.23 | 0.18 | 28.36 | -0.23 | 28.54 | -0.23 | 28.3 | -2.87 | 28.59 | -22.21 | -0.06 | -38.79 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0
9.376 | -10.55 | 0 | -35.16 | -0.06 | -63.52 | -0.06 | 0 | 0 | 21.5 | 0.12 | 21.39 | 0.53 | 18.69 | -0.94 | -0.06 | -33.05 | 0.06 | 0.06 | 0.18 | 28.48 | -0.18 | 28.48 | -0.23 | 28.71 | -2.05 | 28.65 | -19.63 | -0.06 | -37.33 | 0 | 0 | -0.18 | 0 | 0 | 0 | -0.06
9.408 | -9.2 | 0 | -32.64 | -0.06 | -48.87 | -0.06 | -12.66 | 0 | 21.39 | 0.12 | 21.39 | 0.47 | 21.5 | -0.7 | -0.06 | -35.98 | 0 | 0 | 0.12 | 28.54 | -0.23 | 28.65 | -0.18 | 28.65 | -1.41 | 28.65 | -17.87 | 0 | -35.33 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06
9.44 | -7.68 | -0.06 | -30.29 | -0.18 | -40.14 | 0 | -62.46 | 0 | 21.39 | 0.18 | 21.39 | 0.29 | 21.39 | -0.23 | 0 | -33.46 | 0.06 | 0.06 | 0.12 | 28.48 | -0.23 | 28.36 | -0.23 | 28.65 | -1 | 28.89 | -15.94 | 0 | -33.69 | 0 | -24.02 | 0 | 0 | 0 | 0 | 0.12
9.472 | -6.39 | 0 | -28.36 | -0.06 | -38.67 | -0.06 | -48.4 | -0.06 | 21.39 | 0.06 | 21.39 | 0.47 | 21.39 | -0.18 | -0.06 | -29.3 | 0.12 | 0 | 0.06 | 28.59 | -0.18 | 28.54 | -0.18 | 28.48 | -0.7 | 28.89 | -14.3 | -0.06 | -31.23 | 0 | -78.23 | -0.18 | 0 | 0 | -3.98 | 0
9.504 | -5.1 | 0.06 | -26.78 | 0 | -35.68 | -0.06 | -39.85 | -0.06 | 21.39 | 0.06 | 21.39 | 0.35 | 21.33 | -0.18 | -0.06 | -27.6 | -0.06 | 0 | 0.18 | 28.65 | -0.18 | 28.48 | -0.23 | 28.42 | -0.47 | 28.24 | -12.13 | 0 | -29.12 | 0 | -53.03 | 0 | -11.02 | 0 | -89.3 | 0
9.536 | -2.93 | 0 | -24.08 | -0.12 | -34.28 | -0.06 | -38.5 | -0.06 | 21.39 | 0 | 21.39 | 0.53 | 21.39 | -0.06 | -0.06 | -26.6 | 0.06 | 0 | 0.23 | 28.83 | -0.18 | 0.29 | -0.18 | 28.42 | -0.35 | 28.65 | -9.43 | -0.06 | -27.77 | 0 | -48.87 | -0.06 | -89.12 | 0 | -64.81 | -0.06
9.568 | -1.88 | 3.4 | -22.09 | -0.06 | -32.52 | -0.06 | -35.92 | -0.06 | 21.33 | 0 | 21.39 | 0.47 | 21.33 | 0 | 6.27 | -25.84 | 0 | 0 | 0.12 | 28.54 | -0.18 | 0.12 | -0.18 | 28.71 | -0.47 | 28.3 | -7.27 | 0 | -26.13 | -0.06 | -42.95 | -0.18 | -62.11 | -0.18 | -36.86 | 0
9.6 | -1.35 | 9.49 | -20.39 | -0.06 | -30.12 | -0.06 | -34.69 | -0.06 | 21.39 | 0.06 | 21.5 | 0.53 | 21.39 | 0.12 | 12.07 | -23.03 | -0.06 | 0 | 0.23 | 0.18 | -0.06 | 0 | -0.06 | 28.48 | -0.41 | 28.48 | -4.75 | -0.06 | -24.14 | 0 | -38.5 | -0.06 | -35.51 | -0.06 | -32.29 | 0
9.632 | -0.64 | 15.06 | -18.28 | -0.06 | -27.77 | -0.06 | -32.05 | 0 | 0.12 | 0 | 21.39 | 0.35 | 21.39 | 0.06 | 17.52 | -22.03 | 0.06 | 0 | 0.18 | 0.06 | -0.06 | 0.06 | -0.18 | 28.18 | -0.23 | 28.36 | -3.46 | -0.06 | -22.62 | 0 | -38.32 | -0.06 | -32.7 | -0.06 | -41.13 | 0
9.664 | -0.18 | 20.51 | -16.7 | -0.18 | -26.08 | -0.06 | -30.59 | -0.06 | 0 | 0 | 21.45 | 0.29 | 21.39 | 0.18 | 21.5 | -20.04 | 0 | 0 | 0.12 | 0 | -0.06 | 0 | -0.18 | 28.59 | -0.23 | 28.54 | -1.82 | 0.76 | -19.75 | -0.06 | -38.67 | -0.06 | -41.84 | 0 | -42.48 | 0
9.696 | -0.18 | 21.5 | -14.47 | -0.06 | -23.56 | -0.06 | -28.07 | -0.06 | 0 | -0.06 | 21.27 | 0.23 | 21.45 | 0.12 | 21.39 | -18.16 | 0 | 0 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.23 | -0.29 | 28.42 | -0.76 | 10.9 | -18.28 | 0 | -36.68 | -0.06 | -41.54 | -0.06 | -38.38 | 0
9.728 | -0.18 | 21.45 | -13.07 | 0 | -21.86 | -0.06 | -26.13 | 0.06 | 0.06 | -0.06 | 21.39 | 0.12 | 21.21 | 0.12 | 21.39 | -16.82 | -0.06 | 0 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0.06 | -0.29 | 28.65 | -0.7 | 18.28 | -16.29 | -0.18 | -35.16 | -0.18 | -37.5 | -0.06 | -34.1 | -0.06
9.76 | -0.23 | 21.45 | -11.48 | -0.06 | -18.99 | -0.18 | -23.97 | 0.06 | 0 | 0 | 21.27 | 0.18 | 21.39 | 0.12 | 21.39 | -14.36 | -0.06 | -24.9 | 0.12 | 0 | -0.06 | 0 | -0.18 | 0 | -0.29 | 28.48 | -0.18 | 28.71 | -14.47 | -0.06 | -33.52 | -0.18 | -33.93 | 0 | -32.58 | -0.06
9.792 | -0.35 | 21.45 | -9.9 | 0 | -17.23 | -0.06 | -21.86 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.39 | 0.12 | 21.39 | -12.83 | -0.06 | -73.89 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.23 | 28.54 | -0.23 | 28.65 | -12.48 | -0.18 | -30.7 | -0.06 | -32.52 | -0.06 | -32.34 | 0
9.824 | -0.35 | 21.5 | -9.02 | 0 | -15.82 | 0 | -20.1 | 0.06 | 0 | 0.06 | 21.5 | 0.12 | 21.27 | 0.06 | 21.39 | -11.13 | 0 | -52.33 | 0.23 | 0 | -0.06 | 0 | 0 | 0 | -0.23 | 28.65 | 0.23 | 28.42 | -10.02 | 0 | -29.47 | -0.06 | -32.46 | -0.06 | -31.23 | -0.06
9.856 | -0.29 | 21.45 | -6.86 | -0.06 | -14.12 | -0.06 | -17.4 | 0.06 | 0 | -0.06 | 21.39 | 0 | 21.39 | 0.18 | 21.09 | -9.43 | -0.06 | -31.82 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.35 | 28.65 | -7.56 | 0 | -27.31 | -0.23 | -31.23 | -0.06 | -29.24 | -0.23
9.888 | -0.23 | 21.45 | -4.86 | -0.06 | -12.77 | 0 | -15.76 | 0 | 0 | -0.06 | 21.39 | 0 | 21.39 | 0.23 | 21.15 | -7.68 | -0.06 | -28.65 | 0.18 | 0 | 0 | 0 | -0.06 | 0 | -0.23 | 28.36 | 0.18 | 28.65 | -5.45 | -0.06 | -25.67 | -0.18 | -29.06 | 0 | -27.6 | -0.18
9.92 | -0.18 | 21.45 | -3.93 | -0.06 | -10.61 | -0.18 | -14.24 | 0 | 0 | -0.06 | 21.56 | -0.06 | 21.39 | 0.06 | 21.21 | -5.45 | -0.06 | -34.75 | 0.06 | 0 | 0 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.23 | 28.48 | -3.46 | -0.18 | -24.08 | -0.18 | -26.72 | -0.18 | -26.02 | -0.06
9.952 | -0.23 | 21.45 | -2.46 | -0.06 | -9.02 | -0.06 | -12.6 | -0.06 | 0 | -0.06 | 21.15 | 0.06 | 21.15 | 0.12 | 21.39 | -3.93 | 0.06 | -34.75 | 0.12 | 0 | 0 | -14.94 | -0.06 | 0 | -0.23 | 28.36 | 0.35 | 28.42 | -2.4 | 0 | -21.27 | -0.06 | -26.54 | 0 | -25.02 | -0.18
9.984 | -0.29 | 21.45 | -1.64 | 6.68 | -7.68 | -0.06 | -11.31 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.21 | 0.12 | 21.27 | -2.81 | -0.06 | -31.88 | 0.06 | 0 | -0.06 | -78.17 | -0.06 | 0 | -0.06 | 28.65 | 0.41 | 28.42 | -0.82 | 0 | -19.63 | -0.18 | -24.84 | -0.18 | -23.03 | -0.18
10.016 | -0.23 | 21.45 | -0.59 | 12.48 | -6.5 | 5.33 | -9.38 | -0.06 | 0 | 0 | 0.23 | -0.06 | 21.33 | 0.18 | 21.39 | -1.88 | -0.06 | -28.89 | 0.12 | -31.47 | 0 | -55.61 | -0.06 | 0 | -0.06 | 28.95 | 0.41 | 28.48 | -0.64 | -0.06 | -17.4 | -0.18 | -22.38 | -0.18 | -21.15 | -0.06
10.048 | -0.29 | 21.45 | -0.35 | 18.11 | -4.86 | 11.25 | -7.85 | -0.06 | -26.13 | 0 | 0.06 | 0 | 21.33 | 0 | 21.21 | -1.46 | 0.06 | -27.42 | 0.06 | -72.72 | 0 | -44.71 | -0.06 | 0 | -0.18 | 28.65 | 0.35 | 28.3 | -0.18 | -0.06 | -15.94 | -0.06 | -20.92 | 0 | -19.69 | -0.06
10.08 | -0.18 | 21.45 | -0.06 | 21.8 | -3.16 | 16.7 | -6.56 | -0.06 | -65.92 | 0 | 0 | -0.18 | 21.39 | 0.06 | 21.39 | -0.41 | 0 | -26.66 | 0.12 | -51.15 | -0.06 | -44.53 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.54 | -0.06 | -0.06 | -14.24 | -0.18 | -19.75 | -0.18 | -17.99 | -0.18
10.112 | -0.29 | 21.45 | -0.18 | 21.5 | -1.99 | 21.91 | -4.86 | -0.06 | -46.76 | 0 | 0 | 0 | 21.39 | 0 | 21.39 | -0.29 | 0.7 | -24.79 | 0 | -44.83 | -0.06 | -41.84 | -0.06 | -6.21 | -0.06 | 0.23 | 0.23 | 28.59 | 0.12 | 0.82 | -11.54 | -0.18 | -17.64 | -0.18 | -16.17 | -0.06
10.144 | -0.18 | 21.45 | -0.18 | 21.5 | -1.23 | 21.45 | -3.75 | -0.06 | -41.43 | 0 | 0 | 0 | 21.27 | 0.06 | 21.15 | -0.23 | 8.56 | -22.79 | 0.12 | -43.89 | -0.06 | -41.19 | -0.18 | -75.82 | -0.06 | 0 | 0.23 | 28.65 | 0.35 | 10.55 | -9.32 | -0.06 | -16.29 | -0.18 | -14.53 | -0.06
10.176 | -0.18 | 21.39 | -0.18 | 21.45 | -0.59 | 21.45 | -2.23 | 0.06 | -35.16 | -0.06 | 0 | 0.06 | 19.57 | -0.06 | 21.33 | -0.06 | 15.65 | -21.33 | 0 | -40.9 | 0 | -38.97 | -0.18 | -52.56 | -0.06 | 0 | 0.29 | 28.36 | 0.35 | 18.05 | -7.56 | 9.84 | -14.3 | 1.64 | -13.01 | -0.06
10.208 | -0.23 | 21.39 | -0.41 | 21.45 | -0.23 | 21.45 | -1.23 | -0.06 | -31.29 | -0.06 | 0 | 0.06 | 0.12 | -0.06 | 21.39 | -0.06 | 21.27 | -19.51 | 0.06 | -40.37 | -0.06 | -37.03 | -0.06 | -46.17 | -0.18 | 0 | 0.12 | 28.42 | 0.35 | 25.61 | -5.39 | 17.05 | -12.77 | 11.25 | -11.48 | 2.11
10.24 | -0.18 | 21.45 | -0.35 | 21.45 | -0.06 | 21.45 | -0.47 | -0.06 | -32.52 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | -17.7 | 0.12 | -38.67 | 0 | -33.99 | 0 | -44.24 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.65 | -3.16 | 24.73 | -11.02 | 18.87 | -9.14 | 11.72
10.272 | -0.18 | 21.45 | -0.29 | 21.45 | -0.06 | 21.45 | -0.06 | 0.06 | -32.4 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | -15.53 | -0.06 | -36.21 | 0 | -32.76 | 0 | -42.66 | -0.18 | 0 | 0.06 | 28.48 | 0.35 | 28.65 | -1.88 | 28.71 | -9.38 | 26.37 | -7.68 | 19.28
10.304 | -0.23 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -0.06 | -0.06 | -30.82 | 0.06 | 0 | 0 | 0 | 0.06 | 21.45 | 0.29 | 21.45 | -13.95 | 0.06 | -34.81 | 0 | -30.7 | -0.06 | -40.78 | -0.18 | 0 | 0.06 | 28.42 | 0.29 | 28.71 | -1 | 28.48 | -7.44 | 28.48 | -5.63 | 27.25
10.336 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.45 | -0.18 | -0.06 | -29.24 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -12.31 | 0 | -31.82 | 0 | -28.95 | -0.06 | -37.85 | -0.06 | 0 | 0.06 | 28.42 | 0.23 | 28.59 | -0.7 | 28.36 | -5.68 | 28.54 | -4.45 | 28.42
10.368 | -0.23 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 0 | -26.49 | -0.06 | 0 | 0 | 0 | -0.06 | 21.21 | 0.23 | 21.39 | -10.9 | 0.06 | -30.35 | 0 | -27.6 | -0.06 | -37.21 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.42 | -0.06 | 28.59 | -4.22 | 28.65 | -2.93 | 28.48
10.4 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.29 | 5.22 | -23.91 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.21 | -8.56 | 0.06 | -28.95 | -0.06 | -24.61 | -0.06 | -34.45 | -0.18 | 0 | 0 | 28.65 | 0.18 | 28.42 | -0.06 | 28.59 | -2.75 | 28.36 | -1.88 | 28.42
10.432 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.35 | 11.25 | -21.56 | -0.06 | -24.67 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | -6.86 | 0 | -25.72 | 0 | -23.15 | 0 | -32.93 | -0.18 | 0 | 0 | 28.59 | 0.23 | 28.89 | 0.12 | 28.65 | -1.7 | 28.65 | -1 | 28.54
10.464 | -0.18 | 0.59 | -0.29 | 21.45 | -0.23 | 21.45 | -0.41 | 16.7 | -19.98 | -0.06 | -77.05 | 0.06 | 0 | 0.06 | 0.12 | 0.23 | 21.39 | -5.1 | 3.16 | -23.85 | 0 | -20.63 | -0.18 | -31.29 | 0 | 0 | 0.06 | 28.59 | 0.12 | 28.71 | 0.23 | 28.36 | -1.11 | 28.42 | -0.59 | 28.59
10.496 | -0.18 | 0.06 | -0.29 | 21.39 | -0.23 | 21.5 | -0.29 | 21.5 | -17.87 | -0.06 | -51.39 | 0 | 0 | -0.06 | 0 | 0.18 | 21.39 | -3.57 | 11.19 | -22.03 | -0.06 | -18.87 | 0 | -28.65 | 0 | 0 | -0.06 | 28.59 | 0.06 | 28.59 | 0.23 | 28.65 | -0.53 | 28.65 | -0.35 | 28.48
10.528 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.29 | 21.45 | -16.23 | -0.18 | -42.6 | 0 | 0 | -0.06 | 0 | 0.18 | 21.39 | -2.52 | 16.7 | -20.33 | 0.12 | -17.4 | 0 | -26.66 | -0.18 | -58.89 | 0 | 28.83 | 0.06 | 28.48 | 0.29 | 28.59 | -0.53 | 28.59 | -0.47 | 28.42
10.56 | -0.06 | 0.06 | -0.23 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -13.54 | -0.06 | -35.63 | 0.06 | 0 | 0 | 0 | 0.29 | 21.39 | -1.29 | 21.62 | -17.99 | 0 | -14.82 | -0.06 | -24.61 | -0.06 | -75.47 | 0 | 0.29 | 0 | 28.54 | 0.23 | 28.71 | -0.18 | 28.65 | -0.06 | 28.42
10.592 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -11.13 | -0.06 | -31.82 | 0 | 0 | -0.06 | 0 | 0.29 | 21.39 | -0.94 | 21.39 | -16.17 | 0 | -13.54 | 0 | -23.15 | 0 | -51.68 | 0 | 0.12 | 0.06 | 28.3 | 0.35 | 28.24 | -0.06 | 28.42 | -0.06 | 28.59
10.624 | 0 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | -9.43 | -0.06 | -32.58 | -0.06 | -63.69 | 0 | 0 | 0.23 | 21.33 | -0.53 | 21.45 | -14.53 | 2.99 | -12.07 | 0 | -21.09 | 0 | -47.52 | -0.06 | 0 | 0 | 28.3 | 0.29 | 28.42 | 0.06 | 28.42 | 0.06 | 28.42
10.656 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -6.86 | 0 | -32.64 | -0.18 | -61.76 | 0 | 0 | 0.12 | 21.39 | -0.29 | 21.39 | -12.89 | 11.19 | -10.78 | 0.06 | -17.99 | -0.18 | -40.08 | 0 | 0.06 | 0.12 | 28.48 | 0.23 | 28.54 | 0.06 | 28.54 | 0.06 | 28.48
10.688 | -0.18 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.33 | -5.22 | -0.06 | -31.41 | -0.18 | -44.12 | -0.06 | 0 | 0.23 | 21.09 | -0.29 | 21.39 | -11.37 | 18.87 | -9.32 | 8.97 | -16.41 | -0.06 | -36.51 | -0.06 | 0 | 0 | 28.36 | 0.18 | 28.3 | 0.06 | 28.65 | 0.06 | 28.48
10.72 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.5 | -0.18 | 21.39 | -2.75 | -0.06 | -28.48 | -0.06 | -39.44 | 0 | 0 | 0.23 | 21.39 | 0.06 | 21.33 | -10.55 | 29.01 | -7.62 | 16.11 | -15 | -0.06 | -37.44 | 0 | 0 | -0.06 | 28.3 | 0.18 | 28.48 | 0.18 | 28.36 | 0.12 | 28.54
10.752 | 0 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.29 | 21.39 | -1.64 | 0 | -26.84 | -0.06 | -33.87 | 0 | 0 | 0.23 | 21.39 | 0 | 21.27 | -9.32 | 29.01 | -6.45 | 26.6 | -13.83 | 0 | -38.09 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.42 | 0.06 | 28.71 | 0.06 | 28.42
10.784 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | -0.18 | 21.39 | -1.11 | 4.16 | -24.9 | 0 | -32.23 | 0 | 0 | 0.23 | 21.39 | 0.12 | 21.39 | -7.21 | 28.48 | -5.63 | 28.42 | -11.95 | 0 | -36.97 | 0 | 0 | 0 | 28.36 | 0 | 28.54 | 0.12 | 28.59 | 0.12 | 28.13
10.816 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 10.43 | -23.03 | -0.06 | -32.58 | 0 | 0 | 0.29 | 21.45 | 0.12 | 21.39 | -6.15 | 28.48 | -4.39 | 28.3 | -10.55 | 0 | -35.1 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.48 | 0.18 | 28.3 | 0.06 | 28.59
10.848 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 15.59 | -21.33 | -0.18 | -32.17 | -0.06 | 0 | 0.23 | 21.39 | 0.18 | 21.33 | -5.1 | 28.65 | -3.69 | 28.59 | -9.08 | -0.18 | -32.93 | -0.18 | 0 | -0.23 | 28.36 | 0 | 28.48 | 0.12 | 28.48 | 0.23 | 28.42
10.88 | -0.06 | -9.32 | -0.18 | 0.18 | -0.18 | 21.39 | -0.23 | 21.39 | 0.35 | 21.39 | -18.87 | -0.06 | -30.7 | 0 | -43.01 | 0.23 | 21.33 | 0.06 | 21.33 | -4.22 | 28.42 | -2.52 | 28.65 | -7.79 | 0.06 | -30.76 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.3 | 0.18 | 28.77 | 0.12 | 28.48
10.912 | -0.06 | -62.05 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.45 | 0.29 | 21.56 | -16.7 | -0.06 | -28.13 | -0.06 | -65.98 | 0.12 | 21.39 | 0.23 | 21.33 | -3.22 | 28.36 | -1.82 | 28.65 | -6.5 | 0 | -29.24 | -0.06 | 0 | -0.18 | 28.36 | 0.06 | 28.65 | 0.23 | 28.36 | 0.18 | 28.48
10.944 | -0.06 | -50.1 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.45 | 0.29 | 21.45 | -14.41 | -0.06 | -25.37 | -0.18 | -42.48 | 0.06 | 21.39 | 0.12 | 21.39 | -2.46 | 28.3 | -1.46 | 28.77 | -5.22 | -0.18 | -27.19 | 0 | 0 | 0 | 28.59 | 0 | 28.48 | 0.18 | 28.3 | 0.12 | 28.42
10.976 | -0.18 | -40.37 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -12.13 | -0.06 | -23.56 | -0.06 | -27.48 | 0.06 | 21.39 | 0.23 | 21.39 | -2.17 | 28.3 | -1.23 | 28.65 | -4.04 | 0 | -24.9 | -0.06 | -1.41 | -0.06 | 28.42 | -0.18 | 28.48 | 0.12 | 28.48 | 0.06 | 28.48
11.008 | 0 | -38.73 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.53 | 21.39 | -10.49 | -0.06 | -21.56 | 0 | -32.52 | 0 | 0.23 | 0.12 | 21.39 | -1.76 | 28.59 | -0.94 | 28.48 | -3.46 | -0.06 | -23.38 | 0 | -86.84 | -0.06 | 20.57 | -0.06 | 28.48 | 0 | 28.54 | 0.12 | 28.48
11.04 | -0.18 | -36.39 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -7.62 | -0.06 | -19.98 | 0 | -35.68 | 0 | 0.06 | 0.23 | 21.39 | -1.29 | 28.42 | -0.59 | 28.59 | -2.75 | -0.06 | -21.39 | 0 | -57.78 | -0.18 | 0.12 | -0.06 | 26.31 | 0 | 28.48 | 0.06 | 28.42
11.072 | -0.18 | -34.4 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.47 | 21.45 | -5.33 | -0.06 | -17.81 | -0.06 | -33.69 | 0.06 | 0 | 0.29 | 21.39 | -1 | 28.65 | -0.35 | 28.3 | -2.29 | 0.06 | -19.92 | 0 | -48.75 | -0.06 | 0 | -0.06 | 0.18 | 0 | 13.89 | 0.06 | 28.42
11.104 | -0.06 | -32.52 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 21.39 | 0.47 | 21.39 | -3.34 | 0 | -15.59 | -0.06 | -29.77 | -0.18 | 0 | 0.18 | 21.39 | -0.7 | 28.48 | -0.35 | 28.59 | -1.46 | -0.06 | -17.75 | 0 | -44.83 | 0 | 0.06 | -0.06 | 0 | -0.06 | 0.12 | 0 | 3.22
11.136 | -0.06 | -30.41 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.88 | -0.06 | -13.54 | -0.06 | -27.66 | 0 | 0 | 0.23 | 21.39 | -0.59 | 28.59 | -0.41 | 28.48 | -0.82 | 9.96 | -15.35 | 0 | -38.32 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.12 | 0.06
11.168 | -0.06 | -27.83 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -1.41 | 3.34 | -11.07 | -0.06 | -26.9 | 0.06 | 0 | 0.18 | 21.33 | -0.53 | 28.83 | -0.35 | 28.36 | -0.59 | 17.17 | -13.3 | 0 | -36.39 | -0.06 | 0 | 0 | 0 | -0.06 | 0.06 | 0.12 | 0
11.2 | -0.06 | -26.13 | -0.18 | 0 | 0 | 0 | -0.06 | 21.45 | 0.41 | 21.45 | -0.53 | 9.55 | -9.38 | -0.18 | -25.72 | -0.18 | 0 | 0.23 | 21.39 | -0.47 | 28.36 | -0.29 | 28.42 | -0.47 | 25.08 | -11.6 | -0.06 | -37.74 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | 0.12 | 0.06
11.232 | -0.06 | -23.85 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.35 | 21.39 | 0.18 | 16.76 | -6.97 | -0.06 | -23.67 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | 28.83 | -0.35 | 28.42 | -0.35 | 28.77 | -9.08 | 0 | -37.38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11.264 | 0 | -22.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.18 | 21.74 | -3.93 | 3.34 | -22.09 | 0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.83 | -0.35 | 28.65 | -0.41 | 28.65 | -7.03 | -0.06 | -36.45 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
11.296 | -0.06 | -19.34 | -0.06 | -35.16 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | 0.29 | 21.39 | -2.23 | 11.13 | -19.86 | -0.06 | 0 | 0.12 | 21.33 | -0.41 | 28.65 | -0.23 | 28.95 | -0.29 | 28.48 | -4.28 | 0 | -34.75 | 0.06 | 0 | 0 | 0 | -0.18 | 0 | 0 | 0
11.328 | -0.06 | -17.46 | -0.18 | -56.78 | -0.06 | -29.06 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | 21.39 | -1.46 | 16.82 | -18.34 | 0 | 0 | 0.12 | 21.33 | -0.35 | 28.65 | -0.29 | 28.48 | -0.29 | 28.48 | -2.64 | 0 | -32.58 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
11.36 | -0.06 | -15.53 | -0.18 | -42.95 | -0.06 | -62.58 | -0.06 | 0 | 0.12 | 21.33 | 0.47 | 21.33 | -0.76 | 21.68 | -16.7 | 0 | 0 | 0.18 | 0.12 | -0.29 | 28.65 | -0.18 | 28.42 | -0.23 | 28.48 | -1.7 | -0.18 | -29.94 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 0
11.392 | -0.06 | -14.36 | -0.06 | -38.79 | -0.06 | -44.94 | -0.18 | 0 | 0.12 | 21.21 | 0.47 | 21.39 | -0.41 | 21.39 | -15.06 | -0.06 | 0 | 0.12 | 0 | -0.23 | 28.48 | -0.23 | 28.54 | -0.29 | 28.3 | -1 | -0.06 | -27.83 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
11.424 | -0.06 | -12.77 | -0.06 | -37.38 | -0.06 | -40.31 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -0.18 | 21.45 | -13.13 | 0 | -1.17 | 0.18 | 0 | -0.35 | 28.65 | -0.18 | 28.65 | -0.23 | 28.65 | -0.41 | -0.18 | -26.31 | 0 | -0.47 | -0.06 | 0 | -0.06 | 0 | 0 | 0
11.456 | -0.06 | -10.9 | -0.06 | -35.57 | 0 | -38.67 | -0.06 | 0 | 0.06 | 21.5 | 0.41 | 21.33 | 0.12 | 21.39 | -11.25 | -0.06 | -73.13 | 0.18 | 0 | -0.23 | 28.3 | -0.18 | 28.83 | -0.23 | 28.77 | -0.06 | 0 | -24.61 | 0 | -84.38 | -0.18 | 0 | -0.18 | 0 | 0 | 0
11.488 | 3.4 | -9.49 | -0.06 | -33.99 | -0.06 | -35.92 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | 0.06 | 21.39 | -9.61 | 0.06 | -51.27 | 0.12 | 0 | -0.29 | 28.83 | -0.18 | 28.71 | -0.29 | 28.48 | -0.06 | -0.06 | -23.32 | -0.06 | -67.56 | 0 | -77.29 | -0.06 | -0.53 | -0.06 | 0
11.52 | 9.55 | -8.03 | -0.18 | -31.11 | 0 | -33.63 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | 0.29 | 21.39 | -7.79 | 0.06 | -30.76 | 0.12 | 0 | -0.23 | 0.29 | -0.23 | 28.59 | -0.18 | 28.18 | 0.12 | 3.34 | -20.39 | 0.06 | -48.75 | -0.06 | -72.6 | 0 | -86.43 | -0.06 | -8.91
11.552 | 14.88 | -6.33 | -0.06 | -28.89 | -0.06 | -31.7 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.33 | 0.41 | 21.39 | -5.92 | -0.06 | -28.07 | 0.18 | 0 | -0.18 | 0.12 | -0.18 | 13.95 | -0.23 | 29.01 | 0.29 | 11.6 | -18.05 | -0.06 | -44.24 | -0.06 | -41.9 | -0.06 | -68.73 | 0 | -87.89
11.584 | 20.39 | -5.04 | -0.06 | -26.84 | -0.06 | -30 | -0.18 | 0 | -0.06 | 21.45 | 0.41 | 21.39 | 0.41 | 21.39 | -4.22 | -0.06 | -34.16 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0.12 | -0.18 | 28.3 | 0.35 | 19.22 | -16.7 | -0.06 | -38.97 | -0.18 | -29.24 | -0.18 | -38.85 | 0 | -63.05
11.616 | 21.5 | -3.52 | -0.18 | -24.9 | 0 | -27.54 | 0 | 0 | -0.06 | 21.33 | 0.35 | 21.39 | 0.41 | 21.27 | -2.23 | 0.06 | -34.92 | 0.18 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.18 | 28.42 | 0.29 | 29.06 | -14.53 | 0 | -37.09 | -0.06 | -40.31 | -0.18 | -31.58 | 0 | -31.76
11.648 | 21.5 | -2.17 | -0.06 | -22.79 | 0 | -25.14 | -0.06 | 0 | 0 | 21.15 | 0.29 | 21.33 | 0.41 | 21.39 | -2.17 | 0.06 | -32.58 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0.23 | 28.54 | -12.83 | 0 | -37.79 | -0.18 | -43.01 | -0.06 | -40.55 | 0 | -31.47
11.68 | 21.45 | -1.17 | -0.06 | -20.98 | 0.06 | -23.38 | 0 | 0 | -0.06 | 0.23 | 0.18 | 21.39 | 0.35 | 21.33 | -1.11 | 0 | -28.77 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.42 | -10.25 | 0.06 | -37.21 | -0.06 | -38.85 | -0.06 | -41.78 | -0.06 | -41.54
11.712 | 21.5 | -0.59 | -0.06 | -18.22 | -0.06 | -21.39 | -0.06 | -38.67 | 0 | 0 | 0.23 | 21.39 | 0.29 | 21.33 | -0.76 | 0.06 | -27.36 | 0.18 | 0 | -0.18 | 0 | 0 | 0 | -0.18 | 28.59 | 0.35 | 28.48 | -8.09 | 0.06 | -36.68 | 0 | -34.86 | -0.06 | -38.32 | 0 | -42.13
11.744 | 21.45 | -0.18 | -0.06 | -16.23 | 0 | -18.99 | -0.06 | -72.54 | -0.06 | 0 | 0.12 | 21.39 | 0.29 | 21.33 | -0.41 | -0.06 | -26.13 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.18 | 0.35 | 28.65 | -5.8 | -0.06 | -34.16 | -0.06 | -33.69 | -0.06 | -34.63 | -0.06 | -37.97
11.776 | 21.39 | -0.18 | -0.06 | -14.94 | -0.06 | -17.05 | -0.06 | -45 | -0.06 | 0 | 0.12 | 21.33 | 0.29 | 21.33 | -0.18 | 0 | -24.55 | 0.12 | -36.21 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.35 | 28.65 | -3.87 | 0.06 | -32.76 | -0.18 | -32.93 | 0 | -32.7 | 0 | -33.17
11.808 | 21.45 | -0.18 | -0.06 | -13.36 | -0.06 | -15.18 | 0 | -39.32 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | -0.06 | 0 | -22.21 | 0.12 | -70.49 | -0.18 | 0 | 0 | 0 | -0.06 | 28.42 | 0.35 | 28.48 | -2.29 | 2.99 | -31.11 | 0 | -32.23 | -0.06 | -32.81 | 0 | -32.58
11.84 | 21.45 | -0.23 | -0.06 | -12.01 | 0 | -13.77 | -0.06 | -34.86 | -0.18 | 0 | 0.06 | 21.33 | 0.23 | 21.45 | -0.06 | 0 | -20.74 | 0.12 | -45.76 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.23 | 28.36 | -1 | 11.25 | -28.01 | -0.06 | -29.94 | -0.06 | -31.76 | -0.06 | -32.81
11.872 | 21.45 | -0.41 | -0.06 | -10.37 | -0.06 | -12.07 | -0.06 | -32.05 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.06 | -0.06 | -19.45 | 0.12 | -29.71 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.65 | -0.82 | 18.69 | -26.31 | 0 | -27.89 | -0.06 | -29.53 | 0 | -32.05
11.904 | 21.45 | -0.53 | 6.15 | -8.85 | -0.06 | -10.72 | -0.06 | -32.52 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.45 | 0.23 | 0 | -18.16 | 0.06 | -29.3 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.48 | -0.23 | 28.83 | -24.32 | 0 | -27.01 | 0 | -27.48 | -0.06 | -29.24
11.936 | 21.45 | -0.35 | 13.71 | -7.27 | -0.06 | -9.14 | -0.06 | -31.88 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | 0.23 | 5.39 | -15.7 | 0.06 | -35.68 | -0.18 | -4.16 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.36 | -0.29 | 28.65 | -22.97 | 0 | -25.02 | 0 | -25.72 | 0 | -27.01
11.968 | 21.45 | -0.23 | 19.4 | -5.57 | 0 | -7.44 | -0.18 | -30.7 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.15 | 0.18 | 11.31 | -14.24 | 0.06 | -34.75 | 0 | -73.83 | -0.06 | 0 | -0.18 | 28.48 | 0.06 | 28.48 | 0 | 28.65 | -21.21 | -0.06 | -23.2 | 0 | -24.43 | 0 | -25.61
12 | 21.45 | -0.29 | 21.5 | -4.39 | 0.06 | -6.09 | -0.18 | -28.77 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.45 | 0.29 | 18.63 | -12.6 | 0.06 | -31.58 | -0.06 | -52.74 | -0.06 | -45.76 | -0.18 | 25.84 | 0 | 28.36 | 0.06 | 28.65 | -18.52 | 0 | -21.86 | -0.06 | -23.15 | -0.06 | -24.67
12.032 | 21.39 | -0.18 | 21.5 | -2.64 | 3.05 | -4.63 | -0.06 | -26.25 | 0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.04 | 0.23 | 21.5 | -10.84 | 0.12 | -29.06 | -0.18 | -45.29 | 0 | -61.41 | -0.06 | 0.18 | 0.12 | 28.59 | 0.18 | 28.54 | -16.7 | 0 | -19.98 | 0 | -21.91 | -0.06 | -22.97
12.064 | 21.45 | -0.18 | 21.45 | -1.7 | 9.14 | -3.57 | -0.06 | -24.2 | -0.06 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0.23 | 21.39 | -8.61 | 0.12 | -27.42 | -0.06 | -44.77 | -0.06 | -45.24 | -0.06 | 0 | 0.12 | 28.54 | 0.18 | 28.59 | -14.71 | -0.06 | -18.87 | -0.18 | -19.86 | -0.18 | -21.27
12.096 | 21.39 | -0.23 | 21.45 | -0.94 | 14.59 | -1.82 | 2.99 | -21.39 | 0 | -25.08 | -0.18 | 0 | -0.06 | 21.21 | 0.23 | 21.39 | -6.62 | 0.12 | -26.31 | -0.06 | -42.54 | -0.06 | -45.53 | -0.18 | 0 | 0 | 28.54 | 0.29 | 28.42 | -12.89 | -0.06 | -16.52 | -0.18 | -17.75 | -0.06 | -19.34
12.128 | 21.45 | -0.18 | 21.5 | -0.41 | 20.16 | -1 | 9.26 | -19.81 | 0.06 | -76.76 | 0 | 0 | 0 | 21.33 | 0.29 | 21.09 | -5.1 | 0.06 | -24.43 | -0.18 | -41.54 | 0 | -43.3 | -0.18 | 0 | 0.06 | 28.48 | 0.23 | 28.42 | -10.72 | 3.34 | -14.71 | -0.18 | -16.47 | 0 | -17.99
12.16 | 21.39 | -0.29 | 21.45 | -0.18 | 21.5 | -0.47 | 14.59 | -17.7 | -0.06 | -53.15 | -0.06 | 0.06 | 0 | 0.12 | 0.23 | 21.39 | -3.69 | 0.06 | -21.97 | -0.18 | -37.97 | 0 | -41.13 | -0.18 | 0 | 0 | 28.59 | 0.18 | 28.48 | -8.14 | 11.66 | -13.59 | -0.06 | -15 | 1.46 | -16
12.192 | 21.45 | -0.23 | 21.45 | -0.06 | 21.45 | -0.18 | 20.27 | -16.06 | 0 | -41.66 | 0 | 0 | 0.12 | 0 | 0.23 | 21.45 | -2.46 | 0.12 | -20.92 | -0.06 | -36.04 | -0.06 | -38.97 | -0.18 | 0 | 0 | 28.54 | 0.29 | 28.54 | -5.98 | 21.91 | -11.54 | 1.88 | -13.07 | 11.07 | -14.65
12.224 | 21.39 | -0.29 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -14.47 | -0.06 | -35.16 | 0 | 0 | -0.06 | 0 | 0.29 | 20.98 | -1.46 | 0.12 | -18.99 | -0.18 | -35.22 | -0.06 | -37.21 | 0 | 0 | 0 | 28.24 | 0.29 | 28.42 | -3.93 | 28.95 | -10.14 | 11.25 | -11.72 | 18.4 | -12.77
12.256 | 21.45 | -0.29 | 21.39 | -0.29 | 21.45 | -0.18 | 21.45 | -11.72 | 0.06 | -31.52 | 0 | 0 | 0 | 0 | 0.18 | 21.39 | -0.82 | -0.06 | -17.4 | -0.18 | -33.28 | -0.18 | -35.74 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | -2.23 | 28.54 | -7.56 | 18.87 | -9.08 | 26.25 | -10.66
12.288 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | -0.35 | 21.39 | -8.91 | -0.06 | -32.64 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.53 | -0.06 | -15.59 | -0.18 | -31.41 | 0 | -34.22 | -0.06 | 0 | -0.18 | 28.3 | 0.12 | 28.65 | -1.46 | 28.48 | -6.33 | 26.43 | -7.56 | 28.48 | -9.38
12.32 | 21.39 | -0.18 | 21.45 | -0.41 | 21.39 | -0.35 | 21.45 | -7.38 | -0.06 | -32.76 | 0 | 0 | 0 | 0 | 0.18 | 21.33 | -0.29 | 0.12 | -13.59 | -0.18 | -28.65 | -0.06 | -32.23 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.65 | -0.7 | 28.42 | -4.51 | 28.48 | -5.98 | 28.65 | -7.85
12.352 | 21.39 | -0.23 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -4.92 | -0.06 | -31.06 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.39 | -0.29 | -0.06 | -12.42 | 0 | -27.01 | 0 | -29.06 | -0.06 | 0 | -0.18 | 28.42 | 0.06 | 28.59 | -0.41 | 28.01 | -3.16 | 28.48 | -4.75 | 28.42 | -6.09
12.384 | 0.76 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.39 | -2.99 | -0.06 | -28.89 | 0 | 0 | -0.06 | 0 | 0.06 | 21.21 | 0.12 | 0 | -10.61 | -0.06 | -24.38 | -0.06 | -27.72 | 0 | 0 | -0.18 | 28.24 | 0 | 28.65 | 0 | 28.48 | -2.4 | 28.3 | -3.16 | 28.48 | -4.51
12.416 | 0.06 | -0.18 | 21.45 | -0.35 | 21.45 | -0.29 | 21.45 | -1.46 | -0.06 | -26.78 | 0 | 0 | 0.06 | 0 | 0.06 | 21.33 | 0.12 | 0 | -8.14 | 0 | -22.74 | -0.06 | -25.55 | 0 | 0 | -0.06 | 0.23 | 0.12 | 28.65 | -0.06 | 28.83 | -1.17 | 28.59 | -2.11 | 28.59 | -2.64
12.448 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.94 | -0.18 | -24.55 | -0.06 | 0 | -0.06 | 0 | 0 | 21.21 | 0.06 | 2.23 | -6.33 | -0.06 | -20.92 | -0.06 | -23.5 | 0 | -60.06 | -0.06 | 0.12 | 0.18 | 28.65 | -0.06 | 28.48 | -0.59 | 28.48 | -1.17 | 28.42 | -2.11
12.48 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.18 | 21.39 | -0.41 | -0.06 | -22.85 | 0 | -5.33 | -0.06 | 0 | 0.12 | 21.39 | 0 | 9.55 | -4.63 | -0.06 | -18.16 | 0 | -21.86 | 0.06 | -63.81 | -0.06 | 0 | 0 | 28.65 | 0.41 | 28.42 | -0.53 | 28.54 | -0.82 | 28.48 | -1.29
12.512 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.35 | 21.45 | -0.35 | -0.06 | -20.39 | -0.06 | -72.07 | -0.23 | 0 | 0.18 | 21.39 | 0.23 | 14.88 | -3.22 | -0.06 | -16.64 | 0.06 | -19.63 | 0 | -45.94 | -0.06 | 0.06 | 0 | 28.65 | 0.29 | 28.71 | -0.35 | 28.54 | -0.47 | 28.65 | -0.76
12.544 | 0 | -0.06 | 21.45 | -0.29 | 21.5 | -0.35 | 21.45 | 0.06 | -0.06 | -18.22 | -0.06 | -48.4 | -0.18 | 0 | 0.18 | 21.39 | 0.23 | 20.68 | -2.17 | 0.06 | -15.18 | -0.06 | -17.05 | -0.06 | -45 | -0.06 | 0 | -0.18 | 28.71 | 0.47 | 28.59 | -0.18 | 28.71 | -0.35 | 28.59 | -0.47
12.576 | 0 | 0 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | 0.23 | 0 | -16.23 | -0.06 | -42.6 | -0.06 | -32.52 | 0.23 | 21.33 | 0.23 | 21.45 | -1.35 | -0.06 | -13.95 | 0 | -15.7 | -0.06 | -43.6 | -0.18 | 0 | 0 | 28.42 | 0.29 | 28.42 | -0.06 | 28.36 | -0.18 | 28.65 | -0.23
12.608 | 0 | 0.06 | 21.45 | -0.29 | 21.5 | -0.23 | 21.39 | 0.35 | -0.06 | -14.53 | 0.06 | -35.86 | -0.06 | -68.26 | 0.06 | 21.39 | 0.23 | 21.39 | -0.64 | 0 | -12.25 | -0.06 | -14.3 | -0.06 | -41.84 | -0.06 | 0 | 0.06 | 28.48 | 0.47 | 28.65 | -0.06 | 28.54 | 0 | 27.89 | -0.06
12.64 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.29 | -0.06 | -12.83 | -0.06 | -32.58 | -0.06 | -46.76 | 0.06 | 21.39 | 0.23 | 21.39 | -0.53 | -0.06 | -10.31 | 3.93 | -12.72 | 0.06 | -39.85 | -0.06 | 0 | 0 | 28.65 | 0.41 | 28.59 | 0.06 | 28.54 | 0.18 | 28.24 | 0
12.672 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.23 | 21.5 | 0.41 | -0.06 | -9.84 | 0 | -32.52 | -0.06 | -29.06 | -0.06 | 21.39 | 0.18 | 21.04 | -0.23 | -0.06 | -9.02 | 12.25 | -11.13 | 0.06 | -37.5 | -0.06 | 0 | -0.06 | 28.59 | 0.23 | 28.65 | 0.06 | 28.48 | 0.12 | 28.65 | -0.06
12.704 | 0 | -0.06 | 21.45 | -0.29 | 21.5 | -0.29 | 21.39 | 0.47 | 0 | -7.32 | -0.18 | -32.34 | -0.06 | -29.12 | 0.06 | 21.39 | 0.23 | 21.39 | -0.12 | 5.04 | -7.91 | 20.16 | -9.32 | 0 | -36.51 | -0.06 | 0 | 0 | 0.29 | 0.23 | 28.54 | 0.12 | 28.54 | 0.06 | 28.3 | 0.06
12.736 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | 0.47 | 4.16 | -5.39 | 0 | -31.29 | -0.06 | -35.74 | 0 | 21.33 | 0.18 | 21.39 | -0.06 | 13.24 | -6.8 | 27.54 | -8.44 | 0 | -33.93 | 0 | 0 | 0 | 0.12 | 0.18 | 28.42 | 0.12 | 28.48 | 0.18 | 28.59 | 0.18
12.768 | 0 | -0.06 | 15.94 | -0.23 | 21.45 | -0.18 | 21.45 | 0.53 | 10.37 | -3.28 | -0.06 | -29.12 | 0 | -34.22 | -0.06 | 21.33 | 0.18 | 21.33 | 0.12 | 20.8 | -5.57 | 28.77 | -7.03 | 0.06 | -31.23 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.48 | 0.18 | 28.42 | 0.12 | 28.71 | 0.18
12.8 | -5.33 | -0.18 | 0.12 | -0.18 | 21.45 | -0.18 | 21.39 | 0.53 | 17.64 | -2.23 | 0.06 | -26.78 | -0.06 | -31 | 0.12 | 19.92 | 0.18 | 21.39 | 0.12 | 28.65 | -4.28 | 28.65 | -5.86 | 2.99 | -29.77 | 0 | 0 | 0 | 0.06 | 0.23 | 28.48 | 0.12 | 28.42 | 0.29 | 28.54 | 0.18
12.832 | -62.4 | 0 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.41 | 21.68 | -0.94 | 0.06 | -24.55 | -0.06 | -28.59 | -0.06 | 0.12 | 0.29 | 21.33 | 0.12 | 28.83 | -3.63 | 28.83 | -4.63 | 11.31 | -27.6 | 0 | -3.16 | -0.06 | 0 | 0 | 28.59 | 0.18 | 28.36 | 0.18 | 28.13 | 0.29
12.864 | -48.99 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | 21.45 | -0.18 | -0.06 | -22.21 | -0.06 | -27.48 | -0.06 | 0 | 0.23 | 21.33 | 0 | 28.42 | -2.75 | 28.65 | -3.87 | 18.69 | -25.55 | 0 | -93.34 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.48 | 0.18 | 28.54 | 0.18
12.896 | -40.37 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.29 | 21.39 | -0.18 | -0.06 | -20.45 | -0.06 | -26.31 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 28.48 | -2.17 | 28.42 | -2.75 | 29.06 | -22.91 | 0 | -54.03 | -0.18 | 0 | 0.12 | 28.48 | 0 | 28.42 | 0.29 | 28.48 | 0.18
12.928 | -38.67 | 0 | 0 | -0.18 | 4.16 | -0.18 | 21.45 | 0.29 | 21.39 | 0.29 | -0.06 | -18.16 | -0.06 | -23.56 | 0.06 | 0.06 | 0.23 | 21.21 | 0.23 | 28.59 | -1.7 | 28.54 | -2.23 | 28.54 | -20.98 | -0.06 | -47.81 | -0.06 | 0 | 0.06 | 28.42 | 0.06 | 28.59 | 0.18 | 28.42 | 0.12
12.96 | -36.33 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 21.39 | 0.23 | 21.39 | 0.35 | -0.06 | -16.76 | -0.06 | -22.38 | 0 | 0 | 0.23 | 21.39 | 0.12 | 28.65 | -1.23 | 28.3 | -1.64 | 28.65 | -19.1 | -0.06 | -43.71 | 0 | 0 | -0.06 | 28.54 | 0 | 28.42 | 0.23 | 28.59 | 0.12
12.992 | -34.63 | 0 | 0 | -0.18 | 0 | -0.18 | 3.46 | 0.23 | 21.39 | 0.47 | -0.06 | -13.95 | -0.06 | -20.86 | 0 | 0 | 0.23 | 21.39 | 0.29 | 28.65 | -1 | 28.54 | -1.05 | 28.83 | -17.23 | -0.06 | -37.85 | -0.06 | 0 | 0.06 | 28.48 | 0 | 28.59 | 0.29 | 28.42 | 0.18
13.024 | -32.29 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 0.12 | 0.12 | 21.39 | 0.53 | -0.06 | -11.6 | -0.18 | -19.45 | 0 | 0 | 0.23 | 21.33 | 0.23 | 28.65 | -0.82 | 28.42 | -0.76 | 28.42 | -16.23 | 0 | -36.92 | -0.06 | 0 | 0 | 0.23 | 0 | 28.36 | 0.18 | 28.48 | 0.18
13.056 | -29.77 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | -0.06 | -9.79 | -0.18 | -16.7 | 0 | 0 | 0.18 | 21.39 | 0.23 | 28.59 | -0.53 | 28.54 | -0.59 | 28.42 | -14.06 | -0.06 | -37.74 | 0 | 0 | -0.06 | 0.06 | 0 | 28.48 | 0.18 | 14.94 | 0.12
13.088 | -27.72 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.06 | 0.06 | 21.33 | 0.47 | -0.06 | -7.68 | 0 | -15.06 | 0 | 0 | 0.23 | 21.39 | 0.12 | 28.77 | -0.59 | 28.18 | -0.41 | 28.71 | -12.72 | 0 | -37.79 | -0.06 | 0 | 0 | 0 | -0.18 | 13.65 | 0.12 | 0.12 | 0.18
13.12 | -26.13 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 3.63 | -5.92 | -0.06 | -13.59 | -0.06 | 0 | 0.23 | 21.39 | 0.12 | 28.89 | -0.41 | 28.65 | -0.23 | 28.36 | -11.31 | 0 | -35.68 | -0.06 | -3.87 | -0.18 | 0.06 | -0.18 | 0.18 | 0 | 0 | 0.18
13.152 | -23.56 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.21 | 0.47 | 9.96 | -3.16 | 0.06 | -11.89 | -0.06 | 0 | 0.12 | 21.27 | 0.23 | 28.65 | -0.35 | 28.65 | -0.29 | 28.54 | -9.67 | -0.06 | -34.16 | 0 | -89.48 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0.06 | 0.23
13.184 | -21.91 | 0 | -0.06 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 15 | -1.76 | -0.06 | -10.2 | -0.06 | 0 | 0.29 | 21.39 | 0.29 | 27.89 | -0.35 | 28.65 | -0.29 | 28.65 | -8.32 | -0.06 | -31.64 | -0.06 | -55.96 | 0 | 0 | -0.18 | 0.06 | 0 | 0 | 0.23
13.216 | -19.04 | -0.06 | -48.93 | -0.06 | 0 | 0 | 0 | -0.06 | 21.45 | 0.41 | 21.09 | -1 | -0.06 | -7.62 | 0 | 0.06 | 0.12 | 21.33 | 0.12 | 28.83 | -0.35 | 28.36 | -0.23 | 28.95 | -7.09 | 7.5 | -29.24 | 0 | -48.69 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18
13.248 | -17.58 | -0.06 | -52.85 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.41 | 21.45 | -0.59 | -0.18 | -5.98 | -0.06 | -34.75 | 0.12 | 21.39 | 0.23 | 28.24 | -0.35 | 28.59 | -0.35 | 28.42 | -5.68 | 14.82 | -28.01 | 0 | -44.59 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.12
13.28 | -15.7 | -0.06 | -40.49 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.15 | 0.35 | 21.39 | -0.29 | -0.06 | -4.39 | -0.06 | -61.41 | 0.12 | 21.39 | 0.12 | 28.89 | -0.23 | 28.42 | -0.29 | 28.59 | -4.63 | 22.97 | -26.54 | -0.18 | -38.5 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.12
13.312 | -14.24 | -0.06 | -40.14 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | 0.29 | 4.92 | -2.93 | -0.06 | -37.68 | -0.06 | 21.39 | 0.23 | 28.65 | -0.23 | 28.65 | -0.18 | 28.24 | -3.75 | 29.01 | -24.61 | 0 | -36.86 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.18
13.344 | -12.48 | -0.06 | -37.97 | -0.06 | -1.93 | -0.18 | 0 | 0.06 | 21.33 | 0.23 | 21.45 | 0.29 | 12.66 | -2.11 | -0.06 | -27.77 | 0.12 | 0.59 | 0.12 | 28.59 | -0.23 | 28.48 | -0.29 | 28.65 | -3.11 | 28.48 | -22.68 | 0 | -37.62 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.18
13.376 | -10.61 | -0.06 | -35.16 | -0.18 | -60.12 | -0.18 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.35 | 18.16 | -1.11 | 0 | -31.76 | 0.18 | 0.06 | 0.23 | 28.59 | -0.18 | 28.42 | -0.18 | 28.71 | -2.4 | 28.48 | -19.98 | 0 | -37.68 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0
13.408 | -9.14 | -0.18 | -33.05 | -0.06 | -51.04 | -0.18 | -4.45 | 0.06 | 21.56 | 0.12 | 21.39 | 0.35 | 21.62 | -0.64 | 0.06 | -35.27 | 0.06 | 0 | 0.23 | 28.59 | -0.18 | 28.42 | -0.29 | 28.59 | -1.64 | 28.65 | -18.22 | 0 | -36.45 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
13.44 | -7.91 | -0.06 | -30.82 | -0.06 | -40.43 | -0.18 | -58.42 | 0 | 21.39 | 0.23 | 21.39 | 0.53 | 21.39 | -0.41 | -0.06 | -33.22 | -0.06 | 0 | 0.29 | 28.48 | -0.18 | 28.54 | -0.18 | 28.65 | -1 | 28.42 | -16.58 | 0 | -34.16 | -0.18 | -7.97 | -0.18 | 0 | 0 | 0 | 0
13.472 | -6.21 | -0.06 | -29.12 | 0 | -38.79 | -0.18 | -49.46 | 0 | 21.33 | 0.12 | 21.39 | 0.47 | 21.39 | -0.18 | -0.18 | -29.77 | -0.06 | 0 | 0.23 | 28.65 | -0.18 | 28.42 | -0.18 | 28.83 | -0.7 | 28.59 | -15 | 0 | -31.88 | -0.06 | -83.38 | -0.18 | 0 | -0.06 | -0.59 | 0
13.504 | -4.98 | -0.06 | -27.25 | -0.06 | -36.74 | -0.18 | -40.43 | -0.06 | 21.39 | 0.18 | 21.39 | 0.53 | 21.33 | -0.18 | -0.06 | -28.24 | 0 | 0 | 0.23 | 28.83 | -0.18 | 28.54 | -0.18 | 28.77 | -0.53 | 28.71 | -12.6 | 0 | -29.65 | 0 | -54.9 | -0.06 | -2.05 | -0.06 | -86.84 | 0
13.536 | -3.28 | -0.06 | -24.2 | -0.06 | -34.92 | -0.06 | -39.67 | -0.06 | 21.39 | 0.12 | 21.15 | 0.53 | 21.39 | 0.06 | -0.18 | -27.48 | -0.18 | 0 | 0.23 | 28.54 | -0.18 | 0.64 | -0.23 | 28.48 | -0.53 | 28.89 | -9.84 | 0 | -27.77 | -0.06 | -47.99 | -0.06 | -85.55 | 0 | -67.5 | 0.12
13.568 | -2.17 | 2.75 | -22.21 | -0.06 | -32.81 | -0.06 | -36.51 | 0.06 | 21.45 | -0.06 | 21.39 | 0.41 | 21.39 | 0.12 | 5.39 | -26.02 | 0.12 | 0 | 0.23 | 28.48 | -0.06 | 0.12 | -0.18 | 28.65 | -0.29 | 28.42 | -7.85 | 0 | -25.9 | -0.06 | -44.59 | -0.06 | -66.86 | -0.18 | -38.44 | 0
13.6 | -1.35 | 8.73 | -20.22 | -0.06 | -30.12 | -0.06 | -34.75 | 0.06 | 21.39 | 0.12 | 21.39 | 0.53 | 21.33 | 0.12 | 11.13 | -23.15 | -0.06 | 0 | 0.06 | 0.29 | -0.18 | 0 | -0.18 | 28.59 | -0.35 | 28.54 | -5.57 | 0.06 | -24.2 | -0.06 | -37.68 | -0.06 | -38.61 | -0.18 | -30.82 | 0
13.632 | -0.7 | 14.18 | -17.99 | -0.06 | -27.89 | -0.06 | -32.7 | 0.06 | 0.12 | 0.06 | 21.39 | 0.29 | 21.39 | 0.12 | 16.76 | -21.62 | 0.06 | 0 | 0.06 | 0.12 | -0.18 | 0.06 | -0.18 | 28.48 | -0.35 | 28.54 | -3.87 | -0.06 | -22.62 | 0 | -37.91 | -0.06 | -31.47 | -0.18 | -40.9 | 0
13.664 | -0.41 | 19.98 | -16.41 | -0.06 | -26.25 | -0.06 | -30.53 | 0 | 0 | 0.06 | 21.21 | 0.41 | 21.39 | 0.18 | 21.62 | -20.39 | -0.06 | 0 | 0.12 | 0 | -0.18 | 0 | 0 | 28.54 | -0.29 | 28.13 | -2.17 | 2.23 | -20.51 | -0.06 | -37.85 | -0.18 | -41.02 | -0.06 | -42.54 | 0
13.696 | -0.06 | 21.45 | -14.47 | -0.06 | -23.67 | -0.06 | -27.77 | -0.18 | 0 | 0 | 21.45 | 0.29 | 21.39 | 0.12 | 21.39 | -19.04 | -0.06 | 0 | 0.12 | 0.06 | -0.06 | 0 | -0.18 | 0.23 | -0.41 | 28.48 | -1.41 | 9.96 | -18.34 | -0.06 | -37.44 | -0.18 | -42.31 | -0.06 | -38.79 | 0
13.728 | -0.18 | 21.5 | -12.72 | -0.06 | -22.44 | -0.06 | -26.25 | -0.18 | 0.06 | 0 | 21.27 | 0.23 | 21.39 | 0.23 | 21.21 | -16.7 | -0.06 | 0 | 0.23 | 0 | -0.06 | 0 | 0 | 0.06 | -0.29 | 28.65 | -0.76 | 17.17 | -16.47 | -0.18 | -36.1 | -0.06 | -37.68 | -0.06 | -34.63 | 0
13.76 | -0.29 | 21.45 | -11.48 | -0.06 | -18.93 | -0.06 | -23.91 | -0.06 | 0 | 0 | 21.27 | 0.23 | 21.27 | 0.12 | 21.15 | -14.82 | 0 | -11.25 | 0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.59 | -0.06 | 27.54 | -14.47 | -0.06 | -33.69 | 0 | -34.28 | -0.06 | -32.7 | -0.06
13.792 | -0.29 | 21.45 | -10.14 | -0.06 | -17.17 | -0.18 | -22.15 | 0 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.12 | 21.39 | -12.95 | 0.06 | -73.19 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.41 | 28.59 | -0.18 | 28.54 | -12.77 | -0.06 | -30.94 | -0.06 | -33.11 | -0.06 | -32.99 | 0.06
13.824 | -0.29 | 21.45 | -9.08 | -0.06 | -15.94 | 0 | -19.81 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.23 | 21.33 | -11.25 | -0.06 | -55.43 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.12 | 28.65 | -10.14 | 0 | -29.24 | -0.18 | -32.93 | -0.18 | -31.29 | -0.06
13.856 | -0.29 | 21.45 | -7.21 | -0.06 | -14.18 | -0.18 | -17.4 | -0.06 | 0 | 0 | 21.39 | 0 | 21.33 | 0.18 | 21.33 | -9.55 | -0.06 | -33.05 | 0.12 | 0 | -0.06 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0.18 | 28.65 | -8.2 | -0.06 | -27.36 | -0.06 | -31.99 | -0.06 | -29.3 | -0.06
13.888 | -0.35 | 21.39 | -5.39 | -0.06 | -12.77 | -0.18 | -15.65 | -0.18 | 0 | -0.06 | 21.33 | 0.06 | 21.39 | 0.18 | 21.39 | -8.2 | 0 | -27.72 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.23 | 28.65 | 0.23 | 28.65 | -5.86 | -0.18 | -26.13 | -0.18 | -29.71 | -0.18 | -27.77 | -0.06
13.92 | -0.29 | 21.45 | -3.93 | 0 | -10.72 | -0.18 | -14.41 | -0.06 | 0 | -0.06 | 21.21 | 0.06 | 21.39 | 0.23 | 21.21 | -5.74 | -0.18 | -34.51 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.36 | -4.04 | 0.06 | -24.26 | -0.06 | -27.42 | -0.06 | -26.13 | 0
13.952 | -0.23 | 21.45 | -2.75 | -0.06 | -9.61 | -0.06 | -12.42 | 0 | 0 | -0.06 | 21.27 | 0.06 | 21.33 | 0.12 | 21.27 | -4.04 | 0 | -34.92 | 0.06 | 0 | 0.06 | -9.55 | -0.06 | 0 | -0.18 | 28.42 | 0.53 | 28.48 | -2.23 | -0.06 | -21.74 | -0.18 | -25.78 | 0 | -24.9 | 0
13.984 | -0.18 | 21.45 | -1.76 | 5.8 | -8.14 | -0.06 | -11.31 | 0 | 0 | -0.06 | 21.33 | 0 | 21.09 | 0.18 | 21.39 | -2.81 | -0.06 | -32.93 | 0.12 | 0 | -0.18 | -76.7 | -0.06 | 0 | -0.18 | 28.3 | 0.47 | 28.59 | -1.35 | -0.06 | -20.33 | -0.18 | -25.02 | -0.06 | -22.85 | -0.06
14.016 | -0.18 | 21.5 | -0.7 | 11.54 | -6.5 | 5.04 | -9.61 | 0 | 0 | -0.06 | 0.29 | 0.06 | 21.45 | 0.06 | 21.15 | -1.99 | -0.06 | -28.89 | 0.06 | -18.69 | -0.06 | -56.25 | -0.06 | 0 | -0.06 | 28.42 | 0.29 | 28.83 | -0.53 | -0.06 | -18.05 | -0.18 | -22.97 | -0.18 | -21.56 | 0
14.048 | -0.23 | 21.39 | -0.41 | 17.29 | -5.04 | 11.13 | -7.91 | -0.06 | -20.51 | -0.18 | 0.06 | 0 | 21.15 | 0.23 | 21.27 | -1.29 | 0 | -27.42 | 0 | -77.76 | -0.06 | -44.3 | -0.18 | 0 | -0.18 | 28.54 | 0.29 | 28.54 | -0.35 | -0.06 | -16.47 | -0.18 | -21.09 | -0.06 | -19.86 | -0.06
14.08 | -0.29 | 21.45 | -0.06 | 21.8 | -3.57 | 16.41 | -6.86 | -0.06 | -67.27 | -0.18 | 0 | 0.06 | 21.39 | 0 | 21.39 | -0.76 | 0.06 | -27.13 | 0.06 | -54.49 | -0.18 | -45.12 | -0.06 | 0 | -0.18 | 28.42 | 0.35 | 28.48 | -0.06 | -0.06 | -14.06 | -0.18 | -20.22 | -0.06 | -18.05 | -0.06
14.112 | -0.23 | 21.45 | -0.06 | 21.45 | -2.17 | 21.8 | -5.04 | -0.06 | -46.7 | -0.06 | 0.06 | 0 | 21.27 | 0.12 | 21.33 | -0.29 | 0.12 | -24.61 | 0.12 | -44.94 | -0.06 | -41.13 | -0.06 | -3.63 | -0.06 | 0.23 | 0.35 | 28.48 | 0 | 0.53 | -11.48 | -0.06 | -17.99 | -0.06 | -16.76 | 0
14.144 | -0.23 | 21.39 | -0.23 | 21.45 | -1.29 | 21.5 | -3.87 | -0.06 | -40.84 | -0.18 | 0 | -0.06 | 21.27 | 0 | 21.39 | -0.18 | 7.62 | -22.74 | 0.06 | -43.89 | -0.06 | -40.55 | -0.18 | -73.36 | -0.18 | 0.06 | 0.29 | 28.59 | 0.12 | 10.31 | -9.79 | -0.18 | -16.06 | -0.18 | -15 | 0
14.176 | -0.23 | 21.45 | -0.35 | 21.45 | -0.53 | 21.45 | -2.29 | -0.06 | -35.39 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.33 | -0.06 | 14.94 | -20.74 | 0.12 | -41.31 | 0 | -38.85 | -0.06 | -52.74 | -0.06 | 0 | 0.29 | 28.18 | 0.41 | 17.81 | -7.79 | 8.97 | -14.82 | 0.41 | -13.24 | 0
14.208 | -0.18 | 21.45 | -0.41 | 21.45 | -0.18 | 21.45 | -1.52 | 0 | -31.64 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.33 | 0 | 20.63 | -19.75 | 0.12 | -39.96 | 0 | -37.38 | 0 | -45.59 | -0.18 | 0.06 | 0.06 | 28.65 | 0.35 | 25.37 | -5.86 | 16.11 | -13.24 | 10.14 | -11.31 | 2.64
14.24 | -0.18 | 21.45 | -0.41 | 21.45 | -0.29 | 21.45 | -0.76 | -0.06 | -32.4 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | -18.16 | 0.06 | -38.56 | 0 | -34.98 | 0 | -44.94 | 0 | 0 | 0.12 | 28.42 | 0.29 | 28.71 | -3.75 | 23.73 | -11.48 | 17.52 | -9.55 | 10.78
14.272 | -0.23 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -0.35 | 0.06 | -32.29 | -0.18 | 0 | -0.06 | 0 | 0 | 21.45 | 0.18 | 21.39 | -15.76 | 0.06 | -36.45 | 0 | -33.11 | -0.18 | -42.31 | -0.18 | 0 | 0.06 | 28.59 | 0.29 | 28.54 | -1.93 | 28.83 | -9.14 | 25.37 | -7.79 | 18.16
14.304 | -0.18 | 21.45 | -0.29 | 21.39 | -0.23 | 21.45 | -0.06 | 0 | -31 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 21.45 | -14.24 | -0.06 | -34.28 | -0.06 | -31 | -0.06 | -41.43 | 0 | 0 | 0.18 | 28.42 | 0.35 | 28.48 | -1.23 | 28.65 | -7.62 | 28.71 | -5.98 | 25.84
14.336 | -0.18 | 21.45 | -0.18 | 21.5 | -0.35 | 21.45 | -0.06 | 0 | -28.71 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | -12.36 | -0.06 | -32.23 | 0 | -28.77 | -0.06 | -38.91 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.48 | -0.7 | 28.48 | -6.15 | 28.42 | -4.63 | 28.48
14.368 | -0.18 | 21.45 | -0.23 | 21.39 | -0.23 | 21.45 | -0.29 | 0.06 | -26.84 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | -11.13 | 0.12 | -30.53 | -0.06 | -27.6 | -0.06 | -36.92 | -0.06 | 0 | 0.06 | 28.42 | 0.35 | 28.42 | -0.35 | 28.3 | -4.39 | 28.42 | -3.16 | 28.42
14.4 | -0.18 | 21.39 | -0.35 | 21.45 | -0.29 | 21.39 | -0.35 | 4.92 | -23.91 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -8.85 | 0.06 | -29.24 | -0.06 | -24.02 | -0.06 | -35.16 | -0.06 | 0 | 0.06 | 28.65 | 0.18 | 28.71 | -0.18 | 28.48 | -3.34 | 28.24 | -2.05 | 28.48
14.432 | -0.06 | 21.5 | -0.23 | 21.45 | -0.23 | 21.45 | -0.35 | 11.07 | -21.91 | -0.06 | -16.17 | -0.06 | 0 | -0.06 | 21.15 | 0.18 | 21.27 | -7.03 | 0 | -26.54 | 0 | -22.56 | 0 | -33.28 | -0.18 | 0 | 0.06 | 28.48 | 0.18 | 28.48 | 0.06 | 28.36 | -2.11 | 28.42 | -1.46 | 28.48
14.464 | -0.18 | 10.9 | -0.23 | 21.39 | -0.23 | 21.39 | -0.41 | 16.35 | -20.22 | -0.06 | -77 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.39 | -5.33 | 2.99 | -24.08 | -0.06 | -20.92 | 0 | -31.99 | -0.18 | 0 | 0 | 28.48 | 0.12 | 28.59 | 0.06 | 28.48 | -1 | 28.42 | -0.76 | 28.42
14.496 | -0.18 | 0.06 | -0.35 | 21.45 | -0.18 | 21.45 | -0.23 | 21.5 | -18.28 | -0.06 | -53.5 | 0 | 0 | -0.06 | 0 | 0.23 | 21.27 | -3.75 | 11.07 | -22.56 | 0 | -19.04 | -0.06 | -28.83 | -0.06 | 0 | 0.06 | 28.48 | 0.12 | 28.65 | 0.23 | 28.42 | -0.94 | 28.36 | -0.59 | 28.48
14.528 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | -16.29 | -0.18 | -43.13 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -2.64 | 16.41 | -19.98 | 0 | -17.23 | 0 | -26.84 | 0 | -46.12 | 0 | 28.48 | 0.06 | 28.48 | 0.35 | 28.59 | -0.35 | 28.42 | -0.18 | 28.18
14.56 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -13.89 | -0.18 | -36.45 | -0.18 | 0 | 0 | 0 | 0.18 | 21.39 | -1.7 | 21.62 | -18.46 | 0 | -15 | 0 | -25.37 | 0 | -80.69 | 0 | 0.7 | 0.06 | 28.42 | 0.29 | 28.48 | -0.35 | 28.65 | -0.18 | 28.65
14.592 | -0.18 | 0 | -0.23 | 21.45 | -0.12 | 21.45 | -0.29 | 21.39 | -12.07 | -0.06 | -32.17 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.94 | 21.39 | -16.52 | 0.06 | -13.54 | -0.06 | -22.68 | -0.18 | -54.14 | 0 | 0.12 | 0.06 | 28.54 | 0.41 | 28.3 | 0 | 28.65 | 0.06 | 28.3
14.624 | -0.06 | 0 | -0.23 | 21.5 | -0.18 | 21.45 | -0.23 | 21.39 | -9.84 | -0.06 | -32.4 | 0 | -56.37 | 0 | 0 | 0.18 | 21.21 | -0.64 | 21.45 | -14.53 | 0.7 | -12.01 | -0.06 | -20.92 | 0 | -49.16 | 0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.13 | 0 | 28.48 | 0.12 | 28.59
14.656 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.23 | 21.39 | -7.62 | -0.06 | -32.7 | -0.06 | -63.81 | 0 | 0 | 0.18 | 21.39 | -0.41 | 21.45 | -13.24 | 10.72 | -10.61 | -0.06 | -18.46 | 0 | -42.25 | 0 | 0 | -0.06 | 28.48 | 0.29 | 28.3 | 0 | 28.42 | 0 | 28.42
14.688 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | -5.51 | -0.06 | -31.88 | -0.18 | -44.36 | -0.06 | 0 | 0.06 | 21.33 | -0.18 | 21.39 | -12.01 | 18.16 | -9.61 | 8.2 | -16.47 | 0 | -37.15 | 0.06 | 0 | -0.06 | 28.42 | 0.23 | 28.77 | 0.18 | 28.54 | 0.23 | 28.36
14.72 | -0.06 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -2.99 | -0.06 | -29.01 | -0.06 | -39.79 | 0 | 0 | 0.18 | 21.39 | 0 | 21.33 | -10.78 | 28.42 | -7.44 | 15.35 | -15.06 | 0 | -37.03 | 0 | 0 | 0.06 | 28.48 | 0.12 | 28.48 | 0.06 | 28.65 | 0.29 | 28.65
14.752 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | -0.18 | 21.39 | -1.64 | -0.18 | -27.13 | -0.18 | -33.81 | -0.18 | 0 | 0.18 | 21.39 | 0.12 | 21.33 | -9.32 | 28.77 | -6.33 | 25.72 | -13.71 | 0 | -37.38 | -0.06 | 0 | 0 | 28.42 | 0.18 | 28.48 | 0.12 | 28.65 | 0.29 | 28.42
14.784 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -1.35 | 3.34 | -24.55 | -0.18 | -32.46 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 21.39 | -7.27 | 28.54 | -5.51 | 28.65 | -12.42 | 0 | -37.38 | 0.06 | 0 | 0 | 28.59 | 0.06 | 28.48 | 0.23 | 28.48 | 0.23 | 28.65
14.816 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.5 | -0.23 | 21.45 | -0.41 | 9.61 | -22.68 | 0.12 | -32.64 | 0 | 0 | 0.12 | 21.39 | 0.18 | 21.39 | -6.5 | 28.48 | -4.51 | 28.65 | -10.31 | -0.06 | -35.22 | 0 | 0 | 0 | 28.65 | 0 | 28.48 | 0.23 | 28.65 | 0.06 | 28.65
14.848 | -0.06 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -0.35 | 14.94 | -20.86 | -0.06 | -32.76 | -0.18 | 0 | 0.12 | 21.33 | 0.23 | 21.15 | -5.22 | 28.48 | -3.75 | 28.65 | -9.08 | -0.06 | -33.46 | 0 | 0 | 0 | 28.48 | -0.06 | 28.3 | 0.12 | 28.83 | 0.18 | 28.42
14.88 | -0.06 | -2.34 | -0.06 | 0.23 | 0 | 21.39 | -0.06 | 21.39 | 0.18 | 20.45 | -18.4 | 0 | -30.94 | -0.06 | -35.92 | 0.18 | 21.39 | 0.18 | 21.27 | -4.63 | 28.65 | -2.81 | 28.48 | -7.79 | -0.06 | -30.94 | 0.12 | 0 | 0.06 | 28.42 | 0.12 | 28.18 | 0.18 | 28.42 | 0.23 | 28.48
14.912 | -0.06 | -57.37 | -0.06 | 0 | -0.18 | 0.12 | -0.29 | 21.39 | 0.23 | 21.45 | -16.17 | 0 | -28.54 | -0.18 | -69.85 | 0.18 | 21.39 | 0.06 | 21.15 | -3.57 | 28.3 | -2.23 | 28.65 | -6.45 | -0.06 | -29.36 | 0 | 0 | 0 | 28.65 | 0 | 28.65 | 0.18 | 29.01 | 0.12 | 28.42
14.944 | -0.06 | -52.68 | -0.06 | 0 | -0.06 | 0.06 | -0.18 | 21.39 | 0.29 | 21.45 | -14.59 | 0.06 | -26.13 | -0.06 | -46.06 | 0.18 | 21.39 | 0.12 | 21.39 | -2.75 | 28.42 | -1.76 | 28.59 | -5.39 | 0 | -27.48 | 0 | 0 | 0 | 28.3 | 0 | 28.65 | 0.12 | 28.48 | 0.12 | 28.42
14.976 | -0.06 | -40.61 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -12.77 | 0 | -24.26 | -0.06 | -27.19 | 0.12 | 21.33 | 0.29 | 21.39 | -2.23 | 28.48 | -1.11 | 28.59 | -4.22 | -0.06 | -25.14 | -0.06 | 0 | -0.06 | 28.42 | -0.06 | 28.3 | 0.23 | 28.42 | 0.12 | 28.48
15.008 | -0.06 | -39.14 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -10.37 | 0.06 | -21.68 | 0.06 | -31.47 | 0 | 0.47 | 0.18 | 21.39 | -1.7 | 28.48 | -0.94 | 28.54 | -3.22 | -0.06 | -23.38 | 0 | -76.23 | 0 | 22.85 | 0.06 | 28.54 | 0.12 | 28.48 | 0.12 | 28.42
15.04 | -0.06 | -37.79 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -7.62 | 0.12 | -19.92 | -0.06 | -35.45 | 0.12 | 0.06 | 0.23 | 21.33 | -1.35 | 28.48 | -0.64 | 28.89 | -2.58 | 0 | -22.03 | 0 | -61.29 | 0 | 0.12 | 0 | 28.48 | 0 | 28.48 | 0.06 | 28.42
15.072 | -0.06 | -34.86 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | -5.22 | -0.06 | -17.93 | -0.06 | -34.63 | 0.06 | 0 | 0.29 | 21.45 | -1.11 | 28.59 | -0.47 | 28.48 | -2.11 | 0 | -19.81 | 0 | -48.52 | 0 | 0 | -0.06 | 0.23 | 0.06 | 23.26 | 0.06 | 28.42
15.104 | 0 | -32.7 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.45 | 0.47 | 21.45 | -3.63 | -0.06 | -15.47 | 0 | -30.94 | 0.12 | 0 | 0.18 | 21.15 | -0.76 | 28.54 | -0.41 | 28.54 | -1.46 | 0 | -17.93 | 0.06 | -44.83 | -0.06 | 0 | 0 | 0 | 0.06 | 0.18 | 0.12 | 17.93
15.136 | -0.06 | -30.18 | 0 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -2.11 | -0.06 | -13.71 | -0.06 | -28.13 | 0.06 | 0 | 0.23 | 21.45 | -0.53 | 28.59 | -0.23 | 28.54 | -0.94 | 9.02 | -15 | -0.06 | -38.97 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 0.12
15.168 | 0 | -27.95 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.27 | 0.35 | 21.39 | -1.29 | 1.46 | -11.48 | -0.06 | -27.72 | 0 | 0 | 0.29 | 21.27 | -0.59 | 28.3 | -0.29 | 28.48 | -0.64 | 16.11 | -13.36 | 0 | -36.74 | 0 | 0 | 0 | 0 | 0 | 0 | 0.06 | 0
15.2 | 0.06 | -26.13 | 0 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | -0.29 | 8.79 | -9.08 | -0.06 | -25.72 | 0.12 | 0 | 0.23 | 21.39 | -0.59 | 28.89 | -0.29 | 28.48 | -0.47 | 24.08 | -11.31 | 0 | -38.09 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.06 | 0.06
15.232 | 0 | -24.84 | 0 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.29 | 21.27 | -0.18 | 16.06 | -6.97 | -0.06 | -24.08 | 0 | 0 | 0.23 | 21.39 | -0.35 | 28.71 | -0.23 | 28.42 | -0.35 | 28.71 | -9.61 | 0 | -38.79 | 0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.06 | 0
15.264 | 0 | -22.09 | -0.06 | 0 | 0 | 0 | -0.18 | 21.39 | 0.29 | 21.33 | 0.12 | 21.45 | -4.1 | 1.46 | -21.97 | 0 | 0 | 0.23 | 21.33 | -0.41 | 28.59 | -0.23 | 28.59 | -0.29 | 28.59 | -7.32 | 0 | -36.97 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 0
15.296 | -0.06 | -19.34 | -0.06 | -31.58 | -0.06 | 0 | -0.06 | 0.12 | 0.29 | 21.39 | 0.41 | 21.45 | -2.52 | 10.66 | -20.1 | 0 | 0 | 0.18 | 21.27 | -0.35 | 28.48 | -0.23 | 28.59 | -0.23 | 28.65 | -4.34 | -0.06 | -34.63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15.328 | 0.06 | -17.81 | -0.18 | -59.65 | -0.06 | -27.31 | -0.06 | 0 | 0.29 | 21.21 | 0.41 | 21.39 | -1.7 | 16.06 | -18.69 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.65 | -0.23 | 28.48 | -0.23 | 28.54 | -2.87 | -0.06 | -32.23 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0.06 | 0
15.36 | 0 | -15.94 | -0.06 | -44.24 | -0.06 | -62.99 | -0.06 | 0 | 0.18 | 21.21 | 0.23 | 21.39 | -0.76 | 21.68 | -16.7 | 0.06 | 0 | 0.23 | 0.12 | -0.47 | 28.77 | -0.23 | 28.83 | -0.23 | 28.65 | -1.99 | 0 | -30.06 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
15.392 | -0.06 | -14.24 | -0.06 | -40.02 | -0.06 | -45.41 | -0.06 | 0 | 0.12 | 21.33 | 0.53 | 21.39 | -0.41 | 21.33 | -15.41 | -0.06 | 0 | 0.18 | 0.06 | -0.23 | 28.48 | -0.18 | 28.48 | -0.23 | 28.18 | -0.82 | -0.06 | -29.01 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
15.424 | -0.06 | -12.72 | -0.18 | -37.56 | -0.06 | -40.02 | -0.06 | 0 | 0.12 | 21.27 | 0.53 | 21.39 | -0.06 | 21.39 | -13.48 | -0.18 | 0 | 0.12 | 0 | -0.18 | 28.54 | -0.23 | 28.65 | -0.18 | 28.48 | -0.64 | -0.18 | -26.49 | -0.06 | -0.23 | -0.06 | 0 | -0.06 | 0 | 0 | 0
15.456 | 0 | -10.96 | -0.06 | -35.22 | -0.06 | -38.38 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.45 | 0.06 | 21.39 | -11.25 | -0.06 | -73.01 | 0.23 | 0 | -0.29 | 29.3 | -0.18 | 28.89 | -0.29 | 28.36 | -0.23 | 0 | -24.43 | -0.06 | -79.51 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
15.488 | 0.88 | -9.49 | -0.06 | -33.69 | -0.06 | -35.74 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.21 | 0.35 | 21.39 | -9.79 | 0.06 | -54.2 | 0.18 | 0 | -0.23 | 28.42 | -0.23 | 28.48 | -0.18 | 28.42 | 0 | 0 | -22.62 | -0.18 | -68.15 | -0.18 | -61.82 | -0.06 | 0 | 0 | 0
15.52 | 8.67 | -8.32 | 0 | -31.47 | -0.06 | -34.45 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.45 | 0.35 | 21.39 | -7.85 | -0.06 | -32.11 | 0.18 | 0 | -0.23 | 0.29 | -0.18 | 28.65 | -0.23 | 28.59 | 0.12 | 1 | -20.57 | 0 | -48.28 | -0.18 | -76.94 | -0.06 | -79.69 | -0.18 | -1.93
15.552 | 14.18 | -6.91 | 0 | -29.42 | 0 | -31.76 | -0.18 | 0 | 0 | 21.39 | 0.53 | 21.33 | 0.29 | 21.33 | -6.45 | -0.06 | -28.3 | 0.12 | 0 | -0.23 | 0.12 | -0.18 | 20.39 | -0.18 | 28.42 | 0.12 | 10.84 | -18.52 | -0.06 | -44.42 | -0.06 | -46.47 | -0.18 | -71.02 | -0.06 | -81.68
15.584 | 19.98 | -5.39 | -0.06 | -27.01 | 0 | -30.47 | 0 | 0 | -0.06 | 21.27 | 0.53 | 21.45 | 0.35 | 21.33 | -4.63 | 0 | -33.81 | 0.12 | 0 | -0.23 | 0 | -0.12 | 0.18 | -0.23 | 28.42 | 0.47 | 18.4 | -16.88 | 0.06 | -38.32 | -0.06 | -30.18 | -0.18 | -41.25 | 0 | -66.15
15.616 | 21.33 | -3.4 | 0 | -25.02 | -0.06 | -27.6 | -0.06 | 0 | 0.06 | 21.15 | 0.23 | 21.33 | 0.41 | 21.21 | -2.64 | -0.06 | -34.98 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.42 | -14.82 | 0 | -36.62 | -0.18 | -38.15 | -0.18 | -30.94 | -0.06 | -33.17
15.648 | 21.5 | -2.52 | -0.06 | -23.09 | 0 | -25.14 | -0.06 | 0 | 0 | 21.04 | 0.23 | 21.39 | 0.41 | 21.33 | -1.88 | 0.06 | -32.58 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.65 | -12.89 | 0 | -38.15 | -0.18 | -43.13 | -0.18 | -40.2 | -0.06 | -30.18
15.68 | 21.5 | -1.46 | -0.06 | -21.39 | 0 | -23.15 | -0.18 | 0 | -0.18 | 0.64 | 0.29 | 21.39 | 0.47 | 21.39 | -1.17 | -0.06 | -29.77 | 0.18 | 0 | -0.29 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.23 | 28.65 | -10.25 | -0.06 | -37.85 | -0.18 | -39.79 | -0.18 | -42.6 | 0 | -41.13
15.712 | 21.45 | -0.82 | 0 | -18.22 | -0.06 | -20.98 | -0.06 | -36.15 | 0.06 | 0 | 0.18 | 21.39 | 0.35 | 21.27 | -0.82 | -0.06 | -27.54 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.42 | -8.14 | -0.06 | -36.21 | -0.06 | -35.51 | -0.06 | -38.67 | -0.18 | -43.13
15.744 | 21.39 | -0.23 | -0.12 | -16.76 | 0 | -19.81 | -0.18 | -73.83 | -0.18 | 0 | 0.06 | 21.21 | 0.29 | 21.15 | -0.35 | 0 | -27.01 | 0.23 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.42 | -5.98 | -0.06 | -34.04 | 0 | -33.28 | -0.06 | -34.45 | 0 | -38.61
15.776 | 21.45 | -0.06 | 0 | -15.29 | 0 | -17.17 | -0.06 | -46.35 | 0.06 | 0 | 0.18 | 21.21 | 0.47 | 21.33 | -0.29 | 0 | -24.96 | 0.18 | -33.46 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.35 | 28.65 | -4.1 | -0.06 | -31.99 | -0.06 | -33.52 | -0.06 | -33.22 | 0 | -33.4
15.808 | 21.5 | -0.23 | -0.06 | -13.71 | -0.06 | -15.18 | -0.06 | -40.08 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.27 | 0.06 | 0 | -22.74 | 0.06 | -70.49 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.36 | 0.29 | 28.36 | -2.4 | 0.94 | -30.29 | -0.06 | -32.23 | -0.18 | -32.58 | -0.18 | -33.11
15.84 | 21.45 | -0.41 | 0 | -12.31 | -0.06 | -13.89 | -0.06 | -35.16 | -0.06 | 0 | 0.12 | 21.33 | 0.23 | 21.39 | 0.06 | 0.06 | -20.74 | 0.18 | -47.64 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.42 | -1.7 | 10.61 | -28.13 | -0.06 | -30.18 | -0.06 | -31.93 | -0.06 | -32.93
15.872 | 21.45 | -0.41 | 0.06 | -10.02 | -0.06 | -12.31 | -0.06 | -31.99 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.33 | -0.06 | -0.06 | -19.75 | 0.12 | -31.06 | -0.18 | 0 | -0.18 | 0 | 0 | 28.42 | 0.23 | 28.48 | -0.76 | 18.16 | -26.66 | -0.18 | -28.3 | -0.18 | -29.65 | -0.06 | -31.99
15.904 | 21.45 | -0.35 | 5.92 | -8.61 | 0.06 | -10.84 | -0.06 | -32.29 | 0 | 0 | -0.06 | 21.27 | 0.12 | 21.33 | 0.06 | -0.06 | -18.34 | 0.06 | -29.12 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.36 | 0.18 | 28.48 | -0.23 | 28.77 | -24.43 | -0.06 | -27.13 | 0 | -27.6 | 0 | -29.53
15.936 | 21.45 | -0.35 | 13.42 | -7.38 | -0.06 | -9.2 | -0.06 | -32.29 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.33 | 0.06 | 4.92 | -15.76 | 0.12 | -35.57 | -0.18 | -1 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.3 | 0 | 28.65 | -23.15 | -0.18 | -25.31 | 0 | -25.9 | 0 | -26.84
15.968 | 21.5 | -0.35 | 19.04 | -5.92 | -0.06 | -7.62 | -0.06 | -30.53 | 0 | 0 | 0 | 21.33 | 0.06 | 21.39 | 0.23 | 10.84 | -14.24 | 0.12 | -34.75 | -0.18 | -68.91 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.18 | 0 | 28.24 | -21.27 | -0.06 | -23.5 | -0.06 | -25.14 | 0 | -26.37
16 | 21.39 | -0.23 | 21.62 | -4.22 | 0 | -6.33 | -0.06 | -28.71 | -0.06 | 0 | 0 | 21.39 | 0.06 | 20.98 | 0.18 | 18.28 | -12.42 | 0.06 | -31.47 | 0 | -53.44 | -0.06 | -41.43 | -0.18 | 28.54 | 0 | 28.65 | 0.18 | 28.36 | -18.4 | 0 | -22.21 | 0 | -23.44 | 0 | -24.67
16.032 | 21.45 | -0.35 | 21.5 | -2.93 | 2.75 | -4.45 | 0 | -25.72 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.21 | 0.12 | 21.5 | -10.9 | 0.12 | -28.48 | 0 | -44.83 | -0.18 | -62.64 | -0.06 | 0.18 | 0 | 28.59 | 0.18 | 28.71 | -16.58 | -0.18 | -20.45 | -0.06 | -21.45 | 0 | -22.97
16.064 | 21.45 | -0.23 | 21.5 | -1.88 | 8.56 | -3.57 | -0.06 | -24.38 | -0.06 | 0 | 0 | 0.23 | 0.12 | 21.39 | 0.18 | 21.39 | -9.02 | 0.06 | -27.36 | 0 | -44.77 | -0.06 | -45.18 | 0 | 0 | 0 | 28.48 | 0.47 | 28.59 | -14.65 | 0 | -18.87 | -0.06 | -20.1 | -0.06 | -21.39
16.096 | 21.45 | -0.23 | 21.45 | -1.17 | 14.18 | -1.88 | 2.64 | -22.09 | -0.06 | -12.31 | -0.06 | 0.06 | -0.06 | 21.21 | 0.23 | 21.39 | -6.45 | 0.06 | -26.37 | -0.18 | -42.01 | 0 | -46.23 | -0.18 | 0 | 0 | 28.48 | 0.29 | 28.3 | -13.07 | -0.06 | -16.99 | -0.06 | -18.11 | -0.06 | -19.51
16.128 | 21.39 | -0.23 | 21.45 | -0.41 | 19.69 | -1.05 | 8.73 | -19.63 | -0.06 | -76.88 | 0 | 0 | 0.12 | 21.39 | 0.18 | 21.45 | -5.22 | 0 | -24.96 | 0 | -41.19 | 0 | -44.3 | -0.18 | 0 | 0.06 | 28.42 | 0.35 | 28.42 | -10.9 | 2.93 | -14.82 | -0.06 | -16.64 | -0.06 | -18.16
16.16 | 21.45 | -0.23 | 21.39 | -0.18 | 21.5 | -0.47 | 14 | -17.99 | 0 | -56.19 | 0 | 0 | -0.18 | 0.23 | 0.12 | 21.39 | -3.57 | 0 | -22.21 | 0 | -38.09 | -0.06 | -41.13 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.59 | -8.85 | 11.13 | -13.48 | -0.06 | -15 | 1.11 | -16.41
16.192 | 21.45 | -0.18 | 21.45 | -0.29 | 21.5 | -0.06 | 19.86 | -16 | 0 | -42.48 | 0.06 | 0 | -0.06 | 0.06 | 0.12 | 21.27 | -2.46 | 0.06 | -20.74 | 0 | -36.62 | 0 | -39.38 | 0 | 0 | 0 | 28.54 | 0.29 | 28.42 | -6.21 | 21.27 | -12.01 | 2.46 | -13.65 | 10.72 | -14.47
16.224 | 21.39 | -0.18 | 21.39 | -0.23 | 21.5 | 0 | 21.5 | -13.95 | 0.06 | -35.74 | -0.06 | 0 | 0 | 0 | 0.18 | 21.27 | -1.7 | -0.06 | -19.45 | -0.06 | -34.98 | 0 | -37.44 | 0 | 0 | 0 | 28.48 | 0.29 | 28.59 | -4.04 | 28.65 | -10.55 | 10.25 | -11.6 | 18.22 | -13.01
16.256 | 21.39 | -0.23 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | -11.25 | -0.06 | -32.34 | -0.06 | 0 | -0.18 | 0 | 0.23 | 21.33 | -0.94 | -0.06 | -17.23 | -0.06 | -32.99 | -0.06 | -36.45 | -0.18 | 0 | -0.06 | 28.65 | 0.18 | 28.71 | -2.46 | 28.71 | -7.91 | 17.7 | -9.84 | 26.02 | -10.9
16.288 | 21.39 | -0.23 | 21.5 | -0.41 | 21.39 | -0.41 | 21.45 | -9.2 | -0.06 | -32.23 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.45 | -0.76 | 0 | -15.47 | 0 | -30.7 | -0.06 | -34.28 | -0.06 | 0 | 0 | 28.65 | 0.12 | 28.54 | -1.7 | 28.54 | -6.86 | 25.43 | -8.09 | 28.65 | -9.38
16.32 | 21.45 | -0.23 | 21.45 | -0.29 | 21.5 | -0.35 | 21.45 | -7.38 | -0.18 | -32.23 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | -0.06 | 0 | -13.95 | 0 | -29.06 | -0.06 | -32.64 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.48 | -0.76 | 28.48 | -5.33 | 28.95 | -6.15 | 28.65 | -7.38
16.352 | 21.39 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | -4.75 | -0.18 | -30.76 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.18 | 0.06 | -12.36 | 0 | -26.37 | -0.06 | -30 | -0.18 | 0 | -0.06 | 28.65 | 0 | 28.36 | -0.41 | 28.59 | -3.63 | 29.01 | -4.8 | 28.65 | -5.68
16.384 | 10.78 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.45 | -3.16 | -0.06 | -29.01 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 0.06 | -10.49 | -0.06 | -24.84 | -0.06 | -27.36 | 0 | 0 | 0 | 28.65 | 0 | 28.48 | -0.06 | 28.42 | -2.29 | 28.48 | -3.11 | 28.42 | -4.51
16.416 | 0.06 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | -1.7 | -0.06 | -27.13 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.45 | 0 | 0 | -7.97 | -0.06 | -23.15 | -0.06 | -25.72 | -0.18 | 0 | 0 | 0.29 | 0.06 | 28.65 | 0 | 28.71 | -1.29 | 28.65 | -2.4 | 28.42 | -2.75
16.448 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.76 | -0.06 | -25.2 | -0.18 | 0 | 0.06 | 0 | 0.18 | 21.33 | 0.12 | 1.99 | -6.62 | -0.06 | -20.86 | -0.06 | -23.5 | -0.18 | -48.4 | -0.06 | 0.12 | 0 | 28.83 | 0.18 | 28.59 | -0.94 | 28.65 | -0.76 | 28.59 | -2.11
16.48 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -0.53 | -0.06 | -22.97 | -0.06 | -1.17 | 0 | 0 | 0.12 | 21.39 | 0.12 | 9.14 | -4.63 | -0.06 | -18.52 | 0 | -22.09 | -0.06 | -67.68 | -0.18 | 0 | 0 | 28.54 | 0.12 | 28.54 | -0.53 | 28.83 | -0.94 | 28.65 | -1.29
16.512 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -0.18 | 0 | -20.63 | 0 | -74.65 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 14.41 | -3.52 | 0 | -16.47 | -0.06 | -19.28 | -0.06 | -48.46 | -0.06 | 0.06 | 0 | 28.54 | 0.23 | 29.06 | -0.29 | 28.36 | -0.64 | 28.65 | -0.7
16.544 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | 0 | -0.06 | -18.87 | -0.18 | -49.57 | 0 | 0 | 0.06 | 21.39 | 0.18 | 20.04 | -2.17 | -0.06 | -15.06 | 0 | -17.46 | -0.06 | -46.23 | -0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.3 | -0.29 | 28.48 | -0.29 | 28.65 | -0.59
16.576 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | 0.23 | -0.06 | -16.93 | -0.06 | -42.36 | 0 | -24.14 | 0.18 | 21.33 | 0.18 | 21.5 | -1.17 | -0.06 | -13.77 | -0.06 | -15.76 | -0.06 | -43.36 | 0 | 0 | -0.06 | 28.48 | 0.29 | 28.42 | 0.06 | 28.54 | -0.18 | 28.59 | -0.18
16.608 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | 0 | -14.53 | -0.06 | -37.15 | -0.06 | -71.19 | 0.06 | 21.21 | 0.18 | 21.09 | -0.82 | 0 | -12.54 | -0.06 | -14.24 | -0.06 | -41.13 | -0.06 | 0 | -0.06 | 28.65 | 0.29 | 28.24 | 0 | 28.36 | -0.06 | 28.42 | -0.06
16.64 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | 0.29 | -0.06 | -12.77 | -0.18 | -32.23 | -0.06 | -50.51 | 0.12 | 21.39 | 0.18 | 21.45 | -0.53 | 0 | -10.37 | 1.05 | -12.95 | -0.06 | -39.79 | -0.06 | 0 | 0 | 28.71 | 0.41 | 28.36 | 0.12 | 28.42 | 0 | 28.48 | -0.18
16.672 | 0 | 0 | 21.39 | -0.23 | 21.45 | -0.23 | 21.5 | 0.41 | 0 | -9.79 | -0.06 | -32.23 | -0.06 | -30.35 | 0 | 21.21 | 0.23 | 21.39 | -0.41 | -0.06 | -9.14 | 11.43 | -11.13 | -0.06 | -37.91 | -0.06 | 0 | -0.06 | 28.36 | 0.29 | 28.83 | 0.06 | 28.42 | 0 | 28.3 | 0.06
16.704 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | -0.06 | -7.68 | -0.18 | -32.76 | -0.18 | -28.59 | 0.12 | 21.33 | 0.18 | 21.33 | 0.06 | 3.81 | -7.79 | 18.87 | -9.32 | -0.06 | -35.92 | -0.06 | 0 | -0.06 | 0.29 | 0.18 | 28.59 | 0.18 | 28.48 | 0.06 | 28.42 | 0
16.736 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.18 | 21.5 | 0.47 | 3.34 | -5.45 | -0.06 | -31.58 | 0 | -35.74 | -0.06 | 21.39 | 0.12 | 21.5 | -0.18 | 12.36 | -6.86 | 26.43 | -8.32 | -0.06 | -34.16 | -0.18 | 0 | 0 | 0.12 | 0.18 | 28.42 | 0.12 | 28.3 | 0.06 | 28.65 | 0.12
16.768 | 0 | -0.18 | 17.52 | -0.18 | 21.45 | -0.18 | 21.45 | 0.41 | 9.55 | -3.57 | -0.06 | -29.42 | -0.06 | -34.86 | 0.12 | 21.33 | 0.23 | 21.39 | 0.06 | 19.86 | -5.8 | 28.65 | -7.09 | -0.06 | -32.17 | 0 | 0 | 0 | 0 | 0.18 | 28.48 | 0.12 | 28.71 | 0.18 | 28.54 | 0.18
16.8 | -0.53 | -0.06 | 0.12 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | 16.7 | -2.23 | -0.06 | -27.07 | 0.06 | -31.64 | 0 | 21.45 | 0.18 | 21.27 | 0.06 | 27.54 | -4.45 | 28.3 | -6.21 | 2.64 | -29.53 | 0 | 0 | -0.18 | 0 | 0.18 | 28.42 | 0.06 | 28.3 | 0.12 | 28.48 | 0.18
16.832 | -57.25 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.35 | 21.74 | -1.41 | 0.06 | -24.32 | -0.06 | -28.48 | 0.06 | 0.12 | 0.29 | 21.39 | 0.06 | 28.65 | -3.52 | 28.36 | -4.34 | 10.72 | -28.01 | 0 | -0.18 | -0.18 | 0 | 0 | 28.83 | 0.12 | 28.42 | 0.12 | 28.65 | 0.12
16.864 | -51.45 | -0.06 | 0 | -0.12 | 21.45 | -0.18 | 21.45 | 0.29 | 21.45 | -0.18 | -0.18 | -22.33 | 0 | -27.42 | 0.06 | 0 | 0.23 | 21.33 | -0.06 | 28.54 | -2.7 | 28.24 | -3.52 | 18.16 | -26.72 | 0 | -90.24 | -0.18 | 0 | 0.12 | 28.48 | 0.06 | 28.54 | 0.06 | 28.59 | 0.12
16.896 | -40.14 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.39 | -0.23 | -0.06 | -20.74 | -0.06 | -26.31 | -0.06 | 0 | 0.12 | 21.33 | 0.12 | 28.36 | -2.11 | 28.36 | -3.11 | 28.89 | -23.61 | -0.18 | -56.25 | -0.18 | 0 | 0 | 28.59 | 0.06 | 28.18 | 0.12 | 28.48 | 0
16.928 | -39.14 | -0.06 | 0 | -0.18 | 11.19 | -0.18 | 21.45 | 0.23 | 21.45 | 0.18 | -0.06 | -18.87 | -0.06 | -24.32 | 0 | 0 | 0.18 | 21.39 | 0.23 | 28.42 | -1.46 | 28.59 | -2.23 | 28.83 | -21.91 | -0.06 | -48.28 | -0.06 | 0 | 0.06 | 28.83 | 0 | 28.42 | -0.06 | 28.42 | 0.23
16.96 | -36.21 | -0.06 | 0 | -0.06 | 0.06 | -0.18 | 21.45 | 0.18 | 21.39 | 0.18 | 0 | -17.05 | 0.06 | -22.33 | 0.06 | 0 | 0.12 | 21.45 | 0.18 | 28.3 | -1.17 | 28.54 | -1.93 | 28.77 | -19.4 | -0.06 | -44.24 | -0.18 | 0 | 0.06 | 28.59 | 0.12 | 28.42 | 0.06 | 28.42 | 0.12
16.992 | -34.98 | -0.06 | 0 | -0.18 | 0 | -0.18 | 11.48 | 0.18 | 21.33 | 0.35 | -0.06 | -14.71 | 0 | -20.68 | -0.06 | 0 | 0.23 | 21.39 | 0.23 | 28.95 | -0.94 | 28.71 | -1.11 | 28.36 | -17.17 | 0 | -38.56 | -0.18 | 0 | 0 | 28.42 | -0.06 | 28.18 | 0.06 | 28.48 | 0.12
17.024 | -32.93 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.18 | 21.39 | 0.35 | -0.06 | -11.89 | -0.06 | -19.51 | 0.12 | 0 | 0 | 21.45 | 0.18 | 28.36 | -0.7 | 29.01 | -0.76 | 28.65 | -15.76 | -0.06 | -36.45 | -0.06 | 0 | 0 | 0.29 | 0 | 28.48 | 0.06 | 28.48 | 0.06
17.056 | -30 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.27 | 0.59 | 0 | -10.49 | 0 | -17.58 | -0.06 | 0 | 0.18 | 21.27 | 0.23 | 28.42 | -0.59 | 28.65 | -0.59 | 28.59 | -14 | 0 | -38.2 | -0.06 | 0 | 0.12 | 0.12 | 0 | 28.42 | 0.12 | 17.46 | 0.06
17.088 | -28.3 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.53 | 0 | -8.03 | -0.06 | -15.18 | 0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.59 | -0.53 | 28.83 | -0.59 | 28.42 | -12.6 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0 | -0.06 | 23.15 | 0.12 | 0.12 | 0.18
17.12 | -26.31 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.53 | 3.34 | -5.98 | -0.06 | -13.48 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 28.36 | -0.35 | 28.59 | -0.41 | 28.48 | -11.31 | -0.06 | -35.98 | -0.06 | -1.05 | 0 | 0.06 | -0.18 | 0.18 | 0.06 | 0 | 0.18
17.152 | -24.08 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.47 | 9.55 | -3.4 | -0.06 | -12.07 | 0 | 0 | 0.12 | 21.39 | 0.18 | 28.54 | -0.35 | 28.36 | -0.18 | 28.83 | -9.9 | 0.06 | -34.45 | 0 | -86.14 | 0 | 0 | 0.06 | 0 | 0 | 0.06 | 0.06
17.184 | -22.38 | -0.06 | -0.06 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 14.94 | -1.76 | -0.18 | -10.14 | 0 | 0 | 0.18 | 21.39 | 0.12 | 28.71 | -0.29 | 28.54 | -0.18 | 28.65 | -8.32 | 0 | -32.05 | -0.06 | -57.89 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.12
17.216 | -19.86 | -0.06 | -45.7 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 20.45 | -1.41 | 0.06 | -7.91 | -0.06 | 0 | 0.12 | 21.21 | 0.12 | 28.54 | -0.29 | 28.65 | -0.35 | 28.3 | -6.86 | 6.27 | -30 | -0.06 | -48.81 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12
17.248 | -17.52 | -0.06 | -53.15 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.5 | 0.41 | 21.5 | -0.41 | 0 | -6.33 | -0.06 | -29.47 | 0.12 | 21.39 | 0.18 | 28.3 | -0.29 | 28.59 | -0.29 | 28.83 | -5.8 | 14.06 | -27.95 | -0.18 | -44.59 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.12
17.28 | -16.23 | -0.06 | -41.37 | -0.06 | 0 | -0.06 | 0 | 0 | 21.5 | 0.29 | 21.39 | -0.41 | -0.18 | -4.98 | 0 | -63.87 | 0.18 | 21.21 | 0.06 | 28.48 | -0.29 | 28.42 | -0.18 | 28.48 | -4.45 | 21.97 | -26.37 | -0.06 | -38.85 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.23
17.312 | -14.77 | 0 | -39.61 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.23 | 21.39 | 0.23 | 4.69 | -3.11 | 0 | -39.55 | 0.18 | 21.39 | 0.06 | 28.71 | -0.29 | 28.59 | -0.29 | 28.54 | -3.87 | 29.06 | -24.43 | -0.06 | -36.45 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06
17.344 | -12.77 | -0.06 | -38.38 | -0.06 | -0.7 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.39 | 0.29 | 12.42 | -2.17 | 0 | -27.6 | 0.12 | 6.8 | 0.12 | 28.36 | -0.18 | 28.65 | -0.23 | 28.36 | -3.16 | 28.48 | -22.97 | 0 | -37.91 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12
17.376 | -11.25 | -0.06 | -35.22 | -0.06 | -57.83 | -0.06 | 0 | -0.06 | 21.27 | 0.18 | 21.39 | 0.35 | 18.05 | -1.46 | 0 | -30.94 | 0.18 | 0.12 | 0.06 | 28.36 | -0.18 | 28.59 | -0.18 | 28.77 | -2.11 | 28.48 | -20.33 | -0.18 | -37.97 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06
17.408 | -9.61 | -0.06 | -33.52 | -0.06 | -51.97 | -0.06 | -1.46 | -0.18 | 21.39 | 0.12 | 21.39 | 0.41 | 21.56 | -0.53 | 0.06 | -35.22 | 0.12 | 0 | 0.12 | 28.59 | -0.18 | 28.54 | -0.18 | 28.36 | -1.7 | 28.36 | -18.46 | 0 | -36.39 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0
17.44 | -8.09 | -0.06 | -31.23 | -0.06 | -39.96 | -0.06 | -54.14 | -0.06 | 21.15 | 0.18 | 21.39 | 0.41 | 21.39 | -0.47 | 0 | -34.04 | 0.18 | 0.06 | 0.23 | 28.59 | -0.18 | 28.54 | -0.23 | 28.65 | -1.17 | 28.59 | -16.41 | -0.06 | -34.1 | -0.06 | -3.52 | -0.18 | 0 | -0.06 | 0 | 0.06
17.472 | -6.62 | -0.06 | -29.06 | -0.06 | -38.85 | 0 | -51.39 | -0.06 | 21.15 | 0.12 | 21.33 | 0.53 | 21.33 | -0.18 | 0 | -30 | 0.12 | 0 | 0.12 | 28.42 | -0.18 | 28.89 | -0.18 | 28.42 | -0.76 | 28.13 | -14.71 | -0.18 | -32.23 | -0.06 | -88.13 | -0.06 | 0 | -0.06 | -0.29 | 0.12
17.504 | -5.1 | 0 | -27.01 | -0.06 | -36.68 | -0.06 | -40.84 | -0.06 | 21.39 | 0.06 | 21.39 | 0.53 | 21.15 | 0 | 0.12 | -28.07 | 0.12 | 0 | 0.23 | 28.65 | -0.18 | 28.54 | -0.18 | 28.36 | -0.53 | 28.24 | -12.95 | -0.18 | -29.83 | -0.06 | -55.84 | -0.06 | -0.06 | 0 | -85.14 | 0.06
17.536 | -3.63 | -0.06 | -24.61 | -0.06 | -35.16 | -0.06 | -38.91 | -0.06 | 21.39 | 0.12 | 21.39 | 0.35 | 21.39 | -0.06 | -0.06 | -27.25 | 0 | 0 | 0.23 | 28.54 | -0.18 | 13.83 | 0 | 28.36 | -0.59 | 28.59 | -9.9 | -0.06 | -27.83 | 0 | -48.69 | -0.18 | -77.7 | -0.06 | -69.2 | 0
17.568 | -2.29 | 0.41 | -22.33 | -0.06 | -33.17 | -0.06 | -36.86 | -0.06 | 21.33 | -0.06 | 21.39 | 0.35 | 21.33 | 0.12 | 4.92 | -25.72 | 0 | 0 | 0.18 | 28.42 | -0.18 | 0.12 | -0.18 | 28.54 | -0.29 | 28.42 | -7.85 | -0.06 | -26.49 | -0.18 | -44.3 | -0.18 | -70.78 | -0.06 | -39.02 | -0.18
17.6 | -1.41 | 8.03 | -20.39 | 0 | -30.06 | 0 | -35.39 | -0.06 | 21.39 | 0.06 | 21.45 | 0.47 | 21.39 | 0.06 | 10.9 | -23.91 | 0.06 | 0 | 0.12 | 0.76 | -0.18 | 0 | -0.18 | 28.36 | -0.29 | 28.48 | -5.74 | -0.06 | -25.08 | 0 | -38.15 | -0.18 | -41.43 | -0.18 | -31.29 | -0.06
17.632 | -0.7 | 13.83 | -18.69 | -0.06 | -28.36 | -0.06 | -32.46 | -0.06 | 0.12 | 0 | 21.39 | 0.47 | 21.39 | 0.29 | 16.17 | -21.27 | 0.06 | 0 | 0.12 | 0.12 | -0.06 | 0 | 0 | 28.36 | -0.29 | 28.48 | -3.98 | -0.06 | -23.03 | -0.06 | -37.38 | -0.18 | -30.53 | -0.06 | -41.13 | -0.18
17.664 | -0.41 | 19.45 | -16.58 | -0.06 | -26.49 | -0.06 | -30.94 | -0.06 | 0 | -0.06 | 21.45 | 0.29 | 21.21 | 0.29 | 21.74 | -20.1 | 0.06 | 0 | 0.12 | 0 | -0.06 | 0 | -0.18 | 28.65 | -0.18 | 28.13 | -1.99 | 0 | -20.22 | -0.18 | -37.79 | -0.06 | -40.55 | 0.12 | -42.31 | -0.06
17.696 | -0.23 | 21.45 | -14.65 | -0.06 | -23.91 | -0.06 | -28.36 | 0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.18 | 21.39 | -19.04 | 0 | 0 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0.29 | -0.29 | 28.54 | -1.17 | 8.97 | -18.22 | -0.18 | -37.44 | -0.23 | -42.31 | -0.06 | -38.2 | 0
17.728 | -0.29 | 21.5 | -12.95 | -0.06 | -22.62 | -0.06 | -25.84 | 0 | 0 | -0.18 | 21.45 | 0.23 | 21.39 | 0.29 | 21.39 | -17.23 | 0 | 0 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0.12 | -0.29 | 28.36 | -0.94 | 16.11 | -16.23 | 0.06 | -35.98 | -0.18 | -38.44 | -0.18 | -34.86 | -0.06
17.76 | -0.29 | 21.45 | -11.48 | -0.06 | -19.4 | -0.06 | -23.97 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.33 | 0.18 | 21.33 | -14.82 | 0 | -3.98 | 0.18 | 0 | -0.06 | 0 | -0.23 | 0 | -0.29 | 28.54 | -0.23 | 26.43 | -14.71 | 0 | -33.69 | -0.06 | -34.28 | -0.18 | -33.17 | -0.18
17.792 | -0.35 | 21.45 | -10.08 | -0.06 | -17.75 | -0.06 | -22.38 | 0.06 | 0 | -0.06 | 21.45 | 0.18 | 21.39 | 0.18 | 21.33 | -13.13 | 0 | -70.96 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.65 | -0.06 | 28.48 | -12.95 | 0 | -31.47 | -0.06 | -33.34 | -0.18 | -32.76 | 0
17.824 | -0.35 | 21.45 | -8.73 | -0.06 | -15.76 | -0.06 | -20.16 | 0.06 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.23 | 21.27 | -12.01 | -0.06 | -57.25 | 0.23 | 0 | -0.06 | 0 | -0.06 | 0 | -0.29 | 28.54 | 0 | 28.48 | -10.31 | 0 | -29.77 | -0.06 | -33.11 | -0.06 | -31.47 | 0.06
17.856 | -0.47 | 21.45 | -6.97 | -0.06 | -14.12 | 0 | -17.99 | 0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.33 | 0.18 | 21.39 | -10.08 | 0 | -34.22 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0.18 | 28.36 | -8.2 | -0.06 | -28.3 | -0.18 | -31.93 | -0.18 | -29.3 | 0
17.888 | -0.35 | 21.39 | -5.22 | -0.06 | -12.72 | -0.06 | -15.76 | -0.06 | 0 | 0 | 21.27 | 0.06 | 21.39 | 0.18 | 21.33 | -8.32 | 0.06 | -27.6 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.23 | 28.42 | -6.09 | -0.18 | -26.02 | -0.18 | -29.65 | -0.06 | -27.72 | 0.06
17.92 | -0.35 | 21.45 | -4.04 | -0.18 | -11.13 | -0.18 | -14.71 | -0.06 | 0 | 0.06 | 21.5 | 0.18 | 21.39 | 0.23 | 21.45 | -5.63 | -0.06 | -34.22 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.23 | 28.54 | -4.16 | -0.06 | -23.91 | -0.18 | -28.13 | -0.06 | -26.25 | 0
17.952 | -0.18 | 21.39 | -2.46 | -0.06 | -9.49 | -0.06 | -13.24 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.33 | 0.23 | 21.21 | -4.28 | -0.06 | -35.27 | 0.18 | 0 | 0 | -1.41 | -0.06 | 0 | -0.23 | 28.65 | 0.29 | 28.54 | -2.29 | 0 | -22.21 | -0.18 | -26.08 | -0.06 | -25.14 | 0
17.984 | -0.23 | 21.39 | -1.7 | 4.98 | -7.62 | -0.06 | -11.37 | 0 | 0 | 0 | 21.45 | 0.06 | 21.5 | 0.06 | 21.33 | -2.99 | -0.06 | -33.05 | 0.06 | 0 | 0 | -69.96 | -0.18 | 0 | -0.06 | 28.71 | 0.47 | 28.48 | -1.17 | -0.06 | -20.39 | -0.18 | -24.67 | 0 | -23.32 | -0.06
18.016 | -0.18 | 21.39 | -0.76 | 11.13 | -6.5 | 4.16 | -9.61 | -0.06 | 0 | 0 | 0.53 | -0.06 | 21.21 | 0.12 | 21.27 | -2.11 | -0.18 | -29.65 | 0.12 | -7.56 | -0.06 | -60 | -0.06 | 0 | -0.06 | 28.54 | 0.29 | 28.48 | -0.94 | -0.06 | -17.99 | -0.06 | -23.44 | -0.06 | -21.21 | 0
18.048 | -0.35 | 21.45 | -0.18 | 16.41 | -5.51 | 10.49 | -8.14 | -0.06 | -7.32 | 0 | 0.06 | -0.18 | 21.21 | 0.18 | 21.27 | -1.23 | -0.18 | -27.6 | 0.06 | -78.58 | -0.06 | -44.83 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 28.18 | -0.06 | -0.06 | -16.47 | -0.06 | -21.74 | -0.18 | -19.75 | 0
18.08 | -0.29 | 21.45 | -0.18 | 21.74 | -3.52 | 15.7 | -6.97 | -0.06 | -71.9 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | -0.59 | -0.06 | -26.54 | 0 | -56.6 | 0 | -45.35 | -0.06 | 0 | -0.18 | 28.65 | 0.35 | 28.36 | -0.23 | 0 | -14.41 | -0.06 | -20.16 | -0.06 | -17.75 | 0
18.112 | -0.23 | 21.45 | -0.06 | 21.45 | -2.29 | 21.39 | -5.22 | -0.06 | -48.75 | 0 | 0 | -0.18 | 21.33 | 0 | 21.33 | -0.53 | -0.06 | -24.79 | 0.12 | -44.94 | 0 | -42.54 | -0.06 | -0.76 | -0.18 | 0.23 | 0.23 | 28.48 | 0.23 | 0 | -11.72 | -0.06 | -18.05 | -0.06 | -16.29 | -0.06
18.144 | -0.29 | 21.39 | -0.18 | 21.45 | -1.46 | 21.5 | -4.04 | 0 | -42.42 | 0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.21 | -0.18 | 6.68 | -23.15 | 0.06 | -44.06 | 0 | -41.25 | -0.06 | -68.44 | -0.06 | 0.06 | 0.23 | 29.12 | 0.23 | 9.43 | -10.08 | -0.18 | -16.76 | 0 | -14.47 | -0.06
18.176 | -0.35 | 21.45 | -0.29 | 21.45 | -0.59 | 21.5 | -2.52 | 0 | -36.62 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | 0 | 14.18 | -21.68 | 0.06 | -42.13 | -0.06 | -39.14 | 0 | -54.08 | -0.18 | 0 | 0.29 | 28.48 | 0.35 | 16.7 | -8.14 | 7.85 | -14.65 | -0.12 | -13.24 | 0
18.208 | -0.29 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -1.52 | -0.06 | -31.93 | -0.06 | 0 | -0.06 | 0.18 | 0.06 | 21.33 | 0.18 | 19.86 | -19.92 | 0.06 | -40.37 | -0.06 | -37.62 | -0.06 | -45.7 | -0.06 | 0 | 0.23 | 28.65 | 0.29 | 24.43 | -5.92 | 15.12 | -12.89 | 9.32 | -11.13 | 2.29
18.24 | -0.18 | 21.45 | -0.29 | 21.5 | -0.06 | 21.45 | -0.76 | -0.06 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0 | 21.27 | 0.12 | 21.45 | -18.22 | -0.06 | -38.56 | -0.06 | -35.27 | -0.06 | -45.35 | -0.18 | 0 | 0.06 | 28.24 | 0.35 | 28.83 | -3.93 | 23.03 | -11.6 | 16.7 | -9.2 | 9.84
18.272 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -0.41 | 0 | -31.99 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.33 | -16.17 | 0 | -36.97 | 0 | -33.46 | -0.06 | -43.48 | 0 | 0 | 0.12 | 28.48 | 0.29 | 28.54 | -1.99 | 29.01 | -9.2 | 24.38 | -8.14 | 16.99
18.304 | -0.23 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | -0.06 | -31.17 | 0 | 0 | 0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | -14.65 | 0.06 | -34.28 | 0 | -31.7 | -0.18 | -40.9 | 0 | 0 | 0.06 | 28.42 | 0.41 | 28.65 | -1.46 | 28.83 | -7.85 | 28.71 | -6.15 | 24.84
18.336 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | -0.18 | -0.06 | -29.53 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -12.77 | 0 | -33.28 | -0.06 | -29.47 | 0 | -39.08 | -0.06 | 0 | 0.06 | 28.65 | 0.29 | 28.65 | -0.35 | 28.36 | -6.33 | 28.54 | -4.86 | 28.54
18.368 | -0.06 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | -0.29 | -0.06 | -27.42 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -10.96 | 0.06 | -31.11 | 0 | -27.72 | -0.06 | -36.86 | -0.06 | 0 | 0.12 | 28.83 | 0.23 | 28.59 | -0.35 | 28.42 | -4.69 | 28.54 | -3.16 | 28.65
18.4 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.5 | -0.29 | 4.22 | -24.38 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | -9.08 | 0 | -28.36 | -0.06 | -25.14 | -0.06 | -35.1 | -0.06 | 0 | -0.06 | 28.71 | 0.23 | 28.18 | 0.12 | 28.65 | -3.22 | 28.48 | -2.17 | 28.42
18.432 | -0.18 | 21.45 | -0.23 | 21.5 | -0.23 | 21.5 | -0.47 | 10.37 | -21.91 | -0.18 | -11.02 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | -7.15 | 0.18 | -26.9 | 0 | -22.62 | 0 | -32.81 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.42 | 0.18 | 28.36 | -2.23 | 28.65 | -1.05 | 28.42
18.464 | -0.23 | 15.94 | -0.29 | 21.45 | -0.18 | 21.39 | -0.35 | 15.65 | -20.22 | -0.06 | -75.82 | 0 | 0 | -0.06 | 0.12 | 0.29 | 21.33 | -5.39 | 0.88 | -24.26 | -0.06 | -20.98 | -0.06 | -31.88 | 0 | 0 | 0 | 28.48 | 0.06 | 28.36 | 0.29 | 28.42 | -1.05 | 28.71 | -0.76 | 28.59
18.496 | -0.18 | 0.06 | -0.29 | 21.45 | -0.18 | 21.33 | -0.41 | 21.68 | -18.46 | -0.06 | -53.56 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | -3.93 | 10.2 | -22.21 | 0 | -19.04 | -0.06 | -29.24 | 0.06 | 0 | -0.06 | 28.36 | 0.18 | 28.3 | 0.35 | 28.77 | -0.7 | 28.59 | -0.47 | 28.54
18.528 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.35 | 21.39 | -16.7 | 0.06 | -43.71 | -0.06 | 0 | 0.06 | 0 | 0.12 | 21.39 | -2.75 | 15.59 | -20.1 | -0.06 | -17.4 | -0.06 | -26.84 | 0 | -34.75 | 0 | 28.89 | 0.12 | 28.42 | 0.35 | 28.42 | -0.47 | 28.3 | -0.35 | 28.59
18.56 | -0.18 | 0 | -0.23 | 21.45 | -0.18 | 21.5 | -0.35 | 21.56 | -13.89 | -0.06 | -36.62 | 0 | 0 | 0 | 0 | 0.18 | 21.09 | -1.82 | 21.27 | -18.52 | 0 | -15.18 | -0.06 | -25.37 | -0.06 | -85.96 | 0 | 13.77 | 0 | 28.42 | 0.29 | 28.59 | -0.18 | 28.59 | -0.18 | 28.36
18.592 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.56 | -11.95 | -0.06 | -32.34 | -0.18 | 0 | -0.06 | 0 | 0.35 | 21.33 | -1.11 | 21.39 | -17.23 | 0 | -14.06 | -0.06 | -23.56 | -0.06 | -55.96 | -0.18 | 0.12 | 0.12 | 28.42 | 0.35 | 28.54 | -0.06 | 28.3 | -0.12 | 28.42
18.624 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.39 | -10.02 | 0 | -32.23 | 0 | -52.09 | -0.06 | 0 | 0.23 | 21.27 | -0.35 | 21.33 | -14.82 | 2.46 | -12.54 | 0 | -21.09 | 0 | -48.93 | 0 | 0 | 0 | 28.54 | 0.18 | 28.13 | -0.18 | 28.48 | 0 | 28.48
18.656 | -0.06 | 0 | -0.23 | 21.39 | -0.29 | 21.45 | -0.29 | 21.39 | -7.79 | -0.06 | -32.64 | -0.18 | -65.51 | -0.06 | 0.06 | 0.23 | 21.09 | -0.29 | 21.39 | -13.18 | 10.55 | -11.13 | 0.12 | -18.28 | 0 | -41.84 | 0 | 0 | 0 | 28.65 | 0.35 | 28.36 | 0.06 | 28.42 | 0.06 | 28.42
18.688 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.29 | 21.39 | -5.57 | -0.06 | -31.7 | 0 | -45.47 | -0.06 | 0 | 0.18 | 21.15 | -0.12 | 21.27 | -12.07 | 17.81 | -9.38 | 7.5 | -17.05 | 0 | -36.86 | 0.06 | 0 | 0 | 28.83 | 0.18 | 28.59 | 0.12 | 28.3 | 0.12 | 28.59
18.72 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | -3.16 | -0.06 | -29.06 | -0.06 | -40.2 | 0 | 0 | 0.12 | 21.27 | 0.12 | 21.21 | -10.61 | 28.13 | -7.97 | 15.06 | -15.23 | -0.06 | -36.68 | 0.06 | 0 | -0.06 | 28.42 | 0.23 | 28.48 | 0.12 | 28.48 | 0.18 | 28.48
18.752 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -1.52 | -0.06 | -26.84 | 0.06 | -34.45 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 21.27 | -9.49 | 28.59 | -6.86 | 25.31 | -13.71 | -0.06 | -37.74 | 0.06 | 0 | 0.06 | 28.24 | 0.18 | 28.59 | 0.06 | 28.48 | 0.18 | 28.42
18.784 | -0.06 | 0 | -0.23 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | -1.05 | 2.23 | -24.84 | -0.06 | -32.7 | 0 | 0 | 0.12 | 21.39 | 0.06 | 21.21 | -7.79 | 28.36 | -5.74 | 28.65 | -12.36 | -0.06 | -37.38 | -0.06 | 0 | -0.18 | 28.71 | 0.18 | 28.42 | 0.12 | 28.36 | 0.12 | 28.59
18.816 | -0.06 | 0 | -0.23 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | -0.59 | 9.08 | -22.62 | -0.06 | -32.7 | -0.06 | 0 | 0.23 | 21.39 | 0 | 21.33 | -6.33 | 28.48 | -4.86 | 28.65 | -10.61 | -0.06 | -35.27 | -0.06 | 0 | 0.06 | 28.65 | 0.12 | 28.65 | 0.12 | 28.83 | 0.23 | 28.48
18.848 | -0.18 | 0 | -0.18 | 21.5 | -0.18 | 21.5 | -0.18 | 21.39 | -0.06 | 14.47 | -20.86 | -0.06 | -32.76 | -0.18 | 0 | 0.18 | 21.39 | 0.18 | 21.15 | -5.22 | 28.54 | -3.75 | 28.65 | -9.02 | -0.06 | -33.4 | 0 | 0 | 0 | 28.48 | 0.12 | 28.59 | 0.18 | 28.42 | 0.18 | 28.42
18.88 | -0.06 | -0.47 | -0.18 | 0.12 | -0.06 | 21.5 | -0.23 | 21.39 | -0.06 | 20.04 | -18.4 | -0.06 | -31.23 | -0.18 | -32.87 | 0.12 | 21.39 | 0.29 | 21.33 | -4.63 | 28.3 | -2.87 | 28.24 | -8.2 | -0.06 | -31.06 | 0 | 0 | 0.06 | 28.48 | 0 | 28.42 | 0.18 | 28.42 | 0.18 | 28.48
18.912 | -0.18 | -52.5 | -0.18 | 0 | -0.06 | 0.23 | -0.23 | 21.27 | 0.35 | 21.39 | -16.41 | -0.06 | -28.48 | 0 | -71.02 | 0.12 | 21.15 | 0.29 | 21.15 | -3.75 | 28.71 | -2.05 | 28.3 | -6.68 | -0.18 | -29.83 | -0.18 | 0 | -0.18 | 28.71 | 0.12 | 28.48 | 0.12 | 28.42 | 0.06 | 28.48
18.944 | -0.06 | -53.5 | -0.18 | 0 | -0.18 | 0.06 | -0.23 | 21.39 | 0.35 | 21.39 | -14.53 | -0.06 | -25.78 | -0.18 | -46.64 | 0.12 | 21.33 | 0.18 | 21.39 | -2.87 | 28.54 | -1.52 | 29.01 | -5.57 | -0.06 | -27.6 | 0 | 0 | -0.06 | 28.95 | -0.06 | 28.42 | 0.12 | 28.48 | 0.12 | 28.42
18.976 | -0.06 | -40.02 | 0 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -12.77 | -0.18 | -23.85 | 0 | -27.13 | 0.18 | 21.45 | 0.23 | 21.33 | -2.4 | 28.42 | -1.17 | 28.42 | -4.22 | -0.06 | -25.14 | 0 | 0 | 0 | 28.42 | 0.12 | 28.48 | 0.12 | 28.48 | 0.23 | 28.48
19.008 | -0.06 | -39.61 | -0.06 | 0 | -0.06 | 0.06 | -0.23 | 21.39 | 0.35 | 21.33 | -10.72 | 0 | -21.86 | -0.06 | -31.29 | 0.12 | 10.66 | 0.29 | 21.27 | -1.7 | 28.36 | -0.94 | 28.48 | -3.52 | 0 | -24.2 | 0 | -60.76 | 0 | 28.42 | -0.06 | 28.36 | 0.06 | 28.48 | 0.23 | 28.48
19.04 | -0.18 | -36.74 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.47 | 21.39 | -7.68 | 0 | -19.98 | 0.06 | -35.92 | 0.23 | 0.06 | 0.18 | 21.33 | -1.41 | 28.83 | -0.53 | 28.59 | -2.99 | 0.06 | -22.33 | 0 | -66.1 | 0 | 0.12 | 0.06 | 28.3 | 0.06 | 28.59 | 0.06 | 28.54
19.072 | -0.06 | -34.69 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.21 | 0.59 | 21.21 | -5.57 | 0 | -18.16 | -0.06 | -34.86 | 0.23 | 0 | 0.18 | 21.39 | -0.94 | 28.54 | -0.41 | 28.42 | -2.11 | 0 | -20.45 | 0 | -48.52 | -0.06 | 0 | -0.06 | 0.23 | 0 | 28.48 | 0.06 | 28.48
19.104 | -0.06 | -32.81 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.53 | 21.39 | -3.75 | 0.06 | -15.59 | -0.06 | -30.94 | 0.06 | 0 | 0.12 | 21.39 | -0.76 | 28.42 | -0.47 | 28.48 | -1.52 | -0.06 | -17.87 | 0 | -46.29 | -0.06 | 0 | 0 | 0.06 | 0 | 0.12 | 0 | 26.31
19.136 | -0.06 | -30.53 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.47 | 21.39 | -2.05 | 0.06 | -13.54 | -0.06 | -28.07 | -0.06 | 0 | 0.18 | 21.39 | -0.76 | 28.65 | -0.35 | 28.65 | -0.94 | 7.73 | -15.94 | -0.06 | -39.9 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0.12
19.168 | 0 | -28.48 | -0.06 | 0 | 0 | 0 | -0.06 | 21.33 | 0.41 | 21.39 | -1.46 | 2.34 | -11.48 | 0 | -27.01 | 0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.36 | -0.29 | 28.59 | -0.7 | 15.06 | -13.89 | -0.18 | -36.39 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
19.2 | -0.06 | -26.6 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.35 | 21.39 | -0.76 | 8.03 | -9.67 | 0 | -26.02 | 0.12 | 0 | 0.18 | 21.39 | -0.35 | 28.65 | -0.35 | 28.65 | -0.47 | 22.97 | -11.95 | -0.06 | -37.97 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
19.232 | -0.06 | -24.26 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.09 | 0.29 | 21.39 | -0.18 | 15.12 | -7.15 | -0.06 | -23.91 | 0.06 | 0 | 0.23 | 21.33 | -0.35 | 28.65 | -0.23 | 28.83 | -0.41 | 29.06 | -9.84 | 0 | -37.97 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
19.264 | -0.06 | -22.44 | -0.06 | 0 | 0 | 0 | -0.18 | 21.33 | 0.29 | 21.39 | 0.12 | 20.74 | -4.34 | 2.29 | -22.09 | 0 | 0 | 0.23 | 21.39 | -0.47 | 28.65 | -0.29 | 28.48 | -0.29 | 28.65 | -8.09 | 0.06 | -36.86 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
19.296 | -0.06 | -19.86 | -0.06 | -22.85 | 0 | 0 | -0.06 | 0.12 | 0.29 | 21.45 | 0.18 | 21.39 | -2.75 | 9.96 | -19.98 | 0 | 0 | 0.23 | 21.39 | -0.47 | 28.54 | -0.29 | 28.65 | -0.29 | 28.42 | -4.86 | 0 | -34.86 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
19.328 | -0.06 | -17.93 | -0.06 | -62.76 | -0.06 | -14.94 | -0.18 | 0 | 0.23 | 21.39 | 0.12 | 21.33 | -1.88 | 15.29 | -18.52 | 0 | 0 | 0.12 | 21.39 | -0.29 | 28.59 | -0.29 | 28.24 | -0.23 | 28.42 | -3.16 | -0.06 | -32.58 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0 | 0
19.36 | 0 | -16.41 | -0.18 | -46.82 | -0.06 | -63.23 | -0.18 | 0 | 0.12 | 21.39 | 0.29 | 21.39 | -0.94 | 20.98 | -16.82 | 0.06 | 0 | 0.18 | 0.23 | -0.29 | 28.54 | -0.23 | 28.59 | -0.35 | 28.95 | -2.05 | -0.06 | -30.53 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
19.392 | -0.06 | -14.77 | -0.06 | -40.2 | -0.06 | -47.87 | 0 | 0 | 0.12 | 21.45 | 0.41 | 21.39 | -0.76 | 21.45 | -15.7 | -0.06 | 0 | 0.23 | 0.06 | -0.29 | 28.48 | -0.18 | 28.59 | -0.18 | 28.65 | -1.11 | -0.06 | -28.3 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
19.424 | -0.06 | -13.01 | -0.06 | -37.85 | -0.06 | -40.14 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.33 | 0 | 21.39 | -13.48 | 0.06 | 0 | 0.18 | 0 | -0.29 | 28.36 | -0.18 | 28.42 | -0.29 | 28.42 | -0.53 | -0.06 | -26.72 | -0.18 | 0.06 | -0.18 | 0 | -0.06 | 0 | 0 | 0
19.456 | -0.06 | -10.72 | -0.06 | -35.92 | -0.06 | -38.44 | -0.06 | 0 | 0 | 21.39 | 0.53 | 21.15 | 0 | 21.33 | -11.31 | -0.06 | -69.38 | 0.12 | 0 | -0.23 | 28.77 | -0.18 | 28.65 | -0.29 | 28.65 | -0.23 | 0.06 | -24.96 | 0 | -66.51 | -0.06 | 0 | -0.06 | 0 | 0 | 0
19.488 | 0.29 | -9.55 | -0.06 | -34.28 | -0.06 | -35.39 | -0.06 | 0 | -0.06 | 21.45 | 0.41 | 21.33 | 0.29 | 21.33 | -9.61 | -0.06 | -56.72 | 0.23 | 0 | -0.35 | 28.48 | -0.23 | 28.65 | -0.23 | 28.48 | 0 | -0.06 | -23.2 | -0.18 | -72.48 | -0.06 | -52.5 | -0.06 | 0 | -0.06 | 0
19.52 | 7.68 | -8.14 | -0.06 | -31.99 | -0.06 | -34.28 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | 0.29 | 21.33 | -8.09 | -0.06 | -34.63 | 0.18 | 0 | -0.23 | 0.29 | -0.23 | 28.54 | -0.23 | 28.59 | 0.12 | 0.18 | -20.74 | -0.06 | -48.46 | -0.06 | -79.81 | -0.18 | -68.56 | 0 | 0
19.552 | 13.42 | -6.86 | -0.06 | -29.42 | -0.06 | -32.58 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | 0.47 | 21.33 | -6.5 | -0.06 | -27.42 | 0.23 | 0 | -0.29 | 0.12 | -0.18 | 23.32 | -0.18 | 28.36 | 0.23 | 9.96 | -18.52 | 0.06 | -45.88 | 0 | -49.69 | -0.18 | -75.65 | 0 | -68.85
19.584 | 19.1 | -5.39 | -0.06 | -26.9 | -0.06 | -30.35 | -0.06 | 0 | 0 | 21.45 | 0.47 | 21.33 | 0.35 | 21.39 | -4.63 | -0.06 | -32.76 | 0.18 | 0 | -0.23 | 0 | -0.18 | 0.18 | -0.23 | 28.89 | 0.29 | 17.17 | -17.23 | -0.06 | -40.14 | -0.06 | -30.47 | -0.06 | -44.71 | 0 | -71.66
19.616 | 21.39 | -3.69 | 0 | -25.14 | 0 | -28.36 | -0.06 | 0 | 0 | 21.45 | 0.35 | 21.39 | 0.53 | 21.39 | -2.93 | 0 | -35.51 | 0.18 | 0 | -0.29 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.23 | 27.42 | -15.35 | -0.06 | -36.97 | 0 | -36.97 | -0.06 | -29.65 | 0 | -35.86
19.648 | 21.5 | -2.29 | -0.06 | -22.91 | -0.06 | -25.78 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.27 | 0.53 | 21.39 | -2.05 | 0.06 | -32.58 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.3 | 0.29 | 28.77 | -13.3 | -0.06 | -37.56 | -0.29 | -43.3 | -0.06 | -38.85 | -0.18 | -29.53
19.68 | 21.5 | -1.46 | -0.06 | -21.15 | -0.06 | -23.44 | -0.06 | 0 | -0.06 | 5.39 | 0.23 | 21.39 | 0.35 | 21.39 | -1.23 | 0 | -29.42 | 0.18 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.29 | 28.54 | -10.66 | -0.06 | -37.62 | 0 | -39.79 | 0 | -43.13 | 0 | -39.85
19.712 | 21.5 | -0.76 | -0.06 | -19.16 | 0 | -21.45 | -0.06 | -26.25 | -0.06 | 0.06 | 0.23 | 21.33 | 0.29 | 21.39 | -0.64 | -0.06 | -27.6 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.35 | 28.54 | -8.61 | 0.06 | -36.27 | -0.06 | -35.74 | 0 | -39.26 | -0.06 | -42.78
19.744 | 21.39 | -0.29 | -0.06 | -17.23 | -0.06 | -19.51 | -0.06 | -77.11 | -0.06 | 0 | 0.23 | 21.39 | 0.41 | 21.33 | -0.7 | -0.06 | -27.31 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.42 | -6.86 | 0 | -34.75 | -0.18 | -33.75 | 0 | -35.27 | 0 | -39.2
19.776 | 21.5 | -0.23 | -0.06 | -15.35 | 0 | -17.23 | -0.18 | -47.99 | -0.06 | 0 | 0.18 | 21.39 | 0.41 | 21.33 | -0.06 | 0 | -25.37 | 0.06 | -24.79 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.41 | 28.36 | -4.45 | 0 | -32.52 | -0.06 | -32.99 | -0.06 | -33.52 | 0 | -34.22
19.808 | 21.45 | -0.18 | 0 | -13.83 | 0 | -15.59 | -0.18 | -41.02 | -0.18 | 0 | 0.06 | 21.27 | 0.29 | 21.33 | -0.06 | 0.06 | -23.03 | 0.18 | -73.66 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.29 | 28.48 | -2.64 | 0.41 | -30.53 | 0 | -32.23 | -0.06 | -33.17 | 0 | -32.99
19.84 | 21.45 | -0.18 | -0.06 | -12.54 | 0 | -14.3 | -0.18 | -36.45 | -0.18 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | -0.06 | 0.06 | -21.15 | 0.18 | -49.92 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.35 | 28.3 | -1.29 | 10.37 | -27.83 | -0.06 | -30.76 | 0 | -31.64 | -0.18 | -32.64
19.872 | 21.39 | -0.35 | 0 | -10.37 | 0 | -12.07 | -0.06 | -32.05 | -0.06 | 0 | 0.06 | 21.5 | 0.23 | 21.39 | 0.23 | -0.06 | -19.98 | 0.18 | -31.7 | -0.18 | 0 | 0 | 0 | -0.06 | 28.54 | 0.35 | 28.59 | -1.11 | 17.81 | -26.78 | -0.06 | -28.3 | 0 | -30.06 | 0 | -32.23
19.904 | 21.45 | -0.35 | 4.98 | -8.91 | 0 | -11.25 | -0.18 | -32.52 | -0.18 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | 0.23 | -0.06 | -18.46 | 0.06 | -29.42 | 0 | 0 | 0 | 0 | -0.06 | 28.42 | 0.18 | 28.77 | -0.18 | 28.24 | -24.61 | -0.06 | -27.01 | 0 | -28.13 | 0 | -30
19.936 | 21.5 | -0.35 | 12.66 | -7.85 | -0.06 | -9.49 | -0.06 | -31.99 | 0 | 0 | 0 | 21.27 | 0.18 | 21.27 | 0.23 | 4.63 | -16.23 | 0.12 | -35.16 | -0.18 | -0.47 | -0.18 | 0 | -0.18 | 28.36 | 0.23 | 28.42 | -0.18 | 28.65 | -23.79 | -0.06 | -24.96 | -0.18 | -26.13 | 0 | -26.9
19.968 | 21.45 | -0.23 | 18.28 | -5.8 | 0 | -7.97 | -0.06 | -30.94 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.45 | 0.12 | 10.61 | -14.47 | 0.12 | -34.86 | -0.06 | -66.68 | 0 | 0.06 | 0 | 28.48 | 0.12 | 28.42 | 0 | 28.54 | -21.09 | 0 | -23.5 | -0.18 | -24.84 | 0 | -26.25
20 | 21.45 | -0.29 | 21.68 | -4.86 | 0 | -6.39 | -0.06 | -29.06 | 0.06 | 0 | 0 | 21.21 | 0 | 21.45 | 0.29 | 18.05 | -12.54 | 0 | -31.99 | -0.06 | -54.96 | 0 | -38.97 | -0.18 | 28.3 | 0.12 | 28.59 | 0.18 | 28.65 | -18.69 | -0.06 | -22.21 | -0.06 | -23.32 | -0.06 | -25.02
20.032 | 21.5 | -0.23 | 21.45 | -3.16 | 2.4 | -4.75 | -0.06 | -26.49 | -0.06 | 0 | -0.06 | 21.09 | 0.06 | 21.27 | 0.18 | 21.56 | -11.31 | -0.06 | -28.48 | -0.06 | -44.42 | 0 | -63.58 | -0.18 | 0.29 | 0.12 | 28.42 | 0.35 | 28.42 | -16.88 | -0.06 | -20.51 | -0.06 | -21.68 | 0.06 | -23.38
20.064 | 21.5 | -0.29 | 21.45 | -2.05 | 8.32 | -3.87 | -0.06 | -24.43 | 0.06 | 0 | -0.06 | 0.23 | 0 | 21.39 | 0.18 | 21.39 | -9.55 | -0.06 | -26.72 | -0.06 | -45.06 | 0 | -45.24 | -0.18 | 0.06 | 0.06 | 28.36 | 0.35 | 28.48 | -14.94 | 0 | -19.4 | -0.18 | -20.39 | 0.06 | -21.68
20.096 | 21.45 | -0.29 | 21.45 | -1.23 | 13.89 | -2.17 | 0.41 | -21.91 | 0.06 | -3.81 | -0.06 | 0.12 | -0.06 | 21.39 | 0.23 | 21.27 | -7.03 | 0.06 | -25.9 | 0 | -42.42 | 0 | -46.12 | -0.06 | 0 | 0 | 28.42 | 0.29 | 28.42 | -13.54 | -0.06 | -17.29 | -0.06 | -18.22 | 0.06 | -19.98
20.128 | 21.45 | -0.23 | 21.45 | -0.64 | 19.4 | -1 | 8.32 | -19.98 | 0 | -75 | 0 | 0 | 0.06 | 21.39 | 0.35 | 21.33 | -5.39 | 0.06 | -24.84 | -0.18 | -41.31 | 0 | -44.01 | -0.06 | 0 | 0 | 28.89 | 0.35 | 28.48 | -11.07 | 2.7 | -15.53 | -0.06 | -16.93 | 0 | -18.4
20.16 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -0.59 | 13.95 | -17.87 | -0.06 | -58.3 | -0.06 | 0 | 0 | 0.29 | 0.23 | 21.45 | -4.04 | 0 | -22.56 | 0 | -37.91 | -0.06 | -41.13 | -0.18 | 0 | 0.12 | 28.42 | 0.35 | 28.65 | -8.14 | 10.78 | -13.77 | -0.06 | -15.18 | 0.29 | -16.58
20.192 | 21.39 | -0.29 | 21.5 | -0.29 | 21.5 | -0.35 | 19.4 | -15.7 | -0.06 | -41.6 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.33 | -2.52 | 0.06 | -20.74 | -0.06 | -37.03 | -0.06 | -39.79 | 0 | 0 | 0 | 28.71 | 0.47 | 28.42 | -6.33 | 20.74 | -12.01 | 0.18 | -13.18 | 9.96 | -15.65
20.224 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.5 | -14.18 | -0.06 | -35.68 | 0 | 0 | -0.06 | 0 | 0.29 | 21.39 | -2.17 | 0 | -19.22 | -0.18 | -34.98 | 0 | -36.92 | 0 | 0 | 0 | 28.42 | 0.23 | 28.42 | -4.1 | 28.83 | -10.78 | 9.32 | -11.6 | 17.17 | -13.18
20.256 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | -11.48 | -0.06 | -33.17 | -0.06 | 0 | -0.06 | 0.06 | 0.23 | 21.39 | -0.76 | 0 | -17.4 | -0.06 | -32.93 | -0.06 | -35.92 | 0 | 0 | -0.06 | 28.42 | 0.18 | 28.36 | -2.4 | 28.54 | -8.67 | 16.7 | -9.67 | 24.9 | -11.07
20.288 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -9.43 | 0 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.7 | -0.06 | -15.7 | -0.06 | -31.64 | -0.06 | -34.04 | -0.06 | 0 | 0 | 28.65 | 0.06 | 28.59 | -1.7 | 28.71 | -6.86 | 24.26 | -8.09 | 28.54 | -9.49
20.32 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | -0.47 | 21.45 | -7.27 | 0.06 | -32.29 | 0 | 0 | -0.06 | 0 | 0.18 | 21.27 | -0.06 | -0.06 | -14.24 | 0 | -28.95 | -0.06 | -32.17 | 0 | 0 | -0.18 | 28.42 | 0.23 | 28.65 | -0.7 | 28.48 | -5.04 | 28.71 | -6.39 | 28.42 | -8.2
20.352 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.47 | 21.45 | -5.22 | 0.06 | -30.94 | 0 | 0 | -0.18 | 0 | 0.18 | 21.33 | -0.06 | 0.06 | -12.31 | 0 | -27.25 | -0.06 | -29.94 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.48 | -0.47 | 28.59 | -3.69 | 28.65 | -4.75 | 28.42 | -6.56
20.384 | 17.87 | -0.18 | 21.39 | -0.35 | 21.45 | -0.41 | 21.45 | -3.4 | 0 | -29.71 | 0 | 0 | -0.18 | 0 | 0.12 | 21.27 | -0.18 | 0 | -10.66 | 0 | -24.84 | 0 | -27.01 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 28.65 | -0.18 | 28.95 | -2.23 | 28.36 | -3.22 | 28.48 | -4.92
20.416 | 0.06 | -0.06 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | -1.7 | 0.06 | -27.01 | -0.06 | 0 | 0 | 0 | 0.23 | 21.45 | 0.06 | -0.06 | -8.2 | -0.06 | -23.09 | 0 | -25.55 | -0.06 | 0.06 | -0.06 | 0.7 | 0.06 | 28.36 | 0.12 | 28.36 | -1.46 | 28.36 | -1.88 | 28.36 | -3.11
20.448 | 0 | -0.06 | 21.45 | -0.23 | 21.5 | -0.29 | 21.39 | -0.82 | 0.06 | -25.37 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | 0.18 | 1.99 | -6.68 | -0.06 | -21.39 | 0 | -23.5 | 0 | -41.02 | -0.06 | 0.12 | 0.12 | 28.48 | 0.12 | 28.3 | -0.82 | 28.42 | -1.64 | 28.83 | -2.46
20.48 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.59 | 0.06 | -23.2 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 8.85 | -5.04 | -0.06 | -18.52 | -0.06 | -21.74 | -0.06 | -70.96 | 0 | 0 | 0 | 28.65 | 0.35 | 28.48 | -0.53 | 28.59 | -0.94 | 28.3 | -1.46
20.512 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.33 | -0.18 | -0.06 | -21.21 | -0.06 | -76.82 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 14.18 | -3.75 | -0.06 | -16.93 | -0.06 | -19.45 | -0.06 | -48.69 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.65 | -0.41 | 28.65 | -0.47 | 28.42 | -0.7
20.544 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | 0 | 0.06 | -18.87 | -0.06 | -52.68 | 0 | 0 | 0.12 | 21.33 | 0.23 | 20.1 | -2.46 | 0 | -15.23 | -0.06 | -17.23 | -0.06 | -45.59 | 0 | 0 | 0 | 28.48 | 0.29 | 29.06 | -0.18 | 28.48 | -0.35 | 28.42 | -0.64
20.576 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.23 | 21.45 | 0.18 | -0.06 | -16.99 | -0.06 | -42.6 | 0 | -7.91 | 0.06 | 21.33 | 0.23 | 21.39 | -1.46 | 0 | -14.06 | -0.06 | -15.94 | -0.06 | -43.3 | 0 | 0 | 0 | 28.59 | 0.29 | 28.54 | 0 | 28.54 | -0.23 | 28.54 | -0.23
20.608 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.23 | 0 | -14.71 | -0.12 | -36.68 | 0 | -73.66 | 0.06 | 21.33 | 0.18 | 21.39 | -0.76 | -0.06 | -12.54 | 0 | -14.47 | 0.06 | -41.31 | -0.18 | 0 | 0 | 28.36 | 0.29 | 28.48 | 0 | 28.48 | 0 | 28.13 | -0.23
20.64 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.45 | 0.35 | -0.06 | -12.83 | -0.06 | -32.93 | -0.06 | -53.91 | 0.12 | 21.39 | 0.23 | 21.39 | -0.7 | -0.06 | -11.07 | 0.59 | -12.89 | 0 | -40.55 | -0.18 | 0 | -0.18 | 28.54 | 0.29 | 28.65 | 0 | 28.65 | 0.12 | 28.48 | -0.18
20.672 | 0 | 0 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | 0.47 | -0.06 | -10.49 | 0 | -32.46 | 0 | -32.17 | 0 | 21.33 | 0.23 | 21.39 | -0.29 | 0 | -9.2 | 10.9 | -11.31 | 0 | -38.44 | -0.06 | 0 | 0 | 28.3 | 0.23 | 28.65 | 0.06 | 28.54 | 0.12 | 28.77 | 0.12
20.704 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.47 | 0 | -8.09 | -0.06 | -32.58 | 0.06 | -28.07 | 0.06 | 21.39 | 0.18 | 21.33 | -0.18 | 3.16 | -7.91 | 18.16 | -9.84 | 0 | -36.51 | -0.18 | 0 | -0.06 | 0.29 | 0.18 | 28.3 | 0.12 | 28.54 | 0.12 | 28.48 | 0.18
20.736 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.41 | 2.75 | -5.74 | 0.06 | -31.58 | -0.06 | -35.16 | 0.06 | 21.33 | 0.18 | 21.39 | 0.06 | 11.48 | -6.86 | 25.72 | -8.44 | 0 | -34.28 | 0 | 0 | -0.06 | 0.12 | 0.18 | 28.83 | 0.18 | 28.48 | 0 | 28.42 | 0.12
20.768 | 0 | -0.06 | 21.5 | -0.18 | 21.5 | -0.18 | 21.45 | 0.47 | 8.79 | -3.63 | -0.06 | -29.24 | -0.06 | -34.69 | 0.06 | 21.39 | 0.23 | 21.39 | 0.06 | 19.22 | -5.74 | 28.65 | -7.44 | -0.06 | -32.46 | 0 | 0 | 0 | 0 | 0.12 | 28.48 | 0 | 28.24 | 0.18 | 28.42 | 0.12
20.8 | -0.12 | -0.06 | 0.12 | -0.23 | 21.45 | -0.18 | 21.45 | 0.47 | 16.06 | -2.11 | 0 | -27.6 | -0.06 | -31.82 | 0 | 21.39 | 0.18 | 21.15 | 0.12 | 26.6 | -4.45 | 28.65 | -5.92 | 2.4 | -30.24 | 0 | 0 | -0.06 | 0 | 0.06 | 28.54 | 0.18 | 28.65 | 0.18 | 28.48 | 0.06
20.832 | -48.93 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | 21.5 | -1 | -0.06 | -24.96 | -0.06 | -29.01 | -0.06 | 0.23 | 0.18 | 21.39 | 0.18 | 28.54 | -3.69 | 28.65 | -4.63 | 10.2 | -28.59 | -0.18 | 0 | -0.18 | 0 | 0.06 | 28.59 | 0.12 | 28.18 | 0.18 | 28.3 | 0.18
20.864 | -53.85 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.41 | 21.39 | -0.53 | 0 | -22.44 | 0 | -27.77 | 0 | 0 | 0.12 | 21.33 | 0.23 | 28.83 | -2.93 | 28.54 | -3.69 | 17.81 | -26.49 | 0 | -82.56 | -0.18 | 0 | 0.12 | 28.83 | 0.18 | 28.59 | 0.12 | 28.65 | 0.18
20.896 | -39.61 | 0 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.35 | 21.45 | -0.18 | 0 | -20.98 | -0.06 | -26.31 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 28.65 | -2.4 | 28.3 | -3.16 | 28.24 | -24.08 | -0.06 | -59.65 | -0.06 | 0 | 0 | 28.36 | 0 | 28.48 | 0.12 | 28.42 | 0.18
20.928 | -39.44 | 0 | 0 | -0.18 | 13.59 | -0.18 | 21.45 | 0.35 | 21.39 | -0.06 | 0 | -19.22 | 0 | -24.96 | -0.06 | 0 | 0.06 | 21.45 | 0.23 | 28.59 | -1.93 | 28.36 | -2.4 | 28.77 | -21.62 | -0.06 | -47.81 | -0.06 | 0 | 0.06 | 28.42 | 0.12 | 28.59 | 0.18 | 28.42 | 0.18
20.96 | -37.33 | -0.18 | 0 | -0.18 | 0.06 | -0.18 | 21.45 | 0.35 | 21.39 | 0.18 | -0.06 | -17.05 | -0.06 | -22.62 | 0.06 | 0 | 0.18 | 21.45 | 0.06 | 28.54 | -1.29 | 28.71 | -1.93 | 28.89 | -19.63 | -0.18 | -44.77 | 0 | 0 | 0 | 28.59 | 0 | 28.48 | 0.12 | 28.36 | 0.18
20.992 | -34.63 | -0.06 | 0 | -0.06 | 0 | -0.06 | 13.42 | 0.18 | 21.39 | 0.23 | -0.06 | -14.71 | 0 | -20.86 | -0.18 | 0 | 0.12 | 21.33 | 0.18 | 28.83 | -1.05 | 28.48 | -1.23 | 28.77 | -17.93 | -0.06 | -39.67 | 0 | 0 | 0 | 28.3 | -0.06 | 28.42 | 0.12 | 28.36 | 0.23
21.024 | -32.93 | 0 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.18 | 21.45 | 0.35 | 0.06 | -12.83 | 0.06 | -19.92 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 28.54 | -0.76 | 28.42 | -0.94 | 28.42 | -16.52 | -0.06 | -36.68 | -0.06 | 0 | -0.18 | 0.29 | 0 | 28.42 | 0.18 | 28.42 | 0.18
21.056 | -30.94 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | -0.06 | -10.49 | 0.06 | -17.4 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 28.65 | -0.64 | 28.48 | -0.76 | 28.59 | -14.12 | -0.06 | -37.68 | -0.06 | 0 | 0 | 0.12 | 0 | 28.48 | 0.12 | 25.96 | 0.18
21.088 | -28.59 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.06 | 0.12 | 21.39 | 0.41 | -0.06 | -8.32 | -0.06 | -15.7 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 28.59 | -0.41 | 28.77 | -0.53 | 28.54 | -12.83 | 0 | -37.91 | 0 | 0 | 0 | 0 | -0.06 | 28.48 | 0 | 0.18 | 0.23
21.12 | -26.49 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0.35 | 2.7 | -6.56 | 0 | -14 | 0 | 0 | 0.06 | 21.33 | 0.12 | 28.07 | -0.35 | 28.59 | -0.41 | 28.65 | -11.6 | 0 | -36.27 | -0.06 | -0.29 | -0.06 | 0 | 0 | 0.23 | -0.06 | 0 | 0.18
21.152 | -24.08 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 8.79 | -3.63 | -0.06 | -12.07 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.65 | -0.35 | 28.42 | -0.29 | 28.59 | -9.9 | 0 | -34.45 | 0 | -82.85 | 0 | 0 | 0 | 0 | 0 | 0 | 0.18
21.184 | -22.62 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 14.3 | -2.05 | 0 | -10.84 | -0.06 | 0 | 0.12 | 21.33 | 0 | 28.42 | -0.35 | 28.48 | -0.29 | 28.48 | -8.67 | 0 | -32.46 | -0.06 | -58.95 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.23
21.216 | -20.39 | -0.06 | -37.44 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 19.57 | -1.29 | 0.06 | -8.32 | 0.06 | 0 | 0.12 | 21.39 | 0.23 | 28.77 | -0.35 | 28.42 | -0.41 | 28.89 | -7.03 | 5.16 | -30.18 | 0 | -48.17 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12
21.248 | -17.99 | -0.06 | -57.01 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.5 | -0.76 | -0.06 | -6.21 | -0.18 | -25.02 | 0.12 | 21.27 | 0.18 | 28.48 | -0.23 | 28.42 | -0.18 | 28.48 | -5.92 | 13.24 | -28.83 | -0.06 | -45.47 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
21.28 | -15.94 | -0.06 | -43.36 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.45 | -0.47 | -0.06 | -5.16 | -0.06 | -64.98 | 0.18 | 21.39 | 0.18 | 28.77 | -0.23 | 28.42 | -0.29 | 28.48 | -4.75 | 20.98 | -26.25 | 0 | -38.85 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.18
21.312 | -14.41 | -0.06 | -39.96 | -0.18 | 0 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.39 | 0.35 | 3.75 | -3.57 | 0 | -40.49 | 0.06 | 21.45 | 0.18 | 28.65 | -0.23 | 28.42 | -0.23 | 28.65 | -3.69 | 28.71 | -24.79 | -0.18 | -36.86 | -0.06 | 0 | -0.29 | 0 | 0 | 0 | 0.18
21.344 | -13.07 | 0 | -37.68 | -0.06 | -0.23 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.41 | 11.72 | -2.29 | -0.06 | -28.48 | 0.06 | 10.37 | 0.23 | 28.3 | -0.23 | 28.54 | -0.18 | 28.48 | -3.46 | 28.65 | -23.15 | -0.18 | -37.91 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12
21.376 | -11.31 | 0 | -35.22 | -0.06 | -55.61 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.39 | 0.29 | 17.17 | -1.29 | 0 | -31.06 | 0.12 | 0.06 | 0.12 | 28.89 | -0.18 | 28.48 | -0.23 | 28.42 | -2.4 | 28.42 | -20.68 | 0 | -37.79 | -0.06 | 0 | -0.23 | 0 | 0 | 0 | 0.12
21.408 | -9.55 | -0.06 | -32.99 | -0.06 | -52.44 | -0.18 | -0.47 | 0 | 21.39 | 0.23 | 21.45 | 0.35 | 21.74 | -0.94 | 0 | -35.45 | 0.23 | 0 | 0.12 | 28.59 | -0.23 | 28.42 | -0.18 | 28.65 | -1.64 | 28.83 | -19.16 | 0 | -36.21 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06
21.44 | -8.56 | -0.06 | -31.93 | -0.06 | -40.02 | -0.06 | -50.57 | -0.06 | 21.45 | 0.12 | 21.39 | 0.53 | 21.39 | -0.47 | -0.18 | -33.69 | 0.23 | 0 | 0.18 | 28.42 | -0.18 | 28.48 | -0.18 | 28.83 | -1.11 | 28.48 | -16.88 | -0.06 | -34.51 | 0 | -1.29 | -0.18 | 0 | -0.18 | 0 | 0
21.472 | -6.91 | -0.06 | -29.24 | -0.18 | -39.08 | -0.06 | -52.44 | 0 | 21.27 | 0.18 | 21.39 | 0.53 | 21.39 | -0.18 | -0.18 | -30 | 0.06 | 0 | 0.06 | 28.65 | -0.18 | 28.48 | -0.18 | 28.36 | -0.94 | 28.54 | -15 | -0.06 | -31.99 | -0.06 | -89.89 | -0.18 | 0 | -0.06 | 0 | 0
21.504 | -5.39 | 0 | -27.54 | 0 | -36.86 | -0.18 | -41.08 | -0.06 | 21.33 | 0 | 21.33 | 0.53 | 21.39 | -0.06 | -0.06 | -28.3 | 0.18 | 0 | 0.12 | 28.54 | -0.18 | 28.3 | -0.23 | 28.48 | -0.64 | 28.36 | -13.24 | -0.06 | -29.83 | -0.06 | -57.31 | -0.23 | 0 | -0.06 | -76.47 | 0
21.536 | -3.93 | 0 | -24.61 | -0.06 | -34.51 | -0.06 | -39.08 | -0.06 | 21.39 | 0 | 21.39 | 0.53 | 21.39 | 0 | -0.18 | -27.07 | 0.12 | 0 | 0.12 | 28.48 | -0.23 | 20.45 | -0.18 | 28.54 | -0.41 | 28.42 | -10.55 | 0.06 | -27.72 | 0 | -47.99 | -0.06 | -63.52 | -0.06 | -71.9 | 0
21.568 | -2.46 | 2.29 | -23.09 | 0 | -32.58 | -0.06 | -37.33 | -0.06 | 21.39 | 0.06 | 21.5 | 0.47 | 21.39 | 0 | 4.69 | -26.31 | 0.06 | 0 | 0.06 | 28.48 | -0.18 | 0.12 | -0.18 | 27.95 | -0.35 | 28.59 | -8.91 | -0.18 | -26.02 | 0 | -44.77 | -0.06 | -74.12 | 0 | -42.31 | 0.12
21.6 | -1.7 | 7.85 | -20.92 | 0 | -30.12 | -0.06 | -35.68 | -0.06 | 21.39 | -0.06 | 21.39 | 0.53 | 21.33 | 0.12 | 10.72 | -23.5 | 0.06 | 0 | 0.12 | 9.2 | -0.18 | 0 | -0.18 | 28.36 | -0.23 | 28.48 | -6.33 | 0 | -24.26 | -0.06 | -38.67 | -0.18 | -44.65 | -0.18 | -29.59 | -0.06
21.632 | -0.88 | 13.42 | -18.87 | -0.06 | -28.48 | -0.06 | -32.76 | 0.06 | 0.23 | 0.06 | 21.39 | 0.35 | 21.45 | 0.23 | 15.94 | -21.39 | 0.12 | 0 | 0.06 | 0.12 | -0.18 | 0.06 | -0.06 | 28.48 | -0.35 | 28.42 | -4.51 | 0 | -22.79 | 0 | -37.44 | 0 | -29.12 | -0.06 | -38.32 | 0
21.664 | -0.41 | 19.04 | -17.05 | -0.18 | -26.72 | -0.06 | -31.41 | -0.06 | 0.06 | 0 | 21.27 | 0.35 | 21.45 | 0.29 | 21.5 | -20.1 | -0.06 | 0 | 0.12 | 0 | -0.18 | 0 | -0.06 | 28.36 | -0.29 | 28.42 | -2.58 | -0.06 | -20.45 | 0 | -38.2 | -0.18 | -38.91 | -0.06 | -42.72 | 0
21.696 | -0.23 | 21.62 | -14.77 | -0.06 | -23.97 | -0.06 | -28.48 | -0.18 | 0 | 0.12 | 21.33 | 0.29 | 21.15 | 0.29 | 21.39 | -19.22 | 0.12 | 0 | 0.18 | 0.06 | -0.18 | 0 | -0.18 | 0.29 | -0.29 | 28.48 | -1.52 | 8.2 | -18.34 | 0.06 | -37.33 | -0.06 | -42.72 | 0 | -39.32 | 0
21.728 | -0.18 | 21.45 | -13.13 | -0.06 | -22.15 | 0 | -26.43 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.23 | 21.45 | -17.4 | 0.06 | 0 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0.12 | -0.29 | 28.77 | -0.7 | 15.29 | -16.7 | -0.18 | -35.92 | -0.18 | -39.61 | 0 | -35.16 | 0
21.76 | -0.18 | 21.45 | -11.72 | -0.06 | -19.4 | 0.06 | -23.97 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | 0.18 | 21.33 | -14.94 | 0.12 | -1.23 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.29 | 28.42 | -0.47 | 25.78 | -14.36 | -0.18 | -34.1 | -0.18 | -35.74 | 0 | -33.22 | 0
21.792 | -0.35 | 21.45 | -10.55 | -0.06 | -17.7 | -0.06 | -22.33 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.23 | 21.39 | -13.18 | 0.12 | -68.62 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.59 | -0.18 | 28.65 | -12.72 | -0.06 | -31.17 | -0.06 | -33.22 | -0.06 | -32.81 | -0.06
21.824 | -0.53 | 21.5 | -8.85 | -0.18 | -16.17 | -0.06 | -20.27 | -0.06 | 0 | -0.06 | 21.21 | 0 | 21.39 | 0.18 | 21.33 | -11.37 | 0.06 | -58.65 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.35 | 28.54 | 0 | 28.65 | -10.37 | -0.06 | -30.06 | -0.06 | -32.99 | -0.06 | -32.05 | -0.06
21.856 | -0.35 | 21.5 | -7.32 | -0.06 | -14.53 | -0.06 | -17.58 | 0 | 0 | 0.06 | 21.27 | 0.06 | 21.5 | 0.23 | 21.5 | -10.31 | 0 | -35.16 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.83 | 0.23 | 28.65 | -8.56 | 0 | -28.01 | -0.18 | -31.99 | -0.18 | -30.29 | -0.06
21.888 | -0.29 | 21.45 | -5.74 | -0.18 | -13.48 | -0.06 | -16.41 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.33 | 0.12 | 21.39 | -8.32 | 0.06 | -27.31 | 0.06 | 0 | 0 | 0 | -0.06 | 0 | -0.23 | 28.42 | 0.23 | 28.36 | -6.09 | 0 | -26.13 | -0.18 | -29.77 | 0 | -28.07 | 0
21.92 | -0.35 | 21.45 | -4.45 | -0.06 | -10.96 | -0.06 | -14.47 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.23 | 21.33 | -5.92 | 0.06 | -34.63 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.23 | 28.48 | 0.35 | 28.42 | -4.16 | -0.06 | -24.38 | -0.06 | -27.6 | -0.06 | -26.25 | 0
21.952 | -0.23 | 21.45 | -2.75 | -0.18 | -9.55 | 0 | -13.01 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.27 | 0.12 | 21.15 | -4.45 | 0 | -35.22 | 0.06 | 0 | -0.06 | -0.41 | -0.18 | 0 | -0.18 | 28.3 | 0.23 | 28.42 | -2.64 | -0.18 | -22.21 | -0.06 | -26.72 | 0 | -25.25 | -0.06
21.984 | -0.18 | 21.39 | -1.76 | 4.45 | -8.32 | 0 | -11.54 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.45 | 0.18 | 21.15 | -3.22 | -0.06 | -33.22 | 0.12 | 0 | -0.06 | -66.15 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.65 | -1.46 | 0 | -20.74 | -0.18 | -24.9 | -0.06 | -23.09 | -0.06
22.016 | -0.18 | 21.39 | -0.76 | 10.72 | -6.97 | 3.34 | -9.49 | -0.06 | 0 | -0.18 | 10.72 | -0.06 | 21.39 | 0.18 | 21.39 | -2.11 | 0.06 | -29.47 | 0.12 | -2.93 | 0 | -61.7 | -0.18 | 0 | -0.06 | 28.48 | 0.29 | 28.65 | -0.59 | -0.06 | -18.4 | -0.06 | -23.79 | -0.18 | -21.86 | 0
22.048 | -0.35 | 21.39 | -0.23 | 15.94 | -5.33 | 9.55 | -8.14 | -0.06 | -0.82 | -0.06 | 0.06 | 0 | 21.39 | 0.12 | 21.33 | -1.46 | -0.06 | -27.19 | 0.18 | -76.82 | 0 | -45.47 | 0 | 0 | 0 | 28.48 | 0.35 | 28.42 | -0.41 | -0.06 | -16.58 | 0 | -21.62 | -0.18 | -19.92 | -0.06
22.08 | -0.35 | 21.45 | -0.23 | 21.5 | -3.63 | 14.88 | -6.56 | 0 | -75.47 | -0.06 | 0 | -0.06 | 21.15 | 0.12 | 21.15 | -0.7 | 0 | -26.84 | 0.06 | -58.19 | -0.06 | -45.7 | -0.06 | 0 | -0.18 | 28.65 | 0.41 | 28.65 | -0.18 | -0.06 | -14.36 | -0.18 | -20.22 | -0.06 | -18.46 | -0.06
22.112 | -0.35 | 21.45 | -0.06 | 21.5 | -2.29 | 20.63 | -5.22 | 0.06 | -51.33 | 0 | 0 | -0.06 | 21.27 | 0.06 | 21.27 | -0.47 | 1.99 | -25.37 | 0.12 | -44.47 | 0 | -43.3 | -0.06 | -0.12 | -0.18 | 0.29 | 0.29 | 28.42 | 0.06 | -0.06 | -12.01 | -0.18 | -18.22 | -0.06 | -16.76 | -0.06
22.144 | -0.23 | 21.45 | -0.06 | 21.5 | -1.46 | 21.5 | -4.1 | 0 | -42.48 | 0 | 0 | -0.06 | 21.33 | 0 | 21.33 | -0.35 | 6.09 | -22.97 | 0 | -44.77 | -0.06 | -40.9 | -0.06 | -64.57 | -0.18 | 0.12 | 0.23 | 28.59 | 0.18 | 8.44 | -9.96 | -0.18 | -16.88 | -0.18 | -14.82 | -0.06
22.176 | -0.23 | 21.45 | -0.18 | 21.45 | -0.59 | 21.5 | -2.17 | 0.12 | -37.21 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | -0.06 | 13.71 | -21.21 | 0.06 | -42.01 | 0 | -39.61 | -0.06 | -55.49 | -0.18 | 0 | 0.18 | 28.59 | 0.12 | 15.76 | -8.09 | 6.33 | -15.29 | -0.06 | -13.36 | -0.06
22.208 | -0.23 | 21.45 | -0.47 | 21.45 | -0.41 | 21.5 | -1.29 | 0 | -32.46 | -0.06 | 0 | -0.06 | 0.18 | 0.06 | 21.33 | -0.06 | 19.28 | -19.81 | 0.06 | -40.78 | -0.06 | -37.44 | -0.06 | -44.88 | -0.06 | 0.06 | 0.12 | 28.42 | 0.35 | 23.38 | -6.21 | 14.12 | -13.3 | 8.5 | -11.48 | -0.06
22.24 | -0.18 | 21.45 | -0.41 | 21.5 | -0.06 | 21.5 | -0.64 | -0.06 | -32.64 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 21.39 | -18.4 | 0.18 | -38.38 | -0.06 | -35.22 | -0.06 | -45.12 | -0.18 | 0 | 0.18 | 28.65 | 0.29 | 28.65 | -3.87 | 21.86 | -11.6 | 15.76 | -9.08 | 9.26
22.272 | -0.18 | 21.45 | -0.41 | 21.45 | -0.06 | 21.45 | -0.53 | 0 | -32.46 | 0.06 | 0 | -0.06 | 0 | -0.06 | 21.15 | 0.06 | 21.39 | -16.47 | 0 | -36.15 | -0.06 | -33.46 | -0.06 | -43.3 | -0.18 | 0 | 0.06 | 28.42 | 0.35 | 28.42 | -2.23 | 28.89 | -9.9 | 23.38 | -8.03 | 16.47
22.304 | -0.18 | 21.39 | -0.41 | 21.45 | -0.29 | 21.39 | -0.29 | 0 | -31.29 | -0.12 | 0 | -0.18 | 0 | -0.18 | 21.45 | 0.12 | 21.33 | -14.36 | 0.06 | -34.28 | 0.12 | -31.29 | -0.06 | -40.84 | -0.06 | 0 | 0.12 | 28.59 | 0.35 | 28.71 | -1.29 | 28.42 | -8.14 | 28.83 | -6.33 | 24.08
22.336 | -0.06 | 21.39 | -0.29 | 21.45 | -0.06 | 21.45 | -0.29 | 0 | -29.53 | 0 | 0 | 0 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | -13.01 | 0.12 | -32.64 | 0 | -29.77 | 0.06 | -39.08 | -0.06 | 0 | 0.12 | 28.59 | 0.35 | 28.65 | -0.7 | 28.65 | -6.33 | 28.59 | -4.63 | 28.77
22.368 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 0 | -27.31 | 0 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.18 | 21.39 | -11.6 | 0.12 | -31 | -0.06 | -27.36 | 0 | -37.15 | -0.06 | 0 | 0.06 | 28.59 | 0.23 | 28.48 | -0.47 | 28.59 | -4.75 | 28.54 | -3.57 | 28.42
22.4 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.41 | 3.34 | -24.43 | 0 | 0 | -0.06 | 0 | -0.06 | 21.15 | 0.18 | 21.33 | -9.02 | 0.12 | -29.3 | 0 | -24.9 | -0.06 | -35.86 | -0.06 | 0 | 0.06 | 28.3 | 0.18 | 28.65 | -0.35 | 28.71 | -3.28 | 28.59 | -2.29 | 28.65
22.432 | -0.06 | 21.45 | -0.35 | 21.45 | -0.29 | 21.39 | -0.35 | 9.73 | -22.74 | 0 | -3.22 | -0.06 | 0 | -0.06 | 21.27 | 0.23 | 21.39 | -7.03 | 0.06 | -27.01 | 0 | -22.68 | -0.06 | -33.52 | 0 | 0 | 0 | 28.48 | 0.23 | 28.42 | 0.06 | 28.48 | -2.17 | 28.48 | -1.41 | 28.65
22.464 | -0.12 | 17.64 | -0.23 | 21.45 | -0.23 | 21.45 | -0.41 | 14.88 | -20.74 | -0.06 | -71.19 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.33 | -5.45 | 2.23 | -24.43 | 0.12 | -21.5 | -0.06 | -31.82 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.83 | 0.06 | 28.95 | -1.46 | 28.42 | -0.64 | 28.77
22.496 | -0.18 | 0.06 | -0.29 | 21.45 | -0.23 | 21.45 | -0.29 | 21.74 | -19.16 | -0.06 | -56.96 | 0 | 0 | -0.06 | 0 | 0.18 | 21.33 | -3.98 | 9.55 | -22.79 | 0 | -19.45 | 0.06 | -29.3 | 0 | 0 | 0.06 | 28.59 | 0.18 | 28.18 | 0.18 | 28.42 | -0.7 | 28.36 | -0.53 | 28.59
22.528 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | -17.05 | 0.12 | -43.48 | 0 | 0 | -0.18 | 0 | 0.18 | 21.39 | -2.75 | 14.88 | -20.74 | 0 | -17.29 | 0 | -27.01 | 0 | -22.85 | 0 | 28.59 | 0.06 | 28.24 | 0.35 | 28.42 | -0.7 | 28.59 | -0.41 | 28.65
22.56 | -0.18 | 0.06 | -0.23 | 21.45 | -0.23 | 21.5 | -0.35 | 21.39 | -14.41 | 0.06 | -36.74 | -0.18 | 0 | -0.06 | 0 | 0.29 | 21.15 | -1.46 | 20.74 | -19.04 | 0 | -15.23 | 0 | -25.84 | -0.06 | -90.94 | 0 | 20.39 | 0.06 | 28.3 | 0.35 | 28.54 | -0.41 | 28.54 | -0.18 | 28.48
22.592 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -12.48 | -0.18 | -31.99 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.94 | 21.56 | -17.23 | 0 | -13.83 | 0 | -23.32 | 0 | -58.89 | 0.12 | 0.18 | 0 | 28.48 | 0.23 | 28.42 | -0.23 | 28.48 | 0 | 28.48
22.624 | -0.18 | 0 | -0.23 | 21.39 | -0.29 | 21.5 | -0.23 | 21.39 | -10.61 | -0.06 | -32.87 | -0.06 | -42.89 | -0.06 | 0 | 0.06 | 21.33 | -0.64 | 21.39 | -14.82 | 0 | -11.95 | 0.06 | -20.92 | -0.18 | -49.04 | 0 | 0 | 0.18 | 28.48 | 0.35 | 28.89 | -0.06 | 28.65 | 0.06 | 28.24
22.656 | -0.06 | 0 | -0.29 | 21.45 | -0.29 | 21.45 | -0.29 | 21.33 | -8.2 | 0 | -32.76 | 0 | -68.32 | -0.06 | 0 | 0.29 | 21.15 | -0.23 | 21.33 | -13.54 | 9.43 | -10.66 | 0 | -18.75 | -0.06 | -43.36 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.48 | 0 | 28.36 | 0.12 | 28.83
22.688 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | -5.92 | 0.12 | -31.99 | 0 | -47.17 | -0.06 | 0 | 0.18 | 21.27 | -0.18 | 21.39 | -12.48 | 16.7 | -9.67 | 6.39 | -17.05 | -0.06 | -37.38 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.42 | -0.06 | 28.48 | 0.12 | 28.3
22.72 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -3.4 | 0 | -29.94 | -0.06 | -41.66 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 21.33 | -11.07 | 26.84 | -7.85 | 14.18 | -15.29 | -0.06 | -37.38 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.48 | 0.23 | 28.54 | 0.18 | 28.36
22.752 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | -1.7 | -0.06 | -26.78 | -0.06 | -35.33 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 21.27 | -9.32 | 28.65 | -7.03 | 24.38 | -13.83 | 0 | -37.44 | 0.12 | 0 | 0.12 | 28.48 | 0.06 | 28.48 | 0.23 | 28.42 | 0.18 | 28.42
22.784 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.39 | -0.23 | 21.45 | -1.64 | 2.75 | -25.25 | -0.18 | -32.29 | -0.06 | 0 | 0.23 | 21.39 | 0.12 | 21.21 | -7.62 | 28.65 | -5.51 | 28.71 | -12.48 | 0.06 | -37.09 | 0 | 0 | -0.18 | 28.54 | 0.06 | 28.59 | 0.18 | 28.65 | 0.23 | 28.48
22.816 | -0.18 | 0 | -0.18 | 21.45 | -0.06 | 21.5 | -0.23 | 21.39 | -0.18 | 8.85 | -23.61 | -0.18 | -32.34 | -0.06 | 0 | 0.12 | 21.21 | 0.23 | 21.39 | -6.86 | 28.59 | -4.45 | 28.36 | -10.49 | -0.06 | -34.98 | -0.06 | 0 | 0.06 | 28.42 | 0.12 | 28.71 | 0.18 | 28.3 | 0.12 | 28.48
22.848 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | -0.23 | 14.24 | -21.27 | -0.18 | -32.46 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 21.33 | -5.33 | 28.77 | -3.75 | 28.48 | -8.91 | 0 | -33.4 | 0 | 0 | 0 | 28.42 | 0 | 28.48 | 0.06 | 28.65 | 0.29 | 28.48
22.88 | 0 | -0.06 | -0.06 | 0.23 | -0.06 | 21.45 | -0.23 | 21.39 | 0.23 | 19.57 | -18.75 | -0.18 | -31.23 | -0.06 | -29.77 | 0.06 | 21.45 | 0.18 | 21.39 | -4.39 | 28.42 | -2.87 | 28.48 | -7.79 | 0 | -31.52 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.65 | 0.18 | 28.54 | 0.18 | 28.18
22.912 | -0.06 | -49.04 | -0.18 | 0 | -0.18 | 0.64 | -0.18 | 21.39 | 0.23 | 21.45 | -16.7 | -0.06 | -28.71 | -0.06 | -71.19 | 0.18 | 21.39 | 0.23 | 21.39 | -3.69 | 28.48 | -2.23 | 28.42 | -6.62 | 0 | -29.77 | -0.18 | 0 | 0 | 28.77 | 0.06 | 28.3 | 0.12 | 28.65 | 0.12 | 28.71
22.944 | 0 | -55.49 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.45 | 0.29 | 21.39 | -14.71 | -0.06 | -25.9 | 0 | -48.81 | 0.12 | 21.33 | 0.29 | 21.39 | -2.99 | 28.65 | -1.7 | 28.48 | -5.39 | -0.06 | -28.13 | 0 | 0 | -0.06 | 28.48 | 0 | 28.48 | 0.12 | 28.48 | 0.18 | 28.42
22.976 | -0.06 | -39.32 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -12.6 | -0.06 | -23.97 | -0.06 | -28.01 | 0.18 | 21.15 | 0.23 | 21.33 | -2.46 | 28.65 | -1.41 | 28.48 | -4.16 | 0 | -26.13 | 0 | 0.06 | -0.18 | 28.48 | -0.06 | 28.48 | 0.18 | 28.42 | 0.12 | 28.48
23.008 | -0.18 | -39.44 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.45 | -11.13 | -0.06 | -22.21 | -0.06 | -30.53 | 0.18 | 15.7 | 0.12 | 21.15 | -1.82 | 28.65 | -1.11 | 28.83 | -3.63 | 0 | -24.32 | -0.06 | -55.84 | 0.06 | 28.13 | 0.06 | 28.13 | 0.12 | 28.48 | 0.23 | 28.54
23.04 | -0.06 | -37.85 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.21 | 0.35 | 21.39 | -8.09 | -0.06 | -20.22 | 0 | -35.45 | 0.18 | 0.12 | 0.12 | 21.33 | -1.35 | 28.65 | -0.76 | 28.48 | -2.7 | 0.06 | -22.38 | -0.06 | -67.56 | -0.06 | 0.18 | -0.18 | 28.42 | 0.12 | 28.48 | 0.12 | 28.42
23.072 | -0.06 | -35.16 | -0.06 | 0 | 0 | 0 | -0.06 | 21.33 | 0.47 | 21.39 | -5.92 | -0.06 | -18.34 | 0.06 | -34.45 | 0.06 | 0 | 0.12 | 21.33 | -1.11 | 28.65 | -0.41 | 28.65 | -2.05 | 0.06 | -20.22 | -0.06 | -49.34 | 0 | 0 | 0 | 0.23 | 0.06 | 28.42 | 0.06 | 28.48
23.104 | -0.18 | -32.7 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.53 | 21.39 | -4.1 | -0.06 | -15.82 | -0.06 | -31.58 | 0.06 | 0 | 0.18 | 21.39 | -0.7 | 28.42 | -0.29 | 28.65 | -1.7 | 0 | -18.52 | -0.06 | -45.94 | 0 | 0 | -0.06 | 0.06 | 0 | 0.23 | 0.12 | 28.42
23.136 | -0.06 | -30.29 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.45 | 0.53 | 21.39 | -1.93 | 0 | -13.89 | 0.06 | -28.54 | 0 | 0 | 0.29 | 21.39 | -0.64 | 28.59 | -0.23 | 28.83 | -0.94 | 6.33 | -15.53 | -0.06 | -39.85 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.12 | 0.12
23.168 | -0.06 | -28.48 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.64 | 0.12 | -11.66 | -0.06 | -27.54 | 0.18 | 0 | 0.23 | 21.39 | -0.64 | 28.59 | -0.29 | 28.65 | -0.59 | 14.12 | -14.12 | 0 | -35.86 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.12 | 0
23.2 | -0.06 | -26.72 | -0.06 | 0 | 0 | 0 | -0.18 | 21.33 | 0.47 | 21.39 | -0.76 | 8.03 | -9.84 | 0 | -25.9 | -0.06 | 0 | 0.23 | 21.33 | -0.59 | 28.59 | -0.23 | 28.59 | -0.53 | 21.74 | -12.25 | 0 | -38.09 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
23.232 | -0.06 | -24.43 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -0.18 | 14.77 | -7.79 | 0 | -24.14 | -0.06 | 0 | 0.29 | 21.39 | -0.47 | 28.89 | -0.23 | 28.42 | -0.35 | 29.18 | -10.2 | 0 | -38.79 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.12 | 0
23.264 | 0 | -22.68 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.33 | 0.12 | 20.39 | -4.75 | 2.11 | -22.33 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.54 | -0.29 | 28.71 | -0.35 | 28.71 | -8.03 | -0.06 | -37.44 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
23.296 | -0.06 | -19.75 | -0.06 | -18.4 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.39 | 0.29 | 21.39 | -2.93 | 9.55 | -20.16 | 0 | 0 | 0.23 | 21.39 | -0.29 | 28.65 | -0.29 | 28.65 | -0.18 | 28.42 | -5.1 | -0.06 | -34.63 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
23.328 | -0.06 | -17.58 | -0.06 | -64.28 | -0.06 | -6.97 | -0.18 | 0.06 | 0.18 | 21.33 | 0.23 | 21.45 | -1.7 | 14.77 | -18.93 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | 28.42 | -0.18 | 28.65 | -0.41 | 28.3 | -3.34 | 0 | -33.34 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
23.36 | -0.06 | -16.29 | -0.18 | -47.11 | -0.06 | -60.71 | -0.18 | 0 | 0.23 | 21.39 | 0.47 | 21.39 | -1 | 20.45 | -16.88 | 0.06 | 0 | 0.23 | 0.59 | -0.23 | 28.54 | -0.23 | 28.42 | -0.35 | 28.48 | -1.88 | 0 | -30.7 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
23.392 | -0.06 | -14.82 | -0.06 | -40.02 | -0.06 | -48.75 | -0.18 | 0.06 | 0.12 | 21.39 | 0.53 | 21.39 | -0.7 | 21.39 | -15.59 | -0.06 | 0 | 0.18 | 0.12 | -0.41 | 28.54 | -0.23 | 28.65 | -0.29 | 28.36 | -1.23 | 0 | -28.59 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
23.424 | -0.06 | -12.95 | -0.18 | -38.32 | -0.06 | -41.19 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | 21.39 | 0.06 | 21.27 | -13.77 | -0.06 | 0.06 | 0.12 | 0 | -0.35 | 28.48 | -0.23 | 28.54 | -0.23 | 28.48 | -0.64 | 0 | -27.07 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
23.456 | 0 | -11.31 | -0.18 | -35.68 | 0 | -38.79 | -0.06 | 0 | 0.06 | 21.27 | 0.53 | 21.39 | 0.12 | 21.33 | -11.54 | -0.06 | -63.93 | 0.18 | 0 | -0.23 | 28.54 | -0.18 | 28.65 | -0.23 | 28.59 | -0.06 | -0.06 | -24.96 | -0.06 | -54.67 | -0.06 | 0 | -0.06 | 0 | 0 | 0
23.488 | 0.12 | -9.55 | -0.06 | -34.81 | -0.06 | -36.33 | -0.18 | 0 | 0.06 | 21.21 | 0.41 | 21.39 | 0.29 | 21.39 | -9.61 | -0.06 | -59.65 | 0.18 | 0 | -0.35 | 28.65 | -0.23 | 28.48 | -0.23 | 28.36 | 0 | 0 | -23.5 | -0.06 | -76.47 | -0.06 | -43.3 | -0.06 | 0.06 | 0 | 0
23.52 | 7.27 | -8.14 | -0.06 | -31.88 | -0.06 | -34.45 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.39 | 0.47 | 21.39 | -8.2 | -0.06 | -35.45 | 0.18 | 0 | -0.23 | 3.87 | -0.18 | 28.48 | -0.23 | 28.36 | 0.18 | 1.99 | -21.5 | -0.06 | -47.87 | -0.06 | -84.2 | -0.06 | -53.97 | -0.06 | 0.06
23.552 | 13.01 | -6.97 | -0.06 | -29.71 | -0.06 | -32.93 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | 0.41 | 21.33 | -6.5 | -0.06 | -27.6 | 0.12 | 0 | -0.29 | 0.12 | -0.18 | 28.65 | -0.18 | 28.3 | 0.29 | 9.02 | -19.45 | 0 | -46.47 | -0.06 | -51.68 | -0.06 | -78.46 | 0 | -63.93
23.584 | 18.52 | -5.45 | -0.06 | -27.83 | 0 | -29.83 | -0.06 | 0 | 0 | 21.33 | 0.41 | 21.39 | 0.41 | 21.45 | -4.63 | -0.06 | -32.05 | 0.18 | 0 | -0.23 | 0 | -0.23 | 0.18 | -0.18 | 28.65 | 0.29 | 16.11 | -17.23 | 0 | -40.37 | -0.18 | -31.88 | -0.06 | -49.51 | 0 | -73.6
23.616 | 21.39 | -3.4 | -0.18 | -26.66 | -0.18 | -28.18 | -0.18 | 0 | 0 | 21.45 | 0.41 | 21.39 | 0.53 | 21.39 | -2.93 | -0.18 | -35.57 | 0.23 | 0 | -0.18 | 0 | -0.23 | 0 | -0.18 | 28.42 | 0.41 | 26.66 | -15.06 | 0 | -36.86 | 0 | -36.45 | -0.18 | -28.83 | 0.06 | -36.8
23.648 | 21.56 | -2.29 | 0 | -22.91 | -0.06 | -26.25 | -0.06 | 0 | 0.06 | 21.33 | 0.35 | 21.39 | 0.47 | 21.39 | -2.23 | -0.06 | -33.17 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.41 | 28.54 | -13.83 | 0 | -38.09 | -0.18 | -43.3 | 0 | -36.21 | -0.06 | -29.83
23.68 | 21.5 | -1.7 | -0.06 | -20.98 | -0.06 | -24.08 | -0.18 | 0 | 0.06 | 10.61 | 0.29 | 21.45 | 0.41 | 21.33 | -0.76 | 0 | -29.65 | 0.23 | 0 | -0.23 | 0 | -0.18 | 0.06 | -0.18 | 28.36 | 0.29 | 28.65 | -10.78 | 0 | -38.15 | 0 | -40.61 | -0.06 | -42.31 | 0 | -39.32
23.712 | 21.45 | -0.64 | -0.06 | -19.04 | -0.06 | -22.09 | -0.06 | -14.47 | 0.06 | 0.06 | 0.29 | 21.39 | 0.47 | 21.33 | -0.82 | -0.06 | -27.77 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.83 | 0.23 | 28.42 | -8.73 | 0 | -37.21 | 0 | -35.92 | 0 | -40.31 | -0.06 | -43.13
23.744 | 21.45 | -0.18 | -0.06 | -16.93 | -0.06 | -19.69 | -0.06 | -76.94 | 0.06 | 0 | 0.23 | 21.39 | 0.41 | 21.39 | -0.59 | -0.06 | -27.25 | 0.18 | 0 | -0.23 | 0 | -0.06 | 0 | -0.06 | 28.71 | 0.35 | 28.42 | -6.39 | -0.06 | -34.75 | -0.06 | -33.69 | -0.06 | -35.92 | -0.18 | -39.32
23.776 | 21.45 | 0 | 0 | -15.59 | -0.06 | -17.81 | -0.06 | -48.87 | -0.06 | 0.06 | 0.12 | 21.39 | 0.41 | 21.15 | -0.23 | 0.06 | -25.08 | 0.18 | -8.44 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.41 | 28.48 | -4.39 | 0 | -32.93 | -0.18 | -33.05 | -0.06 | -33.28 | 0 | -34.63
23.808 | 21.45 | -0.29 | -0.06 | -13.83 | -0.06 | -15.59 | -0.06 | -41.49 | 0 | 0 | 0.12 | 21.39 | 0.41 | 21.33 | -0.18 | -0.06 | -23.38 | 0.18 | -70.96 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.54 | -2.75 | 2.05 | -30.7 | -0.06 | -32.4 | 0 | -32.64 | 0 | -33.28
23.84 | 21.45 | -0.35 | -0.06 | -12.42 | 0.06 | -13.95 | -0.06 | -36.45 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.27 | -0.06 | 0.06 | -21.27 | 0.12 | -53.79 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.48 | -1.82 | 9.49 | -28.54 | -0.06 | -30.94 | -0.06 | -31.82 | -0.18 | -32.99
23.872 | 21.39 | -0.41 | -0.06 | -10.61 | 0 | -12.6 | 0 | -32.4 | 0.12 | 0 | 0.06 | 21.56 | 0.18 | 21.33 | -0.06 | 0 | -19.92 | 0.12 | -33.99 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.36 | 0.35 | 28.54 | -0.7 | 16.76 | -26.54 | 0 | -28.77 | 0 | -30.53 | -0.06 | -32.05
23.904 | 21.45 | -0.35 | 4.16 | -9.08 | -0.06 | -11.07 | -0.06 | -31.99 | 0 | 0 | 0.12 | 21.39 | 0.18 | 21.39 | 0 | 0 | -18.28 | 0.12 | -27.89 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.12 | 28.59 | -0.53 | 27.07 | -24.9 | -0.06 | -27.36 | -0.06 | -27.77 | -0.06 | -29.65
23.936 | 21.45 | -0.29 | 12.07 | -7.97 | -0.06 | -9.43 | -0.06 | -32.34 | -0.06 | 0 | 0.12 | 21.33 | 0.12 | 21.33 | 0.12 | 3.75 | -16.23 | 0.06 | -34.51 | -0.18 | 0 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.71 | -0.23 | 28.65 | -23.61 | -0.06 | -25.55 | -0.06 | -26.37 | -0.06 | -27.54
23.968 | 21.45 | -0.29 | 17.52 | -6.09 | -0.18 | -7.91 | 0 | -31.17 | -0.06 | 0 | 0.06 | 21.45 | 0.12 | 21.45 | 0.06 | 9.96 | -14.53 | 0.12 | -35.39 | -0.06 | -55.2 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.42 | 0 | 28.71 | -21.68 | -0.06 | -23.85 | 0 | -25.31 | 0 | -26.72
24 | 21.45 | -0.18 | 21.74 | -4.8 | -0.06 | -6.33 | -0.06 | -29.24 | 0 | 0 | 0.06 | 21.56 | 0.12 | 21.33 | 0.06 | 17.11 | -13.07 | 0.06 | -32.76 | -0.18 | -57.89 | 0 | -31.35 | 0 | 28.54 | 0.18 | 28.71 | 0.06 | 28.36 | -19.45 | 0 | -22.44 | -0.06 | -23.44 | -0.06 | -25.14
24.032 | 21.45 | -0.23 | 21.5 | -3.4 | 0.18 | -4.98 | -0.06 | -26.84 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.33 | 0.23 | 21.74 | -11.07 | 0.06 | -29.3 | -0.06 | -44.59 | -0.06 | -67.33 | -0.06 | 0.29 | 0 | 28.95 | 0.29 | 28.59 | -17.64 | 0 | -20.51 | -0.06 | -21.97 | -0.18 | -23.32
24.064 | 21.45 | -0.18 | 21.45 | -2.05 | 7.56 | -3.57 | -0.06 | -24.96 | -0.18 | 0 | -0.06 | 0.76 | 0.06 | 21.39 | 0.18 | 21.39 | -9.49 | 0.12 | -27.42 | -0.18 | -45.47 | 0 | -45.94 | -0.18 | 0.12 | 0.06 | 28.48 | 0.23 | 28.59 | -15.06 | -0.18 | -18.87 | -0.18 | -20.74 | -0.06 | -21.86
24.096 | 21.45 | -0.23 | 21.45 | -1.17 | 13.13 | -2.11 | 0.18 | -22.97 | -0.06 | -1.58 | 0.12 | 0.06 | 0 | 21.39 | 0.23 | 21.33 | -7.38 | 0.18 | -26.84 | 0 | -42.6 | -0.06 | -46.23 | -0.06 | 0 | 0.12 | 28.59 | 0.29 | 28.59 | -13.36 | 0 | -16.99 | 0 | -18.46 | 0 | -20.16
24.128 | 21.39 | -0.23 | 21.45 | -0.64 | 18.63 | -1.17 | 7.62 | -20.45 | 0 | -73.07 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.39 | -5.39 | -0.06 | -25.08 | -0.06 | -40.67 | -0.06 | -43.89 | -0.18 | 0 | 0 | 28.83 | 0.35 | 28.48 | -11.54 | 0.18 | -15.47 | -0.06 | -16.76 | -0.18 | -18.52
24.16 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | -0.59 | 13.13 | -18.28 | -0.06 | -58.77 | -0.06 | 0 | -0.06 | 1.05 | 0.12 | 21.33 | -4.1 | 0 | -22.38 | -0.06 | -38.61 | -0.06 | -41.6 | -0.18 | 0 | 0.06 | 28.65 | 0.29 | 28.48 | -9.02 | 9.84 | -13.71 | -0.06 | -15.18 | -0.06 | -16.88
24.192 | 21.45 | -0.18 | 21.5 | -0.18 | 21.45 | -0.35 | 18.69 | -16.52 | -0.06 | -42.89 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.5 | -2.99 | -0.06 | -20.74 | -0.06 | -37.03 | -0.18 | -40.02 | 0 | 0 | -0.06 | 28.48 | 0.29 | 28.48 | -6.97 | 19.86 | -12.07 | 0 | -13.89 | 9.32 | -15.18
24.224 | 21.45 | -0.29 | 21.39 | -0.18 | 21.45 | -0.18 | 21.62 | -14.65 | -0.06 | -37.79 | -0.06 | 0 | 0 | 0 | 0.18 | 21.21 | -1.76 | 0 | -19.75 | -0.06 | -35.45 | -0.06 | -37.68 | -0.06 | 0 | -0.06 | 28.59 | 0.23 | 28.36 | -4.63 | 27.48 | -10.84 | 8.85 | -12.01 | 16.52 | -13.3
24.256 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | -0.18 | 21.45 | -11.89 | -0.06 | -32.76 | -0.06 | 0 | 0 | 0 | 0.12 | 21.39 | -0.82 | 0.06 | -18.16 | -0.06 | -32.93 | -0.06 | -36.39 | -0.06 | 0 | 0 | 28.36 | 0.18 | 28.42 | -2.75 | 28.65 | -8.56 | 16.06 | -9.84 | 24.26 | -11.31
24.288 | 21.45 | -0.23 | 21.45 | -0.41 | 21.45 | -0.23 | 21.45 | -9.84 | -0.06 | -32.17 | -0.06 | 0 | 0 | 0 | 0.12 | 21.33 | -0.59 | -0.06 | -16.11 | -0.06 | -31.64 | -0.18 | -33.52 | -0.06 | 0 | 0.06 | 28.42 | 0.23 | 28.42 | -1.93 | 28.65 | -6.8 | 23.56 | -8.38 | 28.65 | -9.84
24.32 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | -0.41 | 21.45 | -7.44 | 0 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | -0.47 | 0 | -13.89 | -0.18 | -28.83 | 0 | -32.93 | -0.06 | 0 | 0 | 28.3 | 0.12 | 28.59 | -0.94 | 28.48 | -5.74 | 28.83 | -6.91 | 28.59 | -8.38
24.352 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.47 | 21.45 | -5.04 | -0.06 | -31.06 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.29 | -0.06 | -12.72 | 0 | -27.01 | -0.06 | -30.7 | -0.06 | 0 | -0.18 | 28.42 | 0.12 | 28.48 | -0.41 | 28.59 | -3.75 | 28.65 | -5.39 | 28.54 | -6.8
24.384 | 20.98 | -0.18 | 21.45 | -0.41 | 21.45 | -0.35 | 21.45 | -3.63 | -0.18 | -29.47 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | -0.06 | 0 | -11.07 | 0 | -26.08 | -0.06 | -27.83 | -0.06 | 0 | -0.06 | 28.48 | 0 | 28.59 | -0.29 | 28.54 | -2.46 | 28.42 | -3.75 | 28.54 | -4.92
24.416 | 0.06 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | -1.82 | -0.06 | -27.72 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.27 | 0.06 | -0.06 | -8.85 | -0.06 | -23.09 | -0.06 | -25.9 | 0 | 0 | -0.06 | 13.89 | 0.18 | 28.59 | 0.06 | 28.42 | -1.64 | 28.54 | -2.23 | 28.59 | -2.99
24.448 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -0.94 | 0 | -25.31 | -0.06 | 0 | 0 | 0 | 0 | 21.39 | 0.06 | 1.99 | -7.27 | -0.06 | -21.33 | -0.06 | -24.08 | -0.06 | -31.7 | 0 | 0.12 | 0 | 28.42 | 0.12 | 28.54 | -1 | 28.3 | -1.17 | 28.42 | -2.05
24.48 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | -0.47 | 0 | -23.32 | 0 | 0.12 | 0 | 0 | 0.06 | 21.39 | 0.06 | 8.03 | -5.45 | -0.06 | -18.63 | -0.06 | -21.86 | -0.06 | -75.71 | 0 | 0 | 0 | 28.54 | 0.29 | 28.48 | -0.53 | 28.42 | -1.23 | 28.48 | -1.46
24.512 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -0.06 | 0 | -20.98 | 0.06 | -76.12 | 0.06 | 0 | 0.06 | 21.39 | 0.12 | 13.54 | -3.98 | -0.06 | -17.17 | 0 | -19.86 | -0.06 | -51.68 | 0 | 0 | -0.18 | 28.65 | 0.23 | 28.36 | -0.29 | 28.59 | -0.18 | 28.42 | -0.7
24.544 | 0 | -0.18 | 21.45 | -0.29 | 21.5 | -0.23 | 21.45 | 0 | -0.06 | -18.69 | 0 | -53.85 | -0.18 | 0 | 0.06 | 21.33 | 0.06 | 19.04 | -2.64 | -0.06 | -15.59 | 0 | -17.4 | 0 | -46.23 | -0.06 | 0 | 0.06 | 28.42 | 0.29 | 28.36 | -0.29 | 28.36 | -0.47 | 28.54 | -0.64
24.576 | 0 | -0.18 | 21.5 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | -0.06 | -16.52 | -0.18 | -43.07 | -0.18 | -4.16 | 0.06 | 21.33 | 0.29 | 21.5 | -1.35 | 0 | -14 | 0 | -16.23 | -0.06 | -44.3 | 0 | 0 | -0.06 | 28.48 | 0.35 | 28.65 | -0.06 | 28.54 | -0.06 | 28.36 | -0.41
24.608 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.29 | 21.5 | 0.29 | -0.06 | -14.94 | -0.06 | -37.21 | -0.18 | -72.54 | 0.06 | 21.39 | 0.29 | 21.39 | -0.94 | 0.06 | -12.25 | 0 | -14.47 | 0.06 | -41.66 | -0.18 | 0 | 0 | 28.42 | 0.41 | 28.48 | -0.06 | 28.42 | -0.06 | 28.3 | 0
24.64 | 0 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.47 | -0.06 | -13.01 | -0.18 | -32.93 | -0.06 | -55.37 | 0.06 | 21.39 | 0.29 | 21.39 | -0.64 | 0 | -10.61 | 2.4 | -13.24 | -0.06 | -40.9 | 0.06 | 0 | 0 | 28.24 | 0.35 | 28.42 | -0.06 | 28.42 | 0.06 | 28.24 | -0.18
24.672 | 0 | -0.06 | 21.5 | -0.18 | 21.45 | -0.29 | 21.39 | 0.47 | -0.06 | -10.37 | 0 | -32.34 | -0.06 | -32.29 | 0.12 | 21.33 | 0.29 | 21.33 | -0.35 | 0 | -9.9 | 10.25 | -11.07 | -0.06 | -38.44 | 0 | 0 | -0.06 | 28.3 | 0.29 | 28.65 | 0.23 | 28.54 | 0.06 | 28.42 | -0.06
24.704 | 0 | 0 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | 0.47 | -0.06 | -8.09 | 0 | -32.64 | -0.06 | -27.54 | 0.12 | 21.39 | 0.12 | 21.15 | -0.06 | 3.05 | -8.09 | 17.7 | -10.25 | 0 | -36.74 | 0 | 0 | 0 | 3.81 | 0.18 | 28.59 | 0.12 | 28.42 | 0.18 | 28.65 | 0
24.736 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | -0.18 | 21.5 | 0.47 | 1 | -5.92 | -0.06 | -31.23 | -0.18 | -34.86 | -0.06 | 21.39 | 0.23 | 21.15 | 0.12 | 11.02 | -6.91 | 25.43 | -8.61 | 0.06 | -34.63 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 28.42 | 0.18 | 28.59 | 0.18 | 28.42 | 0
24.768 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.53 | 8.44 | -3.63 | -0.06 | -30.29 | -0.06 | -35.39 | -0.06 | 21.39 | 0.35 | 21.21 | 0 | 18.87 | -5.92 | 28.59 | -7.21 | 0 | -32.7 | -0.06 | 0 | 0 | 0 | 0.23 | 28.77 | 0.12 | 28.42 | 0.18 | 28.42 | -0.06
24.8 | -0.06 | -0.06 | 0.12 | -0.18 | 21.45 | 0 | 21.45 | 0.47 | 15.53 | -2.11 | -0.06 | -27.89 | -0.06 | -32.23 | 0.06 | 21.33 | 0.23 | 20.98 | 0.06 | 26.37 | -4.45 | 28.54 | -6.39 | 0.06 | -30.06 | 0 | 0 | -0.18 | 0 | 0 | 28.42 | 0.12 | 28.42 | 0.12 | 28.36 | 0.06
24.832 | -42.89 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.35 | 21.39 | -1 | -0.06 | -24.84 | -0.18 | -29.06 | 0 | 0.12 | 0.35 | 21.15 | 0.06 | 28.65 | -3.46 | 28.42 | -4.86 | 9.38 | -29.06 | 0 | 0.06 | -0.06 | 0 | 0.06 | 28.65 | 0.18 | 28.48 | 0.23 | 28.42 | 0.06
24.864 | -55.61 | 0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.29 | 21.56 | -0.76 | -0.06 | -22.91 | -0.06 | -27.54 | 0 | 0 | 0.23 | 21.39 | 0.29 | 28.65 | -2.93 | 28.59 | -3.87 | 16.7 | -26.6 | 0 | -69.73 | -0.06 | 0 | 0.12 | 28.54 | 0.06 | 28.65 | 0.12 | 28.48 | 0.12
24.896 | -40.55 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.29 | 21.39 | -0.06 | 0.06 | -21.56 | -0.06 | -26.84 | 0.06 | 0 | 0.18 | 21.21 | 0.23 | 28.54 | -2.23 | 28.59 | -3.34 | 27.07 | -23.85 | -0.06 | -62.64 | 0 | 0 | 0.06 | 28.48 | 0.06 | 28.42 | 0.18 | 28.48 | 0.12
24.928 | -39.32 | -0.18 | 0 | -0.23 | 19.86 | -0.06 | 21.39 | 0.23 | 21.5 | 0.12 | 0 | -19.34 | -0.06 | -24.38 | 0 | 0.06 | 0.18 | 21.33 | 0.18 | 28.65 | -1.76 | 28.42 | -2.4 | 28.59 | -22.33 | -0.06 | -47.76 | 0 | 0 | 0 | 28.71 | 0.12 | 28.48 | 0.23 | 28.42 | 0.18
24.96 | -37.38 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.39 | 0.23 | 21.39 | 0.18 | -0.06 | -17.58 | -0.06 | -22.74 | -0.06 | 0 | 0.23 | 21.39 | 0.29 | 28.54 | -1.41 | 28.48 | -1.7 | 28.36 | -20.39 | -0.06 | -45.53 | 0 | 0 | 0 | 28.13 | 0 | 28.42 | 0.12 | 28.48 | 0.12
24.992 | -35.45 | -0.06 | 0 | -0.18 | 0 | -0.18 | 19.75 | 0.29 | 21.39 | 0.41 | -0.06 | -14.94 | -0.18 | -20.98 | 0.06 | 0 | 0.23 | 21.27 | 0.18 | 28.42 | -1 | 28.48 | -1.35 | 28.42 | -18.46 | -0.06 | -40.55 | -0.06 | 0 | -0.06 | 28.48 | 0 | 28.42 | 0.18 | 28.24 | 0.12
25.024 | -33.05 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 21.45 | 0.29 | -0.06 | -12.72 | -0.18 | -19.45 | 0 | 0 | 0.23 | 21.27 | 0.23 | 28.71 | -0.82 | 28.59 | -1 | 28.48 | -16.52 | 0 | -36.8 | 0 | 0 | 0.06 | 0.64 | 0.12 | 28.42 | 0.12 | 28.42 | 0.12
25.056 | -30.59 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | -0.18 | -10.84 | -0.18 | -17.4 | -0.06 | 0 | 0.18 | 21.21 | 0.12 | 28.59 | -0.53 | 28.42 | -0.76 | 28.65 | -14.65 | -0.06 | -37.91 | 0 | 0 | 0.06 | 0.12 | -0.06 | 28.42 | 0.06 | 28.42 | 0.12
25.088 | -28.65 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.53 | 0 | -8.91 | 0 | -15.35 | 0 | 0 | 0.23 | 21.33 | 0.18 | 28.48 | -0.41 | 28.89 | -0.64 | 28.71 | -13.07 | -0.18 | -38.15 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.54 | 0 | 0.23 | 0.12
25.12 | -26.84 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | 0.53 | 2.34 | -6.56 | 0.06 | -14.24 | 0 | 0 | 0.23 | 21.15 | 0.18 | 28.65 | -0.35 | 28.48 | -0.29 | 28.36 | -12.07 | 0 | -35.92 | 0 | 0.12 | 0 | 0 | -0.06 | 0.23 | 0.06 | 0 | 0.12
25.152 | -24.61 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | 8.03 | -3.93 | -0.06 | -12.36 | 0.06 | 0 | 0.18 | 21.39 | 0.18 | 28.65 | -0.29 | 28.59 | -0.29 | 28.42 | -10.2 | 0 | -34.28 | -0.06 | -70.43 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06
25.184 | -22.56 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 13.65 | -2.23 | -0.06 | -10.66 | 0 | 0 | 0.18 | 21.39 | 0.23 | 28.65 | -0.29 | 29.24 | -0.35 | 28.42 | -9.02 | 0 | -32.93 | -0.06 | -62.87 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.06
25.216 | -19.86 | 0 | -32.23 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | 19.04 | -1.29 | -0.06 | -8.56 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.48 | -0.23 | 28.54 | -0.23 | 28.42 | -7.21 | 3.81 | -29.94 | 0 | -48.22 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.06
25.248 | -17.7 | -0.06 | -59.24 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.94 | 0 | -6.62 | -0.06 | -10.08 | 0.12 | 21.33 | 0.06 | 28.42 | -0.29 | 28.42 | -0.35 | 28.42 | -6.39 | 12.25 | -28.24 | -0.06 | -45.35 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18
25.28 | -16.23 | 0 | -44.36 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -0.18 | -0.06 | -5.04 | 0 | -67.8 | 0.12 | 21.39 | 0.18 | 28.54 | -0.23 | 28.48 | -0.23 | 28.42 | -5.1 | 19.92 | -26.6 | -0.06 | -40.31 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18
25.312 | -14.82 | 0 | -39.85 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | -0.18 | 2.46 | -3.16 | -0.06 | -43.65 | 0.18 | 21.39 | 0.06 | 28.59 | -0.35 | 28.36 | -0.23 | 28.42 | -4.16 | 27.31 | -25.31 | 0 | -36.92 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06
25.344 | -13.54 | -0.06 | -38.79 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.29 | 10.9 | -2.29 | 0 | -29.06 | 0.23 | 17.64 | 0.23 | 28.48 | -0.18 | 28.48 | -0.29 | 28.59 | -3.4 | 28.71 | -23.38 | -0.06 | -38.2 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.18
25.376 | -11.31 | 0 | -35.27 | -0.18 | -45.35 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.21 | 0.29 | 16.41 | -1.46 | -0.06 | -30.06 | 0.06 | 0.12 | 0.06 | 28.71 | -0.18 | 28.3 | -0.18 | 28.48 | -2.58 | 28.65 | -20.63 | -0.06 | -37.85 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0
25.408 | -9.79 | -0.06 | -33.99 | -0.18 | -54.49 | 0 | -0.12 | -0.06 | 21.39 | 0 | 21.39 | 0.35 | 21.74 | -0.76 | 0 | -35.1 | 0.12 | 0 | 0.06 | 28.77 | -0.23 | 28.42 | -0.18 | 28.54 | -1.82 | 28.65 | -19.04 | -0.06 | -36.39 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0.12
25.44 | -8.38 | -0.06 | -31.29 | -0.06 | -39.79 | -0.06 | -42.48 | 0.06 | 21.39 | 0.06 | 21.39 | 0.29 | 21.39 | -0.64 | 0.06 | -34.16 | 0.06 | 0.06 | 0.12 | 28.77 | -0.06 | 28.65 | -0.18 | 28.48 | -1.46 | 28.3 | -17.05 | -0.06 | -34.28 | 0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0
25.472 | -7.03 | -0.06 | -29.71 | -0.06 | -39.08 | -0.06 | -54.44 | -0.06 | 21.39 | 0.18 | 21.39 | 0.29 | 21.39 | -0.23 | -0.06 | -30.59 | 0.06 | 0 | 0.06 | 28.65 | -0.06 | 28.42 | -0.18 | 28.65 | -1.05 | 28.48 | -15.18 | -0.06 | -32.05 | -0.18 | -93.34 | 0 | 0 | -0.06 | 0.06 | 0.12
25.504 | -5.63 | -0.06 | -28.13 | -0.06 | -36.97 | -0.06 | -41.72 | 0.06 | 21.33 | 0.06 | 21.39 | 0.29 | 21.45 | -0.23 | -0.06 | -28.3 | 0.06 | 0 | 0.12 | 28.48 | -0.18 | 28.48 | -0.18 | 28.65 | -0.64 | 28.59 | -13.18 | -0.06 | -30.41 | 0.06 | -61.29 | -0.18 | 0 | -0.06 | -68.85 | -0.06
25.536 | -3.63 | -0.06 | -24.9 | -0.06 | -35.27 | -0.06 | -39.14 | -0.06 | 21.45 | 0.06 | 21.39 | 0.41 | 21.21 | 0 | 0 | -26.9 | -0.06 | 0 | 0.12 | 28.54 | -0.18 | 23.38 | -0.18 | 28.48 | -0.41 | 28.65 | -10.78 | -0.18 | -28.59 | -0.18 | -47.93 | -0.18 | -55.14 | -0.06 | -74.59 | 0.06
25.568 | -2.46 | 0 | -22.62 | -0.06 | -33.17 | 0 | -37.21 | 0 | 21.56 | 0.12 | 21.45 | 0.47 | 21.33 | -0.06 | 3.75 | -26.31 | 0.06 | 0 | 0.12 | 28.65 | -0.18 | 0.18 | -0.18 | 28.48 | -0.23 | 28.65 | -8.56 | -0.18 | -26.9 | -0.06 | -46.06 | -0.18 | -76.58 | -0.18 | -43.89 | 0
25.6 | -1.64 | 7.03 | -21.21 | -0.18 | -30.41 | -0.06 | -34.98 | -0.06 | 21.45 | 0.06 | 21.39 | 0.35 | 21.39 | 0.18 | 9.96 | -24.26 | 0.06 | 0 | 0.23 | 13.83 | -0.18 | 0 | -0.18 | 28.3 | -0.23 | 28.48 | -6.27 | -0.06 | -24.84 | 0 | -39.85 | -0.18 | -47.64 | -0.06 | -29.42 | -0.06
25.632 | -0.76 | 12.83 | -18.87 | -0.06 | -28.65 | -0.06 | -33.17 | -0.06 | 0.23 | -0.06 | 21.27 | 0.35 | 21.45 | 0.35 | 15.29 | -21.86 | 0 | 0 | 0.12 | 0.12 | -0.12 | 0 | -0.23 | 28.42 | -0.35 | 28.3 | -4.16 | 0 | -23.03 | 0 | -36.86 | -0.18 | -29.24 | 0 | -37.97 | 0.06
25.664 | -0.35 | 18.4 | -17.23 | -0.18 | -26.66 | -0.06 | -31.47 | 0 | 0.06 | -0.06 | 21.39 | 0.18 | 21.33 | 0.29 | 20.8 | -20.39 | 0 | 0 | 0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | -0.29 | 28.48 | -2.23 | -0.06 | -20.86 | -0.06 | -37.74 | -0.18 | -38.67 | 0 | -42.89 | 0
25.696 | -0.18 | 21.39 | -14.94 | -0.06 | -24.43 | -0.06 | -29.01 | -0.06 | 0 | 0 | 21.15 | 0.18 | 21.39 | 0.23 | 21.45 | -19.45 | 0.12 | 0 | 0.18 | 0.06 | -0.06 | 0 | 0 | 4.63 | -0.29 | 28.48 | -1.29 | 7.79 | -18.99 | -0.06 | -38.32 | -0.18 | -42.6 | -0.06 | -40.08 | -0.06
25.728 | -0.23 | 21.56 | -13.54 | -0.06 | -22.15 | 0 | -26.25 | 0 | 0 | -0.06 | 21.33 | 0.23 | 21.45 | 0.23 | 21.45 | -17.7 | 0.06 | 0 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0.12 | -0.29 | 28.48 | -0.7 | 15.06 | -16.88 | -0.18 | -36.33 | -0.06 | -39.61 | -0.06 | -35.16 | 0
25.76 | -0.29 | 21.45 | -11.6 | 0 | -19.98 | -0.06 | -24.79 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.56 | 0.29 | 21.33 | -15.23 | 0 | 0 | 0.23 | 0 | -0.06 | 0 | -0.06 | 0 | -0.29 | 28.59 | -0.59 | 25.61 | -15.18 | -0.06 | -33.69 | -0.18 | -35.39 | 0 | -33.28 | -0.06
25.792 | -0.35 | 21.39 | -10.61 | -0.06 | -17.75 | 0 | -22.33 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.33 | 0.29 | 21.39 | -13.36 | -0.06 | -60.76 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.23 | 28.42 | -0.06 | 28.71 | -13.18 | 0 | -31.82 | -0.06 | -33.05 | 0.06 | -33.05 | 0
25.824 | -0.41 | 21.45 | -9.32 | -0.18 | -16.29 | -0.06 | -20.92 | -0.06 | 0 | 0.06 | 21.5 | 0 | 21.39 | 0.29 | 21.39 | -12.07 | -0.06 | -61.23 | 0.18 | 0 | 0 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.12 | 28.59 | -10.9 | -0.06 | -29.94 | -0.23 | -32.93 | 0 | -31.82 | -0.18
25.856 | -0.35 | 21.5 | -7.62 | -0.18 | -14.77 | -0.06 | -17.99 | 0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.45 | 0.29 | 21.39 | -10.08 | 0.06 | -37.56 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.23 | 28.48 | 0.18 | 28.65 | -8.5 | 0 | -28.48 | -0.18 | -31.93 | -0.06 | -29.77 | 0
25.888 | -0.35 | 21.45 | -5.63 | -0.06 | -13.3 | -0.06 | -16.23 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.27 | 0.18 | 21.33 | -8.85 | -0.06 | -27.89 | 0.06 | 0 | 0 | 0 | -0.06 | 0 | -0.23 | 28.48 | 0.23 | 28.65 | -6.5 | 0 | -26.66 | -0.06 | -29.42 | 0 | -27.89 | 0
25.92 | -0.47 | 21.45 | -4.16 | -0.06 | -11.13 | 0 | -14.65 | -0.06 | 0 | 0 | 21.39 | 0 | 21.33 | 0.23 | 21.39 | -6.39 | 0 | -32.93 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.65 | -4.39 | 0 | -24.61 | -0.06 | -28.3 | -0.12 | -26.31 | 0
25.952 | -0.29 | 21.39 | -3.16 | -0.06 | -9.84 | 0 | -13.07 | -0.18 | 0 | 0 | 21.39 | 0 | 21.56 | 0.29 | 21.27 | -4.69 | 0.06 | -35.16 | 0.18 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.65 | 0.41 | 28.48 | -2.75 | 0 | -22.38 | -0.18 | -26.43 | -0.06 | -25.14 | -0.06
25.984 | -0.18 | 21.45 | -2.11 | 4.22 | -8.38 | 0 | -11.66 | -0.06 | 0 | 0 | 21.39 | -0.06 | 21.33 | 0.12 | 21.39 | -3.34 | 0 | -33.52 | 0 | 0 | -0.06 | -59.06 | -0.06 | 0 | -0.18 | 28.36 | 0.47 | 28.42 | -1.64 | 0 | -20.57 | -0.06 | -24.55 | -0.06 | -23.44 | 0
26.016 | -0.18 | 21.45 | -1.05 | 10.66 | -7.09 | 2.93 | -9.61 | -0.06 | 0 | -0.06 | 17.7 | 0 | 21.56 | 0.06 | 21.27 | -2.23 | -0.06 | -30.18 | 0.12 | -1.05 | -0.06 | -63.64 | 0 | 0 | -0.18 | 28.42 | 0.47 | 28.42 | -0.59 | 0 | -18.34 | -0.06 | -23.85 | -0.06 | -21.86 | 0
26.048 | -0.35 | 21.45 | -0.59 | 15.82 | -5.68 | 9.02 | -8.56 | -0.06 | -0.23 | -0.06 | 0.12 | -0.06 | 21.33 | 0.18 | 21.27 | -1.11 | 0 | -27.95 | 0.06 | -75 | -0.06 | -45.76 | 0 | 0 | -0.06 | 28.59 | 0.35 | 28.42 | -0.53 | -0.06 | -16.17 | -0.18 | -22.21 | 0 | -20.16 | -0.18
26.08 | -0.23 | 21.5 | -0.29 | 21.45 | -3.87 | 14.59 | -7.44 | -0.06 | -77.52 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | -0.82 | -0.06 | -27.01 | 0.06 | -59.65 | 0 | -45.53 | -0.06 | 0.06 | -0.18 | 28.65 | 0.47 | 28.65 | 0.12 | 0 | -14.65 | -0.18 | -20.45 | -0.18 | -18.4 | -0.06
26.112 | -0.29 | 21.45 | -0.18 | 21.5 | -2.52 | 20.33 | -5.57 | 0 | -51.62 | 0 | 0 | 0 | 21.33 | 0 | 21.62 | -0.59 | -0.06 | -25.08 | 0.12 | -44.36 | -0.06 | -43.24 | -0.06 | 0 | -0.18 | 0.7 | 0.23 | 28.48 | 0 | -0.06 | -12.19 | 0 | -18.4 | 0 | -16.58 | 0
26.144 | -0.18 | 21.45 | -0.06 | 21.5 | -1.7 | 21.5 | -4.22 | -0.06 | -43.01 | -0.06 | 0 | 0 | 21.27 | 0.06 | 21.27 | -0.23 | 5.86 | -22.91 | 0.12 | -44.36 | -0.06 | -41.66 | -0.06 | -51.39 | -0.18 | 0.12 | 0.29 | 28.3 | 0.23 | 7.44 | -10.37 | -0.06 | -16.64 | -0.06 | -15 | -0.06
26.176 | -0.23 | 21.5 | -0.23 | 21.45 | -0.82 | 21.5 | -2.7 | 0 | -37.38 | -0.06 | 0 | -0.18 | 21.39 | 0.06 | 21.33 | 0 | 13.59 | -21.62 | 0.06 | -42.83 | -0.06 | -38.15 | -0.06 | -58.95 | -0.18 | 0 | 0.23 | 28.42 | 0.29 | 14.82 | -8.2 | 5.45 | -15.18 | -0.18 | -13.59 | -0.06
26.208 | -0.18 | 21.39 | -0.41 | 21.39 | -0.23 | 21.5 | -1.52 | -0.06 | -32.76 | -0.18 | 0 | -0.18 | 0.12 | 0 | 21.39 | -0.06 | 19.1 | -20.1 | 0 | -40.43 | -0.06 | -37.15 | 0 | -45.76 | -0.06 | 0.06 | 0.12 | 28.59 | 0.41 | 22.79 | -6.33 | 13.48 | -13.36 | 7.09 | -11.72 | -0.06
26.24 | -0.23 | 21.39 | -0.41 | 21.45 | -0.23 | 21.45 | -0.94 | -0.06 | -32.52 | 0 | 0 | 0 | 0.06 | 0 | 21.27 | 0.06 | 21.39 | -18.63 | 0.12 | -38.91 | 0 | -34.63 | 0 | -45.06 | -0.18 | 0 | 0.29 | 28.89 | 0.29 | 28.65 | -4.1 | 21.27 | -11.84 | 14.71 | -9.49 | 9.2
26.272 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.45 | -0.41 | 0 | -32.05 | 0 | 0 | 0 | 0 | 0 | 21.45 | 0.12 | 21.33 | -16.7 | 0 | -37.21 | 0 | -34.1 | 0 | -43.13 | 0 | 0 | 0.23 | 28.65 | 0.29 | 28.71 | -1.88 | 29.18 | -9.84 | 22.44 | -7.97 | 16.17
26.304 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -0.18 | -0.06 | -31.7 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.39 | 0 | 21.33 | -15.18 | 0.06 | -34.86 | 0.06 | -31.41 | 0 | -41.43 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.54 | -1.7 | 28.48 | -7.91 | 29.12 | -6.8 | 23.85
26.336 | -0.29 | 21.45 | -0.23 | 21.45 | -0.41 | 21.45 | -0.23 | -0.06 | -29.59 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.18 | 21.27 | -12.83 | 0.06 | -33.17 | 0 | -29.47 | -0.06 | -38.79 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.48 | -0.41 | 28.59 | -6.86 | 28.54 | -5.04 | 28.54
26.368 | -0.06 | 21.39 | -0.29 | 21.45 | -0.41 | 21.45 | -0.18 | -0.06 | -27.48 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.21 | -11.37 | 0.06 | -30.82 | 0 | -27.95 | 0 | -37.68 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.71 | -0.47 | 28.54 | -4.86 | 28.48 | -3.28 | 28.65
26.4 | -0.18 | 21.45 | -0.29 | 21.39 | -0.23 | 21.45 | -0.35 | 2.7 | -24.67 | 0 | 0 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.21 | -9.26 | 0 | -28.95 | 0 | -24.84 | -0.06 | -35.51 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.48 | 0.12 | 28.48 | -3.75 | 28.48 | -1.99 | 28.59
26.432 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | -0.47 | 8.79 | -22.91 | 0.06 | -0.18 | -0.06 | 0 | -0.18 | 21.45 | 0.12 | 21.27 | -7.21 | 0 | -27.77 | 0 | -23.2 | 0 | -34.22 | -0.06 | 0 | 0.12 | 28.42 | 0.23 | 28.48 | 0.06 | 28.48 | -2.23 | 28.54 | -1.17 | 28.42
26.464 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.5 | -0.47 | 14.18 | -20.92 | -0.06 | -62.7 | -0.06 | 0 | 0 | 0.29 | 0.23 | 21.27 | -5.63 | -0.06 | -23.85 | -0.06 | -21.39 | -0.06 | -32.05 | -0.18 | 0 | -0.06 | 28.59 | 0.12 | 28.3 | 0.41 | 28.48 | -1.64 | 28.42 | -0.82 | 28.42
26.496 | -0.18 | 0.06 | -0.29 | 21.45 | -0.18 | 21.45 | -0.29 | 21.62 | -19.04 | 0 | -60.12 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | -4.22 | 9.26 | -22.44 | -0.06 | -18.75 | 0 | -29.47 | 0 | 0 | 0 | 28.3 | 0.06 | 29.01 | 0.29 | 28.59 | -0.82 | 28.54 | -0.76 | 28.83
26.528 | -0.06 | 0 | -0.29 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -16.88 | -0.06 | -44.3 | 0 | 0 | 0 | 0 | 0.18 | 21.33 | -2.93 | 14.36 | -20.57 | 0 | -17.29 | 0.06 | -27.07 | 0 | -8.38 | 0.06 | 28.83 | 0.12 | 28.48 | 0.35 | 28.42 | -0.64 | 28.48 | -0.23 | 28.48
26.56 | -0.06 | 0 | -0.23 | 21.45 | -0.18 | 21.5 | -0.29 | 21.39 | -14.77 | -0.06 | -37.56 | 0 | 0 | -0.06 | 0.06 | 0.23 | 21.27 | -1.76 | 20.27 | -18.99 | 0 | -15.53 | 0 | -25.37 | -0.06 | -92.93 | 0 | 22.85 | 0 | 28.54 | 0.47 | 28.48 | -0.23 | 28.36 | -0.23 | 28.59
26.592 | -0.18 | 0 | -0.29 | 21.45 | -0.23 | 21.39 | -0.29 | 21.39 | -12.54 | -0.06 | -32.76 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.39 | -1.05 | 21.45 | -16.82 | 0 | -14 | -0.06 | -23.56 | -0.06 | -61.58 | -0.06 | 0.18 | 0.06 | 28.65 | 0.41 | 28.48 | -0.06 | 28.48 | -0.18 | 28.13
26.624 | -0.18 | 0 | -0.29 | 21.5 | -0.35 | 21.5 | -0.35 | 21.39 | -10.31 | 0 | -32.76 | -0.06 | -35.74 | -0.06 | 0 | 0.12 | 21.21 | -0.76 | 21.56 | -15 | -0.06 | -12.6 | 0 | -21.56 | -0.06 | -48.28 | 0 | 0 | 0.06 | 28.48 | 0.29 | 28.54 | -0.06 | 28.54 | 0 | 28.54
26.656 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | -8.14 | -0.06 | -32.52 | 0 | -72.89 | -0.06 | 0 | 0.12 | 21.27 | -0.29 | 21.45 | -13.3 | 8.26 | -11.07 | 0.06 | -19.34 | 0 | -43.71 | 0.06 | 0 | 0 | 28.65 | 0.29 | 28.59 | 0.06 | 28.65 | -0.06 | 28.48
26.688 | -0.06 | 0 | -0.23 | 21.5 | -0.18 | 21.45 | -0.29 | 21.39 | -5.92 | 0.06 | -31.99 | -0.06 | -49.57 | -0.18 | 0 | 0.18 | 21.33 | -0.23 | 21.15 | -12.07 | 15.41 | -9.9 | 5.1 | -16.99 | 0 | -37.62 | 0 | 0 | 0 | 28.48 | 0.29 | 28.42 | 0.23 | 28.54 | 0.12 | 28.48
26.72 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | -3.63 | -0.06 | -29.47 | -0.06 | -41.9 | 0 | 0 | 0.12 | 21.27 | -0.06 | 21.27 | -11.25 | 26.02 | -8.03 | 13.24 | -15.18 | 0 | -36.39 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.48 | 0.12 | 28.48 | -0.18 | 28.36
26.752 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.35 | 21.33 | -1.82 | 0 | -27.54 | -0.06 | -36.51 | 0 | 0 | 0.12 | 21.33 | -0.06 | 21.39 | -9.43 | 28.77 | -6.62 | 23.44 | -14.06 | 0 | -37.56 | -0.18 | 0 | -0.06 | 28.42 | 0.12 | 28.48 | 0.12 | 28.42 | 0.06 | 28.71
26.784 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | -0.23 | 21.39 | -1.35 | 2.34 | -25.72 | -0.06 | -32.29 | -0.06 | 0 | 0.06 | 21.33 | -0.06 | 21.27 | -7.79 | 28.48 | -5.63 | 28.95 | -12.36 | -0.06 | -37.44 | -0.06 | 0 | 0 | 28.77 | 0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.36
26.816 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | -0.41 | 8.03 | -23.67 | 0 | -32.93 | 0 | 0 | 0.12 | 21.33 | 0.06 | 21.33 | -6.56 | 28.59 | -4.86 | 28.54 | -10.9 | -0.06 | -35.57 | -0.06 | 0 | -0.06 | 28.54 | 0.18 | 28.24 | 0.06 | 28.48 | 0.06 | 28.42
26.848 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -0.29 | 13.83 | -21.86 | -0.06 | -32.7 | 0.06 | 0 | 0.18 | 21.33 | 0.18 | 21.21 | -5.57 | 28.48 | -3.87 | 28.59 | -9.32 | 0 | -33.93 | -0.06 | 0 | 0 | 28.48 | 0 | 28.42 | 0.12 | 28.24 | 0.12 | 28.24
26.88 | -0.23 | 0 | -0.06 | 0.23 | -0.18 | 21.39 | -0.18 | 21.27 | -0.06 | 18.99 | -19.04 | -0.06 | -31.52 | -0.06 | -17.29 | 0.12 | 21.33 | 0.29 | 21.15 | -4.45 | 28.59 | -2.81 | 28.48 | -8.44 | 0.06 | -31.99 | 0 | 0 | -0.06 | 28.3 | 0 | 28.59 | 0.23 | 28.59 | 0.12 | 28.36
26.912 | -0.18 | -39.96 | -0.06 | 0 | -0.18 | 7.68 | -0.18 | 21.39 | 0.23 | 21.56 | -16.99 | -0.06 | -29.06 | -0.18 | -73.25 | 0.12 | 21.39 | 0.12 | 21.39 | -3.87 | 28.54 | -1.99 | 28.48 | -6.86 | 0 | -29.94 | 0 | 0 | -0.06 | 28.36 | 0 | 28.42 | 0.12 | 28.48 | 0.18 | 28.59
26.944 | -0.06 | -57.25 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 21.39 | 0.29 | 21.45 | -14.59 | -0.06 | -26.13 | 0 | -51.86 | 0.06 | 21.39 | 0.18 | 21.39 | -2.75 | 28.54 | -1.76 | 28.65 | -5.74 | -0.06 | -28.01 | 0 | 0 | -0.06 | 28.3 | -0.06 | 28.48 | 0.18 | 28.54 | 0.12 | 28.42
26.976 | -0.18 | -40.55 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | -12.77 | -0.06 | -24.14 | -0.06 | -28.71 | 0.12 | 21.39 | 0.12 | 21.39 | -2.23 | 28.48 | -1.46 | 28.48 | -4.45 | 0.06 | -26.13 | 0 | 0 | 0 | 28.48 | -0.06 | 28.3 | 0.06 | 28.42 | 0.12 | 28.42
27.008 | -0.18 | -39.2 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.27 | 0.47 | 21.39 | -10.72 | 0.06 | -22.62 | -0.06 | -30.12 | 0.12 | 17.46 | 0.18 | 21.39 | -1.76 | 28.65 | -1 | 28.83 | -3.46 | 0 | -24.43 | -0.06 | -51.04 | -0.06 | 28.59 | 0 | 28.54 | 0.06 | 28.48 | 0 | 28.42
27.04 | 0 | -37.79 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.33 | 0.53 | 21.45 | -8.38 | -0.06 | -20.39 | 0 | -35.1 | 0 | 0.12 | 0.18 | 21.33 | -1.46 | 28.71 | -0.7 | 28.83 | -2.93 | 0 | -22.21 | -0.18 | -69.61 | -0.06 | 0.18 | 0 | 28.42 | 0.06 | 28.48 | 0.06 | 28.48
27.072 | -0.06 | -35.39 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.47 | 21.39 | -6.09 | 0 | -18.22 | -0.06 | -34.92 | 0 | 0 | 0.29 | 21.15 | -1.11 | 28.59 | -0.35 | 28.59 | -2.29 | -0.06 | -20.39 | 0 | -47.87 | -0.06 | 0 | 0 | 0.29 | 0 | 28.48 | 0.18 | 28.48
27.104 | -0.18 | -33.17 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -4.22 | 0.06 | -16.17 | -0.06 | -31.17 | 0.12 | 0 | 0.23 | 21.33 | -0.76 | 28.65 | -0.29 | 28.36 | -1.52 | 0 | -19.22 | 0 | -47.23 | -0.06 | 0 | 0.06 | 0.12 | 0 | 0.29 | 0.12 | 28.48
27.136 | -0.06 | -30.53 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -2.11 | 0.06 | -13.89 | 0 | -28.3 | 0.06 | 0 | 0.23 | 21.33 | -0.7 | 28.48 | -0.23 | 28.48 | -0.82 | 5.22 | -16.47 | 0 | -41.37 | -0.06 | 0.06 | -0.06 | 0 | 0 | 0 | 0.12 | 0.18
27.168 | -0.18 | -29.24 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | -1.76 | 2.05 | -11.84 | -0.06 | -27.54 | 0.12 | 0 | 0.23 | 21.33 | -0.64 | 28.95 | -0.18 | 28.65 | -0.76 | 13.18 | -14.41 | -0.06 | -36.68 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
27.2 | -0.06 | -27.01 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | -0.41 | 7.21 | -9.79 | -0.18 | -26.6 | 0.18 | 0 | 0.23 | 21.27 | -0.59 | 28.95 | -0.23 | 29.12 | -0.53 | 21.09 | -12.13 | -0.06 | -37.79 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
27.232 | -0.06 | -23.97 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.29 | 21.5 | -0.18 | 14.36 | -7.62 | -0.06 | -24.96 | -0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.54 | -0.23 | 28.65 | -0.23 | 28.71 | -10.37 | 0 | -38.5 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
27.264 | -0.06 | -23.15 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.27 | -0.06 | 20.45 | -4.92 | -0.06 | -22.56 | 0.06 | 0 | 0.23 | 21.15 | -0.35 | 28.77 | -0.29 | 28.42 | -0.23 | 28.59 | -8.03 | 0 | -36.51 | -0.23 | 0 | -0.06 | 0 | -0.18 | 0 | 0 | 0
27.296 | -0.18 | -20.22 | -0.06 | -14.53 | -0.06 | 0 | -0.18 | 0.23 | 0.29 | 21.39 | 0 | 21.5 | -2.99 | 9.43 | -20.74 | 0.18 | 0 | 0.18 | 21.27 | -0.35 | 28.59 | -0.23 | 28.59 | -0.23 | 28.54 | -5.22 | 0 | -34.63 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
27.328 | -0.06 | -17.93 | -0.06 | -64.46 | -0.06 | -3.16 | -0.23 | 0.06 | 0.12 | 21.39 | 0.23 | 21.45 | -1.82 | 14.53 | -18.87 | 0 | 0 | 0.12 | 21.33 | -0.35 | 28.48 | -0.23 | 28.42 | -0.18 | 28.65 | -3.87 | 0 | -32.76 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
27.36 | 0 | -16.58 | -0.06 | -47.64 | -0.18 | -57.07 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -1.23 | 20.27 | -17.05 | 0.06 | 0 | 0.18 | 6.68 | -0.41 | 28.54 | -0.18 | 28.59 | -0.23 | 28.36 | -2.11 | -0.18 | -30.53 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
27.392 | 0 | -15 | -0.06 | -39.85 | -0.06 | -50.86 | -0.06 | 0 | 0.18 | 21.39 | 0.41 | 21.39 | -0.64 | 21.39 | -15.59 | 0.18 | 0 | 0.18 | 0.06 | -0.29 | 28.48 | -0.18 | 28.48 | -0.29 | 28.36 | -1.29 | 0 | -28.65 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | -0.06 | 0
27.424 | -0.06 | -13.07 | -0.06 | -38.56 | -0.06 | -40.43 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 21.39 | -0.18 | 21.27 | -13.59 | -0.18 | 0 | 0.18 | 0 | -0.35 | 28.59 | -0.18 | 28.59 | -0.23 | 28.36 | -0.7 | -0.18 | -27.01 | -0.06 | 0.06 | 0 | 0 | -0.06 | 0 | -0.06 | 0
27.456 | -0.06 | -11.31 | -0.06 | -36.1 | -0.06 | -39.55 | -0.18 | 0 | 0.12 | 21.27 | 0.35 | 21.39 | 0.06 | 21.27 | -11.25 | -0.06 | -59.12 | 0.12 | 0.06 | -0.29 | 28.59 | -0.18 | 28.83 | -0.23 | 28.48 | -0.41 | -0.06 | -25.37 | -0.06 | -47.46 | 0 | 0 | 0 | 0 | 0 | 0
27.488 | 0 | -9.79 | -0.06 | -34.63 | -0.06 | -36.21 | -0.06 | 0 | 0.06 | 21.45 | 0.47 | 21.21 | 0.23 | 21.39 | -9.79 | -0.06 | -60.3 | 0.18 | 0 | -0.35 | 28.13 | -0.18 | 28.42 | -0.29 | 28.65 | -0.06 | 0 | -23.56 | 0 | -78.99 | -0.18 | -38.03 | -0.23 | 0 | 0 | 0
27.52 | 7.09 | -8.67 | 0 | -32.05 | -0.06 | -34.69 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.39 | 0.23 | 21.33 | -8.32 | -0.06 | -37.91 | 0.18 | 0 | -0.18 | 18.34 | -0.18 | 28.65 | -0.18 | 28.42 | -0.06 | 0 | -21.15 | -0.06 | -48.28 | -0.06 | -84.85 | -0.06 | -45.47 | 0 | 0
27.552 | 12.95 | -6.91 | -0.06 | -29.47 | 0 | -32.46 | -0.06 | 0 | 0 | 21.45 | 0.47 | 21.39 | 0.41 | 21.33 | -6.56 | 0 | -27.6 | 0.18 | 0 | -0.18 | 0.18 | -0.06 | 28.59 | -0.29 | 28.48 | 0.18 | 8.14 | -19.1 | -0.06 | -46.23 | -0.06 | -54.73 | -0.06 | -82.68 | 0 | -58.89
27.584 | 18.28 | -5.74 | -0.06 | -27.48 | -0.06 | -30.47 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.53 | 21.39 | -4.98 | -0.06 | -31.52 | 0.18 | 0 | -0.29 | 0 | -0.18 | 0.23 | -0.18 | 28.65 | 0.12 | 15.35 | -17.4 | -0.06 | -40.78 | 0 | -33.22 | 0 | -51.68 | -0.18 | -74.3
27.616 | 21.62 | -3.75 | -0.06 | -26.31 | 0 | -28.71 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.15 | 0.41 | 21.33 | -2.81 | 0 | -35.63 | 0.18 | 0 | -0.23 | 0 | -0.18 | 0.06 | -0.06 | 28.42 | 0.29 | 25.84 | -15.47 | 0 | -36.21 | -0.18 | -34.86 | -0.06 | -28.89 | 0 | -37.79
27.648 | 21.45 | -2.52 | -0.06 | -23.38 | -0.06 | -26.37 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | 0.35 | 21.27 | -2.17 | 0 | -33.34 | 0.06 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.77 | 0.41 | 29.3 | -13.54 | 0 | -36.92 | -0.06 | -42.83 | 0 | -35.33 | -0.06 | -28.54
27.68 | 21.45 | -1.52 | -0.06 | -21.5 | 0 | -24.38 | -0.06 | 0 | -0.06 | 17.64 | 0.29 | 21.33 | 0.35 | 21.33 | -1.17 | -0.06 | -30 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 28.42 | 0.35 | 28.65 | -11.13 | -0.18 | -38.15 | -0.06 | -41.43 | -0.06 | -42.42 | -0.06 | -38.44
27.712 | 21.45 | -0.64 | 0 | -19.22 | -0.06 | -22.15 | -0.18 | -2.46 | -0.06 | 0.06 | 0.18 | 21.33 | 0.47 | 21.27 | -0.76 | -0.06 | -27.6 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.29 | 28.54 | -9.08 | -0.06 | -37.56 | -0.06 | -36.27 | 0 | -41.02 | -0.06 | -42.54
27.744 | 21.45 | -0.35 | -0.06 | -16.47 | -0.06 | -19.98 | -0.06 | -70.73 | -0.06 | 0 | 0.18 | 21.27 | 0.35 | 21.15 | -0.53 | 0 | -27.19 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.65 | 0.23 | 28.71 | -6.5 | 0 | -35.1 | -0.06 | -33.63 | 0.06 | -36.39 | 0 | -39.2
27.776 | 21.5 | -0.18 | 0.06 | -15.41 | 0 | -17.99 | 0 | -51.27 | -0.06 | 0 | 0.29 | 21.21 | 0.29 | 21.33 | -0.18 | 0 | -25.43 | 0.12 | -5.1 | -0.18 | 0 | 0 | 0 | -0.06 | 28.65 | 0.35 | 28.65 | -4.39 | 0 | -32.76 | -0.06 | -32.81 | -0.06 | -33.87 | 0 | -34.86
27.808 | 21.45 | -0.18 | -0.06 | -14.12 | 0 | -15.82 | -0.06 | -42.66 | 0 | 0 | 0.06 | 21.27 | 0.35 | 21.33 | -0.06 | -0.06 | -23.32 | 0.12 | -69.2 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.42 | -2.7 | -0.06 | -30.82 | 0 | -32.7 | -0.18 | -33.11 | 0 | -33.17
27.84 | 21.45 | -0.18 | -0.06 | -12.6 | 0 | -14.36 | -0.06 | -37.62 | -0.06 | 0 | 0.12 | 21.39 | 0.29 | 21.04 | 0.12 | -0.06 | -21.21 | 0.18 | -55.26 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.35 | 28.48 | -1.82 | 8.5 | -28.54 | -0.06 | -31.11 | -0.06 | -32.34 | 0 | -32.52
27.872 | 21.5 | -0.35 | 0 | -10.66 | 0.06 | -12.6 | -0.06 | -32.76 | -0.06 | 0 | 0.12 | 21.21 | 0.18 | 21.39 | 0.18 | -0.18 | -19.75 | 0.18 | -34.69 | 0 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.29 | 28.3 | -0.7 | 15.65 | -27.25 | -0.06 | -28.77 | -0.06 | -30.35 | -0.18 | -31.99
27.904 | 21.45 | -0.47 | 3.63 | -9.14 | -0.06 | -11.48 | -0.06 | -32.34 | -0.18 | 0 | 0 | 21.45 | 0.18 | 21.39 | 0.12 | -0.06 | -18.46 | 0.06 | -27.72 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.48 | -0.59 | 26.19 | -24.9 | -0.18 | -27.07 | -0.06 | -28.48 | 0 | -29.94
27.936 | 21.45 | -0.18 | 11.72 | -7.97 | 0 | -9.67 | -0.06 | -32.64 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | 0.12 | 3.22 | -16.23 | 0.12 | -34.75 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.77 | 0.06 | 28.3 | -0.18 | 28.71 | -23.97 | -0.06 | -25.55 | -0.18 | -27.01 | -0.06 | -27.77
27.968 | 21.45 | -0.23 | 16.93 | -6.39 | -0.06 | -8.03 | -0.06 | -31.23 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.27 | 0.18 | 9.55 | -14.77 | 0.12 | -35.57 | -0.18 | -44.47 | -0.06 | 0 | -0.06 | 28.65 | 0.12 | 28.42 | -0.18 | 28.71 | -22.15 | 0.06 | -23.79 | -0.06 | -25.61 | -0.06 | -26.37
28 | 21.39 | -0.29 | 21.8 | -4.98 | 0.12 | -6.86 | -0.06 | -29.71 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.15 | 0.23 | 16.64 | -12.95 | 0.12 | -33.17 | -0.06 | -61.64 | -0.06 | -18.4 | -0.18 | 28.65 | 0.18 | 28.48 | 0.12 | 28.59 | -19.28 | 0 | -22.56 | -0.06 | -23.67 | 0 | -24.84
28.032 | 21.45 | -0.29 | 21.45 | -3.11 | 1.99 | -5.33 | -0.18 | -27.01 | 0 | 0 | 0 | 21.5 | 0.06 | 21.39 | 0.29 | 21.56 | -11.31 | 0.12 | -29.36 | -0.18 | -45.35 | -0.06 | -69.91 | -0.06 | 0.29 | 0.06 | 28.48 | 0.29 | 28.54 | -17.7 | 0 | -20.92 | 0 | -22.21 | 0 | -23.67
28.064 | 21.45 | -0.29 | 21.5 | -1.7 | 7.03 | -3.75 | -0.18 | -25.2 | -0.06 | 0 | 0 | 2.99 | 0.12 | 21.33 | 0.23 | 21.39 | -9.61 | 0 | -27.6 | -0.06 | -44.88 | -0.06 | -47.81 | -0.06 | 0.12 | 0 | 28.65 | 0.18 | 28.13 | -15.59 | 0 | -19.63 | -0.18 | -20.45 | 0 | -21.45
28.096 | 21.45 | -0.23 | 21.45 | -1.11 | 12.6 | -2.29 | -0.06 | -22.44 | 0 | 0 | -0.18 | 0.06 | 0.12 | 21.21 | 0.23 | 21.33 | -7.32 | -0.06 | -26.25 | 0 | -43.77 | -0.06 | -46.23 | -0.06 | 0 | 0.06 | 28.48 | 0.23 | 28.42 | -13.36 | 0.06 | -17.52 | -0.18 | -18.99 | 0 | -19.63
28.128 | 21.45 | -0.18 | 21.39 | -0.59 | 18.16 | -1.46 | 6.8 | -20.57 | -0.06 | -66.74 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.33 | -5.51 | 0 | -25.2 | -0.06 | -41.66 | -0.06 | -44.42 | 0 | 0.06 | 0.12 | 28.42 | 0.23 | 28.65 | -11.54 | -0.06 | -15.94 | -0.06 | -17.05 | 0 | -18.28
28.16 | 21.45 | -0.18 | 21.45 | -0.23 | 21.5 | -0.76 | 12.36 | -18.87 | 0 | -61.41 | -0.18 | 0 | -0.06 | 2.4 | 0.18 | 21.62 | -3.87 | 0.06 | -23.15 | -0.06 | -39.26 | -0.06 | -41.31 | 0 | 0 | 0.06 | 28.77 | 0.29 | 28.59 | -8.79 | 9.08 | -14.18 | 0 | -15.47 | 0 | -16.7
28.192 | 21.45 | -0.23 | 21.45 | -0.06 | 21.5 | -0.18 | 18.05 | -16.82 | 0 | -42.72 | -0.06 | 0 | -0.18 | 0.06 | 0.29 | 21.39 | -2.58 | 0 | -21.56 | -0.18 | -37.62 | -0.06 | -40.08 | -0.18 | 0 | 0 | 28.59 | 0.23 | 28.42 | -6.97 | 18.87 | -12.48 | -0.06 | -13.71 | 8.91 | -15.41
28.224 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | -0.06 | 21.68 | -15.18 | -0.06 | -37.85 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.09 | -1.88 | 0 | -19.69 | -0.06 | -35.27 | 0 | -37.56 | -0.18 | 0 | 0 | 28.48 | 0.18 | 28.48 | -4.63 | 26.37 | -11.02 | 8.67 | -12.31 | 16.11 | -13.65
28.256 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | -12.07 | 0 | -33.22 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -1.23 | 0 | -18.16 | -0.06 | -33.69 | 0 | -36.33 | -0.06 | 0 | -0.06 | 28.54 | 0.29 | 28.42 | -2.93 | 28.65 | -8.79 | 15.65 | -10.14 | 23.73 | -11.48
28.288 | 21.39 | -0.23 | 21.45 | -0.35 | 21.45 | -0.29 | 21.45 | -10.2 | 0 | -32.64 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.33 | -0.18 | -0.06 | -16.23 | -0.06 | -32.58 | 0 | -34.98 | -0.06 | 0 | 0 | 28.77 | 0 | 28.42 | -1.93 | 28.48 | -7.38 | 23.38 | -8.56 | 28.83 | -9.67
28.32 | 21.5 | -0.18 | 21.5 | -0.29 | 21.45 | -0.23 | 21.45 | -8.2 | -0.18 | -32.52 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.18 | 0 | -14.41 | -0.06 | -29.01 | 0 | -32.7 | 0 | 0 | 0 | 28.71 | 0 | 28.54 | -0.7 | 28.48 | -5.45 | 28.83 | -6.86 | 28.59 | -8.32
28.352 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | -0.41 | 21.45 | -5.86 | -0.06 | -31.41 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.33 | -0.06 | 0.12 | -12.83 | 0 | -27.77 | -0.06 | -30.24 | -0.06 | 0 | 0 | 28.42 | 0 | 28.48 | -0.47 | 28.3 | -3.75 | 28.65 | -5.39 | 28.59 | -6.86
28.384 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -0.35 | 21.5 | -3.98 | 0 | -30.24 | -0.06 | 0 | -0.12 | 0 | 0.12 | 21.27 | -0.06 | 0.06 | -10.96 | 0 | -24.84 | -0.06 | -28.24 | -0.06 | 0 | 0 | 28.83 | 0.12 | 28.65 | -0.18 | 28.54 | -2.52 | 28.42 | -3.98 | 28.42 | -5.16
28.416 | 0.06 | -0.18 | 21.39 | -0.35 | 21.45 | -0.23 | 21.5 | -1.93 | -0.06 | -28.24 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.27 | 0 | 0 | -8.67 | -0.18 | -23.79 | 0 | -26.37 | -0.06 | 0 | 0 | 20.68 | 0 | 28.42 | 0 | 28.54 | -1.64 | 28.42 | -2.17 | 28.48 | -3.22
28.448 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.5 | -1.35 | 0 | -25.37 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 0.06 | -7.09 | 0 | -21.45 | 0 | -24.43 | 0 | -22.21 | -0.06 | 0.18 | 0.06 | 28.48 | 0.23 | 28.48 | -0.82 | 28.65 | -1.7 | 28.59 | -2.4
28.48 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.23 | 21.45 | -0.53 | -0.06 | -23.38 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.33 | 0.18 | 7.09 | -5.33 | 0 | -19.04 | 0 | -22.79 | -0.18 | -79.51 | 0 | 0 | -0.06 | 28.65 | 0.18 | 28.48 | -0.64 | 28.65 | -0.94 | 28.54 | -1.29
28.512 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.18 | 21.45 | -0.23 | -0.06 | -20.74 | -0.18 | -74.18 | 0 | 0 | 0.06 | 21.39 | 0.23 | 12.89 | -4.34 | 0 | -17.05 | -0.06 | -20.1 | -0.06 | -53.91 | -0.18 | 0 | -0.06 | 28.71 | 0.23 | 28.42 | -0.29 | 28.54 | -0.64 | 28.65 | -1.17
28.544 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -0.06 | -0.06 | -18.93 | -0.06 | -54.79 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 18.34 | -2.58 | 0 | -15.18 | 0 | -17.87 | -0.06 | -45.94 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.59 | -0.18 | 28.48 | -0.47 | 28.36 | -0.35
28.576 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | 0.12 | -0.06 | -17.4 | 0 | -42.48 | 0 | -1.99 | 0 | 21.33 | 0.23 | 21.56 | -1.29 | -0.06 | -14.36 | -0.06 | -16.06 | 0.06 | -45.06 | -0.18 | 0 | -0.06 | 28.65 | 0.53 | 28.59 | 0 | 28.83 | 0 | 28.59 | -0.23
28.608 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.29 | 0.12 | -15 | -0.06 | -38.15 | -0.06 | -71.43 | 0.12 | 21.33 | 0.23 | 21.5 | -1.29 | 0 | -12.72 | 0.06 | -14.71 | -0.06 | -41.78 | -0.18 | 0 | 0 | 28.54 | 0.18 | 28.36 | 0 | 28.36 | -0.06 | 28.77 | -0.06
28.64 | 0 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.35 | -0.06 | -13.54 | -0.06 | -33.28 | -0.18 | -56.66 | 0.06 | 21.33 | 0.12 | 21.39 | -0.53 | -0.06 | -10.96 | 2.05 | -13.3 | -0.06 | -41.08 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.59 | 0.06 | 28.48 | -0.06 | 28.48 | 0
28.672 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.35 | 0.06 | -10.61 | -0.06 | -32.99 | -0.06 | -33.4 | 0 | 21.33 | 0.29 | 21.39 | -0.35 | -0.06 | -9.67 | 9.38 | -11.66 | 0 | -39.38 | -0.06 | 0 | -0.18 | 28.65 | 0.23 | 28.18 | 0 | 28.36 | 0.06 | 28.42 | 0.12
28.704 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | -0.23 | 21.45 | 0.29 | -0.06 | -8.56 | -0.06 | -32.99 | 0 | -27.6 | 0 | 21.33 | 0.29 | 21.39 | -0.23 | 2.4 | -8.2 | 16.64 | -9.9 | 0 | -36.45 | -0.06 | 0 | 0 | 18.28 | 0.23 | 28.48 | 0.18 | 28.54 | 0 | 28.42 | 0.12
28.736 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.41 | 0.35 | -6.15 | -0.06 | -31.88 | -0.06 | -35.1 | 0 | 21.33 | 0.23 | 21.39 | 0 | 10.31 | -7.21 | 24.43 | -8.91 | 0 | -34.63 | 0 | 0 | 0 | 0.12 | 0.23 | 28.59 | 0.12 | 28.59 | 0.18 | 28.59 | 0.18
28.768 | 0 | 0.06 | 21.39 | -0.18 | 21.39 | -0.18 | 21.45 | 0.35 | 7.97 | -3.98 | -0.06 | -29.77 | -0.06 | -35.16 | 0.06 | 21.33 | 0.35 | 21.39 | -0.06 | 17.81 | -5.86 | 28.54 | -7.38 | -0.06 | -32.46 | -0.06 | 0 | 0 | 0 | 0.18 | 28.65 | 0.12 | 28.59 | 0.23 | 28.65 | 0.23
28.8 | -0.06 | -0.06 | 0.12 | -0.18 | 21.45 | -0.18 | 21.45 | 0.41 | 15.35 | -2.23 | -0.18 | -27.77 | -0.06 | -32.23 | -0.06 | 21.33 | 0.23 | 21.39 | 0.23 | 25.43 | -4.51 | 28.89 | -6.5 | -0.06 | -30.82 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.36 | 0.12 | 28.59 | 0.29 | 28.59 | 0.23
28.832 | -40.55 | -0.06 | 0 | -0.06 | 21.39 | -0.06 | 21.39 | 0.29 | 20.8 | -1.11 | 0.12 | -25.37 | -0.06 | -29.01 | -0.06 | 0.12 | 0.18 | 21.09 | 0.12 | 28.71 | -3.57 | 28.48 | -4.98 | 8.73 | -28.95 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.54 | 0.18 | 28.54 | 0.12 | 28.59 | 0.06
28.864 | -55.67 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.35 | 21.39 | -0.94 | 0 | -23.03 | 0 | -27.6 | -0.06 | 0 | 0.23 | 21.33 | 0.23 | 28.59 | -2.99 | 28.48 | -4.04 | 15.94 | -27.54 | -0.06 | -60.76 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.59 | 0.06 | 28.65 | 0.23
28.896 | -40.08 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.41 | 21.39 | -0.06 | 0 | -21.15 | 0 | -26.72 | 0.06 | 0 | 0.23 | 21.33 | 0.23 | 28.54 | -2.46 | 28.48 | -3.22 | 26.37 | -24.32 | -0.06 | -66.1 | 0 | 0 | 0.12 | 28.54 | 0.06 | 28.07 | 0.06 | 28.42 | 0.12
28.928 | -39.67 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | 0.06 | 0 | -18.99 | -0.06 | -24.67 | 0 | 0 | 0.18 | 21.27 | 0.23 | 28.65 | -1.82 | 28.59 | -2.7 | 28.59 | -22.38 | 0 | -48.28 | -0.06 | 0 | 0 | 28.36 | 0.06 | 28.54 | 0.12 | 28.48 | 0.23
28.96 | -37.38 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 21.45 | 0.29 | 21.39 | 0.12 | 0.06 | -17.29 | 0.06 | -22.85 | 0.06 | 0 | 0.23 | 21.15 | 0.29 | 28.48 | -1.41 | 28.65 | -2.17 | 28.71 | -20.16 | -0.18 | -45.59 | -0.18 | 0 | -0.06 | 28.13 | 0.18 | 28.48 | 0.06 | 28.42 | 0.29
28.992 | -35.16 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.33 | 0.35 | 21.39 | 0.41 | -0.06 | -14.59 | -0.06 | -20.98 | -0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.89 | -0.94 | 28.59 | -1.11 | 28.42 | -18.69 | 0 | -40.9 | 0 | 0 | 0 | 28.36 | 0.12 | 28.18 | 0.12 | 28.42 | 0.12
29.024 | -33.28 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0.12 | 0.35 | 21.33 | 0.47 | -0.06 | -12.25 | 0 | -19.98 | 0 | 0 | 0.23 | 21.21 | 0.18 | 28.59 | -0.76 | 28.71 | -0.76 | 28.07 | -16.76 | 0 | -36.51 | 0 | 0 | 0 | 13.89 | 0.12 | 28.48 | 0.23 | 28.65 | 0.18
29.056 | -30.06 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.23 | 21.15 | 0.53 | -0.06 | -10.55 | 0.06 | -17.7 | -0.06 | 0 | 0.29 | 21.21 | 0.23 | 28.65 | -0.64 | 28.42 | -0.59 | 28.71 | -14.82 | -0.06 | -37.97 | -0.18 | 0 | -0.06 | 0.12 | 0 | 28.48 | 0.18 | 28.48 | 0.18
29.088 | -28.48 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | -0.06 | -8.56 | 0.06 | -15.94 | -0.06 | 0 | 0.18 | 21.27 | 0.29 | 28.71 | -0.47 | 28.42 | -0.23 | 28.48 | -13.24 | -0.06 | -37.85 | 0 | 0 | 0 | 0 | 0 | 28.48 | 0 | 0.23 | 0.06
29.12 | -27.01 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.27 | 0.41 | 0.18 | -6.56 | 0 | -14.3 | -0.06 | 0 | 0.18 | 21.39 | 0.23 | 28.59 | -0.29 | 28.54 | -0.35 | 28.89 | -12.13 | -0.18 | -36.62 | -0.06 | 0.06 | -0.06 | 0 | -0.06 | 0.18 | 0.06 | 0 | 0.18
29.152 | -24.2 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.29 | 7.38 | -3.75 | 0 | -12.42 | -0.06 | 0 | 0.18 | 21.45 | 0.23 | 28.65 | -0.35 | 28.59 | -0.23 | 29.06 | -10.61 | -0.18 | -34.98 | -0.06 | -59.83 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.29
29.184 | -22.79 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 13.3 | -2.11 | -0.06 | -10.72 | 0 | 0 | 0.18 | 21.27 | 0.18 | 28.13 | -0.29 | 28.95 | -0.35 | 28.36 | -8.79 | 0 | -32.52 | -0.06 | -66.33 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0
29.216 | -19.45 | 0 | -30.35 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 18.52 | -1.7 | 0 | -8.67 | 0.06 | 0 | 0.23 | 21.33 | 0.18 | 28.18 | -0.35 | 28.95 | -0.35 | 28.48 | -7.38 | 1.52 | -30.18 | -0.06 | -48.63 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06
29.248 | -18.16 | -0.06 | -60.41 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.62 | -0.53 | -0.06 | -6.8 | 0 | -5.98 | 0.18 | 21.39 | 0.18 | 28.3 | -0.29 | 28.54 | -0.23 | 28.48 | -6.09 | 11.31 | -28.71 | -0.06 | -45.82 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.06
29.28 | -16.23 | -0.06 | -44.83 | -0.18 | 0 | -0.06 | 0 | 0 | 21.33 | 0.29 | 21.39 | -0.35 | -0.06 | -5.22 | 0 | -69.08 | 0.12 | 21.39 | 0.23 | 28.42 | -0.35 | 28.42 | -0.29 | 28.42 | -5.33 | 18.87 | -27.13 | -0.06 | -40.84 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
29.312 | -15 | -0.06 | -40.02 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0 | 2.64 | -3.34 | -0.18 | -45.35 | 0.23 | 21.33 | 0.23 | 28.48 | -0.23 | 28.48 | -0.18 | 28.48 | -4.34 | 26.37 | -25.2 | -0.06 | -36.45 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0
29.344 | -13.48 | -0.06 | -38.97 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.18 | 10.55 | -2.64 | -0.06 | -29.88 | 0.12 | 21.33 | 0.18 | 28.54 | -0.18 | 28.65 | -0.23 | 28.59 | -3.34 | 28.48 | -23.15 | -0.18 | -37.21 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06
29.376 | -11.25 | 0 | -35.74 | -0.06 | -40.61 | -0.06 | 0 | -0.06 | 21.27 | 0.35 | 21.27 | 0.12 | 15.94 | -1.35 | -0.06 | -30.29 | 0.12 | 0.12 | 0.23 | 28.48 | -0.18 | 28.54 | -0.23 | 28.48 | -2.29 | 28.48 | -21.09 | 0 | -38.09 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0
29.408 | -9.84 | -0.06 | -33.81 | -0.06 | -56.37 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.39 | 0.18 | 21.45 | -0.94 | 0 | -35.16 | 0.12 | 0 | 0.23 | 28.3 | -0.23 | 28.36 | -0.18 | 28.36 | -1.7 | 28.54 | -19.28 | -0.06 | -36.27 | -0.06 | 0 | 0 | 0 | 0 | 0 | -0.06
29.44 | -8.38 | -0.06 | -32.4 | -0.18 | -40.67 | -0.06 | -35.51 | -0.06 | 21.39 | 0.23 | 21.39 | 0.53 | 21.5 | -0.41 | -0.06 | -33.69 | 0.12 | 0 | 0.23 | 28.42 | -0.18 | 28.65 | -0.18 | 28.65 | -1.41 | 28.77 | -17.4 | -0.06 | -34.75 | 0 | 0 | -0.06 | 0 | 0 | 0 | -0.06
29.472 | -6.97 | -0.06 | -30 | 0 | -39.2 | -0.18 | -57.19 | -0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.33 | -0.29 | -0.06 | -31.11 | 0.12 | 0 | 0.06 | 28.71 | -0.18 | 28.59 | -0.18 | 28.59 | -0.76 | 28.36 | -15.53 | -0.18 | -32.76 | 0 | -92.23 | -0.06 | 0 | 0 | 0 | -0.06
29.504 | -5.68 | -0.06 | -27.95 | -0.06 | -37.21 | 0 | -43.13 | -0.18 | 21.27 | 0.12 | 21.39 | 0.47 | 21.39 | -0.18 | 0 | -27.95 | 0.12 | 0 | 0.12 | 28.42 | -0.18 | 28.48 | -0.23 | 28.89 | -0.7 | 28.48 | -13.48 | 0 | -30.53 | -0.06 | -65.45 | 0 | 0 | 0.06 | -63.93 | -0.18
29.536 | -3.87 | 0 | -25.02 | -0.06 | -35.33 | -0.06 | -39.61 | 0 | 21.39 | 0.18 | 21.45 | 0.53 | 21.39 | 0 | -0.06 | -27.01 | 0 | 0 | 0.06 | 28.54 | -0.18 | 28.36 | -0.23 | 28.48 | -0.47 | 28.36 | -10.61 | 0 | -28.71 | 0 | -47.81 | 0 | -49.28 | -0.06 | -76.64 | -0.06
29.568 | -2.99 | -0.06 | -23.15 | -0.06 | -33.46 | -0.06 | -37.44 | 0 | 21.45 | 0.18 | 21.39 | 0.53 | 21.39 | 0.06 | 3.22 | -26.13 | 0.12 | 0 | 0.12 | 28.54 | -0.18 | 0.18 | -0.18 | 28.48 | -0.29 | 28.3 | -8.56 | -0.18 | -26.84 | -0.23 | -46.12 | -0.06 | -77.87 | 0 | -46 | -0.06
29.6 | -1.64 | 6.21 | -21.21 | -0.06 | -30.47 | -0.06 | -35.63 | -0.06 | 21.33 | -0.06 | 21.45 | 0.41 | 21.39 | 0 | 9.55 | -24.08 | 0.06 | 0 | 0.12 | 23.09 | -0.18 | 0 | -0.06 | 28.65 | -0.23 | 28.65 | -6.33 | 0 | -25.31 | -0.18 | -39.67 | -0.18 | -50.92 | 0 | -29.65 | -0.18
29.632 | -0.82 | 12.07 | -19.16 | -0.06 | -28.77 | -0.06 | -33.81 | 0 | 0.29 | -0.06 | 21.39 | 0.35 | 21.33 | 0 | 14.71 | -21.74 | 0.12 | 0 | 0.18 | 0.18 | -0.18 | 0 | -0.06 | 28.3 | -0.18 | 28.65 | -4.63 | 0 | -23.32 | -0.06 | -37.44 | 0 | -28.54 | 0 | -38.15 | -0.18
29.664 | -0.53 | 17.52 | -17.17 | -0.18 | -26.66 | -0.06 | -31.52 | -0.06 | 0.06 | -0.06 | 21.27 | 0.29 | 21.39 | 0.18 | 20.57 | -20.74 | -0.06 | 0 | 0.18 | 0 | -0.18 | 0 | -0.18 | 28.71 | -0.29 | 28.89 | -2.75 | -0.06 | -21.09 | -0.18 | -38.32 | -0.06 | -37.5 | -0.06 | -42.78 | -0.18
29.696 | -0.35 | 21.5 | -15.35 | -0.06 | -24.79 | -0.06 | -29.06 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.21 | 0.29 | 21.21 | -18.69 | -0.06 | 0 | 0.18 | 0.06 | -0.18 | 0 | 0 | 14.94 | -0.18 | 28.42 | -1.88 | 6.33 | -18.99 | 0 | -37.91 | -0.18 | -42.54 | -0.18 | -39.67 | -0.06
29.728 | -0.18 | 21.45 | -13.71 | -0.18 | -22.56 | -0.06 | -26.78 | -0.06 | 0 | 0 | 21.45 | 0.18 | 21.39 | 0.18 | 21.39 | -17.46 | -0.06 | 0 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0.18 | -0.29 | 28.59 | -0.53 | 14.18 | -17.23 | -0.06 | -36.68 | -0.18 | -39.26 | -0.06 | -35.45 | 0
29.76 | -0.23 | 21.45 | -12.07 | -0.06 | -20.22 | 0 | -24.61 | -0.06 | 0 | -0.18 | 21.39 | 0.18 | 21.33 | 0.23 | 21.27 | -15.18 | -0.06 | 0 | 0.12 | 0 | -0.06 | 0 | 0 | 0 | -0.35 | 28.48 | -0.53 | 24.2 | -15.23 | -0.18 | -34.69 | -0.18 | -36.04 | -0.06 | -32.7 | -0.18
29.792 | -0.53 | 21.5 | -10.72 | -0.06 | -17.7 | -0.06 | -23.2 | -0.18 | 0 | 0.06 | 21.33 | 0.18 | 21.39 | 0.18 | 21.5 | -13.13 | 0 | -51.68 | 0.06 | 0 | 0 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0 | 28.59 | -13.3 | -0.06 | -32.05 | -0.06 | -33.34 | 0 | -32.76 | 0
29.824 | -0.41 | 21.45 | -9.55 | -0.06 | -16.47 | -0.06 | -20.68 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.33 | 0.18 | 21.33 | -12.13 | -0.06 | -63.81 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | -0.06 | 28.3 | -10.96 | -0.18 | -30.29 | -0.06 | -33.05 | -0.18 | -32.52 | 0
29.856 | -0.35 | 21.45 | -7.56 | -0.06 | -15.06 | -0.18 | -18.34 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | 0.06 | 21.39 | -10.14 | 0.06 | -40.49 | 0.23 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.12 | 28.59 | -8.79 | -0.18 | -28.36 | -0.06 | -32.52 | -0.18 | -29.94 | 0
29.888 | -0.35 | 21.45 | -5.98 | -0.06 | -13.3 | -0.06 | -16.58 | -0.06 | 0 | -0.06 | 21.21 | 0.12 | 21.39 | 0.18 | 21.33 | -8.56 | -0.06 | -28.65 | 0.12 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.54 | 0.23 | 28.59 | -6.45 | 0 | -27.01 | -0.29 | -30.82 | -0.18 | -27.36 | -0.18
29.92 | -0.35 | 21.45 | -4.22 | -0.18 | -11.37 | 0 | -15 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.45 | 0.12 | 21.33 | -6.09 | -0.06 | -32.76 | 0.18 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.48 | 0.35 | 28.65 | -4.86 | -0.06 | -24.9 | -0.06 | -28.59 | -0.06 | -26.43 | -0.06
29.952 | -0.29 | 21.39 | -2.93 | -0.06 | -9.61 | -0.06 | -13.3 | 0 | 0 | 0.06 | 21.39 | -0.06 | 21.21 | 0.18 | 21.21 | -4.75 | -0.06 | -35.57 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.89 | 0.23 | 28.65 | -2.7 | -0.06 | -22.56 | -0.06 | -26.25 | -0.18 | -25.78 | -0.06
29.984 | -0.18 | 21.45 | -2.11 | 3.34 | -8.32 | -0.18 | -11.89 | 0 | 0 | 0 | 21.39 | -0.06 | 21.39 | 0.29 | 21.33 | -3.22 | -0.06 | -33.69 | 0.12 | 0 | 0 | -50.45 | 0 | 0 | -0.23 | 28.42 | 0.35 | 28.65 | -1.7 | -0.06 | -21.09 | -0.18 | -25.14 | -0.06 | -23.79 | -0.06
30.016 | -0.29 | 21.45 | -1 | 9.67 | -6.86 | 0.76 | -9.9 | 0 | 0 | -0.18 | 20.98 | -0.06 | 21.39 | 0.23 | 21.15 | -2.11 | -0.06 | -30.35 | 0.12 | 0 | 0 | -66.04 | -0.06 | 0 | -0.18 | 28.59 | 0.29 | 28.59 | -1 | -0.06 | -18.69 | -0.18 | -23.67 | -0.06 | -22.03 | -0.06
30.048 | -0.18 | 21.45 | -0.41 | 14.94 | -5.39 | 8.79 | -8.56 | -0.06 | -0.06 | -0.06 | 0.06 | -0.06 | 21.27 | 0.12 | 21.39 | -1.7 | -0.06 | -28.36 | 0.06 | -67.68 | 0 | -46.64 | -0.06 | 0 | -0.23 | 28.48 | 0.23 | 28.59 | -0.47 | -0.06 | -17.23 | -0.23 | -21.91 | 0 | -20.1 | 0
30.08 | -0.29 | 21.45 | -0.23 | 20.39 | -3.87 | 14.36 | -7.62 | 0.06 | -77.29 | 0.12 | 0 | -0.06 | 21.33 | 0.12 | 21.33 | -0.64 | 0 | -26.9 | 0.06 | -62.99 | -0.06 | -45.59 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.65 | -0.18 | -0.18 | -15.06 | -0.06 | -20.57 | -0.06 | -18.52 | 0
30.112 | -0.23 | 21.45 | -0.06 | 21.45 | -2.75 | 20.1 | -5.51 | 0 | -53.5 | 0.06 | 0 | -0.06 | 21.27 | 0.12 | 21.39 | -0.53 | -0.06 | -25.84 | 0.12 | -44.42 | -0.06 | -43.48 | -0.06 | 0.06 | -0.06 | 14.06 | 0.23 | 28.59 | 0 | 0 | -12.72 | -0.18 | -18.28 | -0.06 | -16.88 | 0
30.144 | -0.18 | 21.45 | -0.35 | 21.45 | -1.46 | 21.5 | -4.51 | -0.18 | -42.48 | 0 | 0 | 0 | 21.5 | 0 | 21.04 | -0.18 | 5.04 | -23.61 | 0.06 | -44.24 | 0 | -41.19 | 0 | -47.23 | -0.06 | 0.12 | 0.18 | 28.65 | 0.12 | 7.09 | -10.96 | -0.06 | -16.64 | -0.23 | -15.06 | -0.06
30.176 | -0.23 | 21.45 | -0.41 | 21.5 | -0.94 | 21.5 | -2.52 | -0.18 | -37.74 | -0.18 | 0 | -0.06 | 21.21 | 0.06 | 21.39 | -0.18 | 12.95 | -21.62 | 0.12 | -43.24 | 0 | -39.44 | -0.06 | -60.41 | -0.18 | 0 | 0.12 | 28.48 | 0.12 | 14.65 | -8.85 | 4.98 | -15.47 | 0 | -13.48 | 0
30.208 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.45 | -1.46 | -0.06 | -33.05 | -0.06 | 0 | 0 | 0.23 | 0.12 | 21.39 | 0.06 | 18.34 | -20.63 | 0.12 | -40.55 | -0.06 | -37.62 | -0.06 | -45.53 | -0.18 | 0.06 | 0.18 | 28.42 | 0.29 | 22.44 | -6.8 | 13.18 | -13.48 | 5.57 | -11.95 | 0
30.24 | -0.12 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | -0.94 | 0 | -32.99 | 0 | 0 | 0 | 0.06 | 0.06 | 21.33 | -0.06 | 21.56 | -18.52 | 0.06 | -38.61 | -0.06 | -34.92 | 0 | -46 | -0.18 | 0 | 0.12 | 28.59 | 0.29 | 28.83 | -4.45 | 20.86 | -11.95 | 13.77 | -10.2 | 7.79
30.272 | -0.06 | 21.39 | -0.35 | 21.45 | -0.18 | 21.5 | -0.53 | -0.18 | -32.76 | 0 | 0 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | -16.29 | 0.06 | -36.74 | -0.06 | -34.16 | -0.06 | -44.06 | -0.18 | 0 | 0 | 28.59 | 0.47 | 28.65 | -2.17 | 28.36 | -9.84 | 21.15 | -8.2 | 15.18
30.304 | -0.18 | 21.39 | -0.35 | 21.45 | -0.18 | 21.45 | -0.18 | 0 | -31.64 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.33 | -14.82 | 0.06 | -34.69 | -0.06 | -31.7 | 0 | -41.25 | -0.06 | 0 | 0.12 | 28.59 | 0.29 | 28.42 | -1.76 | 28.65 | -8.14 | 28.83 | -6.5 | 23.03
30.336 | -0.18 | 21.39 | -0.18 | 21.5 | -0.29 | 21.39 | -0.23 | -0.06 | -29.71 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | -13.01 | 0.06 | -33.22 | 0.06 | -29.83 | 0 | -39.08 | -0.18 | 0 | 0.06 | 28.59 | 0.35 | 28.3 | -0.53 | 28.36 | -6.68 | 28.42 | -5.16 | 28.59
30.368 | -0.23 | 21.45 | -0.23 | 21.5 | -0.29 | 21.39 | -0.29 | 0.12 | -27.36 | 0 | 0 | 0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | -11.6 | 0.06 | -31.23 | 0 | -28.13 | 0 | -37.03 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.36 | -0.53 | 28.59 | -5.04 | 28.48 | -3.63 | 28.65
30.4 | -0.06 | 21.39 | -0.23 | 21.45 | -0.23 | 21.5 | -0.23 | 2.58 | -24.43 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.27 | -9.02 | -0.06 | -29.06 | 0 | -25.37 | 0 | -35.98 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.48 | -0.06 | 28.83 | -3.69 | 28.59 | -2.29 | 28.36
30.432 | -0.18 | 21.45 | -0.35 | 21.39 | -0.23 | 21.45 | -0.29 | 8.97 | -22.68 | 0.12 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.04 | -7.56 | 0.06 | -27.25 | 0 | -23.5 | 0.06 | -34.28 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.59 | 0.18 | 28.36 | -2.23 | 28.54 | -1.52 | 28.83
30.464 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | -0.41 | 13.54 | -20.57 | 0.06 | -55.61 | -0.12 | 0 | 0 | 1 | 0.29 | 21.39 | -5.68 | -0.06 | -25.25 | -0.06 | -21.56 | 0.06 | -32.17 | -0.06 | 0 | 0 | 28.59 | 0.23 | 28.48 | 0.12 | 28.48 | -1.29 | 28.59 | -0.7 | 28.48
30.496 | -0.18 | 0.12 | -0.29 | 21.45 | -0.23 | 21.45 | -0.53 | 21.39 | -18.75 | -0.06 | -62.11 | -0.18 | 0 | 0.06 | 0 | 0.29 | 21.33 | -3.93 | 8.97 | -22.74 | -0.06 | -19.81 | 0 | -29.24 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | 0.29 | 28.48 | -1 | 28.59 | -0.53 | 28.42
30.528 | -0.18 | 0 | -0.41 | 21.45 | -0.23 | 21.45 | -0.41 | 21.39 | -17.23 | -0.06 | -45.47 | -0.06 | 0 | -0.06 | 0 | 0.35 | 21.33 | -2.99 | 14.41 | -20.86 | 0 | -17.81 | -0.06 | -27.19 | -0.06 | -1.05 | -0.06 | 28.59 | 0.06 | 28.65 | 0.35 | 28.42 | -0.41 | 28.48 | -0.41 | 28.36
30.56 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.39 | -0.23 | 21.39 | -14.41 | -0.06 | -38.32 | 0 | 0 | -0.06 | 0 | 0.18 | 21.33 | -1.7 | 19.86 | -19.22 | 0 | -15.7 | 0.12 | -25.84 | -0.18 | -92.76 | 0.06 | 28.36 | 0.12 | 28.65 | 0.29 | 28.48 | -0.53 | 28.48 | -0.18 | 28.36
30.592 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -12.54 | -0.06 | -33.05 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -1 | 21.39 | -17.23 | -0.06 | -14.24 | 0.06 | -22.97 | -0.06 | -65.8 | -0.06 | 0.18 | 0.18 | 28.59 | 0.35 | 28.48 | -0.18 | 28.54 | -0.18 | 28.24
30.624 | -0.06 | 0 | -0.29 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -10.72 | -0.06 | -32.46 | 0 | -30.06 | -0.06 | 0 | 0.23 | 21.33 | -0.59 | 21.39 | -15.18 | -0.06 | -12.89 | -0.06 | -21.86 | -0.06 | -47.7 | -0.06 | 0 | 0 | 28.3 | 0.35 | 28.48 | -0.18 | 28.36 | 0 | 28.42
30.656 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -8.44 | -0.06 | -32.93 | -0.06 | -75.94 | -0.06 | 0 | 0.23 | 21.33 | -0.47 | 21.39 | -13.89 | 7.32 | -11.31 | -0.06 | -19.16 | -0.06 | -44.36 | -0.18 | 0 | 0.12 | 28.48 | 0.29 | 28.65 | 0.06 | 28.71 | 0 | 28.42
30.688 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | -6.39 | 0 | -32.58 | 0.06 | -50.63 | -0.18 | 0 | 0.23 | 21.33 | -0.18 | 21.15 | -12.72 | 14.94 | -9.84 | 3.93 | -17.29 | -0.18 | -38.56 | 0 | 0 | 0 | 28.3 | 0.18 | 28.48 | 0.12 | 28.48 | 0.12 | 28.3
30.72 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.18 | 21.39 | -3.28 | -0.06 | -29.65 | 0 | -42.13 | 0 | 0 | 0.23 | 21.33 | -0.18 | 21.39 | -10.9 | 25.31 | -8.67 | 12.25 | -15.82 | -0.06 | -37.15 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.48 | 0.18 | 28.59 | 0.18 | 28.48
30.752 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | -0.18 | 21.33 | -1.7 | 0 | -27.72 | 0 | -36.68 | -0.06 | 0 | 0.18 | 21.21 | -0.06 | 21.5 | -9.61 | 28.83 | -7.09 | 22.33 | -13.89 | 0 | -36.86 | 0 | 0 | 0 | 28.36 | 0.06 | 28.48 | 0.12 | 28.48 | 0.18 | 28.42
30.784 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | -0.23 | 21.39 | -1.64 | 2.11 | -25.61 | -0.06 | -31.82 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 21.33 | -8.32 | 28.59 | -5.57 | 29.12 | -12.89 | 0 | -37.68 | -0.06 | 0 | 0.12 | 28.48 | 0.23 | 28.48 | 0.12 | 28.48 | 0.23 | 28.42
30.816 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.5 | -0.23 | 21.39 | -0.06 | 7.44 | -23.15 | -0.18 | -32.29 | -0.06 | 0 | 0.23 | 21.33 | 0.12 | 21.39 | -6.8 | 28.65 | -4.51 | 28.83 | -10.84 | 0 | -36.1 | 0 | 0 | 0 | 28.42 | 0.06 | 28.3 | 0.12 | 28.24 | 0.18 | 28.42
30.848 | -0.06 | 0 | -0.06 | 21.5 | -0.23 | 21.45 | -0.23 | 21.39 | -0.41 | 12.89 | -21.68 | -0.06 | -32.52 | 0.06 | 0 | 0.29 | 21.39 | 0.18 | 21.39 | -5.68 | 28.59 | -3.93 | 28.59 | -9.79 | 0 | -34.45 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.48 | 0 | 28.59 | 0.23 | 28.24
30.88 | -0.06 | 0 | -0.18 | 0.59 | -0.18 | 21.33 | -0.18 | 21.39 | 0.23 | 18.63 | -19.16 | 0 | -31.58 | -0.06 | -10.49 | 0.06 | 21.39 | 0.18 | 21.5 | -4.51 | 28.65 | -2.7 | 28.65 | -8.32 | 0 | -31.17 | 0 | 0 | -0.06 | 28.18 | 0 | 28.54 | 0.06 | 28.36 | 0.18 | 28.42
30.912 | -0.06 | -33.69 | -0.18 | 0 | -0.06 | 10.55 | -0.23 | 21.33 | 0.23 | 21.5 | -17.23 | -0.06 | -29.24 | 0 | -71.72 | 0.12 | 21.33 | 0.06 | 21.39 | -3.87 | 28.54 | -2.23 | 28.42 | -6.91 | -0.06 | -30.24 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | 0.18 | 28.59 | 0.23 | 28.42
30.944 | -0.06 | -60.53 | -0.18 | 0 | -0.06 | 0.12 | -0.23 | 21.33 | 0.29 | 21.39 | -15.47 | 0.06 | -26.37 | 0 | -53.15 | 0.18 | 21.39 | 0.06 | 21.39 | -2.99 | 28.48 | -1.64 | 28.54 | -5.92 | 0 | -28.07 | -0.06 | 0 | -0.06 | 28.3 | 0 | 28.01 | 0.12 | 28.59 | 0.23 | 28.42
30.976 | -0.06 | -40.9 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -13.3 | 0.12 | -24.61 | 0.06 | -29.24 | 0.18 | 21.39 | 0.18 | 21.33 | -2.29 | 28.48 | -1.35 | 28.89 | -4.63 | 0 | -25.78 | 0 | 0 | 0 | 28.07 | 0 | 28.42 | 0.12 | 28.48 | 0.35 | 28.48
31.008 | -0.06 | -39.55 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.45 | 0.47 | 21.39 | -11.25 | -0.06 | -22.68 | -0.06 | -29.77 | 0.18 | 21.33 | 0.12 | 21.39 | -1.7 | 28.48 | -0.94 | 28.42 | -3.75 | 0 | -24.61 | -0.06 | -42.19 | 0 | 28.24 | -0.06 | 28.18 | 0.06 | 28.48 | 0.12 | 28.42
31.04 | -0.06 | -38.15 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.53 | 21.39 | -8.56 | -0.06 | -20.39 | 0 | -34.86 | 0.18 | 0.18 | 0.12 | 21.33 | -1.29 | 28.54 | -0.7 | 28.36 | -2.87 | 0 | -22.97 | -0.06 | -72.72 | 0 | 0.18 | -0.06 | 28.48 | 0 | 28.48 | 0.12 | 28.42
31.072 | -0.06 | -35.39 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.53 | 21.39 | -6.39 | -0.06 | -18.46 | -0.18 | -34.81 | 0.12 | 0 | 0.12 | 21.33 | -1.05 | 28.18 | -0.47 | 28.3 | -2.17 | 0 | -20.92 | 0 | -48.69 | 0 | 0 | -0.06 | 0.29 | 0 | 28.48 | 0.12 | 28.42
31.104 | -0.06 | -34.22 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.53 | 21.39 | -4.34 | -0.06 | -16.47 | -0.06 | -32.58 | 0.12 | 0 | 0.18 | 21.21 | -0.76 | 28.48 | -0.53 | 28.48 | -1.7 | 0 | -18.93 | -0.06 | -47.23 | 0 | 0 | 0 | 0.12 | 0 | 0.29 | 0.12 | 28.42
31.136 | -0.06 | -31.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.27 | 0.41 | 21.33 | -2.23 | -0.06 | -14.24 | 0 | -28.59 | 0.06 | 0 | 0.18 | 21.33 | -0.64 | 28.13 | -0.29 | 28.48 | -1 | 3.87 | -16 | -0.18 | -40.72 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 0.18
31.168 | -0.06 | -29.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -1.99 | 1.99 | -12.07 | 0 | -27.31 | 0.06 | 0 | 0.18 | 21.33 | -0.35 | 28.48 | -0.29 | 28.65 | -0.76 | 12.25 | -14.24 | 0.06 | -36.39 | 0 | 0 | 0 | 0.06 | 0.06 | 0 | 0.18 | 0
31.2 | -0.06 | -27.01 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -0.76 | 6.21 | -10.08 | 0 | -26.25 | 0.06 | 0 | 0.18 | 21.27 | -0.41 | 28.3 | -0.23 | 28.48 | -0.53 | 19.98 | -12.6 | -0.18 | -38.67 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
31.232 | 0.06 | -25.67 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.33 | -0.06 | 13.83 | -8.32 | 0 | -24.14 | 0 | 0 | 0.12 | 21.33 | -0.53 | 28.48 | -0.29 | 28.36 | -0.35 | 27.42 | -10.55 | 0 | -39.08 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
31.264 | -0.06 | -22.91 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -0.06 | 19.4 | -5.22 | 0.12 | -22.38 | 0 | 0 | 0.23 | 21.33 | -0.41 | 28.42 | -0.23 | 28.3 | -0.29 | 28.48 | -8.14 | -0.06 | -37.03 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 0
31.296 | -0.06 | -20.22 | -0.06 | -4.22 | -0.06 | 0 | -0.18 | 0.18 | 0.18 | 21.33 | 0.29 | 21.56 | -3.34 | 8.5 | -20.39 | 0 | 0 | 0.06 | 21.27 | -0.47 | 28.42 | -0.29 | 28.42 | -0.29 | 28.54 | -5.39 | -0.18 | -34.86 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
31.328 | -0.06 | -18.34 | -0.06 | -62.52 | 0 | -1.46 | -0.18 | 0.06 | 0.18 | 21.33 | 0.29 | 21.39 | -1.88 | 13.89 | -19.04 | 0 | 0 | 0.18 | 21.33 | -0.29 | 28.65 | -0.29 | 28.48 | -0.23 | 28.65 | -3.69 | -0.06 | -32.76 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
31.36 | 0 | -16.88 | -0.06 | -50.74 | -0.06 | -54.14 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -1.17 | 19.45 | -17.52 | 0 | 0 | 0.23 | 10.25 | -0.35 | 28.3 | -0.23 | 28.59 | -0.18 | 28.83 | -2.05 | -0.18 | -31 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0
31.392 | -0.06 | -14.82 | -0.06 | -40.9 | -0.06 | -52.09 | -0.18 | 0 | 0.06 | 21.45 | 0.41 | 21.5 | -0.59 | 21.56 | -15.53 | 0.06 | 0 | 0.18 | 0.06 | -0.23 | 28.3 | -0.23 | 28.77 | -0.23 | 28.36 | -1.46 | -0.06 | -29.01 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | -0.06 | 0
31.424 | -0.06 | -13.13 | -0.06 | -38.5 | 0 | -40.37 | -0.18 | 0 | 0 | 21.27 | 0.53 | 21.27 | -0.18 | 21.5 | -13.83 | 0.06 | 0 | 0.18 | 0 | -0.23 | 28.36 | -0.18 | 28.65 | -0.23 | 28.48 | -0.76 | -0.18 | -27.6 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0
31.456 | -0.06 | -11.66 | -0.18 | -36.51 | -0.06 | -39.44 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 21.39 | 0 | 21.39 | -11.54 | 0 | -49.92 | 0.12 | 0 | -0.23 | 28.54 | -0.23 | 28.48 | -0.29 | 28.42 | -0.35 | 0 | -25.25 | 0.06 | -45.24 | 0 | 0 | -0.06 | 0 | 0 | 0
31.488 | -0.06 | -9.9 | -0.06 | -34.63 | -0.06 | -36.33 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 21.39 | 0.06 | 21.27 | -10.37 | -0.06 | -62.64 | 0.12 | 0 | -0.29 | 28.18 | -0.18 | 28.48 | -0.29 | 28.71 | -0.06 | 0 | -23.85 | -0.06 | -80.45 | -0.18 | -33.58 | -0.06 | 0 | 0.06 | 0
31.52 | 6.27 | -9.32 | -0.18 | -32.46 | 0 | -34.69 | -0.06 | 0 | -0.06 | 21.39 | 0.53 | 21.45 | 0.29 | 21.27 | -8.14 | 0.06 | -40.72 | 0.12 | 0 | -0.18 | 23.91 | -0.18 | 28.3 | -0.18 | 28.42 | 0.06 | -0.06 | -21.21 | 0 | -49.51 | 0 | -89.48 | -0.18 | -34.69 | 0 | 0.06
31.552 | 12.13 | -6.97 | 0.06 | -30 | -0.06 | -32.76 | -0.06 | 0 | -0.06 | 21.33 | 0.53 | 21.39 | 0.35 | 21.39 | -6.86 | 0 | -27.48 | 0.18 | 0 | -0.23 | 0.18 | -0.06 | 28.3 | -0.18 | 28.42 | 0.29 | 7.68 | -19.1 | 0 | -47.29 | 0 | -56.13 | -0.06 | -86.49 | 0 | -47.81
31.584 | 17.7 | -6.09 | -0.06 | -27.6 | -0.06 | -30.82 | -0.06 | 0 | 0 | 21.21 | 0.53 | 21.39 | 0.35 | 21.39 | -4.86 | 0.12 | -30.94 | 0.23 | 0 | -0.18 | 0 | -0.06 | 0.29 | -0.18 | 28.42 | 0.29 | 15.06 | -17.81 | 0 | -40.9 | -0.06 | -34.1 | -0.06 | -56.25 | -0.06 | -78.23
31.616 | 21.45 | -3.75 | -0.06 | -26.02 | -0.06 | -28.36 | -0.18 | 0 | -0.06 | 21.33 | 0.41 | 21.39 | 0.41 | 21.39 | -2.93 | 0 | -36.62 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0.06 | -0.18 | 28.59 | 0.23 | 25.2 | -15.59 | 0 | -36.45 | -0.06 | -34.4 | -0.06 | -30.29 | -0.18 | -41.66
31.648 | 21.45 | -2.7 | 0 | -23.91 | 0 | -26.49 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.45 | 0.35 | 21.39 | -2.46 | 0.06 | -33.69 | 0.23 | 0 | -0.18 | 0.06 | -0.06 | 0 | -0.18 | 28.65 | 0.29 | 28.71 | -13.83 | -0.06 | -37.09 | -0.06 | -42.42 | 0 | -34.1 | -0.06 | -28.54
31.68 | 21.5 | -1.76 | -0.06 | -22.15 | -0.06 | -24.38 | -0.06 | 0 | 0 | 20.8 | 0.35 | 21.45 | 0.41 | 21.39 | -1 | -0.18 | -30.47 | 0.12 | 0 | -0.18 | 0 | 0 | 0.06 | -0.18 | 28.65 | 0.23 | 28.65 | -11.31 | 0 | -38.44 | 0 | -41.31 | -0.06 | -42.66 | 0 | -37.03
31.712 | 21.5 | -1.05 | 0 | -19.16 | -0.06 | -22.15 | -0.06 | -0.76 | 0 | 0.06 | 0.23 | 21.39 | 0.47 | 21.33 | -1 | -0.06 | -28.24 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.59 | -9.32 | 0 | -36.62 | 0 | -36.97 | -0.18 | -41.43 | 0 | -42.36
31.744 | 21.45 | -0.23 | -0.06 | -17.17 | -0.06 | -20.39 | -0.18 | -67.62 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.33 | -0.53 | -0.06 | -27.07 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.77 | 0.29 | 28.54 | -7.15 | 0.06 | -34.98 | -0.06 | -33.93 | 0 | -35.74 | 0 | -40.08
31.776 | 21.45 | -0.29 | 0 | -15.65 | 0 | -17.58 | -0.18 | -52.03 | -0.06 | 0 | 0.18 | 21.5 | 0.29 | 21.39 | -0.23 | -0.06 | -25.61 | 0.18 | -2.93 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.35 | 28.48 | -4.86 | 0 | -33.69 | 0 | -33.05 | -0.06 | -33.34 | -0.06 | -34.75
31.808 | 21.45 | -0.18 | -0.18 | -14.18 | -0.06 | -16 | -0.06 | -41.6 | 0 | 0 | 0.06 | 21.45 | 0.29 | 21.33 | -0.06 | -0.06 | -23.38 | 0.18 | -66.68 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.35 | 28.65 | -3.28 | 1.99 | -31 | -0.06 | -32.4 | -0.06 | -33.17 | -0.06 | -33.28
31.84 | 21.45 | -0.18 | -0.06 | -12.6 | -0.18 | -14.53 | -0.06 | -37.91 | -0.06 | 0 | 0 | 21.27 | 0.29 | 21.27 | 0.12 | 0 | -21.91 | 0.06 | -57.01 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.77 | 0.35 | 28.48 | -2.11 | 7.44 | -28.89 | -0.06 | -30.76 | -0.18 | -32.81 | -0.06 | -32.76
31.872 | 21.45 | -0.35 | -0.06 | -10.66 | -0.06 | -12.77 | -0.06 | -33.17 | -0.06 | 0 | 0.18 | 21.39 | 0.23 | 21.33 | 0.06 | 0.06 | -20.16 | 0.12 | -35.86 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.29 | 28.3 | -1.05 | 14.88 | -27.01 | 0 | -28.95 | -0.06 | -31.47 | -0.06 | -32.23
31.904 | 21.45 | -0.29 | 3.4 | -9.49 | -0.06 | -11.48 | -0.06 | -32.76 | 0 | 0 | 0.12 | 21.39 | 0.12 | 21.15 | 0.12 | -0.06 | -18.69 | 0.12 | -27.72 | 0 | 0 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.24 | -0.41 | 25.2 | -25.55 | -0.06 | -27.66 | -0.06 | -28.65 | 0 | -30.47
31.936 | 21.45 | -0.41 | 11.48 | -7.62 | -0.06 | -9.61 | -0.06 | -32.52 | 0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.29 | 2.46 | -16.7 | 0.12 | -34.45 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.29 | 28.42 | -0.23 | 28.77 | -23.91 | -0.18 | -25.61 | -0.18 | -27.13 | -0.06 | -27.6
31.968 | 21.45 | -0.29 | 16.82 | -6.68 | 0 | -8.09 | -0.06 | -31.29 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.33 | 0.23 | 9.2 | -14.71 | 0 | -35.27 | -0.06 | -39.61 | -0.18 | 0 | -0.18 | 28.36 | 0.29 | 28.3 | -0.23 | 28.48 | -21.68 | 0 | -24.43 | -0.06 | -25.78 | -0.06 | -26.72
32 | 21.39 | -0.23 | 21.91 | -5.04 | 0 | -7.03 | -0.06 | -29.83 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.23 | 16.47 | -13.24 | 0.06 | -32.64 | -0.18 | -63.23 | 0 | -5.1 | 0 | 28.54 | 0.06 | 28.48 | 0.06 | 28.54 | -19.22 | 0 | -22.91 | -0.06 | -24.55 | 0 | -25.02
32.032 | 21.45 | -0.35 | 21.45 | -3.16 | 1.99 | -5.33 | -0.06 | -27.31 | -0.18 | 0 | 0 | 21.39 | 0.12 | 21.45 | 0.23 | 21.56 | -11.48 | 0.06 | -29.83 | -0.18 | -45.29 | -0.06 | -74.07 | -0.06 | 0.7 | 0.06 | 28.18 | 0.29 | 28.71 | -17.23 | -0.06 | -20.86 | -0.06 | -22.97 | -0.18 | -23.97
32.064 | 21.45 | -0.29 | 21.45 | -2.11 | 6.68 | -4.22 | -0.06 | -25.2 | -0.18 | 0 | 0.06 | 14 | 0 | 21.15 | 0.18 | 21.33 | -9.84 | 0.23 | -27.54 | 0 | -45.29 | -0.06 | -49.98 | -0.23 | 0.12 | 0.06 | 28.59 | 0.35 | 28.24 | -15.53 | -0.06 | -19.45 | 0 | -21.09 | 0 | -22.15
32.096 | 21.45 | -0.18 | 21.45 | -1.46 | 12.72 | -2.23 | -0.06 | -22.91 | -0.06 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0.23 | 21.39 | -7.85 | 0 | -26.9 | 0 | -43.71 | 0 | -45.7 | 0 | 0 | 0 | 28.65 | 0.23 | 28.36 | -13.71 | -0.06 | -17.81 | 0 | -18.69 | 0 | -20.16
32.128 | 21.45 | -0.18 | 21.45 | -0.7 | 17.93 | -1.41 | 5.8 | -20.98 | -0.06 | -60.3 | 0 | 0 | 0 | 21.56 | 0.29 | 21.39 | -5.98 | 0.12 | -25.72 | -0.06 | -40.84 | -0.06 | -44.77 | -0.06 | 0.06 | 0 | 28.48 | 0.23 | 28.42 | -11.6 | -0.06 | -16 | -0.06 | -17.23 | -0.06 | -18.75
32.16 | 21.39 | -0.18 | 21.45 | -0.29 | 21.5 | -1.05 | 11.6 | -19.16 | -0.06 | -63.81 | 0.06 | 0 | -0.06 | 13.54 | 0.23 | 21.27 | -4.39 | 0.18 | -23.03 | -0.06 | -39.14 | 0 | -41.95 | -0.06 | 0 | 0.12 | 28.3 | 0.29 | 28.48 | -8.85 | 7.73 | -14.41 | -0.06 | -15.82 | 1.99 | -17.11
32.192 | 21.45 | -0.23 | 21.5 | -0.29 | 21.45 | -0.29 | 17.23 | -17.17 | -0.06 | -42.78 | 0 | 0 | 0 | 0.06 | 0.23 | 21.21 | -3.11 | 0.12 | -21.09 | -0.06 | -37.03 | -0.06 | -39.79 | -0.18 | 0 | 0.06 | 28.54 | 0.35 | 28.42 | -6.62 | 17.81 | -12.48 | -0.06 | -14.36 | 7.85 | -15.7
32.224 | 21.45 | -0.29 | 21.39 | -0.18 | 21.5 | -0.18 | 22.15 | -15 | -0.06 | -37.74 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.33 | -2.11 | 0.12 | -20.1 | -0.06 | -35.51 | -0.06 | -37.91 | 0 | 0 | -0.06 | 28.42 | 0.35 | 28.65 | -4.86 | 25.31 | -11.13 | 6.97 | -12.48 | 15.12 | -13.83
32.256 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -12.54 | -0.06 | -33.05 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.5 | -0.94 | 0.06 | -17.81 | -0.18 | -34.1 | 0 | -37.15 | -0.18 | 0 | 0.12 | 28.18 | 0.35 | 28.42 | -3.11 | 28.77 | -8.85 | 14.65 | -9.9 | 22.97 | -11.78
32.288 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -10.14 | -0.06 | -32.81 | -0.06 | 0 | 0 | 0 | 0.23 | 21.33 | -0.94 | 0.12 | -16.11 | -0.06 | -32.23 | -0.06 | -34.28 | -0.06 | 0 | -0.06 | 28.54 | 0.18 | 28.48 | -1.93 | 28.54 | -7.21 | 22.5 | -8.56 | 28.77 | -10.02
32.32 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.45 | -8.38 | -0.06 | -32.76 | 0 | 0 | 0 | 0 | 0.23 | 21.39 | -0.23 | 0.06 | -14.41 | 0.06 | -29.42 | -0.06 | -33.17 | -0.18 | 0 | 0 | 28.48 | 0.12 | 28.42 | -0.7 | 28.48 | -5.63 | 28.89 | -7.15 | 28.59 | -8.67
32.352 | 21.45 | -0.23 | 21.45 | -0.41 | 21.39 | -0.35 | 21.45 | -6.33 | -0.18 | -31.64 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 0 | -13.24 | -0.06 | -27.72 | -0.06 | -30.18 | 0 | 0 | 0 | 28.59 | 0.06 | 28.83 | -0.53 | 28.48 | -3.87 | 28.48 | -5.57 | 28.65 | -6.86
32.384 | 21.5 | -0.18 | 21.45 | -0.41 | 21.39 | -0.35 | 21.45 | -4.04 | -0.06 | -30 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.18 | 0 | -11.02 | -0.06 | -25.37 | -0.06 | -28.54 | 0 | 0 | 0 | 28.83 | 0.18 | 28.42 | -0.18 | 28.71 | -2.64 | 28.48 | -3.87 | 28.54 | -5.16
32.416 | 0.12 | -0.06 | 21.39 | -0.23 | 21.45 | -0.29 | 21.5 | -1.88 | -0.06 | -28.07 | 0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | -0.06 | -0.06 | -9.02 | 0 | -23.85 | 0 | -26.66 | 0 | 0 | -0.18 | 23.03 | 0.06 | 28.42 | -0.06 | 28.54 | -1.46 | 28.65 | -2.4 | 28.48 | -3.16
32.448 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.5 | -1.23 | -0.06 | -26.08 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.33 | -0.06 | 0.06 | -7.27 | 0 | -22.09 | 0 | -24.08 | 0 | -8.91 | -0.06 | 0.18 | 0 | 28.59 | 0.23 | 28.65 | -1.23 | 28.42 | -1.46 | 28.48 | -2.23
32.48 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | -0.18 | 21.5 | -0.64 | -0.06 | -23.67 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.45 | 0.12 | 6.45 | -5.63 | 0 | -18.87 | 0 | -22.79 | 0 | -76.82 | 0 | 0 | -0.06 | 28.54 | 0.23 | 28.65 | -0.53 | 28.77 | -1.11 | 28.48 | -1.7
32.512 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.23 | 21.45 | -0.41 | 0 | -21.21 | -0.06 | -69.08 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 12.31 | -4.04 | -0.06 | -17.29 | -0.06 | -20.16 | -0.06 | -56.55 | 0 | 0 | -0.06 | 28.59 | 0.29 | 28.42 | -0.35 | 28.54 | -0.76 | 28.42 | -0.82
32.544 | 0 | -0.06 | 21.5 | -0.29 | 21.39 | -0.29 | 21.45 | -0.06 | -0.18 | -18.99 | -0.06 | -58.42 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 17.75 | -3.11 | 0 | -15.65 | 0 | -18.46 | -0.06 | -44.94 | 0 | 0 | -0.06 | 28.48 | 0.35 | 28.42 | -0.18 | 28.59 | -0.35 | 28.07 | -0.35
32.576 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.12 | -0.06 | -17.23 | -0.06 | -43.77 | -0.18 | 0 | 0.18 | 21.21 | 0.18 | 21.5 | -1.99 | -0.06 | -14.53 | 0.06 | -16.29 | -0.06 | -45.76 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.59 | -0.06 | 28.59 | -0.18 | 28.65 | -0.29
32.608 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.23 | 21.45 | 0.23 | -0.06 | -15.23 | -0.06 | -38.32 | -0.06 | -65.8 | 0.18 | 21.33 | 0 | 21.39 | -0.82 | -0.06 | -12.95 | 0.06 | -15.18 | -0.06 | -41.66 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.65 | 0 | 28.42 | -0.06 | 28.89 | -0.18
32.64 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.23 | -0.06 | -13.54 | 0 | -33.28 | 0 | -59.89 | 0.06 | 21.39 | 0.12 | 21.39 | -0.64 | 0 | -10.61 | 1.99 | -13.13 | -0.18 | -41.31 | -0.06 | 0 | 0 | 28.3 | 0.29 | 28.42 | 0.06 | 28.59 | 0 | 28.48 | -0.06
32.672 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.41 | -0.06 | -10.61 | 0 | -32.58 | -0.06 | -35.63 | 0 | 21.33 | 0.06 | 21.33 | -0.53 | -0.06 | -9.84 | 8.56 | -11.48 | 0 | -38.85 | -0.18 | 0 | 0 | 28.89 | 0.29 | 28.59 | 0.12 | 28.77 | 0 | 28.48 | -0.06
32.704 | 0 | 0 | 21.5 | -0.35 | 21.45 | -0.18 | 21.45 | 0.41 | -0.06 | -8.32 | -0.06 | -32.52 | -0.18 | -27.19 | 0 | 21.39 | 0.12 | 21.27 | -0.18 | 0.06 | -8.38 | 15.47 | -10.31 | 0 | -37.21 | -0.06 | 0 | -0.06 | 23.91 | 0.23 | 28.54 | 0 | 28.42 | 0.06 | 28.59 | 0.06
32.736 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.29 | 0 | -6.56 | 0 | -31.99 | 0 | -34.28 | -0.06 | 21.21 | 0.18 | 21.39 | -0.06 | 9.49 | -7.32 | 23.56 | -9.32 | -0.06 | -34.16 | -0.06 | 0 | 0 | 0.18 | 0.18 | 28.42 | 0.23 | 28.83 | 0.18 | 28.48 | 0.12
32.768 | -41.08 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 21.45 | -0.47 | -0.06 | -21.45 | -0.06 | -26.37 | 0 | 0 | 0.18 | 21.39 | 0.18 | 28.42 | -2.46 | 28.42 | -3.11 | 26.02 | -24.43 | -0.06 | -67.62 | -0.18 | 0 | 0.06 | 28.54 | 0.06 | 28.3 | 0.12 | 28.48 | 0.12
32.8 | -39.44 | -0.18 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.29 | 21.39 | -0.06 | -0.06 | -19.63 | -0.06 | -25.2 | -0.06 | 0.06 | 0.18 | 21.5 | 0.12 | 28.24 | -1.82 | 28.65 | -2.58 | 28.36 | -22.97 | 0 | -48.46 | 0 | 0 | 0 | 28.36 | 0 | 28.48 | 0.06 | 28.48 | 0.23
32.832 | -38.5 | 0 | 0 | -0.18 | 0.06 | -0.18 | 21.39 | 0.35 | 21.39 | 0.06 | -0.06 | -17.93 | -0.06 | -23.15 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 28.48 | -1.46 | 28.65 | -1.99 | 28.65 | -20.57 | 0 | -46.06 | -0.18 | 0 | 0.12 | 28.71 | 0.06 | 28.18 | 0.06 | 28.48 | 0.23
32.864 | -34.75 | -0.06 | 0 | -0.18 | 0 | -0.06 | 17.64 | 0.23 | 21.39 | 0.18 | -0.06 | -15.18 | -0.06 | -21.09 | 0 | 0 | 0.18 | 21.39 | 0.18 | 28.65 | -1.05 | 28.42 | -1.29 | 28.48 | -17.99 | -0.06 | -39.08 | 0 | 0 | 0 | 28.3 | 0.06 | 28.77 | 0.12 | 28.48 | 0.18
32.896 | -32.76 | -0.06 | 0 | -0.12 | 0 | -0.06 | 0.12 | 0.29 | 21.39 | 0.29 | -0.06 | -13.13 | 0.06 | -19.45 | 0 | 0 | 0.12 | 21.27 | 0.12 | 28.83 | -0.94 | 28.42 | -0.94 | 28.54 | -16.47 | -0.06 | -36.97 | 0 | 0 | 0 | 23.26 | 0 | 28.13 | 0.23 | 28.48 | 0.18
32.928 | -30.7 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | 0.41 | -0.06 | -11.13 | -0.06 | -17.93 | -0.06 | 0 | 0.23 | 21.45 | 0.06 | 28.59 | -0.7 | 28.65 | -0.64 | 28.42 | -14.77 | -0.12 | -37.85 | 0 | 0 | 0 | 0.18 | -0.06 | 28.3 | 0.06 | 28.48 | 0.12
32.96 | -29.06 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.33 | 0.35 | -0.06 | -9.08 | -0.18 | -16 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.24 | -0.47 | 28.42 | -0.53 | 28.42 | -13.36 | 0 | -37.91 | -0.18 | 0 | 0 | 0 | 0 | 28.48 | 0.06 | 0.23 | 0.12
32.992 | -27.01 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.53 | 2.05 | -6.91 | -0.18 | -14.59 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.48 | -0.35 | 28.59 | -0.29 | 28.48 | -12.07 | -0.06 | -36.51 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.23 | 0.18 | 0 | 0.18
33.024 | -24.2 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 7.21 | -4.16 | -0.06 | -12.36 | 0 | 0 | 0.29 | 21.33 | 0.18 | 28.42 | -0.29 | 28.42 | -0.35 | 28.65 | -10.08 | 0 | -33.69 | -0.18 | -78.23 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.18
33.056 | -22.21 | -0.06 | -0.06 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 12.72 | -2.46 | -0.06 | -10.37 | 0.06 | 0 | 0.18 | 21.39 | 0.23 | 28.42 | -0.29 | 28.3 | -0.29 | 28.48 | -8.67 | -0.18 | -31.93 | -0.06 | -69.49 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.12
33.088 | -20.1 | -0.06 | -27.83 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 18.28 | -1.7 | -0.06 | -9.08 | -0.18 | 0 | 0.12 | 21.39 | 0.23 | 28.65 | -0.35 | 28.89 | -0.18 | 28.65 | -7.97 | 0.88 | -30.24 | -0.18 | -49.22 | 0.06 | 0 | 0 | 0 | 0.18 | 0 | 0.12
33.12 | -18.22 | -0.18 | -61.82 | -0.18 | 0 | -0.18 | 0 | 0 | 21.21 | 0.35 | 21.68 | -0.76 | -0.06 | -6.86 | -0.06 | -3.46 | 0.06 | 21.45 | 0.23 | 28.59 | -0.23 | 28.3 | -0.23 | 28.42 | -6.39 | 10.55 | -28.83 | -0.18 | -46.53 | -0.18 | 0 | -0.18 | 0 | 0.18 | 0 | 0.12
33.152 | -16.11 | 0.06 | -45 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.39 | -0.41 | 0 | -5.39 | -0.06 | -70.78 | 0.06 | 21.39 | 0.35 | 28.3 | -0.23 | 28.83 | -0.23 | 28.48 | -5.33 | 18.11 | -26.9 | -0.18 | -41.54 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06
33.184 | -14.47 | -0.06 | -39.55 | -0.18 | 0 | -0.06 | 0 | 0 | 21.45 | 0.12 | 21.33 | -0.06 | 4.22 | -3.46 | -0.06 | -46.47 | 0.12 | 21.39 | 0.06 | 28.13 | -0.29 | 28.48 | -0.18 | 28.36 | -3.98 | 28.36 | -24.38 | -0.18 | -35.98 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06
33.216 | -13.07 | -0.06 | -38.15 | -0.06 | -0.06 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.15 | 0.12 | 10.25 | -2.4 | 0 | -29.94 | 0.12 | 21.33 | 0.23 | 28.65 | -0.23 | 28.42 | -0.23 | 28.13 | -3.28 | 28.71 | -22.68 | -0.18 | -37.62 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12
33.248 | -11.31 | -0.06 | -35.57 | -0.06 | -36.86 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.12 | 15.7 | -1.29 | -0.06 | -29.65 | 0.12 | 0.12 | 0.18 | 28.65 | -0.18 | 28.59 | -0.29 | 28.59 | -2.7 | 28.54 | -21.09 | -0.06 | -37.62 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06
33.28 | -10.2 | -0.06 | -34.28 | -0.06 | -57.42 | -0.06 | -0.12 | 0 | 21.45 | 0.06 | 21.33 | 0.35 | 21.45 | -0.94 | -0.06 | -34.75 | 0.18 | 0 | 0.23 | 28.48 | -0.06 | 28.42 | -0.18 | 28.48 | -1.88 | 28.65 | -19.28 | -0.06 | -36.68 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06
33.312 | -8.67 | 0.06 | -31.52 | -0.06 | -39.96 | -0.06 | -30.06 | -0.06 | 21.39 | 0.23 | 21.39 | 0.35 | 21.45 | -0.53 | -0.06 | -34.51 | 0.12 | 0 | 0.18 | 28.65 | -0.18 | 28.42 | -0.18 | 28.59 | -1.46 | 28.71 | -17.4 | -0.06 | -35.51 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12
33.344 | -6.45 | -0.06 | -29.53 | -0.18 | -39.44 | -0.06 | -55.26 | 0 | 21.39 | 0.06 | 21.39 | 0.47 | 21.33 | -0.06 | 0.06 | -31.11 | 0.12 | 0 | 0.18 | 28.65 | -0.18 | 28.65 | -0.18 | 28.48 | -0.76 | 28.59 | -14.94 | 0 | -32.34 | -0.06 | -86.84 | -0.23 | 0 | 0 | 0 | -0.06
33.376 | -5.68 | -0.06 | -27.19 | -0.06 | -37.79 | -0.18 | -41.25 | 0.06 | 21.39 | 0.12 | 21.33 | 0.35 | 21.33 | -0.18 | 0 | -28.65 | 0.06 | 0 | 0.12 | 28.65 | -0.18 | 28.42 | -0.18 | 28.36 | -0.53 | 28.65 | -12.77 | -0.06 | -30.24 | 0 | -69.2 | -0.06 | 0.06 | 0 | -51.74 | 0
33.408 | -3.87 | 0 | -25.37 | -0.06 | -34.92 | -0.06 | -39.44 | 0 | 21.39 | 0 | 21.39 | 0.47 | 21.39 | 0.12 | -0.06 | -27.13 | 0.06 | 0 | 0.29 | 28.71 | -0.18 | 28.42 | -0.06 | 28.42 | -0.53 | 28.83 | -10.84 | -0.06 | -28.54 | 0 | -48.28 | -0.06 | -40.43 | -0.18 | -80.39 | 0
33.44 | -2.75 | 1.99 | -23.15 | -0.06 | -33.81 | 0 | -37.79 | 0 | 21.39 | 0.06 | 21.39 | 0.47 | 21.33 | 0.06 | 2.99 | -26.6 | 0.12 | 0 | 0.12 | 28.3 | -0.18 | 0.23 | -0.18 | 28.48 | -0.41 | 28.59 | -8.67 | 0 | -26.66 | 0 | -46.41 | -0.06 | -80.92 | -0.06 | -48.93 | -0.06
33.472 | -1.7 | 5.68 | -21.45 | 0 | -30.41 | -0.06 | -35.92 | -0.06 | 21.39 | 0 | 21.33 | 0.47 | 21.33 | 0.18 | 9.38 | -24.38 | 0.06 | 0 | 0.23 | 28.59 | -0.06 | 0 | -0.18 | 28.89 | -0.23 | 28.77 | -6.62 | -0.18 | -25.14 | 0 | -40.67 | -0.06 | -53.85 | 0.06 | -29.3 | 0
33.504 | -0.7 | 13.07 | -18.87 | -0.06 | -29.3 | -0.06 | -33.52 | -0.06 | 3.52 | 0.06 | 21.39 | 0.35 | 21.39 | 0.18 | 16.35 | -22.33 | 0 | 0 | 0.18 | 0.12 | -0.18 | 0 | -0.18 | 28.42 | -0.29 | 28.48 | -4.04 | -0.18 | -23.2 | -0.06 | -37.62 | -0.18 | -28.71 | 0 | -36.39 | 0
33.536 | -0.23 | 18.99 | -16.29 | -0.06 | -27.36 | -0.06 | -31.29 | 0 | 0.06 | 0 | 21.27 | 0.29 | 21.21 | 0.18 | 21.62 | -20.68 | 0.06 | 0 | 0.12 | 0 | -0.06 | 0 | -0.18 | 28.59 | -0.29 | 28.48 | -2.4 | -0.06 | -21.09 | -0.06 | -37.44 | -0.06 | -36.21 | -0.18 | -42.6 | 0
33.568 | -0.18 | 21.62 | -15.29 | -0.06 | -25.02 | -0.06 | -29.65 | 0.06 | 0 | 0 | 21.33 | 0.35 | 21.21 | 0.18 | 21.39 | -19.22 | 0.06 | 0 | 0.18 | 0 | -0.06 | 0 | -0.06 | 17.23 | -0.35 | 28.42 | -1.52 | 8.44 | -19.51 | -0.18 | -38.15 | -0.18 | -42.6 | 0.06 | -40.37 | 0
33.6 | -0.18 | 21.45 | -13.71 | -0.06 | -22.91 | 0 | -27.01 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.29 | 21.39 | -17.93 | 0.06 | 0 | 0.23 | 0 | -0.18 | 0 | -0.06 | 0.18 | -0.23 | 28.3 | -1 | 15.65 | -17.17 | -0.06 | -36.51 | -0.06 | -40.61 | 0 | -35.74 | -0.18
33.632 | -0.18 | 21.45 | -12.36 | -0.06 | -20.1 | -0.06 | -25.25 | -0.18 | 0 | -0.06 | 21.27 | 0.23 | 21.39 | 0.23 | 21.39 | -15.41 | 0.12 | 0 | 0.29 | 0 | -0.06 | 0 | -0.06 | 0 | -0.29 | 28.48 | -0.47 | 23.32 | -15.23 | -0.06 | -34.69 | 0 | -35.92 | -0.18 | -32.81 | -0.18
33.664 | -0.35 | 21.45 | -10.49 | -0.06 | -17.93 | -0.06 | -22.21 | -0.18 | 0 | -0.06 | 21.39 | 0.12 | 21.33 | 0.23 | 21.15 | -13.54 | 0.06 | -45.53 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0.06 | -0.29 | 28.3 | -0.23 | 28.95 | -13.18 | 0.06 | -31.99 | -0.06 | -33.69 | 0 | -32.93 | 0
33.696 | -0.41 | 21.45 | -8.79 | -0.06 | -16.29 | 0 | -20.63 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.33 | 0.18 | 21.27 | -12.07 | 0.06 | -65.22 | 0.18 | 0 | -0.06 | 0 | 0 | 0 | -0.23 | 28.54 | 0.06 | 28.71 | -10.78 | 0 | -30.53 | -0.18 | -33.11 | 0.06 | -32.64 | -0.06
33.728 | -0.41 | 21.45 | -7.27 | -0.06 | -15.18 | -0.06 | -19.16 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 21.39 | -10.31 | 0.06 | -42.07 | 0.23 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.36 | 0.12 | 28.65 | -9.08 | -0.18 | -28.77 | -0.06 | -32.52 | -0.06 | -30.7 | 0
33.76 | -0.23 | 21.39 | -5.68 | -0.06 | -13.36 | -0.06 | -16.47 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | 0.29 | 21.33 | -8.56 | 0 | -28.83 | 0.23 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.54 | -6.91 | -0.18 | -26.84 | -0.06 | -30.53 | 0 | -28.24 | -0.06
33.792 | -0.23 | 21.45 | -4.63 | -0.06 | -11.48 | 0 | -15.29 | -0.06 | 0 | -0.06 | 21.09 | 0.06 | 21.39 | 0.29 | 21.39 | -6.33 | 0.06 | -32.05 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.54 | 0.29 | 28.83 | -4.86 | -0.06 | -24.84 | -0.06 | -28.71 | 0 | -26.84 | -0.06
33.824 | -0.23 | 21.5 | -3.16 | -0.18 | -9.61 | 0 | -13.54 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.39 | 0.18 | 21.39 | -4.39 | -0.06 | -35.51 | 0.23 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.36 | 0.29 | 28.48 | -2.64 | 0 | -22.97 | -0.18 | -26.54 | -0.06 | -25.61 | -0.18
33.856 | -0.35 | 21.45 | -1.76 | 1.29 | -8.5 | 0 | -11.37 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | 0.23 | 21.45 | -2.99 | 0.06 | -34.28 | 0.06 | 0 | 0 | -41.37 | -0.06 | 0 | -0.29 | 28.54 | 0.29 | 28.65 | -1.7 | 0 | -21.09 | -0.06 | -25.43 | 0.12 | -23.67 | -0.06
33.888 | -0.35 | 21.5 | -1.11 | 8.97 | -7.32 | 0.29 | -10.31 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.09 | 0.23 | 21.39 | -2.23 | 0 | -30.35 | 0.12 | 0 | -0.06 | -68.85 | 0 | 0 | -0.23 | 28.65 | 0.41 | 28.48 | -1.11 | 0 | -18.87 | -0.18 | -24.08 | -0.06 | -22.38 | 0
33.92 | -0.18 | 21.39 | -0.53 | 14.53 | -5.45 | 7.79 | -9.02 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.39 | 0.12 | 21.04 | -1.41 | 0.06 | -28.18 | 0.18 | -58.24 | -0.06 | -48.46 | -0.06 | 0 | -0.06 | 28.65 | 0.23 | 28.3 | -0.59 | 0 | -17.23 | 0 | -21.74 | 0 | -20.51 | 0
33.952 | -0.35 | 21.39 | -0.18 | 20.16 | -3.75 | 13.59 | -7.38 | -0.06 | -75.76 | -0.18 | 0 | 0.12 | 21.39 | 0.12 | 21.33 | -0.82 | -0.06 | -26.72 | 0.18 | -65.8 | 0 | -45.7 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.18 | -0.06 | -0.06 | -15.41 | -0.18 | -20.39 | 0.06 | -18.63 | 0
33.984 | -0.18 | 21.45 | -0.18 | 21.56 | -2.81 | 21.09 | -5.45 | -0.06 | -56.6 | -0.06 | 0 | -0.18 | 21.56 | 0.12 | 21.39 | -0.47 | -0.06 | -25.9 | 0.06 | -44.83 | -0.06 | -43.65 | -0.06 | -0.18 | -0.06 | 0.29 | 0.18 | 28.3 | 0.06 | 0 | -12.83 | -0.06 | -18.99 | 0 | -17.17 | 0
34.016 | -0.23 | 21.45 | -0.18 | 21.45 | -1.52 | 21.5 | -4.16 | -0.06 | -43.07 | -0.06 | 0 | -0.18 | 21.33 | 0.12 | 21.45 | -0.23 | 6.27 | -24.14 | 0.06 | -44.71 | 0 | -41.08 | -0.06 | -72.19 | -0.18 | 0.12 | 0.23 | 28.59 | 0.35 | 5.68 | -10.96 | -0.06 | -16.88 | 0 | -15.65 | -0.06
34.048 | -0.29 | 21.45 | -0.29 | 21.5 | -0.94 | 21.45 | -2.75 | -0.06 | -38.56 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -0.18 | 11.95 | -21.68 | 0 | -43.24 | 0 | -40.02 | -0.18 | -62.17 | -0.06 | 0 | 0.18 | 28.42 | 0.41 | 13.71 | -8.61 | 3.87 | -15.47 | 0 | -14 | 0
34.08 | -0.18 | 21.45 | -0.41 | 21.5 | -0.41 | 21.45 | -1.64 | -0.18 | -33.05 | -0.06 | 0 | -0.06 | 0.23 | 0.06 | 21.27 | -0.18 | 17.52 | -19.98 | 0.06 | -40.37 | 0 | -38.32 | -0.18 | -44.83 | -0.18 | 0 | 0.12 | 28.65 | 0.41 | 21.45 | -6.62 | 12.13 | -13.54 | 4.75 | -12.07 | 0
34.112 | -0.18 | 21.39 | -0.35 | 21.45 | -0.06 | 21.45 | -1.11 | -0.18 | -32.4 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.39 | 0.23 | 21.68 | -18.63 | 0.12 | -38.85 | 0 | -35.86 | 0 | -46.17 | -0.18 | 0 | 0.23 | 28.59 | 0.41 | 28.83 | -4.51 | 19.81 | -12.25 | 13.07 | -10.31 | 6.33
34.144 | -0.18 | 21.39 | -0.35 | 21.45 | -0.18 | 21.5 | -0.47 | -0.06 | -32.76 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.23 | 21.39 | -16.82 | 0.06 | -36.27 | 0 | -34.86 | 0 | -43.07 | -0.06 | 0 | 0.06 | 28.42 | 0.41 | 28.48 | -2.29 | 29.06 | -10.31 | 23.15 | -8.61 | 16.76
34.176 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -0.18 | -0.06 | -31.7 | 0 | 0 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | -14.94 | 0.06 | -35.1 | 0.06 | -32.52 | 0 | -42.25 | 0 | 0 | 0.06 | 28.42 | 0.47 | 28.54 | -1.88 | 28.48 | -8.2 | 29.01 | -6.8 | 24.32
34.208 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | -0.06 | -29.94 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.21 | 0.35 | 21.39 | -12.95 | 0.12 | -33.69 | 0 | -29.94 | 0 | -38.91 | -0.18 | 0 | 0.12 | 28.36 | 0.41 | 28.65 | -0.7 | 28.42 | -6.68 | 28.48 | -5.39 | 28.77
34.24 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.45 | -0.23 | 0 | -27.83 | 0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.29 | 21.39 | -12.07 | 0 | -31.41 | 0 | -27.89 | 0 | -37.33 | 0 | 0 | 0.12 | 28.48 | 0.35 | 28.83 | -0.35 | 28.48 | -5.51 | 28.48 | -3.69 | 28.59
34.272 | -0.06 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | -0.41 | 2.29 | -25.08 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -9.61 | 0 | -29.36 | -0.06 | -25.72 | 0 | -36.04 | -0.06 | 0 | 0 | 28.54 | 0.29 | 28.77 | -0.18 | 28.42 | -3.93 | 28.48 | -2.23 | 28.59
34.304 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | 21.5 | -0.29 | 10.14 | -23.32 | 0.06 | 0.06 | -0.18 | 0 | 0 | 21.39 | 0.29 | 21.45 | -7.38 | 0.06 | -26.84 | 0.06 | -23.61 | 0 | -33.22 | -0.06 | 0 | 0.06 | 28.65 | 0.23 | 28.42 | 0.18 | 28.48 | -2.23 | 28.65 | -1.46 | 28.54
34.336 | -0.18 | 17.87 | -0.23 | 21.45 | -0.29 | 21.45 | -0.35 | 15.18 | -21.33 | -0.06 | -52.68 | -0.18 | 0 | 0.06 | 0.12 | 0.23 | 21.39 | -5.98 | 2.29 | -25.08 | -0.06 | -21.68 | 0 | -31.29 | -0.06 | 0 | -0.06 | 28.65 | 0.18 | 28.83 | 0.12 | 28.65 | -1.46 | 28.48 | -1.17 | 28.48
34.368 | -0.06 | 0.12 | -0.18 | 21.5 | -0.23 | 21.45 | -0.41 | 20.92 | -19.22 | 0.06 | -62.4 | -0.06 | 0 | 0 | 0.06 | 0.23 | 21.45 | -4.28 | 8.03 | -23.2 | -0.06 | -19.51 | -0.06 | -29.59 | 0 | 0 | -0.06 | 28.54 | 0.23 | 28.54 | 0.35 | 28.65 | -1.11 | 28.83 | -0.41 | 28.65
34.4 | -0.18 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -17.23 | 0 | -44.47 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.15 | -2.93 | 13.54 | -21.09 | 0.06 | -17.93 | -0.06 | -27.42 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.48 | 0.35 | 28.3 | -0.7 | 28.42 | -0.41 | 28.42
34.432 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | -14.82 | -0.06 | -37.97 | -0.06 | 0 | 0 | 0 | 0.23 | 21.21 | -1.88 | 19.04 | -19.63 | 0 | -15.35 | 0 | -25.84 | 0 | -88.54 | 0.06 | 28.59 | 0.12 | 28.65 | 0.29 | 28.48 | -0.23 | 28.59 | -0.18 | 28.59
34.464 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | -12.89 | 0 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.27 | -1.11 | 21.45 | -16.82 | -0.06 | -14.47 | -0.06 | -23.32 | -0.06 | -57.66 | -0.06 | 0.29 | 0.06 | 28.48 | 0.41 | 28.59 | -0.18 | 28.59 | -0.06 | 28.65
34.496 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | -10.61 | -0.06 | -32.58 | -0.06 | -24.49 | -0.06 | 0 | 0.23 | 21.39 | -0.59 | 21.45 | -15 | -0.06 | -12.77 | -0.06 | -20.92 | -0.06 | -48.81 | -0.06 | 0.06 | -0.06 | 28.83 | 0.41 | 28.48 | -0.06 | 28.42 | -0.06 | 28.83
34.528 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.29 | 21.39 | -8.14 | 0 | -32.76 | -0.18 | -77.11 | -0.06 | 0 | 0.23 | 21.21 | -0.47 | 21.45 | -14.24 | 7.09 | -11.25 | -0.06 | -18.87 | -0.18 | -44.47 | -0.06 | 0 | -0.06 | 28.59 | 0.29 | 28.71 | -0.06 | 28.65 | 0.06 | 28.59
34.56 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.33 | -6.21 | -0.06 | -32.23 | -0.06 | -52.03 | 0.06 | 0 | 0.18 | 21.21 | -0.18 | 21.39 | -12.54 | 14.88 | -10.08 | 3.22 | -17.46 | -0.06 | -40.02 | 0 | 0 | 0.06 | 28.59 | 0.23 | 28.42 | 0.06 | 28.42 | 0.06 | 28.36
34.592 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -3.34 | -0.06 | -30.29 | -0.06 | -43.01 | -0.06 | 0 | 0.12 | 21.39 | -0.18 | 21.27 | -11.31 | 25.02 | -8.56 | 11.54 | -15.41 | -0.06 | -38.32 | -0.06 | 0 | 0 | 28.59 | 0.18 | 28.54 | 0.06 | 28.3 | 0.23 | 28.48
34.624 | 0 | 0 | -0.23 | 21.45 | -0.18 | 21.45 | -0.23 | 21.33 | -2.11 | -0.06 | -27.42 | -0.06 | -35.27 | -0.06 | 0 | 0.23 | 21.39 | -0.06 | 21.39 | -9.43 | 28.71 | -6.97 | 21.5 | -13.65 | 0 | -39.08 | -0.06 | 0 | -0.18 | 28.71 | 0.12 | 28.3 | 0.23 | 28.24 | 0.29 | 28.54
34.656 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.23 | 21.39 | -1.35 | 1.99 | -25.84 | -0.06 | -31.82 | 0.06 | 0 | 0.18 | 21.33 | 0.18 | 21.33 | -7.97 | 28.71 | -5.92 | 29.18 | -12.13 | 0.06 | -39.26 | -0.18 | 0 | 0.12 | 28.54 | 0.12 | 28.42 | 0.18 | 28.36 | 0.18 | 28.59
34.688 | 0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.76 | 7.09 | -23.61 | -0.06 | -32.93 | 0 | 0 | 0.12 | 21.39 | 0.18 | 21.62 | -6.91 | 28.71 | -5.04 | 28.83 | -10.84 | 0 | -37.38 | -0.18 | 0 | -0.18 | 28.42 | 0 | 28.48 | 0.23 | 28.59 | 0.06 | 28.3
34.72 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 12.66 | -22.09 | 0.06 | -32.64 | -0.06 | 0 | 0.18 | 21.27 | 0.12 | 21.56 | -5.92 | 28.59 | -3.98 | 28.48 | -9.61 | 0 | -35.39 | -0.18 | 0 | 0.06 | 28.59 | 0.12 | 28.13 | 0.12 | 28.42 | 0.12 | 28.65
34.752 | -0.06 | 0 | -0.18 | 7.03 | -0.18 | 21.45 | -0.29 | 21.45 | 0.12 | 18.22 | -19.51 | 0.06 | -31.06 | 0 | -6.21 | 0.06 | 21.39 | 0.06 | 21.33 | -4.63 | 28.24 | -2.93 | 28.48 | -8.56 | 0 | -32.93 | -0.06 | 0 | -0.06 | 28.42 | 0.06 | 28.18 | 0.23 | 28.54 | 0.23 | 28.59
34.784 | -0.18 | -44.47 | -0.18 | 0.06 | -0.18 | 0.23 | -0.29 | 21.39 | 0.18 | 21.56 | -17.23 | 0 | -28.36 | 0 | -69.55 | 0.06 | 21.39 | 0.12 | 21.15 | -3.98 | 28.71 | -2.23 | 28.65 | -6.86 | 0 | -30.12 | 0 | 0 | 0 | 28.65 | 0 | 28.48 | 0.18 | 28.42 | 0.18 | 28.48
34.816 | -0.06 | -54.67 | -0.06 | 0 | -0.18 | 0.06 | -0.23 | 21.39 | 0.29 | 21.39 | -15.47 | 0 | -26.43 | -0.06 | -48.28 | 0.06 | 21.39 | 0.18 | 21.27 | -3.16 | 28.65 | -1.82 | 28.48 | -5.51 | -0.06 | -28.48 | -0.06 | 0 | -0.18 | 28.13 | -0.06 | 28.89 | 0.18 | 28.48 | 0.18 | 28.54
34.848 | -0.06 | -41.08 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | -13.71 | -0.06 | -24.84 | -0.06 | -29.24 | 0.12 | 21.39 | 0.23 | 21.33 | -2.46 | 28.65 | -1.41 | 28.36 | -4.45 | 0 | -26.08 | -0.06 | 0 | 0 | 28.48 | -0.06 | 28.77 | 0.06 | 28.48 | 0.23 | 28.54
34.88 | -0.18 | -40.14 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -11.6 | -0.06 | -22.79 | 0 | -28.36 | 0.12 | 21.39 | 0.18 | 21.39 | -1.88 | 28.71 | -1.23 | 28.42 | -3.63 | 0 | -24.84 | -0.06 | -59.89 | -0.06 | 28.48 | -0.06 | 28.42 | 0.12 | 28.48 | 0.18 | 28.48
34.912 | 0 | -37.91 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.41 | 21.33 | -8.79 | -0.06 | -20.45 | -0.06 | -34.92 | 0.12 | 0.12 | 0.23 | 21.27 | -1.64 | 28.59 | -0.76 | 28.59 | -2.81 | 0 | -22.91 | -0.06 | -78.4 | 0 | 0.29 | -0.06 | 28.48 | 0.12 | 28.65 | 0.23 | 28.48
34.944 | -0.18 | -34.69 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.41 | 21.39 | -6.39 | -0.18 | -18.05 | 0 | -34.63 | 0.12 | 0 | 0.18 | 21.39 | -1.05 | 28.59 | -0.59 | 28.36 | -1.88 | -0.06 | -20.16 | -0.06 | -51.62 | 0 | 0 | -0.06 | 0.23 | 0.06 | 28.48 | 0.18 | 28.48
34.976 | -0.06 | -32.58 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.53 | 21.39 | -4.51 | -0.18 | -16.88 | 0 | -31.82 | 0 | 0.06 | 0.23 | 21.39 | -0.94 | 28.13 | -0.35 | 28.54 | -1.46 | -0.06 | -18.16 | 0 | -48.05 | 0 | 0 | -0.06 | 0.06 | 0 | 0.29 | 0.18 | 28.42
35.008 | -0.06 | -31.23 | -0.06 | 0 | -0.12 | 0 | -0.18 | 21.45 | 0.35 | 21.21 | -2.46 | 0.06 | -14.36 | 0 | -28.89 | -0.06 | 0 | 0.23 | 21.39 | -0.7 | 28.3 | -0.18 | 28.54 | -1 | 2.99 | -15.47 | 0 | -41.02 | -0.06 | 0.06 | 0 | 0 | 0 | 0 | 0.18 | 0.18
35.04 | -0.06 | -29.42 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.45 | -1.88 | -0.18 | -12.48 | 0 | -27.54 | 0.06 | 0 | 0.29 | 21.45 | -0.59 | 28.3 | -0.23 | 28.48 | -0.7 | 11.25 | -13.3 | 0 | -37.09 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
35.072 | -0.06 | -27.19 | -0.06 | 0 | -0.12 | 0 | -0.06 | 21.09 | 0.35 | 21.39 | -0.76 | 5.63 | -10.02 | -0.06 | -26.02 | -0.06 | 0 | 0.12 | 21.21 | -0.53 | 28.54 | -0.18 | 28.65 | -0.53 | 18.81 | -11.31 | -0.06 | -37.91 | -0.06 | 0 | -0.18 | 0 | 0 | 0.06 | 0.06 | 0
35.104 | -0.18 | -24.84 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.27 | 0.35 | 21.39 | -0.29 | 13.42 | -7.56 | -0.06 | -24.26 | -0.06 | 0 | 0.29 | 21.39 | -0.41 | 28.89 | -0.18 | 28.83 | -0.35 | 28.65 | -8.91 | 0.06 | -38.38 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
35.136 | -0.06 | -22.21 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | 0.06 | 19.04 | -5.39 | 0.41 | -22.56 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.77 | -0.23 | 28.3 | -0.23 | 28.65 | -6.8 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
35.168 | -0.06 | -20.63 | -0.06 | -2.4 | -0.06 | 0 | -0.18 | 0.23 | 0.23 | 21.27 | 0.06 | 21.56 | -3.16 | 7.85 | -20.57 | 0.12 | 0 | 0.23 | 21.39 | -0.35 | 28.89 | -0.35 | 28.65 | -0.29 | 28.83 | -4.51 | -0.06 | -34.63 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0.06 | 0
35.2 | 0 | -18.93 | 0 | -60.94 | -0.18 | -0.12 | -0.18 | 0.06 | 0.23 | 21.39 | 0.29 | 21.5 | -2.05 | 13.42 | -19.34 | -0.06 | 0 | 0.29 | 21.39 | -0.41 | 28.59 | -0.23 | 28.65 | -0.29 | 28.42 | -2.7 | 0 | -32.93 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 0
35.232 | -0.06 | -16.52 | -0.18 | -50.69 | -0.06 | -45.12 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.45 | -1.05 | 19.04 | -17.46 | 0.06 | 0 | 0.18 | 17.46 | -0.41 | 28.71 | -0.29 | 28.65 | -0.35 | 28.42 | -1.7 | 0 | -30.76 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
35.264 | -0.06 | -14.71 | -0.18 | -40.37 | -0.06 | -48.75 | -0.06 | 0 | 0.18 | 21.45 | 0.47 | 21.45 | -0.23 | 21.21 | -15.23 | -0.06 | 0 | 0.18 | 0.06 | -0.41 | 28.54 | -0.18 | 28.48 | -0.23 | 28.54 | -0.64 | -0.06 | -28.83 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0
35.296 | -0.06 | -12.77 | -0.06 | -38.85 | -0.06 | -39.85 | -0.06 | 0 | 0 | 21.39 | 0.53 | 21.39 | -0.18 | 21.33 | -14.18 | -0.18 | 0 | 0.23 | 0 | -0.29 | 28.83 | -0.18 | 28.59 | -0.23 | 28.42 | -0.41 | -0.06 | -27.54 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
35.328 | -0.06 | -11.66 | -0.06 | -36.21 | -0.06 | -39.32 | 0 | 0 | 0.06 | 21.45 | 0.53 | 21.33 | 0.23 | 21.39 | -11.89 | -0.06 | -39.79 | 0.18 | 0 | -0.35 | 28.48 | -0.18 | 28.24 | -0.18 | 28.65 | 0.06 | -0.06 | -25.14 | 0 | -31.35 | 0 | 0 | -0.18 | 0 | -0.06 | 0
35.36 | -0.06 | -10.37 | -0.06 | -34.81 | -0.06 | -36.97 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 21.33 | 0.18 | 21.39 | -10.14 | 0 | -65.51 | 0.06 | 0 | -0.29 | 28.42 | -0.18 | 28.36 | -0.18 | 28.18 | 0.12 | 0 | -23.44 | -0.06 | -86.19 | 0.06 | -17.87 | -0.18 | 0 | 0 | 0
35.392 | 5.45 | -8.91 | -0.06 | -32.4 | -0.06 | -35.1 | -0.06 | 0 | 0 | 21.39 | 0.53 | 21.5 | 0.29 | 21.33 | -8.56 | -0.06 | -43.89 | 0.18 | 0 | -0.18 | 25.84 | -0.18 | 28.83 | -0.18 | 28.65 | 0.41 | -0.06 | -22.03 | -0.06 | -50.8 | 0 | -90.36 | 0 | -28.89 | 0 | 0
35.424 | 13.18 | -6.86 | -0.06 | -29.71 | -0.06 | -33.17 | -0.06 | 0 | -0.06 | 21.45 | 0.41 | 21.33 | 0.41 | 21.39 | -5.98 | -0.06 | -27.48 | 0.06 | 0 | -0.23 | 0.18 | -0.18 | 28.65 | -0.06 | 28.48 | 0.41 | 9.43 | -20.1 | -0.18 | -47.76 | 0 | -52.38 | -0.06 | -77.29 | 0 | -63.69
35.456 | 18.57 | -5.22 | -0.06 | -27.89 | -0.06 | -30.12 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.27 | 0.53 | 21.39 | -4.39 | 0.06 | -32.23 | 0.18 | 0 | -0.18 | 0 | -0.23 | 0.23 | -0.18 | 28.36 | 0.35 | 16.7 | -17.93 | -0.06 | -41.31 | -0.18 | -31.11 | -0.06 | -49.75 | 0 | -74.89
35.488 | 21.62 | -4.1 | 0.06 | -26.72 | -0.06 | -28.54 | -0.18 | 0 | 0 | 21.39 | 0.47 | 21.39 | 0.41 | 21.39 | -2.99 | 0 | -35.63 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0.12 | -0.18 | 28.36 | 0.53 | 24.38 | -15.18 | 0 | -36.1 | -0.18 | -32.52 | 0 | -29.59 | -0.18 | -44.47
35.52 | 21.45 | -2.7 | -0.06 | -24.08 | 0 | -26.66 | 0 | 0 | -0.18 | 21.45 | 0.35 | 21.27 | 0.29 | 21.33 | -2.11 | -0.18 | -34.22 | 0.23 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.59 | 28.77 | -13.07 | -0.06 | -37.03 | 0 | -42.31 | -0.06 | -33.52 | -0.18 | -28.77
35.552 | 21.45 | -1.99 | 0 | -21.97 | -0.06 | -24.61 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.47 | 21.21 | -1.46 | -0.18 | -31 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0.06 | -0.06 | 28.48 | 0.59 | 28.48 | -11.48 | -0.06 | -38.38 | -0.06 | -42.13 | -0.18 | -42.48 | -0.06 | -35.04
35.584 | 21.45 | -1 | -0.06 | -19.4 | -0.06 | -21.74 | -0.06 | -13.95 | 0 | 0.06 | 0.23 | 21.15 | 0.35 | 21.39 | -0.7 | -0.18 | -28.07 | 0.12 | 0 | -0.18 | 0 | -0.23 | 0 | -0.18 | 28.42 | 0.53 | 28.65 | -9.32 | -0.06 | -36.92 | -0.06 | -35.92 | -0.06 | -39.96 | 0 | -43.01
35.616 | 21.45 | -0.53 | -0.06 | -17.17 | -0.06 | -19.4 | -0.06 | -77.23 | -0.18 | 0 | 0.23 | 21.33 | 0.41 | 21.21 | -0.59 | -0.18 | -26.84 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.53 | 28.59 | -7.68 | -0.06 | -34.98 | 0 | -33.17 | 0 | -36.27 | -0.18 | -39.79
35.648 | 21.45 | -0.06 | -0.06 | -15.94 | -0.06 | -17.81 | -0.06 | -52.97 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.39 | -0.18 | -0.18 | -25.72 | 0.18 | -0.12 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.41 | 28.65 | -4.86 | 0.06 | -33.52 | -0.18 | -32.76 | -0.06 | -33.22 | -0.06 | -34.45
35.68 | 21.45 | -0.18 | -0.06 | -14.24 | -0.06 | -15.94 | -0.06 | -42.95 | -0.06 | 0 | 0.18 | 21.27 | 0.23 | 21.39 | -0.18 | -0.06 | -23.61 | 0.12 | -57.95 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.47 | 28.48 | -2.75 | 0 | -31.47 | 0.06 | -32.76 | -0.06 | -32.99 | -0.06 | -33.11
35.712 | 21.45 | -0.23 | -0.06 | -12.95 | -0.06 | -14.94 | -0.06 | -37.56 | -0.06 | 0 | 0 | 21.15 | 0.23 | 21.33 | 0.12 | -0.06 | -22.21 | 0.06 | -59.77 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.35 | 28.71 | -1.88 | 6.97 | -29.42 | 0 | -31.06 | -0.06 | -32.93 | -0.06 | -32.93
35.744 | 21.5 | -0.47 | 0 | -11.02 | -0.06 | -12.48 | -0.06 | -31.99 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.06 | -0.06 | -20.22 | 0.12 | -32.34 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.65 | -0.94 | 17.17 | -26.9 | -0.06 | -28.59 | 0 | -30.53 | -0.06 | -32.46
35.776 | 21.45 | -0.35 | 4.63 | -9.49 | 0 | -11.37 | -0.06 | -32.4 | -0.18 | 0 | 0.12 | 21.33 | 0.18 | 21.39 | 0.12 | -0.06 | -18.87 | 0.06 | -28.65 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.36 | -0.82 | 24.96 | -25.37 | 0 | -26.9 | -0.18 | -28.13 | -0.12 | -29.53
35.808 | 21.45 | -0.35 | 10.72 | -7.97 | 0 | -9.79 | -0.06 | -32.29 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.56 | 0.12 | 0.94 | -16.93 | 0.12 | -33.69 | -0.18 | 0.06 | -0.06 | 0 | -0.06 | 28.54 | 0.18 | 28.59 | -0.06 | 28.65 | -23.91 | -0.06 | -26.13 | 0 | -26.72 | -0.06 | -27.89
35.84 | 21.45 | -0.23 | 16 | -6.8 | -0.06 | -8.32 | -0.06 | -31.7 | -0.06 | 0 | 0.06 | 21.39 | 0 | 21.15 | 0.12 | 8.61 | -14.77 | 0.12 | -35.74 | -0.06 | -37.68 | -0.18 | 0 | -0.06 | 28.65 | 0.12 | 28.48 | -0.18 | 28.48 | -22.09 | 0 | -24.32 | 0 | -26.25 | -0.06 | -26.54
35.872 | 21.45 | -0.23 | 21.68 | -5.33 | -0.06 | -7.21 | -0.18 | -29.71 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.33 | 0.23 | 15.65 | -13.36 | 0 | -33.17 | -0.06 | -64.22 | -0.06 | -2.7 | -0.06 | 28.59 | 0.06 | 28.59 | 0.29 | 28.59 | -19.45 | -0.06 | -22.74 | 0 | -24.67 | 0 | -25.9
35.904 | 21.45 | -0.29 | 21.45 | -3.16 | -0.06 | -5.1 | -0.18 | -27.01 | -0.06 | 0 | 0 | 21.39 | -0.06 | 21.39 | 0.23 | 21.33 | -11.31 | -0.06 | -28.48 | -0.06 | -46.17 | 0 | -74.77 | -0.06 | 0.29 | 0.12 | 28.48 | 0.29 | 28.65 | -17.75 | 0 | -20.92 | 0 | -22.09 | -0.06 | -23.79
35.936 | 21.45 | -0.18 | 21.5 | -2.05 | 5.92 | -3.57 | -0.06 | -24.9 | -0.06 | 0 | 0.06 | 18.11 | 0.06 | 21.39 | 0.12 | 21.39 | -9.67 | 0.06 | -27.36 | 0.06 | -45.53 | -0.06 | -50.1 | -0.18 | 0.12 | 0.23 | 28.65 | 0.23 | 28.48 | -15.94 | -0.18 | -19.28 | 0 | -20.22 | -0.18 | -21.68
35.968 | 21.45 | -0.29 | 21.45 | -1.41 | 11.66 | -2.46 | -0.06 | -23.03 | -0.06 | 0 | 0.06 | 0.12 | 0.12 | 21.39 | 0.23 | 21.39 | -7.68 | 0.06 | -26.84 | -0.18 | -43.83 | -0.06 | -45.88 | -0.18 | 0 | 0.06 | 28.42 | 0.47 | 28.42 | -13.83 | 0 | -17.75 | -0.06 | -18.93 | -0.18 | -20.57
36 | 21.45 | -0.23 | 21.5 | -0.76 | 17.29 | -1.7 | 5.27 | -20.39 | 0.06 | -55.72 | -0.18 | 0 | 0 | 21.33 | 0.12 | 21.39 | -5.86 | -0.06 | -25.61 | 0 | -41.08 | -0.06 | -45.06 | 0 | 0.06 | -0.06 | 28.48 | 0.41 | 28.36 | -11.95 | 0 | -16.29 | -0.18 | -17.7 | 0 | -18.69
36.032 | 21.45 | -0.23 | 21.5 | -0.35 | 21.5 | -0.94 | 11.25 | -18.99 | -0.06 | -64.75 | -0.06 | 0 | 0 | 18.11 | 0.18 | 21.39 | -4.63 | 0 | -23.44 | -0.06 | -39.08 | -0.06 | -42.66 | -0.06 | 0 | -0.06 | 28.65 | 0.35 | 28.83 | -9.38 | 6.74 | -14.53 | -0.06 | -16 | -0.06 | -17.23
36.064 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.35 | 18.46 | -16.29 | -0.12 | -42.66 | -0.18 | 0 | -0.06 | 0.06 | 0.12 | 21.5 | -2.64 | 0.06 | -20.68 | 0 | -37.5 | -0.06 | -40.37 | -0.06 | 0 | 0.06 | 28.65 | 0.18 | 28.48 | -7.09 | 17.05 | -12.31 | -0.06 | -13.71 | 6.33 | -15.06
36.096 | 21.5 | -0.23 | 21.45 | -0.06 | 21.45 | -0.23 | 21.74 | -14.18 | -0.18 | -38.79 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.33 | -1.82 | 0 | -19.86 | -0.06 | -35.63 | 0 | -38.2 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.48 | -5.04 | 24.67 | -10.66 | 5.63 | -12.31 | 14.12 | -13.36
36.128 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -12.25 | -0.18 | -34.04 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 21.27 | -1 | -0.06 | -18.69 | -0.18 | -33.63 | 0 | -36.45 | 0 | 0 | 0 | 28.42 | 0.29 | 28.42 | -3.16 | 28.83 | -9.14 | 13.65 | -10.55 | 21.86 | -11.78
36.16 | 21.39 | -0.18 | 21.39 | -0.47 | 21.45 | -0.18 | 21.45 | -10.55 | -0.06 | -32.46 | -0.06 | 0 | 0.06 | 0 | 0.06 | 21.15 | -0.7 | 0.06 | -16.7 | -0.18 | -32.76 | -0.18 | -34.98 | -0.18 | 0 | -0.06 | 28.54 | 0.18 | 28.59 | -1.88 | 28.59 | -7.38 | 21.33 | -9.08 | 29.01 | -10.49
36.192 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.45 | -8.09 | -0.18 | -32.76 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.64 | -0.06 | -14.65 | 0 | -29.77 | 0 | -33.34 | 0 | 0 | -0.18 | 28.65 | 0.18 | 28.42 | -0.53 | 28.89 | -5.86 | 29.12 | -7.21 | 28.59 | -8.91
36.224 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | -0.41 | 21.45 | -5.16 | -0.06 | -31.99 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | -0.18 | -0.06 | -12.77 | -0.06 | -28.18 | 0 | -30.7 | -0.06 | 0 | 0.06 | 28.83 | 0.12 | 28.48 | -0.7 | 28.65 | -3.69 | 28.65 | -5.39 | 28.48 | -6.8
36.256 | 20.04 | -0.23 | 21.45 | -0.35 | 21.45 | -0.35 | 21.39 | -3.16 | -0.06 | -30.06 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | -0.06 | -0.06 | -11.13 | 0 | -26.13 | 0 | -29.06 | 0.06 | 0 | -0.18 | 28.3 | 0.06 | 28.89 | -0.18 | 28.59 | -2.4 | 28.42 | -3.87 | 28.36 | -5.16
36.288 | 0.18 | -0.18 | 21.45 | -0.18 | 21.5 | -0.41 | 21.45 | -1.99 | -0.06 | -28.18 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | 0.06 | 0.06 | -9.32 | -0.06 | -24.43 | 0 | -26.78 | -0.06 | 0 | 0 | 28.54 | 0.12 | 28.77 | -0.06 | 28.48 | -1.64 | 28.65 | -2.4 | 28.07 | -3.75
36.32 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | -0.41 | 21.45 | -0.7 | -0.06 | -25.72 | -0.06 | 0 | -0.18 | 0 | 0.23 | 21.21 | 0 | -0.06 | -7.79 | -0.06 | -21.56 | -0.06 | -24.9 | -0.06 | -2.11 | -0.18 | 0.23 | 0.06 | 28.65 | 0 | 28.65 | -1.35 | 28.42 | -1.7 | 28.42 | -2.52
36.352 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.5 | -0.53 | -0.06 | -23.91 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0 | 6.27 | -5.68 | 0.06 | -19.63 | 0.06 | -22.68 | -0.06 | -71.84 | -0.06 | 0 | 0.18 | 28.42 | 0.18 | 28.59 | -0.41 | 28.54 | -0.82 | 28.83 | -1.7
36.384 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.29 | 21.5 | 0.06 | -0.18 | -21.39 | -0.06 | -63.05 | 0 | 0 | 0.06 | 21.39 | 0.18 | 12.07 | -3.75 | -0.06 | -17.93 | 0 | -20.74 | 0 | -53.26 | 0 | 0 | 0.18 | 28.59 | 0.29 | 28.42 | -0.35 | 28.42 | -0.82 | 28.54 | -0.7
36.416 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | 0 | -0.18 | -19.34 | -0.06 | -59.59 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 17.58 | -2.81 | -0.06 | -16.11 | -0.06 | -18.16 | -0.06 | -45.76 | 0 | 0 | -0.06 | 28.59 | 0.23 | 28.59 | -0.23 | 28.42 | 0 | 28.42 | -0.59
36.448 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | 0.12 | -0.06 | -17.4 | 0 | -44.83 | -0.18 | 0 | 0.18 | 21.33 | 0.06 | 21.56 | -2.05 | 0 | -14.3 | -0.06 | -16.47 | -0.06 | -45.35 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.89 | -0.18 | 28.3 | -0.18 | 28.83 | -0.41
36.48 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.23 | 21.45 | 0.41 | -0.18 | -15.29 | 0 | -39.44 | -0.18 | -61.53 | 0.12 | 21.33 | 0.18 | 21.33 | -1 | 0 | -13.18 | -0.06 | -15.18 | -0.06 | -42.89 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.54 | 0.06 | 28.54 | -0.06 | 28.59 | -0.23
36.512 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | 0.47 | -0.18 | -13.59 | 0 | -34.45 | -0.06 | -61.17 | 0.12 | 21.39 | 0.12 | 21.39 | -0.7 | -0.06 | -11.02 | -0.06 | -13.54 | -0.06 | -41.25 | -0.06 | 0 | 0 | 28.59 | 0.35 | 28.77 | 0.06 | 28.48 | -0.06 | 28.59 | -0.06
36.544 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.41 | -0.18 | -10.84 | -0.06 | -32.46 | -0.06 | -32.52 | 0.06 | 21.33 | 0.18 | 21.33 | -0.18 | -0.06 | -9.61 | 6.97 | -11.84 | 0 | -38.2 | 0 | 0 | 0 | 28.65 | 0.18 | 28.54 | 0 | 28.48 | 0.06 | 28.65 | 0.06
36.576 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.23 | 21.39 | 0.47 | -0.06 | -8.32 | -0.06 | -32.34 | -0.06 | -28.48 | 0 | 21.39 | 0.12 | 21.39 | -0.06 | 1.99 | -8.44 | 14.65 | -10.37 | 0 | -36.27 | -0.06 | 0 | -0.18 | 25.9 | 0.29 | 28.42 | 0.23 | 28.36 | 0.12 | 28.48 | 0
36.608 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | 0.53 | -0.06 | -6.45 | -0.06 | -31.64 | -0.06 | -34.22 | -0.06 | 21.33 | 0.18 | 21.33 | -0.06 | 8.73 | -7.27 | 22.33 | -8.91 | -0.06 | -34.98 | 0.06 | 0 | -0.06 | 0.18 | 0.12 | 28.59 | 0.12 | 28.48 | 0.18 | 28.48 | 0.12
36.64 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | 0.41 | 7.09 | -4.34 | -0.06 | -30.12 | -0.06 | -35.68 | 0.06 | 21.33 | 0.23 | 21.15 | 0 | 15.94 | -6.09 | 29.24 | -7.44 | 0.06 | -33.17 | -0.06 | 0 | 0.12 | 0 | 0.12 | 28.36 | 0.18 | 28.54 | 0.18 | 28.48 | 0.12
36.672 | 0 | -0.06 | 0.12 | -0.18 | 21.39 | -0.18 | 21.39 | 0.47 | 14.53 | -2.46 | -0.06 | -28.48 | -0.06 | -33.28 | 0.06 | 21.39 | 0.18 | 21.33 | 0.06 | 23.73 | -4.75 | 28.48 | -6.62 | -0.06 | -31.64 | 0 | 0 | -0.06 | 0 | 0.12 | 28.65 | 0.12 | 28.36 | 0.12 | 28.48 | 0.18
36.704 | -42.36 | -0.18 | 0 | 0 | 21.39 | -0.18 | 21.39 | 0.41 | 20.33 | -1.41 | 0.06 | -25.43 | -0.18 | -28.71 | 0.12 | 0.23 | 0.18 | 21.33 | 0.12 | 28.71 | -3.87 | 28.54 | -5.16 | 9.9 | -28.77 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.3 | 0 | 28.48 | 0.23 | 28.48 | 0.12
36.736 | -55.96 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.5 | 0.29 | 21.5 | -0.64 | -0.06 | -23.38 | -0.06 | -27.6 | 0.06 | 0.06 | 0.29 | 21.33 | 0.12 | 28.71 | -3.16 | 28.48 | -4.04 | 17.23 | -27.36 | -0.18 | -69.2 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.71 | 0.18 | 28.59 | 0.18
36.768 | -41.31 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.35 | 21.39 | -0.29 | -0.06 | -21.5 | -0.06 | -26.84 | 0 | 0 | 0.23 | 21.39 | 0.06 | 28.54 | -2.46 | 28.65 | -3.46 | 25.02 | -24.84 | 0 | -71.25 | -0.18 | 0 | 0.06 | 28.48 | 0.12 | 28.48 | 0.12 | 28.65 | 0.18
36.8 | -40.14 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.39 | 0.29 | 21.39 | -0.18 | -0.06 | -19.75 | -0.06 | -25.43 | -0.06 | 0.06 | 0.12 | 21.39 | 0.12 | 28.59 | -1.99 | 28.71 | -2.75 | 28.89 | -22.91 | 0.06 | -51.33 | 0 | 0 | 0 | 28.42 | 0.06 | 28.42 | 0.18 | 28.48 | 0.23
36.832 | -38.32 | -0.18 | 0 | -0.06 | 0.12 | -0.18 | 21.39 | 0.23 | 21.39 | 0.29 | 0.12 | -17.81 | -0.06 | -23.38 | 0 | 0 | 0.23 | 21.39 | 0.12 | 28.54 | -1.46 | 28.13 | -2.17 | 28.65 | -21.09 | -0.06 | -47.05 | 0 | 0 | 0 | 28.36 | 0 | 28.3 | 0.23 | 28.59 | 0.12
36.864 | -34.86 | -0.06 | 0 | -0.18 | 0 | -0.18 | 20.33 | 0.23 | 21.33 | 0.12 | -0.06 | -15.7 | -0.06 | -21.09 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 28.48 | -1.17 | 28.59 | -1.35 | 28.83 | -18.63 | -0.18 | -39.38 | 0 | 0 | 0 | 28.48 | 0 | 28.48 | 0.23 | 28.48 | 0.12
36.896 | -33.4 | 0 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.18 | 21.39 | 0.29 | -0.06 | -12.89 | 0.06 | -19.63 | 0.06 | 0 | 0.12 | 21.39 | 0.18 | 28.48 | -0.94 | 28.48 | -1 | 28.77 | -16.47 | 0 | -36.68 | -0.06 | 0 | 0 | 27.66 | 0 | 28.42 | 0.06 | 28.42 | 0.06
36.928 | -31.29 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.33 | 0.41 | -0.06 | -11.25 | -0.06 | -18.22 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 28.48 | -0.7 | 28.48 | -0.7 | 28.65 | -15.18 | -0.06 | -37.56 | 0 | 0 | 0 | 0.18 | 0 | 28.48 | 0.12 | 28.48 | 0.18
36.96 | -29.36 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 21.33 | 0.41 | -0.06 | -8.91 | 0.06 | -16.29 | 0.06 | 0 | 0.06 | 21.21 | 0.12 | 28.89 | -0.47 | 28.95 | -0.53 | 28.48 | -13.36 | -0.06 | -37.91 | -0.06 | 0 | -0.06 | 0 | 0 | 28.65 | 0.12 | 0.29 | 0.35
36.992 | -27.19 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 0 | -7.03 | -0.06 | -14.3 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 28.48 | -0.35 | 28.24 | -0.35 | 28.59 | -12.19 | 0 | -36.68 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.23 | 0.06 | 0 | 0.06
37.024 | -24.43 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 6.21 | -4.28 | -0.06 | -12.36 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 28.59 | -0.23 | 28.42 | -0.35 | 28.48 | -10.61 | 0 | -34.16 | -0.06 | -73.66 | 0 | 0 | 0 | 0 | 0.12 | 0 | 0.06
37.056 | -22.38 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | 11.89 | -2.23 | 0 | -10.61 | 0.06 | 0 | 0.23 | 21.33 | 0.12 | 28.59 | -0.29 | 28.48 | -0.23 | 28.42 | -9.14 | -0.06 | -32.58 | -0.18 | -70.26 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.18
37.088 | -21.21 | -0.06 | -19.92 | -0.18 | 0 | -0.06 | 0 | 0 | 21.27 | 0.29 | 17.52 | -1.99 | -0.06 | -8.85 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 28.48 | -0.23 | 28.48 | -0.35 | 28.36 | -7.85 | 2.46 | -30.35 | 0 | -50.45 | 0 | 0 | 0 | 0 | 0.12 | 0 | 0.12
37.12 | -18.52 | 0 | -63.23 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 21.86 | -0.59 | 0 | -7.09 | -0.06 | -0.23 | 0.12 | 21.39 | 0.12 | 28.48 | -0.29 | 28.59 | -0.23 | 28.65 | -6.39 | 10.31 | -29.24 | 0 | -46.29 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
37.152 | -16.99 | -0.06 | -46.47 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.45 | -0.64 | 0 | -5.57 | -0.18 | -74.07 | 0.06 | 21.33 | 0.23 | 28.3 | -0.23 | 28.42 | -0.29 | 28.3 | -5.63 | 17.64 | -26.84 | 0 | -41.02 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.12
37.184 | -14.71 | -0.06 | -39.85 | 0 | 0 | -0.18 | 0 | -0.18 | 21.33 | 0.35 | 21.39 | 0 | 3.4 | -3.52 | -0.06 | -49.4 | 0.06 | 21.39 | 0.18 | 28.42 | -0.23 | 29.01 | -0.18 | 28.48 | -4.28 | 27.95 | -24.9 | 0 | -36.27 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
37.216 | -13.18 | -0.06 | -38.15 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.45 | 0.29 | 21.39 | 0.23 | 9.61 | -2.29 | 0 | -31 | 0.12 | 21.33 | 0.12 | 28.42 | -0.23 | 28.54 | -0.23 | 28.59 | -3.34 | 28.65 | -23.67 | -0.06 | -37.56 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.18
37.248 | -11.54 | -0.06 | -34.98 | -0.06 | -29.71 | -0.06 | 0 | -0.18 | 21.39 | 0.23 | 21.45 | 0.29 | 14.88 | -1.46 | -0.06 | -28.95 | 0.23 | 0.12 | 0.18 | 28.59 | -0.23 | 28.77 | -0.18 | 28.42 | -2.64 | 28.65 | -20.92 | 0 | -37.91 | -0.18 | 0 | 0 | 0 | 0 | 0 | -0.06
37.28 | -10.61 | -0.06 | -34.28 | -0.06 | -59.89 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | 0.23 | 20.63 | -1 | -0.06 | -34.51 | 0.18 | 0 | 0.12 | 28.65 | -0.18 | 28.42 | -0.18 | 28.36 | -2.05 | 28.48 | -19.63 | -0.06 | -36.74 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06
37.312 | -9.02 | -0.06 | -32.4 | -0.06 | -40.84 | -0.06 | -24.32 | -0.06 | 21.39 | 0.18 | 21.39 | 0.47 | 21.39 | -0.53 | -0.18 | -35.33 | 0.06 | 0 | 0.23 | 28.59 | -0.23 | 28.54 | -0.29 | 28.59 | -1.41 | 28.54 | -17.75 | 0 | -35.1 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0
37.344 | -7.32 | 0 | -29.3 | -0.06 | -39.61 | -0.06 | -50.28 | -0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.21 | -0.18 | 0 | -31.88 | 0.12 | 0 | 0.06 | 28.71 | -0.18 | 28.65 | -0.23 | 28.71 | -0.94 | 28.36 | -15.18 | -0.06 | -32.64 | -0.23 | -80.69 | 0 | 0 | -0.06 | 0 | 0
37.376 | -5.57 | -0.06 | -27.6 | -0.06 | -38.38 | -0.06 | -38.15 | -0.18 | 21.39 | 0.18 | 21.39 | 0.47 | 21.39 | -0.18 | -0.06 | -28.89 | 0.06 | 0 | 0 | 28.36 | -0.18 | 28.42 | -0.18 | 28.59 | -0.59 | 28.71 | -13.01 | 0 | -30.94 | -0.06 | -71.84 | -0.06 | 0 | 0 | -43.07 | 0
37.408 | -4.04 | 0 | -25.43 | -0.06 | -35.68 | -0.06 | -32.4 | -0.06 | 21.39 | 0.06 | 21.39 | 0.41 | 21.33 | 0.23 | 0 | -27.36 | 0.06 | 0 | 0.18 | 28.65 | -0.18 | 28.48 | -0.18 | 28.65 | -0.41 | 28.65 | -11.13 | -0.06 | -29.06 | -0.06 | -48.4 | -0.06 | -25.25 | -0.06 | -83.56 | 0.06
37.44 | -3.34 | -0.06 | -23.2 | -0.06 | -33.52 | -0.06 | -28.48 | -0.06 | 21.21 | 0 | 21.27 | 0.41 | 21.39 | 0.06 | 2.46 | -26.78 | -0.06 | 0 | 0.12 | 28.54 | -0.18 | 0.23 | -0.18 | 28.65 | -0.35 | 28.71 | -8.91 | -0.06 | -27.13 | -0.18 | -46.53 | -0.06 | -85.37 | 0 | -52.21 | 0
37.472 | -1.82 | 5.45 | -21.62 | -0.06 | -30.76 | -0.06 | -26.37 | -0.06 | 21.39 | 0.12 | 21.39 | 0.53 | 21.39 | 0.12 | 8.5 | -24.55 | 0.12 | 0 | 0.18 | 28.65 | -0.18 | 0 | -0.18 | 28.65 | -0.29 | 28.77 | -7.09 | 0 | -25.72 | 0 | -40.61 | -0.18 | -58.77 | 0 | -29.06 | 0
37.504 | -0.94 | 13.18 | -19.22 | -0.06 | -29.24 | -0.06 | -25.72 | -0.06 | 4.1 | 0.06 | 21.39 | 0.41 | 21.39 | 0.18 | 15.7 | -22.33 | 0 | 0 | 0.12 | 0.12 | -0.18 | 0 | -0.18 | 28.65 | -0.35 | 28.48 | -4.39 | -0.06 | -23.15 | -0.06 | -37.03 | -0.06 | -30.29 | -0.06 | -35.45 | 0.12
37.536 | -0.47 | 18.99 | -16.99 | -0.18 | -27.42 | -0.06 | -23.91 | -0.06 | 0.06 | -0.06 | 21.56 | 0.35 | 21.39 | 0.18 | 21.33 | -20.57 | 0.06 | 0 | 0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | -0.23 | 28.54 | -2.46 | 0.06 | -21.33 | -0.06 | -37.68 | -0.06 | -34.63 | 0 | -42.95 | -0.18
37.568 | -0.23 | 21.62 | -15.35 | -0.06 | -24.9 | 0 | -21.68 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | 0.06 | 21.45 | -19.1 | 0 | 0 | 0.06 | 0 | -0.06 | 0 | -0.06 | 25.9 | -0.35 | 28.3 | -1.46 | 7.38 | -19.22 | 0 | -37.5 | -0.18 | -42.31 | 0 | -40.43 | -0.06
37.6 | -0.18 | 21.5 | -13.59 | 0 | -22.85 | -0.18 | -19.34 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.29 | 21.39 | -17.75 | 0.18 | 0 | 0.18 | 0 | 0 | 0 | -0.18 | 0.18 | -0.18 | 28.42 | -1 | 14.88 | -17.4 | -0.18 | -36.04 | -0.12 | -40.67 | -0.06 | -36.45 | -0.18
37.632 | -0.23 | 21.45 | -12.31 | -0.06 | -20.51 | 0 | -17.7 | -0.18 | 0 | -0.18 | 21.33 | 0.29 | 21.33 | 0.29 | 21.39 | -15.35 | 0.06 | 0 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.35 | 28.3 | -0.53 | 22.79 | -15.41 | 0 | -34.63 | 0 | -36.86 | 0 | -32.93 | -0.18
37.664 | -0.35 | 21.45 | -10.61 | -0.06 | -18.87 | -0.06 | -14.82 | -0.06 | 0 | 0 | 21.56 | 0.12 | 21.45 | 0.29 | 21.27 | -13.95 | 0 | -42.31 | 0.23 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.65 | -0.06 | 28.89 | -13.01 | -0.18 | -32.76 | -0.18 | -33.46 | -0.06 | -32.34 | -0.06
37.696 | -0.35 | 21.45 | -8.91 | -0.18 | -16.76 | -0.06 | -12.77 | -0.18 | 0 | 0 | 21.62 | 0.18 | 21.39 | 0.18 | 21.39 | -12.25 | 0.12 | -66.33 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.12 | 28.65 | -11.02 | -0.06 | -31 | -0.18 | -32.76 | 0 | -32.4 | -0.06
37.728 | -0.29 | 21.45 | -7.79 | -0.06 | -15.23 | -0.06 | -10.72 | -0.06 | 0 | 0 | 21.33 | 0.06 | 21.39 | 0.18 | 21.39 | -10.37 | 0 | -43.07 | 0 | 0 | 0 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.12 | 28.24 | -9.43 | -0.06 | -28.71 | -0.06 | -32.7 | 0 | -30.53 | -0.18
37.76 | -0.47 | 21.45 | -5.86 | -0.06 | -13.77 | -0.06 | -8.44 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.45 | 0.29 | 21.39 | -8.79 | 0.06 | -28.65 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.42 | 0.23 | 28.65 | -7.09 | -0.06 | -27.19 | -0.06 | -30.76 | -0.18 | -27.83 | 0
37.792 | -0.29 | 21.5 | -4.75 | -0.18 | -11.25 | -0.06 | -6.33 | -0.06 | 0 | -0.18 | 21.21 | 0.06 | 21.45 | 0.29 | 21.27 | -6.21 | -0.06 | -31.29 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.35 | 28.42 | -5.04 | 0 | -25.37 | -0.18 | -28.89 | 0 | -26.37 | 0.06
37.824 | -0.18 | 21.45 | -3.11 | -0.06 | -10.14 | -0.18 | -3.98 | 0 | 0 | -0.06 | 21.33 | 0.06 | 21.45 | 0.23 | 21.33 | -4.63 | 0 | -35.39 | 0.12 | 0 | -0.06 | 0 | 0 | 0 | -0.18 | 28.42 | 0.35 | 28.36 | -2.7 | 0.06 | -23.2 | -0.18 | -27.25 | -0.06 | -25.61 | -0.06
37.856 | -0.23 | 21.45 | -2.05 | 0.94 | -8.67 | -0.06 | -2.46 | -0.06 | 0 | -0.06 | 21.45 | 0 | 21.39 | 0.12 | 21.39 | -3.46 | -0.06 | -34.63 | 0.12 | 0 | -0.06 | -35.51 | -0.06 | 0 | -0.18 | 28.71 | 0.18 | 28.71 | -1.7 | 0 | -21.33 | -0.06 | -25.61 | -0.06 | -23.91 | -0.06
37.888 | -0.29 | 21.45 | -1.11 | 8.61 | -7.62 | -0.06 | -1.41 | -0.06 | 0 | -0.06 | 21.33 | -0.06 | 21.39 | 0.18 | 21.09 | -2.52 | -0.06 | -30.7 | 0.12 | 0 | -0.06 | -71.78 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.48 | -0.82 | 0 | -18.93 | -0.06 | -24.2 | -0.06 | -22.38 | 0
37.92 | -0.35 | 21.45 | -0.53 | 14.18 | -5.98 | 7.09 | -0.94 | 0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.15 | 0.18 | 21.33 | -1.23 | -0.18 | -28.48 | 0.18 | -50.98 | -0.06 | -49.92 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.42 | -0.7 | -0.06 | -17.46 | -0.06 | -22.09 | -0.06 | -20.92 | -0.18
37.952 | -0.29 | 21.45 | -0.29 | 19.92 | -3.93 | 12.83 | -0.23 | -0.06 | -74.42 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.5 | -1 | -0.06 | -26.66 | 0.06 | -68.03 | -0.18 | -46 | 0 | 0 | -0.06 | 28.42 | 0.35 | 28.48 | -0.06 | -0.06 | -15.41 | -0.06 | -21.09 | 0 | -18.87 | 0
37.984 | -0.29 | 21.45 | -0.18 | 21.5 | -2.75 | 20.16 | -0.06 | -0.06 | -57.42 | -0.18 | 0 | 0 | 21.39 | 0.18 | 21.39 | -0.47 | 0 | -26.13 | 0.12 | -45.41 | -0.06 | -43.3 | 0 | 0 | -0.18 | 0.29 | 0.29 | 28.48 | 0.18 | 0.06 | -13.07 | -0.06 | -18.87 | -0.06 | -17.93 | 0
38.016 | -0.29 | 21.45 | -0.23 | 21.5 | -1.88 | 21.5 | 0.29 | -0.06 | -43.3 | 0.06 | 0 | 0.12 | 21.45 | 0.18 | 21.27 | -0.18 | 5.68 | -24.14 | 0.12 | -44.59 | 0.06 | -41.08 | 0 | -63.75 | -0.18 | 0.12 | 0.23 | 28.65 | 0.12 | 4.45 | -11.13 | -0.06 | -17.29 | -0.18 | -15.53 | -0.06
38.048 | -0.29 | 21.45 | -0.35 | 21.45 | -1.05 | 21.45 | 0.35 | 0 | -37.97 | -0.06 | 0 | 0 | 21.56 | 0 | 21.39 | -0.18 | 11.48 | -21.68 | 0.23 | -43.24 | 0 | -40.08 | -0.18 | -64.46 | -0.18 | 0 | 0.06 | 28.59 | 0.29 | 12.54 | -9.14 | 1.46 | -15.7 | -0.18 | -13.95 | 0
38.08 | -0.23 | 21.5 | -0.23 | 21.45 | -0.53 | 21.45 | 0.47 | -0.06 | -33.22 | -0.06 | 0 | 0 | 0.53 | 0 | 21.39 | -0.06 | 17.05 | -20.57 | 0.12 | -41.08 | 0 | -38.44 | -0.06 | -45.82 | 0 | 0 | 0.12 | 28.59 | 0.47 | 20.27 | -7.38 | 11.19 | -14.06 | 4.39 | -12.19 | -0.06
38.112 | -0.23 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.47 | -0.06 | -32.52 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.33 | 0.06 | 21.62 | -18.28 | 0.06 | -39.61 | -0.06 | -35.98 | 0 | -46.29 | -0.06 | 0 | 0.12 | 28.3 | 0.41 | 29.01 | -5.04 | 18.75 | -12.6 | 12.66 | -10.55 | 4.98
38.144 | -0.18 | 21.45 | -0.41 | 21.45 | -0.06 | 21.45 | 0.59 | 0 | -32.93 | 0 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0 | 21.45 | -16.64 | 0.12 | -36.74 | -0.06 | -33.75 | 0.06 | -43.48 | 0 | 0 | 0.12 | 28.65 | 0.35 | 28.65 | -2.7 | 29.01 | -10.37 | 22.97 | -8.67 | 15.65
38.176 | -0.18 | 21.39 | -0.35 | 21.5 | -0.12 | 21.45 | 0.53 | -0.06 | -32.17 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.39 | -15.06 | 0.06 | -35.98 | 0.06 | -32.93 | 0.06 | -41.31 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.48 | -1.82 | 28.48 | -8.67 | 28.95 | -7.15 | 23.38
38.208 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | 21.39 | 0.47 | 0 | -30 | -0.18 | 0 | -0.06 | 0 | 0 | 21.33 | 0.06 | 21.45 | -13.36 | 0.18 | -33.81 | 0.06 | -30.7 | 0 | -39.67 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.48 | -0.82 | 28.71 | -6.91 | 28.48 | -5.51 | 28.95
38.24 | -0.18 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | 0.53 | -0.06 | -27.95 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | -11.89 | 0.18 | -31.64 | 0.06 | -28.77 | 0 | -37.56 | 0 | 0 | 0.06 | 28.59 | 0.29 | 28.48 | -0.64 | 28.48 | -5.74 | 28.65 | -3.87 | 28.59
38.272 | -0.18 | 21.45 | -0.18 | 21.39 | -0.29 | 21.5 | 0.59 | 2.05 | -24.9 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.33 | -9.79 | 0.06 | -29.77 | 0 | -26.25 | 0 | -36.68 | -0.18 | 0 | 0 | 28.48 | 0.23 | 28.48 | -0.06 | 28.48 | -3.75 | 28.48 | -2.23 | 28.65
38.304 | -0.18 | 21.45 | -0.23 | 21.39 | -0.29 | 21.39 | 0.47 | 9.38 | -22.85 | -0.06 | 0 | 0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.33 | -7.85 | 0.06 | -27.31 | 0 | -23.56 | -0.06 | -34.28 | -0.06 | 0 | 0 | 28.65 | 0.06 | 28.65 | 0.12 | 28.42 | -2.29 | 28.59 | -1.64 | 28.65
38.336 | -0.18 | 19.69 | -0.29 | 21.45 | -0.29 | 21.39 | 0.53 | 14.53 | -21.15 | -0.06 | -42.13 | -0.06 | 0 | -0.06 | 0.23 | 0.35 | 21.39 | -5.98 | 2.05 | -25.43 | 0 | -21.97 | 0 | -31.47 | -0.06 | 0 | 0 | 28.36 | 0.06 | 28.59 | 0 | 28.65 | -1.46 | 28.65 | -0.94 | 28.48
38.368 | -0.18 | 0.12 | -0.29 | 21.45 | -0.23 | 21.45 | 0.47 | 20.39 | -19.1 | -0.06 | -65.74 | 0 | 0 | -0.18 | 0.06 | 0.23 | 21.45 | -4.86 | 7.09 | -23.32 | 0 | -19.81 | -0.06 | -29.47 | -0.06 | 0 | 0.06 | 28.42 | 0.06 | 28.54 | 0.29 | 28.59 | -1 | 28.54 | -0.53 | 28.65
38.4 | -0.06 | 0 | -0.29 | 21.39 | -0.29 | 21.39 | 0.41 | 21.45 | -17.4 | 0 | -47.52 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -3.22 | 12.89 | -21.27 | 0.06 | -17.93 | 0.06 | -27.25 | -0.06 | 0 | 0 | 28.48 | -0.06 | 28.65 | 0.35 | 28.3 | -0.7 | 28.42 | -0.47 | 28.59
38.432 | 0 | 0 | -0.23 | 21.5 | -0.23 | 21.39 | 0.35 | 21.39 | -14.94 | 0 | -38.67 | -0.06 | 0 | 0 | 0 | 0.35 | 21.21 | -2.05 | 18.34 | -19.34 | 0 | -15.47 | 0.06 | -26.31 | -0.06 | -85.96 | 0.06 | 28.42 | 0.18 | 28.36 | 0.35 | 28.3 | -0.47 | 28.42 | -0.23 | 28.65
38.464 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.23 | 21.39 | -12.95 | -0.06 | -33.46 | -0.18 | 0 | -0.06 | 0 | 0.29 | 21.39 | -1.17 | 21.68 | -17.64 | 0 | -14.3 | 0.06 | -23.03 | -0.06 | -59.24 | 0 | 0.29 | 0.06 | 28.42 | 0.47 | 28.59 | -0.18 | 28.48 | -0.06 | 28.83
38.496 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | 0.35 | 21.39 | -10.66 | 0 | -32.05 | -0.18 | -6.91 | -0.06 | 0 | 0.18 | 21.33 | -0.82 | 21.39 | -15.41 | 0.06 | -13.36 | -0.06 | -21.33 | -0.06 | -48.75 | -0.18 | 0.06 | 0.12 | 28.54 | 0.35 | 28.48 | -0.18 | 28.42 | 0 | 28.59
38.528 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.35 | 21.39 | -8.79 | -0.06 | -32.58 | -0.06 | -48.4 | -0.06 | 0 | 0.29 | 21.5 | -0.18 | 21.39 | -13.95 | 5.74 | -11.54 | 0 | -19.69 | 0 | -44.59 | -0.18 | 0 | 0 | 28.48 | 0.29 | 28.48 | 0 | 28.65 | 0 | 28.59
38.56 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.12 | 21.39 | -6.45 | 0 | -32.4 | -0.06 | -43.01 | -0.06 | 0 | 0.12 | 21.15 | -0.29 | 21.39 | -12.48 | 13.71 | -10.55 | 1.17 | -17.4 | 0 | -39.55 | 0 | 0 | 0 | 28.59 | 0.41 | 28.36 | 0.06 | 28.42 | 0.12 | 28.36
38.592 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.18 | 21.39 | -3.93 | -0.06 | -30.24 | -0.06 | -29.24 | -0.06 | 0 | 0.18 | 21.39 | -0.06 | 21.39 | -11.48 | 23.91 | -8.38 | 11.43 | -15.94 | 0 | -37.21 | -0.06 | 0 | 0 | 28.42 | 0.23 | 28.42 | 0.06 | 28.48 | 0.23 | 28.42
38.624 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -2.11 | 0 | -27.83 | -0.06 | -25.72 | -0.06 | 0 | 0.23 | 21.33 | 0 | 21.33 | -9.32 | 28.77 | -7.15 | 21.39 | -14.47 | 0 | -37.79 | 0 | 0 | -0.06 | 28.71 | 0.18 | 28.42 | 0.06 | 28.54 | 0.23 | 28.65
38.656 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0 | 21.39 | -1.17 | 0 | -26.31 | -0.23 | -27.01 | -0.06 | 0 | 0.18 | 21.45 | 0.12 | 21.39 | -8.2 | 28.59 | -5.92 | 29.06 | -12.6 | -0.06 | -37.85 | -0.06 | 0 | 0 | 29.01 | 0.12 | 28.59 | 0.18 | 28.3 | 0.29 | 28.48
38.688 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.18 | 21.27 | -1 | 6.27 | -24.43 | -0.18 | -26.31 | -0.18 | 0 | 0.12 | 21.45 | 0.18 | 21.33 | -6.91 | 28.48 | -4.63 | 28.48 | -11.13 | 0 | -36.27 | -0.18 | 0 | -0.06 | 28.65 | 0.23 | 28.65 | 0.12 | 28.54 | 0.29 | 28.89
38.72 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0 | 21.39 | -0.41 | 12.07 | -22.38 | -0.18 | -25.02 | 0 | 0 | 0.23 | 21.39 | 0.12 | 21.39 | -5.98 | 28.59 | -4.22 | 28.42 | -9.84 | 0 | -34.28 | 0 | 0 | -0.06 | 28.65 | 0.06 | 28.48 | 0 | 28.18 | 0.18 | 28.42
38.752 | -0.06 | 0 | -0.18 | 10.37 | -0.18 | 21.45 | 0.06 | 21.45 | 0.29 | 17.64 | -19.63 | -0.06 | -21.68 | -0.06 | -0.76 | 0.18 | 21.39 | 0.18 | 21.39 | -4.86 | 28.59 | -3.16 | 28.59 | -8.61 | 0 | -31.7 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.54 | 0 | 28.07 | 0.23 | 28.48
38.784 | -0.06 | -38.79 | -0.06 | 0.06 | -0.06 | 4.28 | 0 | 21.39 | 0.18 | 21.5 | -17.64 | -0.18 | -19.45 | 0 | -73.42 | 0.18 | 21.33 | 0.06 | 21.33 | -3.75 | 28.3 | -2.29 | 28.42 | -6.91 | 0 | -29.77 | -0.06 | 0 | 0 | 28.54 | 0 | 28.36 | 0.06 | 28.36 | 0.23 | 28.65
38.816 | -0.06 | -55.14 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.39 | 0.35 | 21.39 | -15.53 | -0.18 | -18.16 | -0.06 | -50.57 | 0.12 | 21.39 | 0.18 | 21.33 | -2.99 | 28.48 | -1.7 | 28.3 | -5.68 | -0.06 | -28.24 | 0 | 0 | 0 | 28.3 | 0 | 28.89 | 0.06 | 28.48 | 0.23 | 28.48
38.848 | -0.06 | -41.9 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.23 | 21.39 | -13.89 | -0.18 | -16.47 | -0.06 | -30.53 | 0.12 | 21.39 | 0.29 | 21.45 | -2.52 | 28.36 | -1.46 | 28.48 | -4.69 | -0.06 | -26.31 | -0.06 | 0.06 | 0 | 28.59 | 0 | 28.13 | 0 | 28.48 | 0.23 | 28.48
38.88 | -0.18 | -40.2 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -11.72 | -0.06 | -14.82 | 0.06 | -28.77 | -0.06 | 21.39 | 0.23 | 21.39 | -1.7 | 28.48 | -1 | 28.36 | -3.75 | 0.06 | -24.61 | -0.18 | -54.32 | 0 | 28.65 | 0.06 | 28.42 | 0 | 28.42 | 0.12 | 28.48
38.912 | -0.06 | -37.68 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -9.14 | 0 | -13.01 | -0.06 | -34.28 | 0.12 | 0.18 | 0.18 | 21.39 | -1.23 | 28.13 | -0.64 | 28.95 | -2.99 | -0.06 | -22.79 | -0.06 | -80.16 | 0 | 0.23 | -0.06 | 28.36 | 0.06 | 28.48 | 0.06 | 28.48
38.944 | -0.06 | -34.51 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.39 | -6.86 | -0.06 | -10.61 | 0.06 | -34.75 | 0.12 | 0 | 0.18 | 21.39 | -1 | 28.83 | -0.59 | 28.54 | -1.99 | 0 | -20.33 | -0.06 | -52.21 | 0.06 | 0.06 | -0.06 | 0.23 | 0.06 | 28.42 | 0.18 | 28.48
38.976 | -0.06 | -32.46 | -0.06 | 0 | -0.06 | 0 | 0 | 21.27 | 0.53 | 21.33 | -4.8 | -0.06 | -8.67 | 0 | -31.88 | 0.06 | 0 | 0.29 | 21.39 | -0.94 | 28.59 | -0.35 | 28.36 | -1.7 | 0 | -18.69 | 0 | -48.46 | 0 | 0 | 0.06 | 0.12 | 0 | 0.23 | 0.12 | 28.48
39.008 | -0.06 | -31.52 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.33 | -2.99 | -0.06 | -6.86 | -0.18 | -28.95 | 0.12 | 0 | 0.29 | 21.21 | -0.53 | 28.18 | -0.29 | 28.42 | -1.29 | 2.64 | -16.76 | 0 | -41.84 | -0.18 | 0.06 | 0 | 0 | -0.06 | 0.06 | 0.12 | 0.29
39.04 | -0.06 | -29.24 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -1.7 | -0.06 | -5.51 | 0 | -27.19 | 0.18 | 0 | 0.23 | 21.33 | -0.53 | 28.59 | -0.35 | 28.59 | -0.82 | 10.72 | -14.94 | 0 | -36.62 | -0.06 | 0 | 0 | 0.06 | 0.06 | 0 | 0.12 | 0.06
39.072 | -0.06 | -27.72 | 0 | 0 | -0.18 | 0 | 0.06 | 21.33 | 0.35 | 21.21 | -0.7 | 5.39 | -3.75 | -0.06 | -26.31 | 0.06 | 0 | 0.18 | 21.27 | -0.41 | 28.65 | -0.23 | 28.54 | -0.64 | 18.11 | -13.01 | 0 | -37.44 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0.06 | 0 | 0
39.104 | -0.06 | -24.43 | -0.06 | 0 | 0 | 0 | 0 | 21.39 | 0.29 | 21.33 | -0.59 | 13.18 | -2.23 | 0 | -24.32 | 0.06 | 0 | 0.23 | 21.33 | -0.35 | 28.71 | -0.23 | 28.65 | -0.35 | 28.65 | -10.31 | -0.18 | -38.2 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
39.136 | -0.06 | -23.03 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.12 | 18.75 | -1.17 | 0.18 | -22.33 | 0 | 0 | 0.29 | 21.15 | -0.41 | 28.59 | -0.23 | 28.3 | -0.29 | 28.48 | -8.14 | -0.06 | -37.15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
39.168 | -0.06 | -20.63 | -0.18 | -1.05 | -0.18 | 0 | 0 | 3.52 | 0.18 | 21.45 | 0.29 | 21.5 | -0.82 | 7.5 | -20.51 | 0 | 0 | 0.12 | 21.39 | -0.35 | 28.59 | -0.35 | 28.36 | -0.29 | 28.59 | -5.74 | -0.18 | -35.22 | -0.18 | 0 | 0.06 | 0 | -0.18 | 0 | 0 | 0
39.2 | -0.18 | -18.69 | -0.06 | -58.6 | -0.06 | 0.06 | -0.06 | 0.12 | 0.06 | 21.39 | 0.18 | 21.39 | -0.76 | 13.18 | -19.63 | 0 | 0 | 0.18 | 21.21 | -0.29 | 28.71 | -0.18 | 28.3 | -0.29 | 28.48 | -3.93 | 0 | -32.93 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
39.232 | 0 | -16.88 | -0.06 | -51.92 | -0.06 | -37.74 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | 21.39 | -0.41 | 18.63 | -17.75 | 0 | 0 | 0.12 | 21.33 | -0.29 | 28.42 | -0.23 | 28.48 | -0.29 | 28.65 | -2.4 | 0 | -30.7 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0
39.264 | -0.06 | -14.77 | -0.06 | -39.85 | -0.06 | -51.39 | -0.06 | 0.06 | 0.18 | 21.33 | 0.41 | 21.33 | -0.06 | 21.62 | -15.47 | 0.06 | 0 | 0.18 | 0.12 | -0.23 | 28.65 | -0.18 | 28.13 | -0.23 | 28.59 | -1.29 | 0.06 | -29.12 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0
39.296 | -0.18 | -13.24 | -0.06 | -38.56 | -0.06 | -40.31 | -0.06 | 0 | 0.06 | 21.45 | 0.41 | 21.21 | 0 | 21.33 | -14 | 0.06 | 0 | 0.29 | 0 | -0.29 | 28.59 | -0.18 | 28.3 | -0.18 | 28.54 | -0.7 | -0.06 | -28.13 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
39.328 | -0.06 | -12.01 | -0.06 | -37.15 | -0.06 | -39.61 | -0.18 | 0 | 0.12 | 21.45 | 0.53 | 21.39 | 0.12 | 21.39 | -12.07 | 0 | -35.51 | 0.18 | 0 | -0.18 | 28.36 | -0.18 | 28.48 | -0.29 | 28.48 | -0.18 | -0.18 | -25.78 | 0.12 | -16.99 | 0 | 0 | -0.06 | 0 | 0 | 0
39.36 | -0.06 | -10.84 | -0.06 | -34.92 | -0.06 | -37.56 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.27 | -0.06 | 21.39 | -10.61 | -0.06 | -67.15 | 0.06 | 0 | -0.35 | 28.42 | -0.23 | 28.59 | -0.23 | 28.48 | -0.18 | -0.06 | -23.56 | -0.23 | -92.05 | -0.06 | -5.74 | 0 | 0 | 0 | 0
39.392 | 4.8 | -9.08 | -0.06 | -32.93 | -0.06 | -34.75 | -0.06 | 0 | 0.06 | 21.45 | 0.41 | 21.33 | 0.29 | 21.33 | -9.02 | 0 | -45.7 | 0.18 | 0 | -0.29 | 28.54 | -0.18 | 28.48 | -0.23 | 28.3 | 0.06 | -0.06 | -21.56 | 0.12 | -52.74 | -0.06 | -87.43 | -0.06 | -23.15 | -0.06 | 0
39.424 | 12.66 | -7.09 | -0.06 | -30.53 | -0.06 | -32.29 | -0.06 | 0 | 0 | 21.45 | 0.41 | 21.33 | 0.29 | 21.39 | -6.39 | 0 | -27.77 | 0.23 | 0 | -0.29 | 0.29 | -0.18 | 28.42 | -0.23 | 28.48 | 0.12 | 8.44 | -19.86 | -0.06 | -48.05 | 0 | -56.37 | -0.18 | -78.69 | -0.06 | -53.44
39.456 | 18.05 | -5.45 | -0.18 | -28.24 | -0.06 | -30.29 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.15 | 0.35 | 21.39 | -4.8 | -0.06 | -31.99 | 0.23 | 0 | -0.29 | 0 | -0.06 | 0.23 | -0.18 | 28.42 | 0.12 | 15.65 | -17.87 | 0 | -42.72 | -0.06 | -34.34 | -0.18 | -51.97 | -0.06 | -77.46
39.488 | 21.74 | -4.28 | -0.06 | -26.02 | -0.06 | -28.83 | 0 | 0 | 0 | 21.21 | 0.41 | 21.33 | 0.35 | 21.27 | -2.99 | -0.06 | -35.39 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0.12 | -0.23 | 28.36 | 0.35 | 23.38 | -15.29 | -0.06 | -37.44 | -0.06 | -34.22 | -0.18 | -29.71 | 0 | -48.87
39.52 | 21.5 | -2.93 | -0.06 | -23.85 | -0.06 | -26.72 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.29 | 21.33 | -2.11 | -0.06 | -34.28 | 0.18 | 0 | -0.29 | 0 | -0.18 | 0 | -0.29 | 28.65 | 0.29 | 29.01 | -13.18 | -0.06 | -37.15 | -0.06 | -43.3 | -0.18 | -32.05 | 0 | -30.47
39.552 | 21.45 | -2.05 | -0.06 | -21.68 | 0 | -24.79 | -0.23 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.41 | 21.39 | -1.64 | 0 | -31.29 | 0.12 | 0 | -0.23 | 0 | -0.06 | 0.06 | -0.23 | 28.36 | 0.23 | 28.65 | -11.66 | -0.18 | -38.79 | -0.06 | -43.48 | -0.06 | -42.36 | -0.06 | -33.69
39.584 | 21.45 | -0.94 | -0.06 | -19.34 | -0.18 | -21.91 | -0.18 | -4.34 | -0.06 | 0.12 | 0.18 | 21.33 | 0.35 | 21.39 | -0.76 | -0.06 | -27.83 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 27.95 | 0.41 | 28.36 | -9.61 | 0 | -37.21 | -0.06 | -37.79 | -0.18 | -39.85 | -0.06 | -42.95
39.616 | 21.39 | -0.35 | -0.06 | -17.99 | -0.18 | -19.98 | -0.18 | -76.35 | 0.06 | 0 | 0.18 | 21.33 | 0.29 | 21.39 | -0.59 | -0.06 | -27.42 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.23 | 27.89 | 0.35 | 28.48 | -7.62 | -0.06 | -36.04 | 0 | -34.92 | -0.06 | -36.33 | -0.18 | -40.2
39.648 | 21.45 | -0.23 | -0.06 | -15.7 | -0.06 | -18.16 | -0.18 | -55.72 | -0.06 | 0 | 0.18 | 21.39 | 0.23 | 21.21 | -0.18 | 0 | -26.43 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 28.71 | 0.29 | 28.54 | -4.8 | 0 | -33.46 | -0.18 | -34.28 | 0 | -33.22 | 0 | -35.86
39.68 | 21.45 | -0.18 | -0.06 | -14.71 | -0.06 | -16 | -0.18 | -43.07 | -0.06 | 0 | 0.12 | 21.21 | 0.29 | 21.27 | -0.18 | -0.06 | -23.67 | 0.18 | -48.52 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 28.42 | -3.11 | -0.06 | -31.29 | -0.18 | -33.75 | -0.06 | -32.76 | -0.18 | -33.4
39.712 | 21.45 | -0.18 | -0.06 | -12.95 | 0 | -14.82 | -0.06 | -38.85 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.39 | 0 | 0.06 | -21.56 | 0.12 | -61.23 | -0.18 | 0 | 0 | 0 | -0.18 | 28.65 | 0.18 | 28.48 | -1.7 | 5.68 | -29.12 | -0.06 | -32.46 | -0.06 | -32.46 | 0 | -32.99
39.744 | 21.45 | -0.29 | -0.06 | -11.07 | -0.06 | -12.6 | -0.06 | -32.58 | -0.06 | 0 | 0.12 | 21.45 | 0.23 | 21.33 | -0.06 | -0.06 | -20.1 | 0.18 | -33.4 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.83 | 0.23 | 28.42 | -1.23 | 16.17 | -27.83 | 0 | -28.89 | -0.06 | -30.24 | -0.18 | -32.93
39.776 | 21.45 | -0.47 | 3.75 | -9.32 | 0 | -11.48 | -0.18 | -32.23 | 0 | 0 | 0 | 21.21 | 0.29 | 21.39 | 0.12 | -0.06 | -18.69 | 0.12 | -27.31 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.18 | 0.23 | 28.48 | -0.7 | 23.79 | -25.9 | 0.06 | -27.54 | -0.06 | -28.3 | 0 | -30.53
39.808 | 21.45 | -0.35 | 9.84 | -8.2 | -0.06 | -9.79 | -0.18 | -32.76 | 0 | 0 | 0.06 | 21.33 | 0.23 | 21.39 | 0.12 | 2.23 | -17.05 | 0.06 | -32.99 | -0.06 | 0.06 | -0.06 | 0 | -0.18 | 28.54 | 0.12 | 28.36 | -0.18 | 28.95 | -24.08 | -0.06 | -26.08 | -0.18 | -26.72 | 0 | -27.95
39.84 | 21.45 | -0.18 | 15.35 | -6.68 | -0.06 | -8.73 | 0 | -31.47 | 0 | 0 | 0.12 | 21.04 | 0.18 | 21.21 | 0.06 | 7.62 | -14.94 | 0.06 | -35.74 | -0.18 | -44.36 | 0 | 0 | -0.06 | 28.42 | 0.12 | 28.65 | -0.06 | 28.65 | -22.38 | 0 | -24.08 | -0.06 | -25.9 | -0.06 | -26.78
39.872 | 21.45 | -0.23 | 20.92 | -5.04 | -0.06 | -7.09 | -0.06 | -30.35 | 0 | 0 | -0.06 | 21.33 | 0.23 | 21.33 | 0.18 | 14.82 | -13.83 | 0.06 | -33.69 | -0.06 | -67.15 | -0.06 | -1 | -0.18 | 28.36 | 0.23 | 28.89 | 0 | 28.48 | -19.75 | -0.06 | -22.21 | 0 | -24.43 | -0.18 | -26.13
39.904 | 21.45 | -0.18 | 21.5 | -3.22 | -0.06 | -5.1 | -0.06 | -27.13 | 0 | 0 | 0 | 21.39 | 0 | 21.39 | 0.23 | 20.63 | -11.48 | 0.12 | -28.59 | 0 | -48.05 | 0 | -75.82 | -0.06 | 0.29 | 0.18 | 28.59 | 0.29 | 29.01 | -18.16 | -0.06 | -19.81 | -0.18 | -22.15 | 0 | -23.91
39.936 | 21.39 | -0.23 | 21.5 | -2.17 | 5.04 | -4.04 | -0.06 | -25.2 | -0.06 | 0 | -0.06 | 19.45 | 0.06 | 21.56 | 0.18 | 21.39 | -9.84 | 0.06 | -27.54 | -0.06 | -45.59 | -0.06 | -51.92 | -0.06 | 0.12 | 0.06 | 28.36 | 0.23 | 28.59 | -15.7 | -0.06 | -18.4 | 0 | -20.68 | 0 | -22.09
39.968 | 21.45 | -0.23 | 21.45 | -1.46 | 11.02 | -2.64 | -0.18 | -23.2 | 0 | 0 | 0.06 | 0.12 | 0.12 | 21.33 | 0.12 | 21.33 | -8.03 | 0.06 | -26.9 | -0.18 | -44.36 | 0 | -45.88 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.59 | -14.06 | -0.06 | -16.58 | -0.12 | -18.87 | -0.06 | -20.68
40 | 21.45 | -0.23 | 21.5 | -0.76 | 16.41 | -1.46 | 4.51 | -21.09 | 0 | -45.06 | 0.06 | 0 | 0.12 | 21.39 | 0.18 | 21.33 | -5.92 | 0 | -26.31 | -0.06 | -41.37 | -0.18 | -45.06 | -0.06 | 0.06 | 0 | 28.48 | 0.35 | 28.54 | -12.07 | -0.06 | -15.18 | -0.06 | -18.05 | -0.06 | -19.04
40.032 | 21.45 | -0.23 | 21.45 | -0.23 | 21.74 | -1 | 10.55 | -19.1 | -0.06 | -67.74 | -0.06 | 0 | 0.06 | 19.98 | 0.18 | 21.39 | -4.8 | 0 | -23.67 | -0.06 | -39.44 | 0 | -42.78 | -0.06 | 0 | 0 | 28.65 | 0.35 | 28.65 | -9.61 | 6.27 | -13.71 | 0 | -16.06 | -0.18 | -17.7
40.064 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -0.41 | 17.81 | -16.7 | 0 | -43.83 | -0.06 | 0 | 0.06 | 0.06 | 0.18 | 21.33 | -2.99 | 0 | -20.92 | 0 | -37.97 | 0.06 | -41.13 | -0.18 | 0 | 0.06 | 28.42 | 0.29 | 28.48 | -7.56 | 16.7 | -11.6 | 0.06 | -13.83 | 5.04 | -15.29
40.096 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.18 | 21.56 | -14.82 | -0.06 | -38.32 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.33 | -2.46 | 0.06 | -20.22 | 0 | -35.92 | -0.06 | -38.32 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.48 | -5.33 | 24.32 | -9.84 | 4.8 | -12.25 | 13.18 | -13.71
40.128 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | -12.77 | 0 | -33.69 | -0.06 | 0 | -0.06 | 0 | 0 | 21.21 | -0.94 | 0.06 | -18.46 | -0.06 | -34.69 | -0.06 | -36.74 | 0 | 0 | 0 | 28.48 | 0.23 | 28.65 | -3.75 | 28.77 | -8.14 | 13.13 | -10.37 | 20.74 | -11.84
40.16 | 21.45 | -0.18 | 21.5 | -0.41 | 21.45 | -0.35 | 21.39 | -10.61 | 0 | -31.99 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | -0.7 | 0.06 | -16.52 | 0 | -32.93 | -0.06 | -34.69 | -0.06 | 0 | 0.06 | 28.65 | 0.18 | 28.54 | -2.11 | 28.54 | -6.68 | 20.51 | -9.2 | 28.89 | -10.49
40.192 | 21.45 | -0.18 | 21.45 | -0.35 | 21.5 | -0.35 | 21.39 | -8.56 | -0.06 | -32.52 | 0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -0.53 | 0.06 | -14.82 | 0 | -29.77 | -0.06 | -33.52 | 0 | 0 | 0 | 28.42 | 0.23 | 28.83 | -1 | 28.48 | -4.75 | 28.77 | -7.27 | 28.59 | -9.02
40.224 | 21.45 | -0.18 | 21.39 | -0.29 | 21.45 | -0.41 | 21.39 | -5.68 | -0.06 | -32.29 | 0.06 | 0 | 0.06 | 0 | 0.12 | 21.27 | -0.18 | 0.06 | -12.25 | -0.06 | -28.18 | -0.06 | -30.76 | 0 | 0 | -0.18 | 28.42 | 0.12 | 28.48 | -0.64 | 28.48 | -3.16 | 28.59 | -5.33 | 28.48 | -6.8
40.256 | 21.45 | -0.06 | 21.39 | -0.35 | 21.45 | -0.35 | 21.39 | -3.4 | -0.06 | -30.47 | 0 | 0 | -0.06 | 0 | 0 | 21.21 | -0.18 | -0.06 | -11.13 | 0 | -25.9 | 0.06 | -28.65 | 0 | 0 | 0 | 28.65 | 0.18 | 28.42 | -0.23 | 28.54 | -2.23 | 28.65 | -3.98 | 28.48 | -5.1
40.288 | 0.12 | -0.06 | 21.45 | -0.35 | 21.45 | -0.29 | 21.39 | -2.05 | 0.12 | -28.59 | 0.12 | 0 | 0 | 0 | 0.18 | 21.33 | -0.06 | 0.06 | -9.32 | -0.06 | -24.26 | -0.06 | -26.6 | 0 | 0 | -0.18 | 28.59 | 0.12 | 28.54 | 0.06 | 28.54 | -1.46 | 28.59 | -2.52 | 28.42 | -3.52
40.32 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | -1.35 | -0.18 | -26.13 | -0.06 | 0 | 0.12 | 0 | 0.12 | 21.33 | 0 | 0 | -7.68 | -0.06 | -22.38 | 0 | -25.25 | 0 | -0.29 | -0.06 | 0.23 | 0.12 | 28.48 | 0.18 | 28.42 | -0.94 | 28.36 | -2.05 | 28.59 | -2.46
40.352 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.18 | 21.39 | -0.47 | 0.06 | -24.08 | 0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 5.39 | -5.68 | 0.06 | -19.28 | 0 | -22.79 | 0 | -61.23 | -0.18 | 0 | 0 | 28.48 | 0.35 | 28.59 | -0.64 | 28.59 | -1.11 | 28.48 | -1.41
40.384 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | -0.35 | 21.39 | -0.18 | -0.06 | -21.33 | -0.06 | -58.89 | 0 | 0 | 0.06 | 21.33 | 0.12 | 11.37 | -3.93 | 0.06 | -17.99 | -0.06 | -20.22 | 0 | -54.44 | -0.18 | 0 | 0.18 | 28.48 | 0.35 | 28.83 | -0.23 | 28.59 | -0.64 | 28.36 | -1
40.416 | 0 | 0 | 21.39 | -0.29 | 21.39 | -0.18 | 21.39 | 0.06 | -0.06 | -19.63 | -0.06 | -61.06 | 0.06 | 0 | 0.18 | 21.45 | 0 | 16.82 | -2.7 | 0 | -15.7 | -0.06 | -18.63 | -0.06 | -44.3 | 0 | 0 | 0.06 | 28.59 | 0.53 | 28.48 | 0 | 28.42 | -0.35 | 28.65 | -0.53
40.448 | 0 | -0.18 | 21.39 | -0.29 | 21.5 | -0.29 | 21.33 | 0.18 | 0 | -17.52 | -0.06 | -44.12 | -0.06 | 0.06 | 0.06 | 21.39 | 0.12 | 21.62 | -2.11 | 0.06 | -14.82 | -0.06 | -16.47 | -0.06 | -42.95 | -0.06 | 0 | -0.06 | 28.42 | 0.53 | 28.42 | -0.06 | 28.48 | -0.18 | 28.42 | -0.35
40.48 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | 0.12 | 0.06 | -15.18 | -0.06 | -39.32 | 0 | -57.66 | 0.06 | 21.33 | 0.18 | 21.39 | -1 | 0 | -13.48 | -0.06 | -14.94 | 0 | -39.44 | -0.18 | 0 | -0.06 | 28.65 | 0.29 | 28.3 | 0 | 28.59 | -0.18 | 28.48 | -0.18
40.512 | 0 | -0.06 | 21.45 | -0.23 | 21.5 | -0.23 | 21.39 | 0.18 | 0 | -13.54 | -0.18 | -34.63 | -0.06 | -62.35 | 0.06 | 21.33 | 0.12 | 21.27 | -0.64 | 0.06 | -11.31 | 0 | -13.77 | 0 | -36.86 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.48 | 0.12 | 28.65 | -0.06 | 28.48 | 0
40.544 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.35 | 21.39 | 0.29 | -0.06 | -10.96 | -0.06 | -32.4 | 0.06 | -32.81 | 0.06 | 21.45 | 0.23 | 21.39 | -0.35 | -0.06 | -10.25 | 5.98 | -11.72 | -0.06 | -33.75 | -0.06 | 0 | -0.06 | 28.42 | 0.29 | 28.36 | 0.06 | 28.59 | 0 | 28.65 | 0.06
40.576 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.33 | 0.35 | -0.18 | -8.73 | -0.06 | -32.64 | 0 | -27.83 | -0.06 | 21.33 | 0.06 | 21.45 | -0.23 | 0.06 | -9.02 | 13.95 | -10.14 | 0 | -32.34 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.48 | 0.12 | 28.48 | 0 | 28.59 | 0.06
40.608 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | -0.06 | -6.91 | -0.06 | -31.99 | 0.06 | -33.63 | -0.06 | 21.45 | 0.23 | 21.39 | 0.06 | 8.38 | -7.62 | 21.68 | -9.2 | -0.06 | -31.76 | -0.06 | 0 | 0 | 0.23 | 0.29 | 28.65 | 0.12 | 28.65 | 0.06 | 28.24 | 0.23
40.64 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | 0.41 | 5.68 | -4.69 | -0.06 | -30.53 | -0.06 | -35.39 | 0.06 | 21.39 | 0.23 | 21.39 | -0.06 | 15.65 | -5.92 | 29.18 | -7.85 | 0.06 | -30.06 | -0.06 | 0 | 0 | 0 | 0.18 | 28.48 | 0.12 | 28.65 | 0.06 | 28.59 | 0.12
40.672 | 0 | -0.06 | 0.23 | -0.23 | 21.5 | -0.18 | 21.33 | 0.35 | 13.42 | -2.81 | -0.06 | -28.48 | 0.06 | -33.52 | 0.06 | 21.21 | 0.12 | 21.45 | -0.06 | 23.09 | -4.69 | 28.48 | -6.8 | -0.06 | -28.01 | 0 | 0 | 0 | 0 | 0.18 | 28.48 | 0.12 | 28.59 | 0.06 | 28.59 | 0.18
40.704 | -38.97 | -0.06 | 0.06 | -0.23 | 21.45 | -0.23 | 21.39 | 0.47 | 19.1 | -1.46 | -0.06 | -25.84 | -0.06 | -29.36 | 0 | 0.59 | 0.12 | 21.27 | 0.12 | 28.77 | -4.16 | 28.54 | -5.22 | 9.02 | -25.43 | -0.06 | 0 | 0 | 0 | 0.12 | 28.54 | 0.12 | 28.48 | 0.18 | 28.48 | 0.18
40.736 | -57.37 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.29 | 21.56 | -0.94 | -0.06 | -23.5 | -0.06 | -27.77 | 0.06 | 0.06 | 0.12 | 21.39 | 0.06 | 28.42 | -3.46 | 28.42 | -4.39 | 16 | -23.2 | -0.18 | -57.25 | 0 | 0 | 0.06 | 28.42 | 0.12 | 28.24 | 0.12 | 28.83 | 0.18
40.768 | -41.49 | -0.06 | 0.06 | -0.18 | 21.5 | -0.18 | 21.33 | 0.23 | 21.5 | -0.35 | -0.06 | -21.74 | -0.06 | -26.49 | 0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.83 | -2.75 | 28.48 | -3.34 | 23.91 | -21.62 | -0.06 | -75.47 | 0 | 0 | 0 | 28.48 | -0.06 | 28.48 | 0.12 | 28.3 | 0.23
40.8 | -39.85 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 21.39 | 0 | -0.06 | -19.63 | -0.06 | -25.37 | -0.06 | 0 | 0.12 | 21.27 | 0.06 | 28.54 | -2.11 | 28.42 | -2.64 | 28.83 | -19.81 | 0 | -52.15 | -0.06 | 0 | 0.12 | 28.24 | 0.12 | 28.42 | 0.06 | 28.48 | 0.23
40.832 | -38.38 | -0.06 | 0 | -0.06 | 0.12 | -0.18 | 21.39 | 0.23 | 21.33 | 0.12 | 0.06 | -17.75 | -0.06 | -23.44 | 0 | 0 | 0.18 | 21.33 | 0.29 | 28.65 | -1.46 | 28.65 | -2.17 | 28.71 | -17.93 | 0 | -47.76 | 0 | 0 | 0 | 28.42 | 0.12 | 28.3 | 0.12 | 28.42 | 0.23
40.864 | -34.98 | -0.06 | 0.06 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.29 | -0.06 | -15.29 | 0 | -21.09 | 0.06 | 0 | 0.18 | 21.45 | 0.23 | 28.48 | -1.41 | 28.65 | -1.41 | 28.3 | -15.53 | -0.06 | -40.9 | 0 | 0 | 0 | 28.59 | -0.06 | 28.24 | 0.12 | 28.48 | 0.23
40.896 | -33.52 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.18 | 21.39 | 0.23 | -0.06 | -13.54 | -0.06 | -20.1 | 0 | 0 | 0.23 | 21.15 | 0.23 | 28.71 | -0.82 | 28.54 | -1 | 28.54 | -13.01 | -0.06 | -36.74 | 0 | 0 | 0 | 28.42 | 0 | 28.48 | 0.12 | 28.48 | 0.12
40.928 | -31 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.47 | -0.06 | -10.78 | -0.06 | -17.99 | 0 | 0 | 0.18 | 21.27 | 0.18 | 28.71 | -0.7 | 28.59 | -0.59 | 28.77 | -11.31 | -0.06 | -37.44 | -0.06 | 0 | 0 | 0.18 | 0 | 28.42 | 0.06 | 28.42 | 0.18
40.96 | -29.83 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | 0.53 | -0.18 | -8.91 | -0.06 | -16.11 | 0 | 0 | 0.06 | 21.15 | 0.18 | 28.59 | -0.53 | 28.36 | -0.41 | 28.42 | -9.38 | -0.06 | -37.74 | -0.06 | 0 | 0.06 | 0.06 | -0.06 | 28.65 | 0.12 | 0.29 | 0.18
40.992 | -28.24 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.45 | 0.41 | -0.06 | -6.8 | 0 | -14.36 | 0 | 0 | 0.18 | 21.33 | 0.29 | 28.3 | -0.35 | 28.48 | -0.47 | 28.54 | -7.56 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0 | -0.18 | 0.29 | 0.12 | 0.06 | 0.12
41.024 | -24.43 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.59 | 5.39 | -4.34 | -0.06 | -12.36 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 28.54 | -0.29 | 28.54 | -0.35 | 28.42 | -5.33 | -0.06 | -34.63 | 0 | -61.06 | 0 | 0 | 0 | 0.06 | 0.18 | 0 | 0.18
41.056 | -22.68 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.53 | 11.19 | -2.46 | -0.06 | -10.25 | 0 | 0 | 0.12 | 21.39 | 0.18 | 28.3 | -0.29 | 28.54 | -0.29 | 28.42 | -3.22 | 0 | -32.99 | 0 | -74.18 | 0 | 0 | -0.18 | 0 | 0.12 | 0.06 | 0.06
41.088 | -20.86 | 0 | -6.68 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.53 | 16.64 | -1.64 | 0 | -8.91 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 28.48 | -0.29 | 28.83 | -0.29 | 28.59 | -2.23 | -0.18 | -30.7 | 0 | -51.39 | 0.12 | 0 | -0.23 | 0.06 | 0.12 | 0 | 0.12
41.12 | -18.87 | -0.06 | -63.46 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.74 | -1 | -0.06 | -7.32 | -0.06 | 0 | 0.18 | 21.39 | 0.23 | 28.36 | -0.23 | 28.59 | -0.29 | 28.48 | -1.23 | 9.49 | -29.36 | 0 | -47.35 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
41.152 | -16.52 | 0 | -49.4 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.45 | -0.23 | -0.06 | -5.39 | 0 | -72.83 | 0.12 | 21.33 | 0.29 | 28.71 | -0.35 | 28.77 | -0.23 | 28.59 | -0.76 | 16.76 | -27.54 | 0 | -42.36 | -0.18 | 0 | 0 | 0 | 0.23 | 0 | 0.18
41.184 | -14.77 | 0 | -40.31 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -0.06 | 2.75 | -3.52 | -0.06 | -52.33 | 0.12 | 21.39 | 0.23 | 28.59 | -0.18 | 28.71 | -0.18 | 28.48 | -0.23 | 27.07 | -25.14 | 0 | -37.21 | 0 | 0 | -0.18 | 0.06 | 0.23 | 0 | 0.06
41.216 | -13.07 | -0.06 | -38.61 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.39 | 0.18 | 8.97 | -2.46 | 0 | -33.52 | 0.12 | 21.39 | 0.18 | 28.59 | -0.18 | 28.83 | -0.23 | 28.54 | 0 | 28.77 | -23.67 | 0 | -37.74 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.06
41.248 | -12.01 | -0.06 | -36.51 | -0.18 | -25.08 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.45 | 0.35 | 14.18 | -1.7 | -0.06 | -28.54 | 0.06 | 0.12 | 0.12 | 28.42 | -0.23 | 28.65 | -0.18 | 28.95 | 0.23 | 28.48 | -21.56 | -0.18 | -38.2 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
41.28 | -10.61 | -0.06 | -33.87 | -0.06 | -62.64 | -0.18 | 0 | -0.06 | 21.45 | 0.35 | 21.39 | 0.35 | 19.81 | -1.11 | 0 | -33.87 | 0.18 | 0 | 0.18 | 28.48 | -0.23 | 28.48 | -0.18 | 28.3 | 0.29 | 28.42 | -19.63 | -0.06 | -36.86 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06
41.312 | -9.02 | 0 | -32.05 | -0.06 | -42.6 | -0.06 | -50.57 | -0.06 | 21.45 | 0.18 | 21.39 | 0.35 | 21.39 | -0.47 | 0 | -35.22 | 0.18 | 0 | 0.12 | 28.42 | -0.23 | 28.54 | -0.18 | 28.77 | 0.29 | 28.59 | -17.4 | 0 | -35.22 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
41.344 | -6.97 | -0.06 | -29.94 | -0.18 | -40.14 | -0.06 | -66.92 | -0.06 | 21.39 | 0.12 | 21.39 | 0.35 | 21.33 | -0.23 | -0.06 | -32.34 | 0 | 0 | 0.18 | 28.48 | -0.18 | 28.48 | -0.18 | 28.42 | 0.41 | 28.42 | -15.53 | -0.06 | -33.05 | -0.18 | -51.74 | -0.18 | 0 | -0.06 | 0 | 0
41.376 | -5.8 | -0.06 | -27.77 | -0.06 | -38.09 | -0.06 | -46.17 | -0.06 | 21.09 | 0.12 | 21.39 | 0.35 | 21.33 | -0.18 | -0.18 | -28.95 | 0.06 | 0 | 0.23 | 28.48 | -0.18 | 28.65 | -0.06 | 28.65 | 0.47 | 28.48 | -13.07 | -0.06 | -31.06 | -0.06 | -60.59 | -0.18 | 0 | -0.06 | -33.22 | 0
41.408 | -4.22 | 0 | -26.25 | -0.18 | -36.62 | -0.06 | -42.01 | -0.06 | 21.33 | 0.12 | 21.39 | 0.35 | 21.27 | 0.12 | -0.18 | -27.42 | 0.06 | 0 | 0.18 | 28.42 | -0.18 | 28.36 | -0.18 | 28.3 | 0.47 | 28.54 | -11.31 | -0.06 | -29.12 | 0 | -37.79 | 0 | -19.86 | 0.06 | -89.07 | 0
41.44 | -2.99 | -0.06 | -23.44 | -0.06 | -34.4 | -0.06 | -35.92 | -0.06 | 21.33 | 0.12 | 21.5 | 0.41 | 21.33 | 0.12 | 0 | -26.37 | 0.12 | 0 | 0.18 | 28.59 | -0.18 | 0.23 | -0.18 | 28.42 | 0.47 | 28.36 | -9.32 | -0.06 | -27.01 | -0.06 | -29.01 | -0.18 | -88.83 | -0.06 | -55.78 | 0
41.472 | -2.17 | 4.69 | -22.38 | -0.06 | -31.29 | -0.18 | -31.76 | -0.06 | 21.39 | 0.06 | 21.33 | 0.41 | 21.45 | 0.12 | 7.68 | -24.79 | 0.06 | 0 | 0.29 | 28.54 | -0.18 | 0 | -0.06 | 28.48 | 0.53 | 28.42 | -7.38 | -0.06 | -25.37 | 0 | -32.7 | -0.06 | -59.89 | 0 | -29.94 | 0
41.504 | -0.76 | 12.48 | -19.22 | -0.06 | -29.24 | 0 | -32.76 | -0.06 | 17.64 | 0.06 | 21.39 | 0.47 | 21.33 | 0.29 | 14.94 | -23.2 | 0 | 0 | 0.23 | 0.12 | -0.18 | 0 | -0.18 | 28.48 | 0.53 | 28.42 | -3.98 | -0.06 | -23.32 | -0.06 | -35.27 | -0.18 | -30.06 | 0 | -33.99 | 0
41.536 | -0.59 | 18.05 | -17.05 | -0.06 | -27.19 | -0.06 | -31.82 | -0.06 | 0.12 | 0 | 21.45 | 0.29 | 21.33 | 0.18 | 20.63 | -21.09 | 0.06 | 0 | 0.12 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.47 | 28.59 | -2.7 | 0 | -21.21 | -0.18 | -32.76 | -0.18 | -34.16 | -0.06 | -42.01 | -0.06
41.568 | -0.18 | 21.68 | -15.41 | 0 | -25.55 | -0.06 | -30.76 | -0.06 | 0 | 0 | 21.45 | 0.23 | 20.98 | 0.23 | 21.39 | -19.4 | -0.06 | 0 | 0.12 | 0 | -0.06 | 0 | -0.06 | 28.54 | 0.53 | 28.65 | -1.7 | 6.97 | -19.4 | 0 | -29.65 | -0.18 | -42.25 | -0.06 | -41.19 | 0
41.6 | -0.18 | 21.45 | -14 | -0.06 | -23.09 | -0.06 | -29.01 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | 0.23 | 21.39 | -17.75 | 0.12 | 0 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.18 | 0.41 | 28.48 | -0.76 | 14.71 | -17.99 | 0 | -27.72 | -0.06 | -41.25 | 0 | -36.45 | 0
41.632 | -0.29 | 21.45 | -12.36 | -0.06 | -20.98 | -0.06 | -26.6 | -0.06 | 0 | -0.06 | 21.45 | 0.18 | 21.33 | 0.29 | 21.33 | -15.7 | -0.06 | 0.06 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | 0.35 | 28.42 | -0.7 | 22.33 | -15.94 | -0.18 | -26.84 | -0.06 | -37.68 | -0.06 | -33.11 | 0
41.664 | -0.41 | 21.45 | -10.78 | -0.06 | -18.69 | -0.06 | -23.67 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.15 | 0.29 | 21.45 | -14 | 0.06 | -36.56 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | 0.35 | 28.89 | -0.06 | 28.95 | -13.24 | -0.06 | -25.43 | -0.12 | -33.69 | -0.06 | -32.81 | -0.06
41.696 | -0.35 | 21.45 | -9.14 | -0.06 | -17.05 | -0.06 | -22.15 | 0 | 0 | 0 | 21.39 | 0.18 | 21.27 | 0.12 | 21.33 | -12.36 | 0 | -69.49 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.41 | 28.48 | 0 | 28.59 | -11.13 | -0.06 | -23.5 | -0.06 | -32.99 | 0 | -32.52 | -0.06
41.728 | -0.41 | 21.45 | -7.91 | 0 | -15.41 | -0.06 | -19.69 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.23 | 21.33 | -10.66 | 0 | -46.88 | 0.12 | 0 | 0 | 0 | -0.23 | 0 | 0.35 | 28.42 | 0.23 | 28.59 | -9.43 | 0 | -21.15 | -0.06 | -32.81 | 0 | -31.29 | 0
41.76 | -0.35 | 21.45 | -6.5 | -0.18 | -13.77 | -0.06 | -17.99 | -0.06 | 0 | -0.06 | 21.33 | 0 | 21.27 | 0.23 | 21.39 | -9.32 | 0.06 | -30.53 | 0.23 | 0 | -0.06 | 0 | -0.06 | 0 | 0.41 | 28.42 | 0.18 | 28.48 | -7.03 | 0 | -19.75 | -0.18 | -31.23 | -0.06 | -29.24 | 0
41.792 | -0.23 | 21.45 | -5.1 | -0.06 | -11.6 | -0.06 | -16.17 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.27 | 0.29 | 21.45 | -6.5 | 0.12 | -31.17 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.29 | 28.54 | 0.35 | 28.3 | -4.98 | -0.06 | -18.52 | 0 | -28.48 | 0.06 | -26.78 | 0
41.824 | -0.41 | 21.45 | -3.34 | -0.06 | -10.37 | -0.06 | -13.54 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.33 | 0.18 | 21.27 | -4.86 | 0 | -35.98 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.42 | 0.35 | 28.3 | -2.7 | -0.06 | -16 | -0.06 | -26.43 | -0.06 | -25.84 | 0
41.856 | -0.23 | 21.5 | -2.05 | 0.29 | -8.91 | -0.18 | -11.6 | 0 | 0 | -0.06 | 21.39 | 0.12 | 21.04 | 0.18 | 21.33 | -3.63 | 0 | -34.63 | 0.18 | 0 | -0.06 | -29.88 | -0.06 | 0 | 0.18 | 28.48 | 0.41 | 28.59 | -1.93 | -0.06 | -14.41 | -0.06 | -26.13 | 0 | -24.14 | -0.18
41.888 | -0.29 | 21.45 | -1.17 | 7.85 | -7.56 | -0.06 | -9.2 | 0.06 | 0 | 0 | 21.5 | -0.06 | 21.39 | 0.23 | 21.33 | -2.11 | -0.06 | -31.52 | 0.18 | 0.06 | -0.06 | -73.66 | 0 | 0 | 0.18 | 28.36 | 0.35 | 28.3 | -0.94 | -0.06 | -12.77 | -0.06 | -24.61 | 0 | -22.68 | 0
41.92 | -0.29 | 21.45 | -0.59 | 13.48 | -5.92 | 6.68 | -7.15 | -0.18 | 0.06 | -0.06 | 0.18 | 0 | 21.04 | 0.18 | 21.39 | -1.93 | 0.18 | -29.01 | 0.23 | -46.47 | 0 | -52.09 | -0.06 | 0 | 0.18 | 28.3 | 0.29 | 27.95 | -0.47 | -0.06 | -11.07 | -0.06 | -22.56 | 0.12 | -20.74 | 0
41.952 | -0.29 | 21.45 | -0.23 | 19.28 | -3.98 | 12.42 | -5.16 | -0.06 | -64.46 | 0 | 0 | -0.06 | 21.27 | 0.18 | 21.33 | -0.76 | 0 | -26.9 | 0.06 | -68.62 | -0.06 | -45.18 | 0 | 0 | 0.06 | 28.59 | 0.41 | 28.3 | -0.35 | -0.18 | -9.43 | -0.06 | -20.92 | 0.06 | -19.1 | 0.06
41.984 | -0.35 | 21.45 | 0 | 21.5 | -3.16 | 19.86 | -2.7 | 0 | -63.11 | -0.06 | 0 | -0.06 | 21.27 | 0 | 21.45 | -0.53 | 0 | -26.31 | 0.12 | -44.77 | -0.06 | -45 | 0 | 0 | 0.06 | 3.87 | 0.29 | 28.42 | -0.06 | 0 | -7.03 | -0.06 | -19.51 | 0 | -17.7 | 0
42.016 | -0.23 | 21.45 | -0.06 | 21.45 | -1.93 | 21.5 | -1.99 | 0 | -44.59 | 0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.04 | -0.29 | 5.45 | -24.08 | 0.06 | -45.35 | 0 | -40.9 | -0.06 | -52.97 | 0.23 | 0.12 | 0.23 | 28.42 | 0 | 3.69 | -5.68 | -0.18 | -17.75 | 0 | -15.82 | 0
42.048 | -0.23 | 21.45 | -0.29 | 21.45 | -1.11 | 21.5 | -0.76 | -0.06 | -39.32 | 0 | 0 | -0.06 | 21.33 | 0.12 | 21.33 | -0.18 | 11.31 | -21.91 | 0.12 | -43.89 | 0 | -40.14 | 0 | -68.44 | 0 | 0 | 0.23 | 28.42 | 0.29 | 12.01 | -3.93 | 0.47 | -15.7 | 0 | -14.06 | 0
42.08 | -0.23 | 21.39 | -0.53 | 21.45 | -0.53 | 21.45 | -0.53 | -0.06 | -34.45 | -0.06 | 0 | 0 | 10.72 | 0.06 | 21.33 | -0.06 | 16.82 | -20.74 | 0 | -40.9 | 0 | -38.15 | 0 | -46.88 | 0.12 | 0 | 0.06 | 28.42 | 0.29 | 19.57 | -2.58 | 10.2 | -14.24 | 3.34 | -12.6 | 0
42.112 | -0.18 | 21.33 | -0.35 | 21.45 | -0.06 | 21.45 | 0.06 | -0.06 | -32.81 | -0.06 | 0 | 0.06 | 0.06 | 0 | 21.33 | 0 | 21.86 | -18.69 | 0.06 | -39.26 | -0.06 | -36.27 | 0 | -46.53 | 0.06 | 0 | 0.12 | 28.42 | 0.35 | 29.24 | -1.64 | 17.7 | -12.89 | 11.6 | -10.78 | 3.93
42.144 | -0.18 | 21.5 | -0.35 | 21.45 | -0.18 | 21.39 | 0.06 | -0.06 | -32.64 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | -17.23 | 0.12 | -36.92 | -0.06 | -34.69 | 0 | -46.23 | 0.12 | 0 | 0.18 | 28.65 | 0.23 | 28.36 | -1.05 | 27.83 | -10.37 | 21.91 | -9.02 | 14.71
42.176 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.29 | -0.06 | -32.29 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -15.47 | -0.06 | -35.63 | 0 | -32.7 | 0 | -43.3 | 0.12 | 0 | 0.06 | 28.48 | 0.23 | 28.65 | -0.53 | 28.65 | -9.08 | 29.18 | -7.27 | 22.38
42.208 | -0.18 | 21.39 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | 0 | -30.35 | 0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | -13.54 | 0.12 | -34.04 | -0.06 | -30.24 | -0.06 | -41.13 | 0.06 | 0 | 0.12 | 28.48 | 0.29 | 28.42 | -0.06 | 28.48 | -7.21 | 28.65 | -5.74 | 28.95
42.24 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.45 | 0.53 | -0.06 | -28.24 | -0.18 | 0 | 0 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | -12.07 | 0 | -31.7 | -0.06 | -28.77 | 0 | -38.97 | 0.06 | 0 | 0.06 | 28.36 | 0.23 | 28.83 | 0 | 28.54 | -5.8 | 28.48 | -4.16 | 28.48
42.272 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.39 | 0.41 | -0.06 | -25.08 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -9.84 | -0.06 | -30 | -0.06 | -26.08 | 0 | -37.33 | 0.12 | 0 | 0 | 28.36 | 0.23 | 28.65 | 0.23 | 28.59 | -3.98 | 28.48 | -2.7 | 28.48
42.304 | -0.18 | 21.45 | -0.29 | 21.45 | -0.41 | 21.39 | 0.47 | 8.67 | -23.44 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | -8.09 | 0.12 | -27.36 | 0 | -23.67 | -0.06 | -34.4 | 0.06 | 0 | 0.12 | 28.65 | 0.18 | 28.65 | 0.23 | 28.36 | -2.4 | 28.42 | -1.76 | 28.42
42.336 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.41 | 14.12 | -21.68 | 0.06 | -36.62 | -0.06 | 0 | -0.06 | 0.59 | 0.29 | 21.33 | -6.45 | 1.99 | -25.72 | -0.06 | -22.09 | -0.06 | -32.99 | 0.06 | 0 | 0 | 28.48 | 0.12 | 28.65 | 0.18 | 28.59 | -1.17 | 28.65 | -0.94 | 28.42
42.368 | -0.06 | 0.12 | -0.23 | 21.45 | -0.41 | 21.45 | 0.47 | 19.63 | -19.51 | -0.06 | -69.2 | 0 | 0 | -0.06 | 0.06 | 0.23 | 21.39 | -4.45 | 6.56 | -23.79 | 0 | -20.33 | 0 | -30.59 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.65 | 0.47 | 28.42 | -1 | 28.48 | -0.76 | 28.07
42.4 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.5 | 0.41 | 21.5 | -18.16 | -0.06 | -49.22 | -0.18 | 0 | 0 | 0 | 0.18 | 21.5 | -3.4 | 12.25 | -21.45 | -0.06 | -18.28 | -0.06 | -28.48 | 0.12 | 0.06 | 0 | 28.83 | 0.12 | 28.48 | 0.29 | 28.59 | -0.64 | 28.59 | -0.41 | 28.01
42.432 | -0.18 | 0 | -0.23 | 21.39 | -0.29 | 21.45 | 0.41 | 21.39 | -15.29 | -0.06 | -39.44 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.15 | -2.17 | 17.81 | -19.34 | 0 | -16.23 | 0 | -26.25 | 0.12 | -75.06 | 0 | 28.48 | 0.06 | 28.42 | 0.41 | 28.59 | -0.47 | 28.36 | -0.23 | 28.65
42.464 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.41 | 21.39 | -13.36 | 0.06 | -33.69 | 0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -1.29 | 21.39 | -17.23 | 0 | -14.71 | 0 | -23.26 | 0 | -62.4 | 0 | 0.23 | 0 | 28.65 | 0.35 | 28.48 | -0.18 | 28.42 | -0.06 | 28.71
42.496 | -0.06 | 0 | -0.29 | 21.45 | -0.29 | 21.45 | 0.35 | 21.39 | -10.78 | 0 | -32.11 | -0.06 | -2.23 | -0.06 | 0 | 0.18 | 21.27 | -0.7 | 21.39 | -15.18 | -0.06 | -12.77 | 0 | -20.92 | 0.06 | -47.93 | 0.06 | 0.06 | 0 | 28.48 | 0.29 | 28.48 | -0.18 | 28.42 | 0 | 28.83
42.528 | -0.18 | 0 | -0.29 | 21.39 | -0.18 | 21.39 | 0.41 | 21.39 | -9.2 | -0.18 | -33.22 | 0 | -67.39 | -0.06 | 0 | 0.18 | 21.33 | -0.59 | 20.98 | -13.95 | 4.34 | -11.84 | 0 | -18.75 | 0 | -45.7 | 0 | 0 | -0.06 | 28.65 | 0.35 | 28.36 | -0.06 | 28.42 | 0.12 | 28.36
42.56 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | 0.29 | 21.39 | -6.86 | 0 | -32.93 | -0.18 | -58.6 | -0.18 | 0 | 0.18 | 21.33 | -0.18 | 21.21 | -12.83 | 12.72 | -10.49 | 2.4 | -16.7 | -0.06 | -40.78 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.48 | 0 | 28.65 | 0 | 28.42
42.592 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.18 | 21.33 | -4.16 | -0.06 | -30.47 | 0 | -34.69 | -0.18 | 0 | 0.23 | 21.21 | -0.06 | 21.27 | -11.31 | 22.91 | -8.79 | 10.31 | -15 | 0.06 | -37.21 | -0.06 | 0 | 0 | 28.54 | 0.18 | 28.36 | 0.06 | 28.48 | 0.18 | 28.59
42.624 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.23 | 21.39 | -2.23 | -0.06 | -28.13 | -0.06 | -29.59 | -0.18 | 0 | 0.12 | 21.39 | -0.06 | 21.33 | -9.84 | 29.12 | -7.21 | 20.33 | -13.59 | 0 | -37.79 | -0.18 | 0 | -0.06 | 28.48 | 0.18 | 28.24 | 0 | 28.18 | 0.12 | 28.42
42.656 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.18 | 21.33 | -1.29 | 0 | -26.43 | -0.06 | -34.98 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 21.39 | -8.09 | 28.65 | -5.92 | 28.01 | -12.01 | 0.06 | -37.5 | 0 | 0 | 0 | 28.59 | 0.12 | 28.42 | 0.12 | 28.54 | 0.12 | 28.42
42.688 | -0.18 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.18 | 21.39 | -0.82 | 5.39 | -24.08 | -0.06 | -35.16 | -0.18 | 0 | 0.18 | 21.33 | 0 | 21.39 | -7.15 | 28.59 | -5.1 | 28.77 | -10.61 | 0.06 | -36.33 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.59 | 0.12 | 28.54 | 0.12 | 28.59
42.72 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0 | 21.21 | -0.35 | 11.19 | -22.15 | -0.18 | -32.64 | -0.18 | 0 | 0.12 | 21.33 | 0.23 | 21.21 | -5.74 | 28.36 | -4.22 | 28.65 | -9.26 | 0.06 | -34.98 | -0.18 | 0 | 0 | 28.65 | 0.06 | 28.83 | 0.18 | 28.36 | 0.06 | 28.59
42.752 | -0.06 | 0 | -0.18 | 17.52 | -0.06 | 21.45 | 0.06 | 21.39 | 0 | 16.7 | -19.69 | 0 | -29.24 | -0.18 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | -4.98 | 28.36 | -3.22 | 28.71 | -7.79 | 0 | -32.93 | -0.06 | 0 | -0.06 | 28.65 | 0.06 | 28.59 | 0.06 | 28.36 | 0 | 28.54
42.784 | -0.18 | -33.93 | -0.06 | 0.06 | -0.18 | 10.31 | 0.06 | 21.39 | 0 | 21.56 | -17.87 | -0.06 | -27.01 | -0.06 | -73.36 | 0.18 | 21.33 | 0.29 | 21.33 | -3.75 | 28.65 | -2.46 | 28.36 | -6.15 | 0 | -30.41 | 0 | 0 | 0.06 | 28.65 | 0.06 | 28.42 | 0.12 | 28.3 | 0.06 | 28.42
42.816 | -0.06 | -58.24 | -0.18 | 0 | -0.18 | 0.06 | 0 | 21.39 | 0.35 | 21.39 | -16.06 | -0.06 | -26.54 | -0.06 | -53.91 | 0.18 | 21.33 | 0.18 | 21.39 | -2.99 | 28.83 | -1.99 | 28.59 | -5.16 | 0.06 | -28.3 | -0.06 | 0 | -0.06 | 28.65 | -0.06 | 28.48 | 0.06 | 28.89 | 0.12 | 29.01
42.848 | -0.06 | -44.12 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -13.89 | 0 | -24.61 | -0.18 | -32.05 | 0.18 | 21.39 | 0.18 | 21.39 | -2.52 | 28.65 | -1.46 | 28.65 | -4.16 | -0.06 | -26.37 | -0.06 | 0 | 0 | 28.65 | 0.12 | 28.48 | 0.06 | 28.59 | 0.18 | 28.54
42.88 | 0 | -39.9 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.33 | 0.35 | 21.39 | -12.07 | 0 | -22.56 | -0.06 | -27.72 | 0.18 | 21.39 | 0.23 | 21.39 | -1.99 | 28.65 | -1.23 | 28.48 | -3.4 | 0 | -24.9 | 0 | -49.16 | 0.06 | 28.3 | -0.06 | 28.48 | 0.06 | 28.83 | 0.06 | 28.48
42.912 | -0.06 | -39.08 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.21 | 0.41 | 21.39 | -9.49 | 0.06 | -21.21 | -0.06 | -34.28 | 0.12 | 0.12 | 0.18 | 21.27 | -1.64 | 29.01 | -0.7 | 28.42 | -2.64 | -0.06 | -23.15 | 0 | -82.09 | 0 | 0.23 | 0.06 | 28.71 | 0.06 | 28.48 | 0.12 | 28.48
42.944 | -0.06 | -35.1 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.33 | -7.38 | 0 | -19.28 | -0.06 | -35.33 | 0.06 | 0 | 0.12 | 21.45 | -1.11 | 28.54 | -0.64 | 28.95 | -1.82 | -0.06 | -20.51 | 0 | -53.62 | 0 | 0.06 | 0 | 0.29 | 0.12 | 28.48 | 0.06 | 28.48
42.976 | -0.06 | -33.11 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -5.1 | 0 | -17.58 | -0.06 | -32.52 | 0.06 | 0 | 0.23 | 21.21 | -0.94 | 28.42 | -0.53 | 28.48 | -1.29 | -0.06 | -18.69 | 0 | -48.93 | 0 | 0 | 0.06 | 0.12 | 0.06 | 0.29 | 0.18 | 28.48
43.008 | -0.06 | -31.29 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.47 | 21.33 | -2.99 | 0 | -15.59 | -0.06 | -28.77 | 0.12 | 0 | 0.12 | 21.39 | -0.64 | 28.59 | -0.41 | 28.54 | -0.94 | 0.47 | -16.88 | 0 | -41.08 | 0 | 0.06 | 0 | 0 | 0 | 0.06 | 0.12 | 0.23
43.04 | -0.06 | -29.47 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -1.93 | -0.06 | -13.95 | -0.06 | -27.54 | 0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.48 | -0.29 | 28.36 | -0.7 | 10.31 | -15.18 | -0.06 | -36.68 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.18 | 0.06
43.072 | -0.06 | -28.01 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.33 | -0.7 | 4.63 | -12.36 | -0.06 | -27.07 | 0.18 | 0 | 0.12 | 21.39 | -0.41 | 28.24 | -0.23 | 28.42 | -0.47 | 17.87 | -12.77 | -0.06 | -37.03 | -0.18 | 0 | 0 | 0 | 0.06 | 0 | 0.06 | 0
43.104 | -0.06 | -25.08 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | -0.35 | 12.42 | -10.02 | -0.06 | -24.43 | 0.12 | 0 | 0.18 | 21.39 | -0.47 | 28.42 | -0.29 | 28.71 | -0.29 | 28.18 | -10.61 | 0 | -37.97 | 0 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0
43.136 | -0.06 | -22.85 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -0.23 | 18.05 | -8.2 | -0.06 | -22.62 | 0.18 | 0 | 0.12 | 21.39 | -0.35 | 28.59 | -0.18 | 28.54 | -0.29 | 28.48 | -8.67 | -0.06 | -36.92 | 0 | 0 | 0 | 0 | 0 | 0 | 0.06 | 0
43.168 | -0.06 | -20.92 | -0.06 | -0.06 | -0.06 | 0 | -0.06 | 11.54 | 0.23 | 21.15 | -0.06 | 21.62 | -6.21 | 6.91 | -21.09 | 0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.3 | -0.23 | 28.65 | -0.23 | 28.36 | -6.21 | -0.06 | -35.33 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.06 | 0
43.2 | -0.06 | -18.52 | -0.18 | -51.04 | -0.06 | -0.06 | -0.18 | 0.06 | 0.23 | 21.39 | 0.35 | 21.45 | -4.92 | 12.66 | -19.45 | 0.06 | 0 | 0.06 | 21.15 | -0.35 | 28.65 | -0.29 | 28.71 | -0.35 | 28.42 | -4.45 | -0.06 | -33.52 | -0.06 | 0 | 0.06 | 0 | 0 | 0 | 0.06 | 0
43.232 | 0.06 | -17.23 | -0.06 | -54.14 | -0.06 | -34.34 | 0 | 0 | 0.18 | 21.39 | 0.41 | 21.39 | -3.4 | 18.22 | -17.93 | -0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.36 | -0.35 | 28.71 | -0.29 | 28.48 | -2.46 | -0.06 | -31.7 | -0.06 | 0 | 0.12 | 0 | 0 | 0 | 0 | 0
43.264 | -0.06 | -15.18 | -0.06 | -40.14 | -0.06 | -52.68 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -2.05 | 21.39 | -15.7 | 0.06 | 0 | 0.12 | 0.06 | -0.29 | 28.48 | -0.29 | 28.48 | -0.18 | 28.77 | -1.35 | -0.06 | -29.24 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
43.296 | -0.18 | -13.36 | 0 | -38.79 | 0 | -40.84 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -1.23 | 21.39 | -13.77 | 0 | 0 | 0.12 | 0 | -0.35 | 28.3 | -0.23 | 28.59 | -0.23 | 28.36 | -0.94 | -0.06 | -27.36 | -0.06 | 0 | 0.06 | 0 | -0.06 | 0 | 0 | 0
43.328 | -0.06 | -11.54 | -0.06 | -37.33 | -0.06 | -39.85 | -0.06 | 0 | 0.06 | 21.33 | 0.47 | 21.39 | -0.76 | 21.39 | -12.07 | -0.06 | -33.4 | 0.23 | 0 | -0.29 | 28.59 | -0.23 | 28.54 | -0.23 | 28.3 | -0.18 | -0.18 | -25.55 | -0.06 | -7.5 | -0.06 | 0 | -0.06 | 0 | 0 | 0
43.36 | -0.06 | -10.55 | -0.06 | -35.16 | -0.06 | -37.44 | 0 | 0 | 0.12 | 21.39 | 0.47 | 21.21 | -0.53 | 21.33 | -10.72 | 0.06 | -67.91 | 0.23 | 0 | -0.23 | 28.59 | -0.18 | 28.24 | -0.23 | 28.59 | -0.18 | 0 | -23.5 | -0.18 | -92.99 | 0.06 | -0.18 | 0 | 0 | -0.06 | 0
43.392 | 4.69 | -9.14 | -0.06 | -33.52 | -0.06 | -35.57 | -0.06 | 0 | 0.12 | 21.39 | 0.59 | 21.27 | -0.23 | 21.33 | -9.08 | 0 | -47.23 | 0.18 | 0 | -0.29 | 28.59 | -0.18 | 28.48 | -0.23 | 28.83 | 0.12 | 0 | -21.86 | 0 | -54.03 | 0 | -76.53 | -0.06 | -7.68 | 0 | 0
43.424 | 12.36 | -7.27 | -0.18 | -30.12 | 0 | -32.7 | -0.06 | 0 | -0.06 | 21.27 | 0.53 | 21.39 | -0.06 | 21.39 | -6.45 | 0.06 | -27.72 | 0.23 | 0 | -0.23 | 0.29 | -0.18 | 28.65 | -0.18 | 28.42 | 0.06 | 6.91 | -20.04 | 0 | -48.22 | -0.06 | -60.12 | -0.18 | -82.85 | 0 | -39.67
43.456 | 18.05 | -5.45 | -0.06 | -28.83 | -0.06 | -30.35 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | 21.33 | 0.06 | 21.39 | -4.92 | -0.06 | -31.47 | 0.12 | 0 | -0.18 | 0.06 | -0.23 | 0.35 | -0.18 | 28.48 | 0.35 | 14.53 | -17.75 | -0.06 | -43.3 | 0 | -35.1 | -0.18 | -54.32 | 0 | -81.04
43.488 | 21.68 | -4.39 | 0 | -26.9 | -0.06 | -28.83 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | 0.12 | 21.27 | -3.28 | 0 | -35.68 | 0.12 | 0 | -0.23 | 0 | -0.06 | 0.12 | -0.18 | 28.48 | 0.29 | 22.44 | -15.65 | -0.06 | -37.79 | 0 | -30.53 | -0.18 | -31.93 | -0.06 | -51.74
43.52 | 21.45 | -3.28 | -0.06 | -24.43 | -0.06 | -27.01 | 0 | 0 | 0 | 21.39 | 0.29 | 21.45 | 0.29 | 21.27 | -2.4 | 0 | -34.4 | 0.18 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.18 | 28.83 | 0.35 | 29.01 | -13.48 | -0.06 | -36.74 | -0.06 | -40.72 | -0.06 | -31.64 | -0.18 | -31.64
43.552 | 21.45 | -2.17 | -0.06 | -22.56 | 0 | -25.02 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.45 | 0.18 | 21.39 | -1.23 | 0 | -32.23 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.3 | 0.23 | 28.65 | -11.48 | -0.18 | -38.32 | 0.06 | -42.13 | -0.06 | -41.66 | -0.06 | -32.64
43.584 | 21.45 | -0.94 | 0 | -19.92 | 0 | -22.38 | -0.18 | -1.05 | -0.06 | 0.12 | 0.12 | 21.39 | 0.29 | 21.27 | -0.94 | 0.06 | -27.89 | 0.23 | 0 | -0.23 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0.47 | 28.54 | -9.38 | -0.06 | -37.62 | -0.06 | -36.97 | -0.06 | -40.55 | 0 | -42.48
43.616 | 21.45 | -0.29 | -0.06 | -17.99 | 0 | -19.92 | -0.06 | -73.54 | 0 | 0 | 0.23 | 21.39 | 0.23 | 21.27 | -0.41 | 0.06 | -27.36 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.23 | 28.48 | 0.29 | 28.48 | -7.56 | -0.06 | -35.68 | 0 | -34.28 | -0.06 | -36.15 | 0 | -41.08
43.648 | 21.45 | -0.23 | -0.18 | -15.82 | -0.06 | -18.52 | -0.18 | -58.13 | -0.06 | 0 | 0.18 | 21.39 | 0.29 | 21.39 | -0.29 | 0 | -26.25 | 0.23 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.36 | 0.35 | 28.42 | -4.69 | -0.18 | -34.28 | 0 | -33.22 | -0.18 | -33.28 | -0.06 | -36.86
43.68 | 21.45 | -0.06 | -0.06 | -14.47 | -0.06 | -16.88 | -0.06 | -43.19 | 0 | 0 | 0.06 | 21.39 | 0.29 | 21.33 | -0.29 | 0 | -23.91 | 0.18 | -65.39 | -0.06 | 0 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.59 | -2.81 | -0.06 | -31.7 | 0 | -32.76 | -0.18 | -32.76 | 0 | -32.76
43.712 | 21.45 | -0.18 | 0 | -13.01 | 0 | -14.82 | -0.06 | -38.44 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | 0 | 0.06 | -21.97 | 0.18 | -76.41 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.29 | 28.36 | -1.93 | 4.39 | -29.12 | 0 | -32.23 | -0.06 | -32.93 | -0.06 | -32.81
43.744 | 21.45 | -0.23 | 0 | -11.13 | -0.06 | -12.6 | -0.18 | -33.05 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.12 | 0.06 | -19.63 | 0.12 | -50.04 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.36 | -0.94 | 15.06 | -27.77 | -0.18 | -29.01 | -0.06 | -30.94 | 0.06 | -32.05
43.776 | 21.39 | -0.35 | 3.22 | -9.43 | 0.06 | -11.37 | -0.06 | -32.7 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | 0.12 | -0.06 | -18.63 | 0.23 | -52.44 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.29 | 28.48 | -0.53 | 22.91 | -25.72 | 0 | -27.36 | -0.06 | -28.71 | -0.18 | -30.7
43.808 | 21.45 | -0.35 | 9.38 | -8.32 | -0.06 | -10.14 | -0.18 | -32.46 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.15 | 0.12 | 1.99 | -17.23 | 0.12 | -49.22 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.36 | 0.12 | 28.77 | -0.23 | 28.83 | -24.08 | -0.06 | -26.37 | 0 | -27.19 | -0.18 | -28.3
43.84 | 21.45 | -0.29 | 15 | -7.21 | -0.06 | -8.38 | -0.06 | -31.82 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.06 | 6.97 | -14.94 | 0.18 | -48.52 | -0.06 | -17.4 | -0.06 | 0 | -0.06 | 28.83 | 0.06 | 28.42 | 0 | 28.83 | -22.56 | -0.06 | -24.9 | 0 | -25.78 | -0.06 | -26.54
43.872 | 21.39 | -0.29 | 20.39 | -5.57 | -0.06 | -7.38 | -0.18 | -30.24 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.12 | 14.47 | -13.71 | 0 | -45.7 | -0.18 | -71.25 | 0 | -0.06 | -0.06 | 28.42 | 0.18 | 28.13 | 0.06 | 28.65 | -19.98 | 0 | -23.32 | -0.06 | -24.67 | -0.18 | -26.31
43.904 | 21.39 | -0.18 | 21.5 | -3.46 | 0 | -5.22 | -0.06 | -27.19 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.39 | 0.06 | 20.16 | -11.48 | 0.12 | -43.01 | -0.18 | -49.63 | 0 | -79.46 | -0.06 | 0.29 | 0.06 | 28.42 | 0.29 | 28.65 | -18.16 | -0.06 | -21.15 | -0.18 | -22.62 | 0 | -23.61
43.936 | 21.45 | -0.18 | 21.45 | -2.46 | 4.16 | -3.93 | -0.18 | -25.14 | 0 | 0 | -0.06 | 21.27 | 0.06 | 21.39 | 0.18 | 21.15 | -9.61 | 0.06 | -42.78 | -0.06 | -45.88 | -0.06 | -53.03 | -0.18 | 0.12 | 0.06 | 28.65 | 0.23 | 28.24 | -16 | -0.06 | -19.4 | -0.06 | -20.57 | 0 | -22.21
43.968 | 21.45 | -0.18 | 21.56 | -1.41 | 10.49 | -2.81 | -0.18 | -23.44 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.39 | 0.23 | 21.39 | -8.09 | 0.06 | -40.08 | -0.06 | -44.83 | -0.06 | -46 | -0.06 | 0 | 0.12 | 28.65 | 0.35 | 28.89 | -14.24 | -0.06 | -17.75 | 0 | -19.45 | -0.06 | -20.74
44 | 21.39 | -0.23 | 21.45 | -0.82 | 15.47 | -1.7 | 4.22 | -21.33 | 0 | -38.38 | 0 | 0 | 0 | 21.39 | 0.12 | 21.33 | -6.15 | 0 | -38.97 | -0.18 | -41.9 | 0 | -45.35 | -0.06 | 0.06 | -0.06 | 28.42 | 0.35 | 28.65 | -12.36 | 0 | -16.58 | -0.06 | -18.16 | -0.06 | -19.4
44.032 | 21.45 | -0.18 | 21.5 | -0.53 | 21.8 | -0.94 | 10.43 | -19.16 | -0.06 | -71.96 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.33 | -4.51 | 0.06 | -37.09 | 0 | -40.2 | 0 | -42.89 | -0.18 | 0 | -0.18 | 28.48 | 0.41 | 28.59 | -9.61 | 5.04 | -15 | -0.06 | -16.29 | -0.06 | -17.7
44.064 | 21.39 | -0.23 | 21.45 | -0.18 | 21.5 | -0.23 | 17.46 | -16.47 | 0 | -45.41 | -0.06 | 0 | 0.06 | 0.06 | 0.23 | 21.39 | -2.99 | 0 | -34.4 | -0.18 | -38.56 | 0 | -40.72 | -0.06 | 0 | 0.06 | 28.59 | 0.29 | 28.42 | -7.68 | 15.59 | -12.89 | -0.06 | -13.83 | 4.22 | -15.7
44.096 | 21.39 | -0.35 | 21.45 | -0.18 | 21.39 | -0.06 | 21.74 | -14.77 | 0.06 | -39.61 | -0.06 | 0 | 0.12 | 0 | 0.18 | 21.39 | -2.05 | 0 | -33.69 | 0 | -36.45 | -0.06 | -38.15 | -0.06 | 0 | -0.18 | 28.65 | 0.35 | 28.48 | -5.51 | 23.26 | -10.78 | 4.39 | -12.48 | 12.42 | -14
44.128 | 21.39 | -0.29 | 21.45 | -0.18 | 21.5 | -0.18 | 21.45 | -12.89 | -0.06 | -34.1 | 0 | 0 | 0 | 0 | 0.12 | 21.39 | -1.17 | 0.06 | -30.76 | -0.06 | -34.28 | 0 | -36.68 | 0 | 0 | 0 | 28.65 | 0.23 | 28.42 | -3.69 | 28.71 | -9.14 | 12.72 | -10.9 | 20.16 | -12.48
44.16 | 21.39 | -0.18 | 21.5 | -0.41 | 21.45 | -0.29 | 21.39 | -10.61 | -0.06 | -31.99 | 0 | 0 | 0 | 0 | 0.12 | 21.39 | -0.82 | 0.12 | -29.53 | -0.06 | -32.76 | -0.06 | -35.45 | -0.06 | 0 | -0.06 | 28.42 | 0.23 | 28.48 | -2.23 | 28.65 | -7.79 | 20.22 | -9.32 | 27.72 | -11.13
44.192 | 21.39 | -0.18 | 21.45 | -0.47 | 21.45 | -0.29 | 21.33 | -8.79 | -0.06 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.33 | -0.41 | 0 | -27.77 | -0.06 | -30.29 | -0.06 | -33.99 | 0 | 0 | 0 | 28.3 | 0.18 | 28.54 | -0.94 | 28.54 | -6.27 | 28.13 | -7.56 | 28.65 | -9.43
44.224 | 21.39 | -0.18 | 21.45 | -0.47 | 21.45 | -0.35 | 21.39 | -5.92 | -0.06 | -32.17 | -0.06 | 0 | 0 | 0 | 0.12 | 21.21 | -0.18 | 0 | -24.55 | -0.06 | -28.77 | -0.06 | -30.94 | 0 | 0 | -0.06 | 28.42 | 0.23 | 28.65 | -0.64 | 28.42 | -4.28 | 28.71 | -5.45 | 28.42 | -7.21
44.256 | 21.45 | -0.06 | 21.45 | -0.35 | 21.5 | -0.29 | 21.39 | -3.93 | 0.06 | -30.47 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | -0.06 | 0 | -23.67 | 0 | -26.6 | -0.06 | -28.77 | 0 | 0 | -0.18 | 28.48 | 0.12 | 28.59 | -0.35 | 28.42 | -2.93 | 28.65 | -4.51 | 28.59 | -5.39
44.288 | 0.12 | -0.06 | 21.5 | -0.35 | 21.39 | -0.29 | 21.39 | -2.23 | -0.06 | -28.83 | -0.18 | 0 | 0 | 0 | 0.12 | 21.33 | 0 | 0 | -21.09 | -0.06 | -24.26 | 0.06 | -26.9 | -0.18 | 0 | -0.06 | 28.48 | 0.06 | 28.59 | -0.18 | 28.54 | -1.88 | 28.36 | -2.75 | 28.48 | -3.98
44.32 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | -1.35 | 0 | -26.13 | -0.06 | 0 | 0 | 0 | 0.06 | 21.33 | 0.12 | 0 | -19.92 | 0.12 | -22.68 | -0.06 | -25.02 | -0.06 | -0.06 | -0.06 | 0.23 | 0 | 28.36 | 0 | 28.36 | -1.17 | 28.36 | -2.11 | 28.65 | -2.4
44.352 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.29 | 21.45 | -0.76 | 0.06 | -24.08 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 4.8 | -17.99 | -0.06 | -19.98 | -0.06 | -23.38 | -0.06 | -71.66 | -0.06 | 0.06 | 0.12 | 28.48 | 0.18 | 28.54 | -0.82 | 28.71 | -1.23 | 28.54 | -2.11
44.384 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.23 | 21.39 | -0.18 | 0 | -21.74 | -0.06 | -47.81 | 0.06 | 0 | 0.18 | 21.39 | 0.12 | 10.72 | -15.53 | 0 | -18.63 | 0 | -20.16 | -0.06 | -57.83 | -0.18 | 0 | 0 | 28.42 | 0.29 | 28.3 | -0.47 | 28.48 | -0.35 | 28.48 | -1.05
44.416 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.29 | 21.45 | -0.06 | -0.06 | -19.63 | 0.06 | -64.16 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 16.29 | -14.3 | 0.06 | -16.41 | 0 | -18.4 | -0.06 | -46.41 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.42 | -0.23 | 28.65 | -0.41 | 28.65 | -0.53
44.448 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.18 | 21.39 | 0.18 | -0.06 | -17.93 | -0.06 | -46.76 | 0.06 | 0 | 0.18 | 21.39 | 0.18 | 21.5 | -12.6 | 0.06 | -14.82 | -0.06 | -16.58 | -0.06 | -43.13 | -0.18 | 0 | 0.06 | 28.59 | 0.35 | 28.3 | 0 | 28.36 | -0.23 | 28.48 | -0.47
44.48 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | 0.35 | -0.06 | -15.7 | 0 | -40.67 | 0 | -50.86 | 0 | 21.39 | 0.12 | 21.39 | -11.72 | -0.06 | -13.71 | -0.06 | -15.23 | -0.18 | -36.51 | 0 | 0 | -0.06 | 28.54 | 0.23 | 28.3 | 0 | 28.65 | -0.06 | 28.48 | -0.23
44.512 | 0 | -0.06 | 21.39 | -0.29 | 21.39 | -0.23 | 21.45 | 0.53 | -0.06 | -14.06 | 0 | -34.75 | -0.06 | -64.4 | -0.06 | 21.5 | 0.18 | 21.39 | -10.49 | -0.06 | -11.54 | 0.06 | -14.24 | -0.06 | -34.45 | 0 | 0 | -0.06 | 28.65 | 0.23 | 28.54 | 0 | 28.54 | -0.18 | 28.71 | -0.06
44.544 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | 0.41 | 0.06 | -11.07 | -0.06 | -31.82 | -0.06 | -34.28 | 0 | 21.39 | 0.29 | 21.27 | -8.56 | 0 | -9.84 | 5.63 | -12.01 | -0.06 | -35.74 | 0 | 0 | 0 | 28.48 | 0.29 | 28.3 | 0.12 | 28.48 | 0.06 | 28.42 | 0.06
44.576 | 0 | 0 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | 0.47 | 0.06 | -9.2 | -0.06 | -32.76 | 0 | -27.19 | 0.06 | 21.39 | 0.18 | 21.33 | -7.62 | -0.06 | -8.91 | 13.89 | -10.84 | -0.06 | -35.16 | 0 | 0 | -0.06 | 28.65 | 0.23 | 28.48 | 0.12 | 28.59 | 0 | 28.65 | 0.23
44.608 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | 0.53 | 0 | -7.03 | -0.18 | -32.05 | -0.06 | -33.46 | 0 | 21.39 | 0.12 | 21.39 | -6.21 | 6.97 | -7.38 | 21.33 | -9.32 | -0.06 | -34.22 | 0 | 0 | 0 | 0.23 | 0.23 | 28.42 | 0.06 | 28.48 | 0.12 | 28.42 | 0.18
44.64 | 0 | 0 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | 0.59 | 5.39 | -4.63 | 0.06 | -30.94 | -0.06 | -35.86 | -0.06 | 21.39 | 0.06 | 21.5 | -5.45 | 14.59 | -6.27 | 29.06 | -7.85 | 0 | -32.05 | -0.18 | 0 | -0.18 | 0.06 | 0.18 | 28.42 | 0 | 28.54 | 0.18 | 28.18 | 0.18
44.672 | 0 | -0.18 | 1.11 | -0.18 | 21.39 | -0.18 | 21.39 | 0.53 | 13.18 | -2.93 | 0.06 | -28.71 | -0.06 | -33.69 | 0 | 21.39 | 0.18 | 21.33 | -4.34 | 22.5 | -4.75 | 28.59 | -6.86 | 0 | -30 | 0 | 0 | -0.06 | 0 | 0.18 | 28.42 | 0.23 | 28.65 | 0.12 | 28.59 | 0.23
44.704 | -36.74 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.33 | 0.53 | 18.75 | -1.76 | -0.06 | -25.55 | -0.18 | -29.01 | -0.06 | 6.74 | 0.12 | 21.39 | -3.46 | 28.65 | -3.93 | 28.48 | -5.22 | 8.2 | -27.25 | 0 | 0.06 | -0.06 | 0 | 0.18 | 28.42 | 0.06 | 28.42 | 0.18 | 28.83 | 0.18
44.736 | -58.95 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.35 | 21.5 | -1.11 | -0.06 | -23.5 | -0.06 | -27.72 | 0.06 | 0.06 | 0.06 | 21.33 | -2.99 | 28.71 | -3.46 | 28.59 | -4.16 | 15.29 | -26.02 | -0.06 | -45.82 | 0 | 0 | 0.06 | 28.42 | 0.06 | 28.42 | 0.23 | 28.48 | 0.23
44.768 | -43.07 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.41 | 21.45 | -0.23 | -0.06 | -21.86 | -0.06 | -26.72 | -0.06 | 0 | 0.18 | 21.33 | -2.4 | 28.54 | -2.64 | 28.83 | -3.57 | 23.26 | -24.2 | -0.06 | -79.98 | -0.18 | 0 | 0.06 | 28.42 | 0.06 | 28.13 | 0.12 | 28.36 | 0.29
44.8 | -39.61 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.35 | 21.45 | 0.12 | 0.06 | -19.81 | -0.06 | -25.84 | 0 | 0 | 0.23 | 21.39 | -2.05 | 28.83 | -2.05 | 28.48 | -2.87 | 28.95 | -21.86 | -0.06 | -53.97 | 0 | 0 | 0.06 | 28.42 | 0 | 28.42 | 0.18 | 28.48 | 0.35
44.832 | -38.15 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 21.39 | 0.29 | 21.39 | 0.18 | 0 | -18.34 | -0.06 | -23.38 | -0.06 | 0 | 0.06 | 21.39 | -1.52 | 28.77 | -1.41 | 28.77 | -2.17 | 28.77 | -19.75 | 0 | -48.93 | 0.06 | 0 | -0.06 | 28.54 | 0.06 | 28.42 | 0.12 | 28.59 | 0.23
44.864 | -35.33 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.18 | 21.39 | 0.12 | -0.06 | -15.53 | -0.06 | -20.74 | 0.06 | 0 | 0.23 | 21.33 | -1.11 | 28.42 | -1.23 | 28.48 | -1.52 | 28.54 | -17.75 | 0 | -40.55 | -0.18 | 0 | 0.06 | 28.59 | -0.06 | 28.3 | 0.18 | 28.42 | 0.23
44.896 | -33.87 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | 0.41 | -0.06 | -13.48 | 0 | -19.45 | -0.06 | 0 | 0.12 | 21.27 | -0.94 | 28.36 | -1 | 28.24 | -1 | 28.42 | -15.23 | -0.06 | -36.62 | -0.06 | 0 | 0 | 28.42 | 0 | 28.42 | 0.18 | 28.48 | 0.23
44.928 | -31.41 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.47 | 0 | -11.31 | -0.06 | -18.28 | -0.06 | 0 | 0.12 | 21.27 | -0.76 | 28.65 | -0.7 | 28.54 | -0.7 | 28.42 | -13.54 | 0 | -37.62 | -0.06 | 0 | -0.06 | 0.23 | -0.06 | 28.48 | 0.18 | 28.48 | 0.18
44.96 | -29.59 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | -0.06 | -9.2 | 0 | -16.47 | 0 | 0 | 0.12 | 21.45 | -0.53 | 28.59 | -0.53 | 28.48 | -0.59 | 28.36 | -11.89 | 0.06 | -38.2 | -0.06 | 0 | 0.06 | 0 | -0.06 | 28.24 | 0 | 0.29 | 0.23
44.992 | -27.54 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.39 | 0.53 | -0.06 | -7.62 | -0.06 | -14.59 | -0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.13 | -0.47 | 28.36 | -0.53 | 28.54 | -9.84 | 0 | -37.56 | -0.06 | 0 | -0.06 | 0 | 0 | 0.29 | 0.12 | 0.06 | 0.23
45.024 | -25.31 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 5.45 | -4.45 | -0.06 | -12.42 | -0.06 | 0 | 0.18 | 21.45 | -0.53 | 28.3 | -0.35 | 28.65 | -0.29 | 28.59 | -6.8 | -0.06 | -34.63 | -0.06 | -48.69 | 0 | 0 | -0.06 | 0.06 | 0.12 | 0 | 0.12
45.056 | -22.62 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 11.13 | -2.46 | -0.06 | -10.72 | -0.06 | 0 | 0.12 | 21.21 | -0.47 | 28.65 | -0.23 | 28.59 | -0.23 | 28.48 | -4.45 | 0 | -33.17 | 0 | -78.23 | 0 | 0 | -0.18 | 0 | 0.18 | 0.06 | 0.12
45.088 | -20.92 | -0.06 | -4.22 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 16.7 | -1.7 | -0.06 | -9.08 | 0 | 0 | 0.18 | 21.15 | -0.47 | 28.48 | -0.23 | 28.83 | -0.23 | 28.36 | -3.28 | 0 | -31.23 | -0.06 | -54.2 | -0.18 | 0 | -0.18 | 0 | 0.18 | 0 | 0.12
45.12 | -18.75 | -0.06 | -62.58 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.35 | 21.74 | -1.29 | -0.06 | -7.38 | 0 | 0.06 | 0.06 | 21.39 | -0.53 | 28.01 | -0.35 | 28.71 | -0.29 | 28.42 | -1.93 | 8.5 | -29.42 | 0 | -47.76 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.18
45.152 | -16.93 | 0 | -50.16 | -0.06 | 0 | -0.06 | 0 | 0 | 21.5 | 0.41 | 21.39 | -0.18 | -0.06 | -5.92 | 0 | -70.49 | 0.18 | 21.39 | -0.53 | 28.54 | -0.29 | 28.42 | -0.29 | 28.42 | -1.29 | 15.65 | -27.36 | 0 | -42.83 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.06
45.184 | -14.59 | -0.06 | -40.72 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | 0 | 0.59 | -3.4 | 0.06 | -55.31 | 0.12 | 21.39 | -0.47 | 28.36 | -0.29 | 28.65 | -0.23 | 28.48 | -0.53 | 25.96 | -25.43 | 0 | -36.45 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
45.216 | -13.24 | -0.06 | -39.08 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | 0.23 | 8.44 | -2.23 | -0.06 | -34.51 | 0.12 | 21.39 | -0.41 | 28.36 | -0.23 | 28.65 | -0.23 | 28.71 | 0.06 | 28.65 | -24.08 | 0 | -37.38 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18
45.248 | -12.31 | -0.06 | -35.74 | -0.18 | -14.71 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.45 | 0.35 | 13.71 | -1.7 | 0 | -27.48 | 0.18 | 0.12 | -0.41 | 28.42 | -0.23 | 28.59 | -0.23 | 28.65 | 0 | 28.71 | -21.74 | 0 | -38.44 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06
45.28 | -10.31 | -0.06 | -34.22 | -0.18 | -63.99 | 0 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.23 | 19.34 | -1.11 | -0.06 | -32.81 | 0.23 | 0.06 | -0.29 | 28.54 | -0.18 | 28.48 | -0.18 | 28.59 | 0.41 | 28.48 | -19.75 | -0.06 | -37.21 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06
45.312 | -9.38 | 0 | -32.52 | -0.18 | -43.42 | -0.06 | -20.63 | -0.06 | 21.39 | 0.23 | 21.39 | 0.53 | 21.39 | -0.59 | -0.06 | -35.45 | 0.06 | 0 | -0.35 | 28.48 | -0.23 | 28.59 | -0.18 | 28.71 | 0.29 | 28.42 | -17.93 | -0.06 | -35.57 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0
45.344 | -6.86 | -0.06 | -30 | -0.18 | -40.2 | -0.06 | -67.33 | -0.06 | 21.56 | 0 | 21.39 | 0.41 | 21.39 | -0.23 | -0.18 | -32.52 | 0.12 | 0 | -0.29 | 28.54 | -0.18 | 28.3 | -0.18 | 28.24 | 0.59 | 28.54 | -15.53 | 0 | -33.11 | -0.06 | -59.83 | -0.18 | 0 | 0 | 0 | -0.06
45.376 | -6.15 | -0.06 | -27.31 | -0.06 | -38.44 | -0.06 | -47.11 | -0.06 | 21.39 | 0.23 | 21.45 | 0.47 | 21.27 | -0.06 | -0.06 | -29.77 | 0.06 | 0 | -0.35 | 28.48 | -0.23 | 28.65 | -0.18 | 28.07 | 0.53 | 28.48 | -13.65 | -0.06 | -30.94 | -0.18 | -78.28 | -0.06 | 0 | 0 | -23.2 | 0
45.408 | -4.45 | -0.06 | -25.9 | -0.06 | -36.62 | -0.18 | -41.9 | -0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.39 | 0 | -0.06 | -27.83 | 0.06 | 0 | -0.41 | 28.42 | -0.06 | 28.59 | -0.06 | 28.59 | 0.53 | 28.48 | -11.54 | -0.06 | -29.24 | 0 | -46.29 | -0.18 | -13.42 | 0 | -89.48 | 0
45.44 | -3.16 | -0.06 | -23.79 | -0.06 | -34.63 | -0.06 | -36.21 | 0.06 | 21.27 | -0.06 | 21.39 | 0.47 | 21.33 | 0.06 | -0.06 | -26.84 | 0.06 | 0 | -0.29 | 28.48 | -0.06 | 0.23 | -0.18 | 28.54 | 0.59 | 28.54 | -9.43 | -0.06 | -27.95 | -0.18 | -30.18 | -0.06 | -90.3 | -0.06 | -59.24 | 0
45.472 | -1.99 | 3.75 | -21.45 | -0.06 | -32.23 | -0.18 | -32.23 | 0 | 21.27 | 0.12 | 21.39 | 0.53 | 21.39 | 0.35 | 6.62 | -24.84 | 0.12 | 0 | -0.35 | 28.59 | -0.18 | 0 | -0.06 | 28.48 | 0.53 | 28.54 | -7.44 | -0.18 | -26.13 | 0 | -35.16 | -0.18 | -61.82 | 0 | -30.41 | 0
45.504 | -0.94 | 11.78 | -19.04 | -0.06 | -29.94 | -0.06 | -32.58 | -0.12 | 19.75 | 0.06 | 21.33 | 0.53 | 21.39 | 0.35 | 14.18 | -22.91 | 0.12 | 0 | -0.35 | 0.18 | -0.18 | 0 | -0.18 | 28.42 | 0.47 | 28.42 | -4.75 | 0 | -23.91 | -0.06 | -42.36 | -0.18 | -31.76 | -0.06 | -32.46 | 0
45.536 | -0.35 | 17.17 | -17.58 | -0.06 | -27.6 | -0.06 | -32.23 | -0.06 | 0.12 | 0.06 | 21.33 | 0.41 | 21.39 | 0.23 | 19.86 | -21.09 | 0.06 | 0 | -0.35 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.53 | 28.65 | -2.93 | 0 | -21.86 | 0 | -39.44 | -0.18 | -33.52 | 0 | -42.07 | -0.06
45.568 | -0.06 | 21.91 | -15.7 | -0.06 | -26.13 | -0.06 | -30.82 | -0.18 | 0 | 0.06 | 21.27 | 0.29 | 21.33 | 0.35 | 21.56 | -20.22 | -0.06 | 0 | -0.29 | 0 | 0 | 0 | 0 | 28.59 | 0.53 | 28.48 | -1.88 | 5.68 | -19.86 | 0 | -35.1 | -0.18 | -41.66 | 0 | -41.19 | -0.06
45.6 | -0.06 | 21.5 | -14.12 | 0 | -23.61 | -0.06 | -28.95 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | 0.18 | 21.33 | -18.52 | 0.06 | 0 | -0.29 | 0 | -0.06 | 0 | -0.06 | 0.23 | 0.53 | 28.59 | -1.23 | 13.71 | -17.87 | -0.06 | -33.4 | -0.06 | -41.66 | 0 | -37.5 | -0.18
45.632 | -0.18 | 21.5 | -12.6 | -0.06 | -20.92 | -0.06 | -26.66 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.45 | 0.18 | 21.39 | -15.94 | 0.12 | 0 | -0.29 | 0 | -0.18 | 0 | -0.18 | 0 | 0.47 | 28.54 | -0.64 | 21.33 | -16.17 | -0.18 | -32.34 | -0.18 | -37.21 | -0.06 | -33.52 | 0
45.664 | -0.41 | 21.45 | -10.72 | -0.06 | -19.1 | -0.06 | -23.91 | 0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.33 | 0.18 | 21.39 | -13.83 | 0 | -32.99 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.41 | 28.54 | -0.18 | 29.01 | -13.36 | 0 | -31.82 | -0.18 | -33.63 | 0 | -32.93 | 0
45.696 | -0.35 | 21.45 | -9.61 | -0.06 | -17.23 | -0.06 | -21.91 | -0.06 | 0 | 0 | 21.5 | 0.18 | 21.27 | 0.29 | 21.39 | -12.6 | -0.06 | -71.31 | -0.29 | 0 | 0 | 0 | -0.06 | 0 | 0.47 | 28.48 | -0.18 | 28.48 | -11.48 | 0 | -29.77 | -0.06 | -33.28 | 0 | -32.52 | -0.06
45.728 | -0.35 | 21.45 | -7.91 | -0.18 | -15.82 | 0 | -20.22 | -0.18 | 0 | -0.06 | 21.39 | 0.18 | 21.21 | 0.29 | 21.33 | -10.96 | 0.06 | -47.46 | -0.23 | 0 | -0.06 | 0 | -0.06 | 0 | 0.35 | 28.71 | 0.12 | 28.71 | -9.84 | -0.06 | -27.77 | -0.18 | -32.93 | 0 | -30.76 | 0
45.76 | -0.29 | 21.45 | -6.56 | -0.06 | -14 | -0.06 | -18.16 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.39 | 0.23 | 21.33 | -9.38 | 0 | -30.41 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.48 | 0.12 | 28.48 | -7.44 | -0.06 | -26.6 | -0.18 | -31 | 0 | -28.77 | -0.06
45.792 | -0.29 | 21.45 | -4.92 | -0.06 | -12.25 | -0.18 | -16.17 | -0.18 | 0 | 0 | 21.39 | -0.06 | 21.39 | 0.18 | 21.45 | -7.32 | 0.06 | -30.29 | -0.23 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.42 | 0.35 | 28.59 | -5.33 | -0.06 | -25.55 | -0.18 | -28.71 | -0.18 | -27.6 | -0.06
45.824 | -0.18 | 21.39 | -3.34 | -0.06 | -10.96 | -0.06 | -13.48 | 0 | 0 | -0.06 | 21.45 | 0.06 | 21.33 | 0.12 | 21.39 | -5.39 | 0 | -35.45 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.23 | 28.42 | 0.29 | 28.48 | -2.75 | 0 | -23.56 | -0.18 | -26.49 | -0.18 | -26.08 | 0
45.856 | -0.18 | 21.39 | -2.29 | -0.06 | -9.2 | -0.06 | -11.48 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.33 | 0.06 | 21.56 | -4.04 | -0.06 | -34.45 | -0.18 | 0 | -0.06 | -25.02 | -0.18 | 0 | 0.18 | 28.42 | 0.35 | 28.59 | -1.93 | 0 | -21.33 | -0.18 | -25.43 | 0 | -24.61 | 0
45.888 | -0.23 | 21.45 | -1.46 | 7.09 | -7.79 | -0.06 | -9.67 | -0.18 | 0 | 0 | 21.5 | 0.06 | 21.21 | 0.12 | 21.09 | -2.4 | -0.06 | -31.41 | -0.29 | 0 | -0.06 | -74.59 | -0.06 | 0 | 0.06 | 28.3 | 0.29 | 28.65 | -0.94 | -0.18 | -20.33 | -0.06 | -24.14 | -0.06 | -23.03 | 0
45.92 | -0.29 | 21.45 | -0.76 | 12.83 | -6.33 | 6.33 | -7.44 | -0.06 | 0 | 0 | 0.12 | 0 | 21.39 | 0.18 | 21.27 | -1.82 | 0 | -28.71 | -0.29 | -36.39 | -0.06 | -52.97 | -0.06 | 0 | 0.06 | 28.59 | 0.35 | 28.36 | -0.41 | 0.06 | -18.69 | -0.18 | -22.62 | -0.18 | -20.74 | 0
45.952 | -0.23 | 21.45 | -0.41 | 18.28 | -4.63 | 12.19 | -5.57 | -0.06 | -60.53 | 0 | 0 | -0.06 | 21.33 | 0.12 | 21.33 | -1.23 | -0.06 | -27.01 | -0.29 | -72.19 | 0 | -46.12 | 0 | 0 | 0.18 | 28.36 | 0.41 | 28.36 | -0.29 | 0 | -17.05 | -0.06 | -21.21 | -0.06 | -19.4 | 0
45.984 | -0.18 | 21.45 | -0.06 | 21.5 | -3.16 | 19.4 | -2.7 | -0.06 | -62.87 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.33 | -0.23 | -0.06 | -26.08 | -0.18 | -46.29 | -0.18 | -44.24 | -0.06 | 0 | 0 | 18.22 | 0.41 | 28.36 | 0.23 | -0.06 | -14.65 | -0.18 | -19.16 | -0.06 | -17.7 | 0
46.016 | -0.23 | 21.5 | -0.23 | 21.5 | -1.7 | 21.62 | -1.76 | -0.06 | -44.06 | 0.06 | 0.06 | -0.06 | 21.33 | 0 | 21.21 | -0.29 | 4.63 | -23.91 | -0.23 | -44.71 | -0.18 | -41.13 | -0.06 | -45.24 | 0.12 | 0.12 | 0.35 | 28.54 | 0.23 | 2.7 | -13.18 | -0.06 | -17.58 | 0 | -15.7 | -0.18
46.048 | -0.23 | 21.45 | -0.35 | 21.45 | -1.23 | 21.45 | -1 | -0.06 | -39.55 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.33 | -0.06 | 10.66 | -21.74 | -0.29 | -43.48 | -0.18 | -40.78 | -0.06 | -70.78 | 0 | 0 | 0.29 | 28.48 | 0.23 | 11.84 | -11.54 | 0.18 | -16.06 | -0.06 | -14.18 | 0
46.08 | -0.23 | 21.45 | -0.41 | 21.39 | -0.59 | 21.39 | -0.47 | -0.18 | -34.45 | 0.06 | 0 | -0.06 | 15.65 | 0 | 21.39 | 0 | 15.94 | -20.57 | -0.18 | -41.25 | 0 | -39.08 | 0 | -47.46 | -0.06 | 0 | 0.23 | 28.48 | 0.29 | 19.4 | -10.02 | 9.73 | -13.89 | 0.76 | -12.72 | 0
46.112 | -0.18 | 21.45 | -0.41 | 21.45 | -0.18 | 21.5 | -0.23 | 0 | -32.46 | -0.06 | 0 | 0.06 | 0.12 | 0.06 | 21.39 | 0.12 | 21.45 | -18.93 | -0.18 | -38.85 | -0.06 | -35.86 | -0.06 | -45.76 | 0 | 0 | 0.23 | 28.48 | 0.41 | 29.01 | -8.2 | 16.93 | -12.6 | 10.55 | -10.72 | 1.88
46.144 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.06 | -0.06 | -32.46 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | -17.4 | -0.18 | -37.38 | 0 | -34.22 | 0.06 | -44.77 | 0.12 | 0 | 0.29 | 28.59 | 0.35 | 28.65 | -6.09 | 27.19 | -10.37 | 20.98 | -8.85 | 13.71
46.176 | -0.18 | 21.39 | -0.41 | 21.5 | -0.29 | 21.5 | 0.23 | 0 | -32.05 | 0 | 0 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | -15.65 | -0.18 | -35.16 | -0.06 | -33.11 | -0.06 | -41.43 | 0.12 | 0 | 0.12 | 28.65 | 0.23 | 28.48 | -4.45 | 28.65 | -8.85 | 28.83 | -7.62 | 21.27
46.208 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.23 | 0 | -30.59 | -0.06 | 0 | 0.06 | 0 | 0 | 21.21 | 0.23 | 21.33 | -14.18 | -0.18 | -34.04 | -0.06 | -31.29 | 0 | -40.08 | 0 | 0 | 0.06 | 28.48 | 0.29 | 28.24 | -2.99 | 28.54 | -7.09 | 28.48 | -5.8 | 29.18
46.24 | -0.18 | 21.39 | -0.29 | 21.5 | -0.35 | 21.5 | 0.41 | 0.06 | -28.18 | -0.06 | 0 | 0.06 | 0 | 0 | 21.15 | 0.29 | 21.39 | -12.42 | -0.18 | -32.23 | -0.06 | -28.83 | 0 | -37.91 | 0 | 0 | 0.12 | 28.48 | 0.29 | 28.59 | -2.29 | 28.65 | -5.57 | 28.59 | -3.98 | 28.48
46.272 | -0.18 | 21.45 | -0.23 | 21.45 | -0.41 | 21.45 | 0.23 | 1.99 | -25.31 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -10.2 | -0.18 | -29.83 | 0 | -25.84 | -0.06 | -36.21 | 0.12 | 0 | 0 | 28.83 | 0.29 | 28.13 | -1.29 | 28.48 | -4.28 | 28.54 | -2.64 | 28.65
46.304 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.45 | 0.41 | 8.44 | -23.5 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -8.14 | -0.18 | -28.13 | -0.06 | -23.85 | -0.06 | -34.22 | -0.06 | 0 | 0.12 | 28.36 | 0.12 | 28.48 | -0.53 | 28.48 | -2.4 | 28.42 | -1.7 | 28.42
46.336 | -0.18 | 21.39 | -0.18 | 21.45 | -0.29 | 21.39 | 0.47 | 13.95 | -21.33 | 0.06 | -36.33 | 0.06 | 0 | -0.18 | 10.72 | 0.29 | 21.21 | -6.09 | -0.18 | -26.08 | 0 | -22.38 | 0 | -31.99 | -0.06 | 0 | 0.06 | 28.65 | 0.06 | 28.48 | -0.29 | 28.59 | -1.99 | 28.24 | -1.23 | 28.54
46.368 | -0.06 | 0.18 | -0.23 | 21.45 | -0.18 | 21.5 | 0.47 | 19.34 | -20.22 | 0 | -69.73 | -0.06 | 0 | -0.18 | 0.12 | 0.23 | 21.5 | -4.63 | 8.67 | -23.79 | -0.06 | -20.45 | 0 | -30.47 | 0 | 0 | 0 | 28.48 | 0.12 | 28.54 | -0.29 | 28.42 | -1.17 | 28.54 | -0.64 | 28.59
46.4 | -0.06 | 0 | -0.23 | 21.5 | -0.23 | 21.45 | 0.53 | 21.5 | -17.93 | 0 | -49.16 | -0.06 | 0 | -0.18 | 0 | 0.23 | 21.09 | -3.28 | 16.11 | -21.68 | -0.06 | -18.22 | 0 | -28.48 | 0.06 | 0 | 0.06 | 28.48 | 0.06 | 28.65 | -0.06 | 28.83 | -0.35 | 28.42 | -0.47 | 28.59
46.432 | -0.06 | 0 | -0.23 | 21.5 | -0.23 | 21.45 | 0.47 | 21.39 | -15.35 | 0 | -39.2 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -1.99 | 24.08 | -20.1 | 0 | -16.23 | -0.06 | -26.43 | 0 | -62.52 | 0 | 28.59 | 0.06 | 28.54 | 0 | 28.48 | -0.29 | 28.59 | -0.18 | 28.48
46.464 | -0.06 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | 0.47 | 21.39 | -13.18 | 0 | -34.1 | -0.06 | 0 | 0 | 0 | 0.29 | 21.39 | -1.11 | 29.42 | -17.05 | 0 | -14.47 | -0.06 | -23.61 | 0 | -65.8 | 0 | 0.23 | 0 | 28.54 | 0.23 | 28.42 | -0.29 | 28.42 | -0.18 | 28.48
46.496 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.5 | 0.41 | 21.39 | -11.31 | -0.06 | -31.7 | -0.06 | -0.53 | -0.06 | 0 | 0.23 | 21.39 | -1.05 | 29.36 | -15.53 | 0 | -13.13 | 0.06 | -21.91 | -0.06 | -48.52 | -0.06 | 0.12 | 0 | 28.59 | 0.18 | 28.54 | 0 | 28.59 | -0.06 | 28.59
46.528 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.41 | 21.39 | -9.55 | -0.06 | -32.81 | -0.06 | -64.98 | -0.06 | 0 | 0.18 | 21.45 | -0.29 | 29.3 | -14.71 | 3.69 | -12.13 | 0 | -19.69 | -0.18 | -46.35 | -0.06 | 0 | 0.06 | 28.83 | 0.23 | 28.36 | 0 | 28.07 | 0.06 | 28.59
46.56 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | 0.23 | 21.39 | -7.21 | 0.06 | -32.58 | -0.06 | -60.88 | -0.06 | 0 | 0.12 | 21.39 | -0.23 | 29.24 | -13.01 | 12.13 | -10.37 | 2.05 | -18.05 | -0.06 | -41.25 | 0 | 0 | 0 | 28.48 | 0.12 | 28.54 | 0 | 27.89 | 0.06 | 28.59
46.592 | 0 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | 0.35 | 21.33 | -4.28 | 0 | -31.29 | -0.06 | -35.74 | -0.06 | 0 | 0.18 | 21.27 | 0 | 29.06 | -11.48 | 22.27 | -8.85 | 9.26 | -16.29 | -0.06 | -36.86 | 0 | 0 | 0 | 28.48 | 0.23 | 28.59 | 0.12 | 28.48 | 0.18 | 28.36
46.624 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0 | 21.45 | -2.52 | 0.12 | -28.36 | -0.06 | -28.48 | -0.06 | 0 | 0.12 | 21.21 | 0.06 | 29.24 | -10.08 | 29.12 | -7.62 | 19.34 | -14.06 | -0.18 | -37.38 | -0.06 | 0 | 0 | 28.54 | 0.12 | 28.48 | 0.12 | 28.48 | 0.12 | 28.13
46.656 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.23 | 21.39 | -1.7 | -0.06 | -26.43 | -0.18 | -34.75 | 0.12 | 0 | 0.18 | 21.39 | 0 | 29.3 | -8.56 | 28.65 | -6.39 | 26.9 | -12.77 | -0.18 | -38.09 | 0 | 0 | 0 | 28.3 | 0.18 | 28.71 | 0.18 | 28.3 | 0.06 | 28.36
46.688 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.5 | 0.06 | 21.45 | -0.64 | 4.57 | -24.43 | -0.06 | -35.63 | 0 | 0 | 0.12 | 21.39 | 0.06 | 29.36 | -7.32 | 28.48 | -5.51 | 28.71 | -11.48 | -0.06 | -36.62 | -0.06 | 0 | 0 | 28.59 | 0.12 | 28.42 | 0.23 | 28.48 | 0.06 | 28.48
46.72 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0 | 21.45 | -0.41 | 10.55 | -22.15 | -0.06 | -32.76 | -0.18 | 0 | 0.29 | 21.15 | 0.06 | 29.24 | -6.39 | 28.65 | -4.39 | 28.71 | -9.9 | 0 | -35.45 | 0.12 | 0 | 0 | 28.36 | 0 | 28.36 | 0.23 | 28.48 | 0.06 | 28.83
46.752 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.12 | 21.39 | 0.06 | 15.94 | -19.86 | -0.06 | -29.24 | -0.06 | 0.06 | 0.18 | 21.21 | 0.23 | 29.3 | -5.1 | 28.65 | -3.16 | 27.89 | -8.61 | 0 | -33.22 | -0.18 | 0 | -0.06 | 28.59 | 0.12 | 28.3 | 0.18 | 28.36 | 0.18 | 28.42
46.784 | -0.18 | -29.47 | -0.06 | 0.06 | -0.18 | 13.48 | -0.06 | 21.39 | 0.06 | 21.5 | -17.7 | -0.06 | -27.77 | -0.06 | -70.2 | 0.12 | 21.21 | 0.18 | 29.36 | -3.98 | 28.59 | -2.29 | 28.83 | -6.91 | -0.06 | -30.47 | -0.18 | 0 | -0.18 | 28.42 | 0 | 28.71 | 0.18 | 28.24 | 0.12 | 28.36
46.816 | -0.06 | -60.3 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 21.39 | 0.23 | 21.39 | -15.94 | -0.18 | -27.01 | 0.12 | -57.42 | 0.23 | 21.39 | 0.18 | 29.42 | -3.34 | 28.65 | -1.46 | 28.65 | -5.98 | -0.06 | -28.77 | -0.18 | 0 | -0.06 | 28.65 | 0.18 | 28.48 | 0.12 | 28.48 | 0 | 28.48
46.848 | -0.06 | -45.53 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.35 | 21.39 | -14 | 0 | -25.37 | 0 | -33.93 | 0.18 | 21.39 | 0.18 | 28.95 | -2.52 | 28.36 | -1.52 | 28.3 | -4.69 | -0.06 | -26.6 | -0.18 | 0 | -0.06 | 27.95 | -0.06 | 28.48 | 0.23 | 28.48 | 0.12 | 28.48
46.88 | -0.06 | -40.55 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | -11.95 | -0.06 | -23.32 | -0.06 | -27.01 | 0.12 | 21.39 | 0.12 | 29.36 | -1.88 | 28.65 | -1.11 | 28.3 | -3.87 | -0.06 | -25.43 | -0.06 | -38.61 | -0.06 | 28.48 | 0.06 | 28.42 | 0.18 | 28.42 | 0.06 | 28.42
46.912 | -0.18 | -38.91 | -0.06 | 0 | -0.18 | 0 | 0 | 21.27 | 0.47 | 21.39 | -9.14 | -0.06 | -21.09 | -0.06 | -33.34 | 0.12 | 0.12 | 0.06 | 29.18 | -1.7 | 28.13 | -0.7 | 28.83 | -3.11 | -0.06 | -23.38 | -0.06 | -86.66 | -0.06 | 0.29 | 0 | 28.48 | 0 | 28.48 | 0.12 | 28.48
46.944 | -0.06 | -35.45 | 0 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.33 | -6.91 | 0 | -19.45 | -0.06 | -35.27 | 0.12 | 0 | 0.29 | 29.3 | -1.11 | 28.36 | -0.59 | 28.59 | -2.17 | -0.06 | -20.51 | 0 | -56.08 | -0.18 | 0.12 | 0.06 | 1.35 | 0.18 | 28.48 | 0.12 | 28.48
46.976 | -0.06 | -33.46 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.47 | 21.33 | -5.1 | -0.06 | -17.99 | -0.18 | -32.23 | 0.12 | 0 | 0.23 | 29.06 | -0.94 | 28.3 | -0.47 | 28.48 | -1.64 | -0.06 | -19.04 | 0 | -49.4 | -0.06 | 0 | 0 | 0.12 | 0.06 | 0.29 | 0.06 | 28.42
47.008 | -0.06 | -31.64 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.53 | 21.33 | -2.7 | 0 | -15.94 | -0.06 | -28.89 | 0.06 | 0 | 0.18 | 29.42 | -0.64 | 28.3 | -0.29 | 28.65 | -1.17 | -0.06 | -16.99 | -0.18 | -42.66 | -0.18 | 0.06 | 0 | 0 | 0 | 0.12 | 0.12 | 0.23
47.04 | -0.06 | -29.65 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -2.05 | -0.06 | -14.36 | -0.06 | -27.6 | 0.12 | 0 | 0.18 | 29.24 | -0.53 | 28.48 | -0.29 | 28.48 | -0.94 | 9.49 | -15.29 | 0 | -37.03 | -0.06 | 0 | 0.06 | 0 | 0 | 0 | 0.18 | 0.06
47.072 | -0.18 | -28.07 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.21 | -0.82 | 3.75 | -12.42 | -0.18 | -26.72 | 0.12 | 0 | 0.06 | 29.47 | -0.47 | 28.42 | -0.35 | 28.65 | -0.64 | 16.64 | -13.13 | -0.06 | -37.44 | -0.18 | 0 | 0 | 0 | -0.18 | 0.06 | 0.06 | 0
47.104 | -0.06 | -25.72 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | -0.29 | 11.48 | -10.25 | -0.06 | -24.79 | 0.06 | 0 | 0.12 | 29.42 | -0.29 | 28.59 | -0.23 | 28.89 | -0.41 | 26.95 | -10.66 | 0 | -37.97 | -0.06 | 0.06 | 0 | 0 | -0.06 | 0 | 0 | 0
47.136 | -0.06 | -22.91 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -0.06 | 17.17 | -8.73 | 0.06 | -22.68 | 0.06 | 0 | 0.23 | 29.24 | -0.35 | 28.59 | -0.47 | 28.89 | -0.35 | 28.71 | -8.38 | -0.18 | -37.21 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
47.168 | -0.06 | -21.09 | -0.06 | 0 | -0.06 | 0 | -0.18 | 13.42 | 0.29 | 21.39 | 0.18 | 21.74 | -6.97 | 6.68 | -20.98 | 0 | 0 | 0.12 | 29.24 | -0.35 | 28.54 | -0.29 | 28.48 | -0.35 | 28.48 | -6.5 | 0 | -35.1 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0
47.2 | 0 | -19.1 | -0.06 | -42.48 | -0.06 | -0.06 | -0.06 | 0.12 | 0.18 | 21.39 | 0.29 | 21.39 | -5.16 | 12.36 | -20.16 | 0.06 | 0 | 0.18 | 29.24 | -0.35 | 28.36 | -0.29 | 28.71 | -0.29 | 28.65 | -4.63 | -0.06 | -33.52 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0.06
47.232 | -0.06 | -17.23 | -0.06 | -56.43 | -0.06 | -32.81 | -0.18 | 0 | 0.23 | 21.33 | 0.35 | 21.45 | -3.52 | 18.05 | -18.28 | 0 | 0 | 0.12 | 29.36 | -0.41 | 28.36 | -0.23 | 28.59 | -0.23 | 28.48 | -2.75 | -0.06 | -31.47 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
47.264 | -0.06 | -15.18 | 0 | -40.43 | 0 | -53.79 | -0.18 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -1.99 | 21.45 | -15.76 | -0.06 | 0 | 0.18 | 0.12 | -0.29 | 28.36 | -0.35 | 28.48 | -0.29 | 28.42 | -1.7 | -0.06 | -29.77 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
47.296 | 0 | -13.71 | 0 | -39.79 | -0.06 | -40.55 | 0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.39 | -1.7 | 21.33 | -14.36 | 0.06 | 0 | 0.23 | 0.06 | -0.35 | 28.59 | -0.18 | 28.65 | -0.23 | 28.48 | -0.7 | -0.06 | -28.01 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 0
47.328 | -0.06 | -12.07 | -0.06 | -37.62 | -0.06 | -39.55 | -0.18 | 0 | 0.06 | 21.15 | 0.47 | 21.39 | -0.64 | 21.39 | -12.77 | 0.06 | -24.38 | 0.18 | 0 | -0.23 | 28.48 | -0.23 | 28.54 | -0.29 | 28.42 | -0.47 | -0.18 | -25.61 | -0.06 | -4.16 | -0.06 | 0 | -0.06 | 0 | 0 | 0
47.36 | -0.06 | -10.49 | 0 | -35.39 | -0.06 | -38.32 | -0.06 | 0 | 0.12 | 21.21 | 0.41 | 21.39 | -0.53 | 21.39 | -10.96 | -0.06 | -70.43 | 0.12 | 0 | -0.35 | 28.42 | -0.18 | 28.48 | -0.23 | 28.65 | -0.18 | -0.06 | -23.91 | 0 | -93.05 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0
47.392 | 3.75 | -9.43 | -0.06 | -33.75 | 0 | -35.45 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.33 | -0.18 | 21.39 | -9.55 | 0.06 | -50.33 | 0.23 | 0 | -0.23 | 28.48 | -0.18 | 28.65 | -0.18 | 28.48 | -0.06 | -0.06 | -22.33 | 0 | -53.21 | 0 | -72.13 | -0.06 | -1.29 | -0.06 | 0
47.424 | 11.89 | -7.56 | -0.06 | -30 | -0.06 | -33.11 | -0.06 | 0 | -0.06 | 21.15 | 0.47 | 21.39 | -0.18 | 21.39 | -7.32 | -0.06 | -28.65 | 0.12 | 0 | -0.29 | 0.23 | -0.06 | 28.54 | -0.18 | 28.71 | 0.23 | 5.68 | -19.92 | -0.06 | -48.87 | -0.06 | -62.4 | -0.18 | -88.01 | -0.06 | -33.05
47.456 | 17.34 | -6.09 | -0.06 | -28.71 | -0.06 | -31.76 | -0.06 | 0 | -0.06 | 21.45 | 0.41 | 21.15 | 0.06 | 21.33 | -5.22 | 0.06 | -31.06 | 0.12 | 0 | -0.23 | 0.06 | -0.18 | 4.57 | -0.18 | 28.18 | 0.12 | 13.83 | -18.16 | -0.06 | -43.89 | -0.06 | -36.68 | -0.06 | -58.6 | -0.06 | -84.32
47.488 | 21.86 | -4.45 | -0.06 | -26.49 | -0.18 | -29.71 | -0.18 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.18 | 21.39 | -3.75 | 0 | -35.74 | 0.12 | 0 | -0.23 | 0 | -0.23 | 0.12 | -0.18 | 28.54 | 0.35 | 21.62 | -15.76 | -0.06 | -37.85 | -0.06 | -30 | -0.06 | -33.87 | -0.18 | -55.84
47.52 | 21.5 | -3.16 | -0.06 | -25.55 | -0.18 | -26.13 | -0.06 | 0 | -0.06 | 21.15 | 0.35 | 21.33 | 0.18 | 21.33 | -2.64 | -0.06 | -35.27 | 0.06 | 0 | -0.23 | 0.06 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 28.83 | -13.77 | -0.06 | -36.74 | 0.06 | -39.61 | -0.06 | -30.7 | -0.18 | -32.64
47.552 | 21.45 | -2.11 | -0.06 | -22.62 | 0 | -24.84 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.33 | 0.06 | 21.39 | -1.64 | 0 | -31.93 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.29 | 28.48 | 0.41 | 28.54 | -12.13 | -0.06 | -38.32 | -0.06 | -43.01 | 0 | -40.78 | -0.18 | -31.41
47.584 | 21.45 | -1.11 | 0 | -19.98 | 0 | -22.21 | -0.18 | -0.29 | 0.12 | 0.12 | 0.18 | 21.39 | 0.12 | 21.33 | -0.94 | 0 | -28.48 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.65 | 0.29 | 28.18 | -10.08 | -0.06 | -37.03 | -0.06 | -37.79 | 0 | -40.84 | -0.23 | -42.78
47.616 | 21.45 | -0.47 | -0.06 | -18.28 | -0.06 | -20.39 | -0.06 | -70.55 | -0.06 | 0 | 0.23 | 21.27 | 0.12 | 21.39 | -0.53 | -0.06 | -27.19 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.29 | 28.59 | -7.56 | -0.18 | -35.63 | -0.18 | -33.99 | -0.06 | -37.15 | -0.18 | -41.31
47.648 | 21.45 | -0.06 | 0 | -15.94 | -0.06 | -18.34 | -0.18 | -58.95 | -0.06 | 0 | 0.12 | 21.33 | 0.23 | 21.27 | -0.23 | -0.06 | -26.08 | 0.12 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.65 | 0.35 | 28.42 | -5.04 | -0.06 | -34.04 | -0.18 | -33.4 | 0 | -33.93 | 0 | -37.44
47.68 | 21.5 | -0.18 | 0 | -14.71 | -0.06 | -16.29 | -0.18 | -44.53 | -0.18 | 0 | 0.18 | 21.33 | 0.29 | 21.39 | -0.06 | 0 | -23.97 | 0.12 | -41.02 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.23 | 28.59 | -3.11 | 0 | -31.64 | -0.18 | -33.52 | -0.06 | -33.4 | -0.18 | -33.28
47.712 | 21.45 | -0.18 | -0.06 | -13.36 | -0.06 | -14.82 | -0.06 | -39.44 | -0.18 | 0 | 0.06 | 21.27 | 0.35 | 21.27 | -0.18 | 0.12 | -22.38 | 0.12 | -68.85 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.36 | -1.88 | 3.69 | -29.77 | 0 | -32.34 | 0 | -32.81 | -0.18 | -32.99
47.744 | 21.39 | -0.35 | 0 | -11.31 | 0 | -13.13 | -0.18 | -32.93 | -0.12 | 0 | 0.06 | 21.39 | 0.29 | 21.45 | 0 | -0.18 | -20.22 | 0.12 | -44.65 | -0.06 | 0 | -0.18 | 0 | 0 | 28.65 | 0.29 | 28.54 | -1.05 | 14.41 | -28.13 | -0.06 | -28.83 | -0.18 | -31.17 | -0.18 | -32.64
47.776 | 21.45 | -0.53 | 2.99 | -9.9 | -0.06 | -11.25 | -0.06 | -32.7 | -0.18 | 0 | 0.06 | 21.27 | 0.23 | 21.33 | 0.06 | -0.06 | -18.87 | 0.06 | -45.59 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.59 | -0.59 | 22.21 | -26.13 | -0.18 | -27.77 | -0.18 | -29.06 | -0.18 | -30.7
47.808 | 21.45 | -0.35 | 9.2 | -8.44 | -0.06 | -10.25 | -0.18 | -32.76 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.33 | 0.12 | -0.06 | -16.99 | 0.18 | -43.77 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.18 | 28.54 | -0.06 | 29.01 | -24.08 | -0.18 | -26.72 | -0.06 | -27.19 | 0 | -28.71
47.84 | 21.45 | -0.35 | 14.71 | -7.03 | -0.06 | -9.02 | -0.06 | -32.34 | -0.06 | 0 | -0.06 | 21.5 | 0.29 | 21.39 | 0.29 | 6.68 | -15.18 | 0.12 | -41.49 | -0.18 | -8.61 | -0.06 | 0 | -0.18 | 28.3 | 0.06 | 28.77 | 0 | 28.65 | -22.62 | -0.06 | -24.79 | 0 | -26.13 | -0.06 | -27.77
47.872 | 21.39 | -0.29 | 20.39 | -5.63 | -0.06 | -7.27 | -0.18 | -30.06 | -0.06 | 0 | -0.18 | 21.33 | 0.06 | 21.15 | 0.29 | 14.3 | -14 | 0.12 | -39.79 | 0 | -73.07 | -0.06 | 0 | -0.18 | 28.36 | 0.18 | 28.65 | 0 | 28.48 | -19.92 | 0 | -23.32 | -0.06 | -24.61 | -0.06 | -26.54
47.904 | 21.39 | -0.18 | 21.5 | -3.75 | -0.06 | -5.51 | -0.06 | -27.25 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.15 | 0.35 | 19.98 | -11.72 | 0.12 | -37.38 | -0.06 | -50.69 | -0.06 | -76.47 | 0 | 4.04 | 0 | 28.36 | 0.23 | 28.48 | -17.99 | 0 | -21.27 | 0 | -22.09 | 0 | -24.32
47.936 | 21.39 | -0.18 | 21.45 | -2.58 | 3.63 | -3.98 | -0.18 | -25.25 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.33 | 0.18 | 21.45 | -9.84 | 0 | -36.15 | -0.06 | -45.94 | -0.18 | -57.19 | 0 | 0.12 | 0.06 | 28.42 | 0.29 | 28.65 | -16.64 | -0.18 | -19.69 | -0.06 | -20.63 | 0 | -22.38
47.968 | 21.39 | -0.18 | 21.45 | -1.46 | 9.84 | -2.93 | -0.18 | -23.5 | 0 | 0 | 0.06 | 0.18 | 0.12 | 21.39 | 0.23 | 21.39 | -8.14 | 0.06 | -33.75 | 0 | -45.35 | 0 | -45.76 | -0.18 | 0 | 0.18 | 28.42 | 0.35 | 28.42 | -14.65 | 0 | -17.99 | -0.06 | -19.69 | -0.06 | -20.86
48 | 21.5 | -0.23 | 21.45 | -0.94 | 15.18 | -1.82 | 3.57 | -21.27 | -0.06 | -28.83 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.33 | -6.21 | 0.06 | -32.46 | 0 | -41.84 | 0 | -45.35 | -0.18 | 0 | 0.12 | 28.36 | 0.35 | 28.48 | -12.36 | 0 | -16.64 | -0.06 | -17.99 | -0.06 | -19.63
48.032 | 21.45 | -0.29 | 21.5 | -0.35 | 21.8 | -0.94 | 9.84 | -19.34 | -0.18 | -74.65 | 0 | 0 | 0.18 | 21.5 | 0.12 | 21.33 | -4.8 | 0.12 | -31.06 | 0 | -40.14 | -0.06 | -43.71 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.59 | -9.61 | 3.81 | -15.18 | -0.06 | -16.47 | -0.06 | -18.22
48.064 | 21.39 | -0.29 | 21.45 | -0.23 | 21.5 | -0.29 | 17.11 | -17.05 | 0 | -47.46 | -0.06 | 0 | 0.06 | 0.12 | 0.23 | 21.21 | -2.87 | 0 | -27.77 | -0.06 | -38.32 | 0 | -41.02 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.42 | -7.62 | 14.71 | -12.6 | -0.06 | -14.06 | 3.87 | -15.94
48.096 | 21.39 | -0.29 | 21.39 | -0.18 | 21.45 | -0.18 | 21.74 | -14.94 | 0 | -41.02 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -2.17 | 0 | -26.13 | -0.06 | -35.92 | -0.06 | -39.44 | -0.06 | 0 | 0 | 28.71 | 0.29 | 28.42 | -5.39 | 22.38 | -11.13 | 3.34 | -12.48 | 12.13 | -14.41
48.128 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | -0.06 | 21.45 | -12.6 | -0.06 | -35.39 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.39 | -1.17 | 0 | -24.55 | -0.18 | -34.69 | -0.06 | -37.21 | 0 | 0 | -0.06 | 28.48 | 0.23 | 28.54 | -3.75 | 28.95 | -9.32 | 11.78 | -10.78 | 19.92 | -12.72
48.16 | 21.39 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -11.13 | 0.06 | -31.88 | -0.06 | 0.06 | 0.06 | 0 | 0.18 | 21.39 | -0.94 | 0 | -22.09 | -0.18 | -33.05 | -0.06 | -35.86 | -0.06 | 0 | 0 | 28.71 | 0.23 | 28.48 | -1.99 | 28.59 | -8.03 | 19.22 | -9.67 | 27.54 | -11.07
48.192 | 21.45 | -0.23 | 21.39 | -0.35 | 21.45 | -0.29 | 21.45 | -9.08 | -0.06 | -32.29 | -0.18 | 0 | 0.06 | 0 | 0.18 | 21.33 | -0.29 | 0 | -21.09 | -0.18 | -30.12 | -0.06 | -33.87 | -0.06 | 0 | 0.06 | 28.65 | 0.12 | 28.71 | -1.46 | 28.65 | -6.33 | 26.95 | -7.62 | 28.42 | -9.2
48.224 | 21.39 | -0.18 | 21.45 | -0.41 | 21.45 | -0.41 | 21.39 | -5.86 | 0.06 | -32.29 | 0 | 0 | 0 | 0 | 0.12 | 21.27 | -0.35 | -0.06 | -17.93 | 0 | -28.95 | 0 | -31.06 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.65 | -0.47 | 28.48 | -3.98 | 28.65 | -5.92 | 28.65 | -7.44
48.256 | 21.45 | -0.18 | 21.45 | -0.29 | 21.5 | -0.29 | 21.39 | -3.75 | -0.06 | -30.7 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.06 | -0.06 | -16.29 | -0.06 | -26.6 | 0 | -29.06 | -0.06 | 0 | -0.06 | 28.54 | 0.12 | 28.59 | -0.18 | 28.54 | -2.75 | 28.65 | -4.04 | 28.42 | -5.74
48.288 | 0.12 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.39 | -2.11 | -0.06 | -29.01 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0.06 | 0.06 | -15.06 | 0 | -24.43 | 0 | -27.77 | -0.06 | 0 | -0.06 | 28.3 | 0.18 | 28.48 | 0 | 28.59 | -1.99 | 28.42 | -2.64 | 28.42 | -4.1
48.32 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -1.46 | 0 | -26.37 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | -0.06 | -13.54 | 0 | -22.56 | 0 | -25.37 | -0.06 | 0.06 | 0 | 0.29 | 0.06 | 28.65 | 0.06 | 28.42 | -1.17 | 28.36 | -2.11 | 28.48 | -2.87
48.352 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.33 | -0.64 | 0 | -24.67 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 4.57 | -12.07 | -0.06 | -20.16 | -0.18 | -23.32 | -0.06 | -65.04 | 0 | 0.12 | 0 | 28.42 | 0.23 | 28.3 | -0.76 | 28.24 | -1.17 | 28.42 | -2.23
48.384 | 0 | -0.06 | 21.39 | -0.29 | 21.39 | -0.23 | 21.39 | -0.06 | 0.06 | -22.33 | -0.06 | -40.61 | -0.18 | 0 | 0.12 | 21.39 | 0.18 | 10.55 | -10.66 | -0.06 | -18.22 | 0 | -21.33 | -0.06 | -65.04 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.42 | -0.53 | 28.48 | -0.76 | 28.24 | -1.17
48.416 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | -0.29 | 21.39 | -0.06 | 0 | -19.75 | -0.06 | -67.15 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 15.82 | -9.02 | -0.06 | -16.76 | 0 | -18.34 | -0.06 | -48.4 | -0.06 | 0 | 0 | 28.42 | 0.47 | 28.42 | -0.23 | 28.83 | -0.41 | 28.48 | -0.7
48.448 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | 0.23 | 0 | -18.52 | -0.06 | -47.64 | -0.06 | 0 | 0.12 | 21.33 | 0 | 21.39 | -7.79 | 0 | -15 | 0 | -17.05 | -0.06 | -46.35 | 0 | 0 | 0 | 28.65 | 0.41 | 28.36 | -0.06 | 28.54 | -0.18 | 28.42 | -0.47
48.48 | 0 | -0.06 | 21.5 | -0.35 | 21.45 | -0.23 | 21.39 | 0.12 | 0.06 | -16.29 | -0.06 | -41.19 | -0.06 | -46.88 | 0.06 | 21.45 | 0.18 | 21.39 | -6.62 | -0.06 | -13.3 | -0.06 | -15.53 | 0 | -40.72 | 0 | 0 | 0 | 27.95 | 0.29 | 28.65 | -0.06 | 28.36 | -0.18 | 28.48 | -0.23
48.512 | 0 | -0.18 | 21.45 | -0.35 | 21.5 | -0.29 | 21.39 | 0.47 | -0.06 | -14.47 | 0.06 | -35.45 | -0.06 | -64.92 | 0 | 21.33 | 0.12 | 21.62 | -5.39 | 0 | -11.13 | 0 | -14.24 | 0 | -36.39 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.54 | 0.06 | 28.48 | 0 | 28.59 | -0.06
48.544 | 0 | -0.06 | 21.45 | -0.29 | 21.5 | -0.23 | 21.39 | 0.35 | -0.06 | -11.37 | 0 | -31.82 | -0.18 | -36.1 | 0.06 | 21.39 | 0.23 | 21.15 | -4.45 | 0 | -9.9 | 4.39 | -12.36 | -0.06 | -37.38 | -0.18 | 0 | -0.18 | 28.89 | 0.35 | 28.71 | 0.06 | 28.59 | 0.12 | 28.36 | 0.12
48.576 | 0 | -0.06 | 21.5 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 0.06 | -9.43 | 0.06 | -32.52 | -0.06 | -27.36 | -0.06 | 21.21 | 0.18 | 21.39 | -3.69 | 0 | -8.91 | 12.72 | -10.49 | -0.06 | -37.97 | 0 | 0 | 0 | 28.48 | 0.29 | 28.42 | 0.12 | 28.42 | 0.06 | 28.59 | -0.06
48.608 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.33 | 0.41 | -0.06 | -7.44 | -0.06 | -32.4 | -0.06 | -32.76 | 0.06 | 21.39 | 0.23 | 21.33 | -2.93 | 5.63 | -7.79 | 20.22 | -9.32 | 0 | -36.86 | 0 | 0 | -0.06 | 0.29 | 0.23 | 28.48 | 0.06 | 28.42 | 0.18 | 28.42 | 0.18
48.64 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 4.57 | -5.1 | -0.06 | -30.94 | -0.06 | -35.74 | 0.06 | 21.39 | 0.23 | 21.39 | -2.29 | 13.83 | -6.45 | 27.83 | -8.03 | -0.06 | -35.22 | -0.06 | 0 | 0 | 0.06 | 0.29 | 28.48 | 0.23 | 28.48 | 0.18 | 28.13 | 0.29
48.672 | 0 | -0.06 | 3.16 | -0.18 | 21.45 | -0.23 | 21.45 | 0.41 | 12.36 | -3.11 | -0.06 | -28.95 | -0.06 | -33.69 | 0.12 | 21.39 | 0.12 | 21.39 | -1.82 | 21.33 | -4.69 | 28.48 | -6.86 | 0 | -33.17 | 0 | 0 | 0 | 0 | 0.06 | 28.54 | 0.06 | 28.48 | 0.18 | 28.65 | 0.23
48.704 | -32.64 | -0.06 | 0.06 | -0.23 | 21.45 | -0.18 | 21.33 | 0.35 | 17.87 | -1.52 | -0.06 | -25.9 | -0.06 | -29.12 | 0 | 10.25 | 0.18 | 21.15 | -1.46 | 28.65 | -4.04 | 28.42 | -5.39 | 7.85 | -30.35 | -0.06 | 0 | -0.06 | 0 | 0.12 | 27.95 | 0 | 28.48 | 0.23 | 28.42 | 0.23
48.736 | -60.94 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.35 | 21.62 | -0.94 | -0.18 | -24.26 | 0 | -27.89 | 0 | 0.12 | 0.12 | 21.39 | -1.17 | 28.65 | -3.34 | 28.36 | -4.45 | 15.06 | -28.3 | -0.06 | -34.75 | 0 | 0 | 0.06 | 28.48 | 0.06 | 28.07 | 0.23 | 28.48 | 0.18
48.768 | -43.54 | -0.06 | 0.06 | -0.06 | 21.39 | -0.18 | 21.39 | 0.41 | 21.39 | -0.59 | -0.06 | -22.33 | -0.06 | -27.01 | -0.06 | 0 | 0.23 | 21.45 | -0.94 | 28.36 | -2.75 | 28.24 | -3.69 | 22.79 | -26.54 | -0.06 | -85.61 | -0.06 | 0 | -0.06 | 28.42 | 0.06 | 28.42 | 0.23 | 28.54 | 0.18
48.8 | -40.14 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.33 | 0.29 | 21.45 | -0.06 | -0.06 | -20.39 | 0 | -25.37 | -0.18 | 0 | 0.18 | 21.33 | -0.76 | 28.65 | -2.05 | 28.65 | -2.75 | 29.01 | -24.96 | 0 | -56.08 | 0 | 0 | 0.18 | 28.48 | 0.06 | 28.18 | 0.18 | 28.24 | 0.23
48.832 | -38.56 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.33 | 0.18 | 21.39 | 0.12 | -0.06 | -17.99 | 0 | -23.91 | -0.06 | 0 | 0.23 | 21.33 | -0.76 | 28.59 | -1.64 | 28.42 | -2.11 | 28.71 | -23.2 | 0 | -48.87 | -0.06 | 0 | -0.06 | 28.3 | 0 | 28.65 | 0.29 | 28.83 | 0.18
48.864 | -35.27 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.27 | 0.29 | 21.45 | 0.23 | -0.06 | -15.82 | -0.06 | -21.68 | -0.06 | 0 | 0.23 | 21.39 | -0.53 | 28.48 | -1.23 | 28.36 | -1.7 | 28.83 | -20.92 | 0 | -41.84 | 0 | 0 | 0 | 28.48 | 0 | 28.42 | 0.23 | 28.42 | 0.12
48.896 | -34.4 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.06 | 21.39 | 0.23 | -0.06 | -13.89 | -0.06 | -19.92 | 0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.77 | -1 | 28.71 | -1.17 | 28.48 | -18.87 | 0.06 | -36.39 | -0.18 | 0 | -0.06 | 28.48 | 0 | 28.48 | 0.23 | 28.42 | 0.23
48.928 | -31.41 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | -0.06 | -11.95 | -0.06 | -18.28 | -0.06 | 0 | 0.12 | 21.39 | -0.41 | 28.3 | -0.7 | 28.42 | -0.76 | 28.36 | -17.23 | 0 | -37.03 | -0.06 | 0 | 0 | 0.23 | 0 | 28.36 | 0.12 | 28.42 | 0.23
48.96 | -30.12 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | -0.06 | -9.55 | 0 | -16.7 | -0.06 | 0 | 0.12 | 21.39 | -0.47 | 28.59 | -0.53 | 28.42 | -0.59 | 28.3 | -15.29 | 0 | -38.44 | 0 | 0 | 0.06 | 0.06 | -0.18 | 28.42 | 0.23 | 0.64 | 0.06
48.992 | -27.19 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.41 | -0.06 | -7.79 | 0 | -15.06 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.18 | -0.35 | 28.42 | -0.35 | 28.24 | -13.54 | 0 | -37.44 | -0.06 | 0 | 0 | 0 | -0.06 | 3.81 | 0.18 | 0.06 | 0.18
49.024 | -24.9 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.47 | 4.04 | -4.92 | -0.06 | -12.72 | -0.06 | 0 | 0.12 | 21.33 | -0.41 | 28.24 | -0.41 | 28.59 | -0.23 | 28.42 | -10.9 | 0 | -35.57 | 0 | -42.36 | -0.06 | 0 | -0.18 | 0.06 | 0 | 0 | 0.18
49.056 | -23.03 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | 10.14 | -2.99 | -0.06 | -10.84 | 0 | 0 | 0.06 | 21.39 | -0.35 | 28.59 | -0.41 | 28.59 | -0.35 | 28.48 | -8.85 | 0 | -33.05 | 0 | -80.92 | 0 | 0 | 0 | 0 | 0.06 | 0.06 | 0.12
49.088 | -21.09 | 0 | -2.99 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 15.59 | -1.93 | -0.06 | -9.02 | -0.06 | 0 | 0.23 | 21.39 | -0.41 | 28.48 | -0.29 | 28.48 | -0.29 | 28.65 | -6.91 | 1.99 | -31.29 | -0.18 | -54.32 | -0.18 | 0 | -0.18 | 0.06 | 0 | 0 | 0.06
49.12 | -18.87 | 0 | -61.64 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.33 | -0.94 | -0.06 | -7.38 | 0 | 0 | -0.06 | 21.39 | -0.23 | 28.54 | -0.23 | 28.48 | -0.29 | 28.65 | -4.45 | 7.38 | -30.29 | -0.06 | -48.52 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.18
49.152 | -17.05 | -0.06 | -50.63 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.45 | -0.64 | 0 | -5.63 | -0.06 | -68.26 | 0.12 | 21.39 | -0.29 | 28.65 | -0.35 | 28.54 | -0.23 | 28.65 | -2.99 | 14.82 | -28.24 | -0.06 | -43.01 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0
49.184 | -15 | -0.06 | -40.31 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.06 | 2.34 | -3.46 | 0 | -56.9 | 0.12 | 21.39 | -0.23 | 28.54 | -0.29 | 28.54 | -0.35 | 28.59 | -1.7 | 25.25 | -25.43 | 0 | -36.62 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
49.216 | -13.48 | 0 | -39.14 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | 0.18 | 8.09 | -2.23 | 0 | -34.92 | 0.23 | 21.39 | -0.23 | 28.65 | -0.18 | 28.48 | -0.23 | 28.42 | -0.76 | 28.71 | -24.08 | 0.06 | -37.68 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12
49.248 | -12.25 | -0.06 | -36.86 | -0.06 | -5.92 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.45 | 0.12 | 13.54 | -1.88 | 0.06 | -27.6 | 0 | 0.23 | -0.18 | 28.54 | -0.18 | 28.48 | -0.18 | 28.71 | -0.53 | 28.54 | -21.97 | -0.18 | -38.44 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0
49.28 | -10.55 | 0 | -35.22 | -0.06 | -63.11 | -0.06 | 0 | -0.06 | 21.21 | 0.18 | 21.21 | 0.47 | 19.04 | -0.82 | 0 | -32.93 | 0 | 0.06 | -0.18 | 28.65 | -0.18 | 28.48 | -0.18 | 28.48 | -0.18 | 28.65 | -19.86 | 0 | -37.74 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0
49.312 | -9.61 | 0 | -32.76 | -0.06 | -44.65 | -0.06 | -12.31 | -0.06 | 21.39 | 0.23 | 21.39 | 0.47 | 21.39 | -0.76 | 0 | -34.75 | 0.06 | 0 | -0.23 | 28.48 | -0.18 | 28.42 | -0.18 | 28.83 | 0.18 | 28.48 | -18.28 | -0.06 | -36.04 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0
49.344 | -7.44 | 0 | -30.18 | -0.06 | -40.37 | -0.06 | -69.44 | -0.06 | 21.45 | 0.23 | 21.39 | 0.41 | 21.21 | -0.23 | 0 | -32.76 | 0.06 | 0 | -0.18 | 28.42 | -0.23 | 28.42 | -0.06 | 28.59 | 0.06 | 28.48 | -15.94 | -0.06 | -33.63 | -0.18 | -48.22 | -0.18 | 0 | -0.06 | 0 | 0.18
49.376 | -5.98 | 0 | -28.18 | -0.06 | -39.08 | -0.06 | -47.11 | -0.06 | 21.21 | 0.12 | 21.45 | 0.47 | 21.33 | -0.18 | 0.06 | -29.94 | 0.06 | 0 | -0.23 | 28.71 | -0.18 | 28.48 | -0.06 | 28.42 | 0.53 | 28.36 | -14 | 0 | -31.11 | -0.06 | -81.51 | -0.18 | 0 | -0.06 | -17.23 | 0
49.408 | -4.63 | 0 | -25.61 | -0.06 | -35.98 | -0.06 | -42.42 | -0.06 | 21.39 | 0.06 | 21.45 | 0.53 | 21.39 | -0.06 | -0.06 | -27.77 | 0.12 | 0 | -0.18 | 28.48 | -0.06 | 28.48 | -0.06 | 28.48 | 0.47 | 28.48 | -12.25 | 0 | -29.65 | -0.06 | -49.87 | -0.06 | -2.87 | -0.18 | -89.36 | 0
49.44 | -3.34 | 0 | -23.44 | -0.06 | -34.04 | -0.06 | -36.04 | 0 | 21.39 | 0.06 | 21.39 | 0.47 | 21.39 | 0.06 | -0.06 | -27.19 | -0.06 | 0 | -0.18 | 28.18 | -0.18 | 0.29 | -0.06 | 28.48 | 0.53 | 28.48 | -10.31 | 0 | -27.54 | -0.06 | -30.47 | -0.18 | -89.01 | -0.06 | -59.83 | 0.06
49.472 | -2.23 | 3.22 | -21.56 | -0.06 | -32.46 | -0.06 | -31.99 | -0.06 | 21.39 | 0.06 | 21.39 | 0.47 | 21.39 | 0.12 | 6.09 | -24.61 | -0.06 | 0 | -0.18 | 28.77 | -0.06 | 0.06 | -0.18 | 28.3 | 0.53 | 29.01 | -7.44 | 0 | -26.13 | -0.12 | -33.52 | -0.06 | -65.92 | 0 | -30.94 | 0.06
49.504 | -0.94 | 11.31 | -19.1 | -0.06 | -29.77 | 0 | -31.99 | -0.18 | 21.39 | -0.06 | 21.39 | 0.47 | 21.27 | 0 | 13.77 | -22.38 | -0.06 | 0 | -0.18 | 0.23 | -0.18 | 0 | -0.18 | 28.65 | 0.47 | 28.48 | -4.86 | 0 | -24.08 | -0.18 | -42.78 | -0.18 | -33.58 | 0 | -32.34 | 0
49.536 | -0.53 | 16.7 | -17.05 | -0.06 | -27.95 | 0 | -32.05 | -0.06 | 0.12 | 0.06 | 21.39 | 0.35 | 21.39 | 0.18 | 19.28 | -21.62 | -0.06 | 0 | -0.23 | 0 | -0.06 | 0.06 | -0.06 | 28.3 | 0.41 | 28.42 | -3.16 | 0 | -22.09 | -0.18 | -40.37 | -0.18 | -32.29 | -0.12 | -41.66 | 0
49.568 | -0.23 | 21.86 | -15.7 | -0.18 | -25.55 | -0.06 | -30.82 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.12 | 21.39 | -19.86 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.41 | 28.65 | -1.93 | 4.39 | -19.92 | 0 | -35.74 | -0.06 | -41.13 | 0 | -41.54 | -0.06
49.6 | -0.23 | 21.56 | -14 | -0.18 | -23.91 | 0 | -29.01 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | 0.23 | 21.39 | -18.52 | 0 | 0 | -0.18 | 0.06 | -0.06 | 0 | -0.18 | 0.23 | 0.41 | 28.48 | -1.23 | 12.72 | -17.58 | 0 | -33.52 | 0 | -41.84 | -0.06 | -37.21 | 0
49.632 | -0.18 | 21.5 | -12.72 | -0.06 | -20.74 | -0.06 | -27.42 | -0.06 | 0 | 0.06 | 21.45 | 0.18 | 21.39 | 0.18 | 21.39 | -16.41 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.41 | 28.65 | -0.59 | 20.33 | -16.29 | -0.18 | -32.76 | -0.18 | -37.62 | -0.06 | -33.17 | 0
49.664 | -0.18 | 21.45 | -10.66 | -0.06 | -19.04 | -0.06 | -24.43 | 0 | 0 | -0.06 | 21.27 | 0.29 | 21.39 | 0.29 | 21.39 | -14.47 | -0.06 | -30.12 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.36 | 0 | 28.3 | -14.06 | -0.06 | -32.23 | 0 | -33.75 | 0 | -32.81 | 0
49.696 | -0.35 | 21.45 | -9.32 | -0.06 | -16.93 | -0.06 | -22.09 | -0.18 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.12 | 21.33 | -12.95 | -0.18 | -71.14 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 29.01 | 0 | 28.48 | -11.6 | -0.18 | -30.53 | -0.18 | -33.28 | -0.06 | -32.23 | -0.06
49.728 | -0.35 | 21.39 | -8.03 | -0.06 | -15.94 | -0.06 | -19.81 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.12 | 21.33 | -11.25 | 0.06 | -48.81 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.29 | 28.48 | 0.06 | 28.59 | -9.84 | -0.06 | -28.71 | -0.06 | -32.76 | -0.06 | -31.47 | -0.06
49.76 | -0.29 | 21.39 | -6.68 | -0.06 | -14.06 | -0.06 | -18.16 | 0 | 0 | 0 | 21.39 | 0.12 | 21.33 | 0.12 | 21.39 | -9.32 | 0.06 | -31.29 | -0.23 | 0 | -0.06 | 0 | -0.18 | 0 | 0.23 | 28.42 | 0.12 | 28.59 | -7.79 | -0.18 | -27.31 | -0.06 | -31.41 | 0 | -28.83 | 0
49.792 | -0.23 | 21.45 | -5.22 | -0.06 | -12.25 | 0.06 | -16.47 | 0 | 0 | -0.06 | 21.27 | 0.06 | 21.27 | 0.12 | 21.39 | -6.97 | 0 | -29.83 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.59 | 0.35 | 28.65 | -5.57 | -0.18 | -25.72 | -0.06 | -29.47 | 0 | -26.9 | -0.06
49.824 | -0.18 | 21.45 | -3.16 | -0.06 | -10.61 | 0.06 | -13.71 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.15 | 0.12 | 21.39 | -5.39 | 0.06 | -35.33 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 28.36 | 0.29 | 28.71 | -2.87 | -0.06 | -23.32 | -0.18 | -27.19 | 0 | -26.43 | -0.06
49.856 | -0.18 | 21.5 | -2.11 | -0.06 | -9.43 | -0.06 | -11.89 | -0.06 | 0 | -0.06 | 21.27 | 0 | 21.39 | 0.06 | 21.39 | -3.75 | 0.06 | -35.1 | -0.18 | 0 | -0.06 | -11.6 | 0 | 0 | 0.12 | 28.48 | 0.35 | 28.54 | -2.05 | -0.06 | -22.33 | -0.18 | -25.9 | -0.06 | -24.96 | 0.06
49.888 | -0.23 | 21.45 | -1.41 | 6.21 | -8.14 | -0.06 | -9.55 | 0.06 | 0 | 0 | 21.39 | -0.06 | 21.39 | 0.18 | 21.33 | -2.64 | -0.06 | -31.99 | -0.06 | 0 | 0 | -78.58 | -0.18 | 0 | 0.12 | 28.48 | 0.35 | 28.48 | -1.17 | -0.06 | -20.57 | -0.18 | -24.61 | -0.06 | -22.68 | 0.06
49.92 | -0.29 | 21.39 | -0.64 | 12.13 | -6.45 | 5.39 | -7.62 | 0 | 0 | -0.06 | 0.23 | -0.06 | 21.21 | 0.06 | 21.39 | -1.41 | -0.18 | -28.59 | -0.18 | -29.83 | 0.06 | -55.96 | 0 | 0 | 0 | 28.48 | 0.23 | 28.71 | -0.53 | -0.18 | -18.46 | -0.18 | -22.91 | -0.06 | -20.86 | 0
49.952 | -0.23 | 21.45 | -0.29 | 17.64 | -4.45 | 11.48 | -5.39 | -0.06 | -50.45 | -0.06 | 0 | -0.18 | 21.68 | 0.06 | 21.39 | -0.7 | 0 | -27.6 | -0.06 | -75.65 | -0.06 | -45.82 | -0.06 | 0 | 0 | 28.36 | 0.29 | 28.48 | -0.23 | -0.06 | -17.4 | -0.18 | -21.33 | -0.06 | -19.45 | 0
49.984 | -0.23 | 21.39 | -0.18 | 21.5 | -3.34 | 18.75 | -2.75 | -0.18 | -66.21 | 0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | -0.59 | 0.06 | -26.37 | -0.18 | -47.58 | -0.06 | -45.35 | 0 | 0 | 0.12 | 23.67 | 0.29 | 28.65 | -0.06 | 0 | -15.53 | -0.18 | -19.4 | -0.18 | -18.16 | 0
50.016 | -0.29 | 21.45 | -0.29 | 21.45 | -2.11 | 21.5 | -2.05 | 0.06 | -46.7 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.33 | -0.29 | 3.69 | -24.38 | -0.06 | -45.18 | -0.06 | -41.66 | 0 | -38.03 | -0.06 | 0.18 | 0.29 | 28.65 | 0.12 | 2.64 | -13.54 | -0.06 | -17.87 | -0.06 | -16.11 | -0.06
50.048 | -0.18 | 21.45 | -0.35 | 21.45 | -1.29 | 21.45 | -0.76 | -0.06 | -40.55 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.27 | -0.23 | 9.84 | -22.33 | -0.18 | -44.01 | 0 | -40.67 | -0.06 | -72.13 | -0.06 | 0 | 0.23 | 28.48 | 0.23 | 10.84 | -12.07 | -0.06 | -16.17 | 0.06 | -14.41 | 0
50.08 | -0.18 | 21.45 | -0.47 | 21.5 | -0.64 | 21.5 | -0.64 | 0.06 | -35.27 | -0.06 | 0 | 0.06 | 17.58 | -0.06 | 21.33 | -0.06 | 15.29 | -21.09 | -0.18 | -41.49 | 0 | -39.08 | -0.06 | -48.81 | 0 | 0 | 0.23 | 28.59 | 0.35 | 18.28 | -10.49 | 9.43 | -14.41 | 0.35 | -12.72 | 0
50.112 | -0.18 | 21.39 | -0.29 | 21.45 | -0.35 | 21.45 | -0.18 | -0.06 | -31.99 | 0 | 0 | -0.18 | 0.12 | 0.06 | 21.33 | 0 | 20.98 | -19.22 | -0.18 | -40.55 | 0 | -35.74 | -0.06 | -46.7 | -0.06 | 0 | 0.18 | 28.83 | 0.29 | 28.48 | -8.56 | 16.7 | -12.95 | 9.73 | -10.49 | 2.64
50.144 | -0.18 | 21.45 | -0.18 | 21.5 | -0.23 | 21.39 | 0 | 0 | -32.52 | -0.18 | 0 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -17.52 | -0.18 | -38.32 | -0.18 | -34.63 | 0 | -44.24 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.65 | -6.56 | 27.01 | -10.14 | 19.81 | -9.02 | 13.13
50.176 | -0.18 | 21.45 | -0.35 | 21.5 | -0.35 | 21.45 | 0.18 | -0.06 | -32.05 | -0.06 | 0 | 0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.45 | -15.47 | -0.18 | -35.27 | 0 | -33.05 | 0 | -41.54 | 0 | 0 | 0.06 | 28.54 | 0.35 | 28.65 | -4.63 | 28.48 | -9.08 | 27.31 | -7.79 | 20.63
50.208 | -0.18 | 21.45 | -0.41 | 21.45 | -0.29 | 21.45 | 0.29 | -0.06 | -30.7 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.33 | -13.54 | -0.06 | -34.63 | -0.06 | -30.7 | -0.06 | -40.78 | 0 | 0 | 0.06 | 28.48 | 0.41 | 28.59 | -3.46 | 28.83 | -7.38 | 28.54 | -6.21 | 28.59
50.24 | -0.18 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | 0.41 | -0.06 | -28.48 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | -12.31 | 0 | -32.4 | 0.06 | -29.01 | 0 | -37.97 | -0.06 | 0 | 0.06 | 28.59 | 0.23 | 28.71 | -2.23 | 28.48 | -5.92 | 28.48 | -4.28 | 28.54
50.272 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.41 | -0.06 | -25.9 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | -9.84 | -0.18 | -30.12 | -0.06 | -26.13 | -0.06 | -36.1 | 0 | 0 | 0.12 | 28.3 | 0.29 | 28.48 | -1.46 | 28.42 | -4.39 | 28.48 | -2.58 | 28.48
50.304 | -0.18 | 21.39 | -0.18 | 21.45 | -0.29 | 21.45 | 0.35 | 7.62 | -23.67 | -0.06 | 0 | 0 | 0 | 0 | 21.33 | 0.18 | 21.21 | -8.09 | -0.06 | -27.89 | 0.06 | -24.38 | 0 | -34.63 | 0 | 0 | 0.06 | 28.59 | 0.12 | 28.54 | -0.64 | 28.48 | -2.52 | 28.65 | -1.7 | 28.42
50.336 | -0.18 | 21.45 | -0.23 | 21.5 | -0.41 | 21.45 | 0.35 | 13.18 | -21.86 | -0.06 | -17.99 | 0 | 0 | -0.06 | 15.59 | 0.29 | 21.33 | -6.86 | -0.06 | -25.9 | 0 | -22.68 | 0.06 | -32.46 | 0.06 | 0 | 0 | 28.54 | 0.12 | 28.42 | -0.53 | 28.65 | -1.88 | 28.42 | -1.23 | 28.36
50.368 | -0.06 | 0.12 | -0.18 | 21.45 | -0.29 | 21.45 | 0.47 | 18.69 | -20.16 | -0.06 | -77.05 | -0.06 | 0 | -0.18 | 0.12 | 0.23 | 21.39 | -4.86 | 6.04 | -24.08 | 0 | -20.86 | 0 | -30.7 | 0 | 0 | 0 | 28.42 | 0.12 | 28.59 | -0.18 | 28.48 | -1.29 | 28.42 | -0.76 | 28.42
50.4 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.35 | 21.56 | -18.22 | -0.06 | -52.38 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -3.46 | 14 | -22.09 | 0 | -18.87 | 0 | -28.13 | 0.06 | 0 | 0 | 28.65 | 0.06 | 28.54 | -0.06 | 28.48 | -0.59 | 28.83 | -0.23 | 28.42
50.432 | 0 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.35 | 21.45 | -15.53 | -0.06 | -41.66 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | -2.4 | 21.74 | -20.22 | 0 | -16.23 | 0 | -26.25 | 0 | -53.73 | 0 | 28.77 | 0.12 | 28.42 | -0.18 | 28.48 | -0.53 | 28.42 | -0.23 | 28.59
50.464 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.41 | 21.39 | -13.59 | 0 | -35.33 | 0.06 | 0 | -0.18 | 0 | 0.18 | 21.21 | -1.29 | 28.59 | -17.4 | 0 | -14.77 | 0 | -24.32 | 0 | -68.5 | -0.06 | 0.29 | 0 | 28.36 | 0.23 | 28.36 | 0 | 28.65 | -0.18 | 28.65
50.496 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.29 | 21.39 | -11.31 | 0 | -32.17 | 0 | -0.06 | -0.06 | 0 | 0.18 | 21.21 | -0.76 | 28.54 | -16.06 | 0.12 | -13.54 | 0 | -22.38 | 0 | -49.87 | -0.06 | 0.12 | 0.06 | 28.59 | 0.06 | 28.42 | -0.06 | 28.18 | 0 | 28.65
50.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | -9.84 | -0.06 | -32.7 | -0.06 | -61.53 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.59 | -14.36 | 3.34 | -12.07 | 0 | -20.22 | -0.06 | -46.23 | -0.06 | 0 | 0.06 | 28.42 | 0 | 28.65 | 0 | 28.36 | 0.12 | 28.59
50.56 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.29 | 21.45 | -7.03 | 0 | -32.46 | -0.06 | -61.41 | -0.06 | 0 | 0.23 | 21.39 | -0.29 | 28.65 | -13.01 | 11.72 | -10.55 | 0 | -17.99 | -0.06 | -41.31 | 0 | 0.06 | 0.06 | 28.48 | 0.18 | 28.42 | 0 | 28.65 | 0.12 | 28.48
50.592 | -0.18 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | 0.23 | 21.39 | -4.22 | -0.18 | -30.65 | -0.06 | -36.56 | -0.06 | 0 | 0.12 | 21.15 | -0.23 | 28.42 | -12.36 | 21.91 | -8.91 | 8.85 | -16.23 | 0 | -36.97 | -0.06 | 0 | 0.06 | 28.65 | 0.12 | 28.42 | 0.18 | 28.65 | 0 | 28.59
50.624 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.12 | 21.21 | -2.64 | 0.06 | -28.89 | 0.06 | -28.18 | 0 | 0 | 0.23 | 21.5 | -0.06 | 28.59 | -10.31 | 29.24 | -7.44 | 18.63 | -14.77 | 0 | -37.27 | -0.06 | 0 | 0 | 28.42 | 0.23 | 28.3 | 0.12 | 28.48 | 0.23 | 28.3
50.656 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | 0 | 21.27 | -1.35 | 0 | -26.72 | 0 | -34.16 | 0 | 0 | 0.23 | 21.15 | 0.12 | 28.48 | -8.91 | 28.59 | -6.8 | 26.37 | -13.13 | -0.06 | -37.68 | 0 | 0 | 0 | 28.48 | 0.18 | 28.59 | 0.12 | 28.42 | 0.29 | 28.65
50.688 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -1.05 | 3.98 | -24.61 | -0.06 | -35.86 | 0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.42 | -7.5 | 28.65 | -5.33 | 28.65 | -11.78 | -0.06 | -36.68 | 0 | 0 | 0 | 28.42 | 0.06 | 28.42 | 0.18 | 28.42 | 0.18 | 28.42
50.72 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0 | 21.39 | -0.41 | 10.2 | -22.62 | -0.06 | -33.17 | 0.06 | 0 | 0.18 | 21.21 | 0.12 | 28.42 | -6.45 | 28.59 | -4.45 | 28.59 | -10.14 | -0.18 | -34.92 | 0 | 0 | -0.06 | 28.54 | 0.18 | 28.48 | 0.29 | 28.65 | 0.12 | 28.65
50.752 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | 0.12 | 21.39 | -0.06 | 15.41 | -19.75 | -0.06 | -29.12 | 0 | 0 | 0.18 | 21.39 | 0.12 | 28.42 | -5.16 | 28.59 | -3.52 | 28.59 | -8.61 | -0.06 | -32.76 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | 0.18 | 28.48 | 0.18 | 28.42
50.784 | -0.06 | -26.9 | -0.18 | 0.12 | -0.18 | 19.92 | 0 | 21.39 | 0.06 | 21.62 | -18.16 | 0 | -27.6 | -0.06 | -66.51 | 0.12 | 21.39 | 0.23 | 28.59 | -3.98 | 28.59 | -2.7 | 28.77 | -6.97 | -0.06 | -30.7 | 0 | 0 | -0.06 | 28.18 | 0 | 28.36 | 0.18 | 28.42 | 0.12 | 28.42
50.816 | 0 | -61.88 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.39 | 0.29 | 21.39 | -16.29 | -0.06 | -27.07 | -0.06 | -59.12 | 0.23 | 21.39 | 0.29 | 28.59 | -3.11 | 28.42 | -1.99 | 28.36 | -6.15 | -0.06 | -28.95 | 0.06 | 0 | 0 | 28.3 | 0 | 28.42 | 0.12 | 28.42 | 0.12 | 28.42
50.848 | -0.06 | -45.76 | -0.18 | 0 | -0.06 | 0 | 0 | 21.27 | 0.29 | 21.39 | -14.47 | -0.06 | -25.55 | 0 | -35.16 | 0.12 | 21.33 | 0.29 | 28.83 | -2.64 | 28.48 | -1.52 | 29.06 | -5.04 | -0.06 | -26.84 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.48 | 0.29 | 28.48 | 0.23 | 28.42
50.88 | -0.06 | -40.14 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -12.72 | -0.06 | -23.03 | -0.06 | -27.01 | 0.06 | 21.39 | 0.12 | 28.3 | -2.05 | 28.36 | -1.11 | 28.65 | -4.16 | -0.06 | -25.72 | 0 | -25.96 | -0.06 | 28.48 | -0.06 | 28.65 | 0.12 | 28.48 | 0.18 | 28.42
50.912 | -0.06 | -38.56 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -10.02 | 0 | -21.39 | -0.06 | -33.17 | 0.18 | 0.29 | 0.23 | 28.3 | -1.7 | 28.48 | -0.82 | 28.65 | -3.28 | -0.06 | -23.44 | 0 | -92.35 | -0.18 | 0.29 | 0 | 28.48 | 0.06 | 28.24 | 0.12 | 28.48
50.944 | -0.06 | -35.74 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.27 | -7.62 | 0.12 | -19.69 | -0.06 | -35.86 | 0.12 | 0 | 0.23 | 28.54 | -1.52 | 28.42 | -0.53 | 28.59 | -2.29 | 0 | -20.98 | -0.18 | -58.24 | 0 | 0.12 | 0 | 3.16 | 0.06 | 28.42 | 0.12 | 28.48
50.976 | -0.06 | -33.75 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.21 | -5.51 | 0 | -18.05 | -0.18 | -32.76 | 0.12 | 0 | 0.29 | 28.54 | -1.05 | 28.65 | -0.35 | 29.06 | -1.88 | -0.06 | -19.45 | 0 | -48.99 | -0.06 | 0 | -0.06 | 0.12 | 0 | 0.88 | 0.12 | 28.42
51.008 | -0.06 | -31.17 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.56 | 0.47 | 21.21 | -3.16 | 0.06 | -16.17 | -0.06 | -29.59 | 0.06 | 0 | 0.23 | 28.42 | -0.76 | 28.65 | -0.23 | 28.3 | -1.29 | 0 | -17.4 | -0.06 | -43.24 | -0.18 | 0.06 | 0 | 0 | 0.12 | 0.12 | 0 | 0.29
51.04 | -0.06 | -30.12 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.99 | 0 | -14.41 | -0.06 | -28.07 | 0.06 | 0 | 0.23 | 28.36 | -0.53 | 28.42 | -0.29 | 28.54 | -0.94 | 8.26 | -15.35 | -0.06 | -37.38 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 0.12
51.072 | -0.18 | -28.18 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -0.76 | 3.4 | -12.6 | 0 | -27.31 | -0.06 | 0 | 0.23 | 28.54 | -0.41 | 28.54 | -0.23 | 28.83 | -0.7 | 15.65 | -13.36 | -0.06 | -37.21 | 0 | 0 | -0.18 | 0 | 0.06 | 0 | 0 | 0
51.104 | 0 | -25.61 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -0.64 | 11.31 | -10.25 | -0.06 | -25.02 | 0.06 | 0 | 0.18 | 28.3 | -0.29 | 28.36 | -0.41 | 28.42 | -0.41 | 26.13 | -10.72 | -0.06 | -37.85 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0.06
51.136 | 0 | -23.32 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | 0 | 16.7 | -8.56 | 0 | -23.09 | 0.06 | 0 | 0.23 | 28.54 | -0.23 | 28.24 | -0.23 | 28.36 | -0.29 | 28.59 | -8.91 | -0.06 | -37.56 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0
51.168 | -0.06 | -20.98 | -0.06 | 0 | -0.06 | 0 | 0 | 19.63 | 0.29 | 21.39 | 0.12 | 21.68 | -6.56 | 5.74 | -21.56 | 0 | 0 | 0.23 | 28.65 | -0.35 | 28.59 | -0.23 | 28.59 | -0.23 | 28.59 | -6.27 | -0.06 | -35.39 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0
51.2 | -0.06 | -19.34 | -0.06 | -34.63 | 0 | 0 | -0.06 | 0.12 | 0.35 | 21.39 | 0.18 | 21.27 | -5.04 | 11.54 | -19.63 | -0.06 | 0 | 0.18 | 28.36 | -0.35 | 28.54 | -0.29 | 28.54 | -0.29 | 28.42 | -4.63 | -0.06 | -33.58 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
51.232 | 0 | -17.7 | -0.18 | -59.18 | -0.06 | -28.36 | -0.06 | 0 | 0.18 | 21.39 | 0.29 | 21.33 | -3.75 | 17.17 | -18.4 | 0 | 0 | 0.12 | 28.54 | -0.23 | 28.48 | -0.29 | 28.65 | -0.29 | 28.59 | -2.75 | -0.06 | -31.7 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0
51.264 | -0.06 | -15.41 | -0.06 | -40.43 | -0.06 | -55.55 | -0.06 | 0 | 0.23 | 21.39 | 0.41 | 21.39 | -2.11 | 21.45 | -16 | 0 | 0 | 0.18 | 0.18 | -0.23 | 28.65 | -0.18 | 28.36 | -0.35 | 28.3 | -1.7 | -0.06 | -30 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0
51.296 | -0.18 | -13.71 | -0.18 | -39.32 | -0.18 | -40.78 | 0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.39 | -1.52 | 21.39 | -14.41 | -0.06 | 0 | 0.18 | 0 | -0.29 | 28.36 | -0.29 | 28.59 | -0.18 | 28.89 | -0.88 | -0.06 | -27.54 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | -0.06 | 0
51.328 | 0.06 | -12.07 | -0.06 | -38.09 | -0.06 | -39.32 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.45 | -1.05 | 21.39 | -12.77 | -0.06 | -10.08 | 0.12 | 0 | -0.41 | 28.54 | -0.06 | 28.77 | -0.29 | 28.42 | -0.29 | 0 | -25.9 | -0.06 | -0.12 | -0.23 | 0 | -0.06 | 0 | -0.06 | 0
51.36 | 0 | -10.31 | -0.06 | -35.27 | -0.06 | -37.97 | -0.18 | 0 | 0.12 | 21.39 | 0.41 | 21.33 | -0.47 | 21.33 | -10.9 | 0.06 | -73.89 | 0.18 | 0 | -0.23 | 28.42 | -0.18 | 28.42 | -0.23 | 28.71 | -0.29 | 0 | -24.14 | -0.06 | -89.48 | -0.18 | 0 | 0 | 0 | -0.06 | 0
51.392 | 3.22 | -9.61 | -0.06 | -33.34 | -0.06 | -35.33 | -0.18 | 0 | 0 | 21.56 | 0.53 | 21.39 | -0.35 | 21.15 | -9.38 | -0.06 | -53.5 | 0.18 | 0 | -0.23 | 28.42 | -0.18 | 28.71 | -0.23 | 28.36 | 0.06 | 0 | -22.38 | -0.18 | -56.49 | -0.06 | -67.44 | -0.06 | 0 | -0.06 | 0
51.424 | 11.25 | -7.44 | -0.06 | -31.23 | -0.06 | -33.34 | -0.06 | 0 | 0.12 | 21.33 | 0.41 | 21.39 | -0.06 | 21.33 | -7.03 | -0.18 | -29.88 | 0.18 | 0 | -0.23 | 0.23 | -0.18 | 28.13 | -0.18 | 28.65 | 0.18 | 4.45 | -20.1 | -0.18 | -48.69 | -0.06 | -63.58 | -0.06 | -89.48 | 0 | -25.67
51.456 | 16.64 | -6.09 | -0.06 | -29.01 | -0.06 | -30.94 | 0.06 | 0 | 0 | 21.33 | 0.35 | 21.39 | -0.06 | 21.39 | -5.1 | -0.06 | -30.53 | 0.06 | 0 | -0.18 | 0.06 | -0.06 | 14.77 | -0.18 | 28.48 | 0.23 | 12.72 | -18.34 | -0.06 | -43.89 | -0.06 | -37.79 | -0.06 | -62.23 | -0.06 | -88.6
51.488 | 21.74 | -4.51 | 0 | -27.01 | -0.06 | -29.36 | -0.06 | 0 | -0.18 | 21.27 | 0.35 | 21.33 | 0 | 21.39 | -3.87 | -0.06 | -35.39 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0.12 | -0.18 | 28.36 | 0.23 | 20.27 | -16.29 | -0.06 | -39.08 | -0.06 | -29.12 | 0 | -35.51 | 0 | -59.77
51.52 | 21.5 | -3.16 | -0.06 | -24.9 | -0.06 | -27.19 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.21 | 0.18 | 21.33 | -2.64 | 0.06 | -34.81 | 0.23 | 0 | -0.23 | 0.06 | -0.18 | 0 | -0.18 | 28.42 | 0.29 | 28.13 | -14 | -0.18 | -36.68 | -0.06 | -38.97 | -0.06 | -29.24 | 0 | -33.75
51.552 | 21.5 | -2.23 | -0.06 | -22.38 | -0.06 | -25.2 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | 0.23 | 21.39 | -1.76 | -0.06 | -32.81 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0.35 | 28.59 | -11.89 | -0.18 | -37.56 | -0.18 | -42.54 | -0.06 | -39.55 | 0 | -30
51.584 | 21.5 | -1.11 | -0.06 | -20.39 | -0.06 | -22.79 | -0.18 | 0.06 | -0.06 | 0.12 | 0.23 | 21.33 | 0.23 | 21.39 | -1.05 | 0 | -29.24 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.35 | 28.36 | -10.55 | 0 | -37.85 | -0.06 | -37.21 | 0 | -41.78 | 0 | -41.25
51.616 | 21.45 | -0.64 | 0 | -18.16 | 0 | -20.51 | -0.06 | -61.29 | -0.06 | 0 | 0.18 | 21.39 | 0.23 | 21.39 | -0.41 | 0 | -26.84 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.3 | -8.44 | -0.06 | -35.98 | -0.06 | -34.98 | 0 | -37.97 | 0 | -41.08
51.648 | 21.45 | -0.23 | -0.06 | -16.58 | 0 | -18.52 | -0.18 | -61.94 | 0 | 0 | 0.23 | 21.39 | 0.18 | 21.27 | -0.29 | -0.06 | -25.9 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.3 | 0.23 | 28.48 | -5.39 | 0 | -33.99 | -0.06 | -33.28 | -0.18 | -34.28 | -0.06 | -37.56
51.68 | 21.45 | -0.23 | 0 | -15.18 | 0 | -16.82 | -0.06 | -44.65 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -0.06 | -0.06 | -25.08 | 0.12 | -30.53 | -0.18 | 0 | 0 | 0 | -0.18 | 28.36 | 0.35 | 28.65 | -3.57 | -0.06 | -31.76 | -0.06 | -32.99 | 0 | -33.52 | -0.06 | -33.75
51.712 | 21.45 | -0.29 | -0.06 | -13.48 | -0.06 | -15.59 | -0.18 | -39.9 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 21.27 | -0.06 | -0.06 | -22.97 | 0.12 | -72.48 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 28.54 | -1.82 | 3.4 | -29.53 | -0.18 | -32.23 | 0 | -32.58 | -0.18 | -32.81
51.744 | 21.45 | -0.18 | -0.06 | -11.54 | -0.06 | -13.3 | -0.06 | -33.17 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.39 | 0.18 | 0.06 | -20.39 | 0.06 | -46.23 | 0 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.23 | 28.48 | -1.76 | 14.18 | -28.24 | -0.18 | -29.24 | -0.06 | -31.23 | -0.06 | -32.05
51.776 | 21.45 | -0.29 | 0.59 | -10.25 | -0.06 | -11.66 | -0.18 | -32.23 | -0.06 | 0 | -0.06 | 21.27 | 0.18 | 21.39 | 0.29 | 0 | -19.4 | 0.12 | -44.53 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.65 | -0.23 | 21.97 | -26.13 | -0.06 | -27.6 | 0 | -28.83 | -0.06 | -30.82
51.808 | 21.45 | -0.35 | 8.26 | -8.67 | 0 | -10.66 | -0.06 | -32.46 | -0.18 | 0 | 0.06 | 21.39 | 0.18 | 21.27 | 0.35 | 1.99 | -17.64 | 0.12 | -43.65 | 0 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.42 | -0.41 | 28.89 | -24.67 | -0.18 | -26.43 | 0 | -27.72 | -0.06 | -29.3
51.84 | 21.39 | -0.35 | 14 | -7.44 | 0 | -8.85 | -0.06 | -31.99 | 0.06 | 0 | 0 | 21.33 | 0.18 | 21.27 | 0.29 | 5.8 | -15.65 | 0.18 | -41.25 | -0.18 | -5.39 | -0.18 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.42 | -23.2 | 0 | -25.37 | -0.06 | -26.13 | 0 | -27.42
51.872 | 21.45 | -0.23 | 19.63 | -5.98 | -0.06 | -7.44 | -0.06 | -30.53 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.33 | 0.29 | 13.59 | -14.24 | 0.18 | -40.37 | 0 | -73.42 | -0.18 | 0.06 | -0.06 | 28.59 | 0.18 | 28.42 | 0.18 | 28.65 | -20.1 | 0 | -23.2 | 0 | -24.84 | 0 | -26.43
51.904 | 21.45 | -0.18 | 21.56 | -4.04 | 0 | -5.74 | -0.06 | -27.83 | -0.18 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | 0.23 | 19.1 | -11.84 | 0.12 | -37.38 | -0.06 | -52.03 | -0.06 | -72.83 | -0.18 | 18.52 | 0 | 28.48 | 0.29 | 28.48 | -18.46 | -0.06 | -21.62 | -0.06 | -22.38 | 0 | -24.55
51.936 | 21.5 | -0.23 | 21.45 | -2.7 | 3.4 | -4.04 | -0.06 | -25.31 | -0.06 | 0 | -0.06 | 21.21 | 0.06 | 21.39 | 0.29 | 21.45 | -10.02 | 0.12 | -35.51 | -0.06 | -45.59 | 0 | -58.77 | -0.18 | 0.12 | 0 | 28.48 | 0.41 | 28.48 | -16.29 | -0.06 | -19.92 | -0.06 | -21.09 | -0.06 | -22.44
51.968 | 21.45 | -0.23 | 21.45 | -1.82 | 9.67 | -2.81 | 0.06 | -23.67 | -0.06 | 0 | -0.18 | 0.12 | 0.23 | 21.21 | 0.29 | 21.39 | -8.56 | 0.12 | -33.81 | -0.06 | -44.59 | -0.06 | -44.42 | -0.06 | 0 | 0 | 28.71 | 0.35 | 28.24 | -14.94 | -0.06 | -18.63 | 0 | -19.86 | -0.06 | -21.39
52 | 21.45 | -0.23 | 21.39 | -1 | 15 | -1.88 | 3.34 | -21.56 | -0.06 | -23.32 | 0 | 0.06 | 0.06 | 21.27 | 0.29 | 21.39 | -6.5 | 0.12 | -31.82 | -0.06 | -42.54 | -0.06 | -46 | -0.06 | 0 | 0.06 | 28.36 | 0.29 | 28.71 | -12.77 | -0.06 | -16.93 | -0.06 | -18.34 | 0 | -19.69
52.032 | 21.45 | -0.23 | 21.45 | -0.53 | 21.91 | -1 | 9.55 | -19.63 | -0.18 | -77.23 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -5.33 | 0.06 | -30.29 | -0.06 | -40.2 | -0.06 | -43.36 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.36 | -10.31 | 3.34 | -15 | -0.06 | -16.7 | 0 | -18.16
52.064 | 21.39 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 16.7 | -17.17 | -0.06 | -47.52 | -0.06 | 0 | 0.06 | 0.12 | 0.18 | 21.39 | -3.52 | 0.12 | -27.48 | -0.18 | -38.79 | -0.06 | -41.25 | 0 | 0 | -0.06 | 28.65 | 0.41 | 28.42 | -8.14 | 14.06 | -12.48 | 0 | -14.41 | 2.99 | -15.76
52.096 | 21.39 | -0.29 | 21.39 | -0.23 | 21.45 | -0.06 | 21.8 | -15.41 | -0.18 | -41.02 | -0.18 | 0 | -0.06 | 0 | 0.29 | 21.39 | -2.23 | 0.12 | -25.72 | -0.06 | -36.68 | -0.06 | -39.55 | 0 | 0 | 0.06 | 28.65 | 0.06 | 28.36 | -5.92 | 21.56 | -11.31 | 2.64 | -12.83 | 11.25 | -14.36
52.128 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | -0.06 | 21.39 | -13.71 | -0.06 | -35.51 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -1.17 | -0.06 | -23.61 | -0.06 | -35.45 | 0 | -37.03 | -0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.36 | -4.39 | 29.24 | -9.61 | 10.72 | -11.02 | 18.69 | -12.89
52.16 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.39 | -11.37 | -0.06 | -32.17 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.45 | -1 | 0 | -21.86 | -0.06 | -32.7 | -0.06 | -35.45 | -0.06 | 0 | 0 | 28.42 | 0.18 | 28.3 | -2.4 | 28.54 | -8.03 | 18.16 | -9.67 | 26.37 | -11.37
52.192 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.41 | 21.39 | -9.2 | -0.18 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | -0.06 | -19.98 | -0.06 | -30.47 | -0.18 | -33.87 | -0.06 | 0 | -0.06 | 28.54 | 0.12 | 28.59 | -1.41 | 28.54 | -6.21 | 25.9 | -7.79 | 28.65 | -9.55
52.224 | 21.45 | -0.06 | 21.45 | -0.35 | 21.39 | -0.41 | 21.39 | -6.21 | -0.06 | -32.58 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -0.29 | 0 | -17.81 | -0.18 | -28.59 | 0 | -31.23 | -0.06 | 0 | 0 | 28.54 | 0.12 | 28.36 | -0.82 | 28.54 | -4.34 | 28.65 | -5.68 | 29.01 | -7.79
52.256 | 21.39 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -4.04 | -0.06 | -30.94 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.33 | -0.06 | 0.06 | -16.06 | -0.06 | -26.31 | -0.06 | -29.24 | -0.06 | 0 | -0.06 | 28.54 | 0.18 | 28.59 | -0.35 | 28.48 | -3.11 | 28.65 | -4.39 | 28.59 | -5.68
52.288 | 0.18 | -0.18 | 21.39 | -0.35 | 21.45 | -0.29 | 21.39 | -2.4 | 0 | -29.18 | 0 | 0 | -0.06 | 0 | 0.23 | 21.33 | 0.12 | 0 | -14.41 | 0 | -24.79 | 0.06 | -27.36 | -0.06 | 0 | 0.06 | 28.59 | 0.12 | 28.65 | -0.23 | 28.83 | -2.23 | 28.42 | -2.75 | 28.48 | -4.34
52.32 | 0.06 | -0.18 | 21.5 | -0.35 | 21.5 | -0.35 | 21.39 | -1.88 | -0.06 | -26.72 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.06 | -0.06 | -13.18 | -0.06 | -23.32 | -0.18 | -25.55 | -0.06 | 0 | -0.06 | 0.29 | 0.06 | 28.42 | 0.06 | 28.36 | -1.41 | 28.42 | -2.05 | 28.48 | -2.64
52.352 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | -0.82 | -0.06 | -24.43 | 0 | 0 | 0 | 0 | 0 | 21.39 | 0 | 3.93 | -11.48 | 0 | -19.81 | 0 | -23.85 | -0.06 | -54.61 | -0.18 | 0.12 | 0 | 28.71 | 0.18 | 28.54 | -0.64 | 28.48 | -1.29 | 28.65 | -2.17
52.384 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -0.35 | -0.06 | -22.09 | -0.06 | -36.74 | -0.18 | 0 | 0 | 21.27 | 0.23 | 10.2 | -9.9 | -0.06 | -17.99 | 0 | -20.74 | 0 | -68.91 | -0.06 | 0 | 0 | 28.59 | 0.12 | 28.65 | -0.53 | 28.65 | -0.64 | 28.42 | -1.17
52.416 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.39 | -0.06 | -0.06 | -20.16 | -0.06 | -70.73 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 15.59 | -8.67 | 0 | -16.47 | 0 | -18.87 | 0 | -48.87 | -0.18 | 0.06 | 0.06 | 28.48 | 0.18 | 28.36 | 0 | 28.42 | -0.59 | 28.42 | -0.47
52.448 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | 0.12 | -0.06 | -18.63 | 0 | -48.22 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | -7.56 | 0 | -15.23 | -0.06 | -17.81 | -0.06 | -46.99 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.65 | -0.18 | 28.54 | -0.35 | 28.65 | -0.41
52.48 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | 0.18 | 0 | -16.29 | 0 | -41.9 | -0.06 | -38.67 | 0.06 | 21.27 | 0.35 | 21.45 | -6.56 | 0 | -13.71 | 0 | -15.76 | -0.06 | -41.25 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.3 | 0 | 28.48 | 0 | 28.48 | -0.18
52.512 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.29 | 21.39 | 0.29 | -0.06 | -14.36 | -0.06 | -36.27 | -0.06 | -68.62 | 0.18 | 21.27 | 0.29 | 21.39 | -5.33 | 0 | -11.66 | -0.18 | -14.24 | 0 | -36.68 | 0 | 0 | 0 | 28.48 | 0.35 | 28.48 | 0.06 | 28.54 | -0.06 | 28.36 | 0
52.544 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.41 | -0.06 | -11.54 | -0.06 | -31.93 | -0.06 | -38.67 | 0.06 | 21.39 | 0.29 | 21.39 | -4.1 | 0 | -9.9 | 3.4 | -11.95 | -0.06 | -37.85 | 0 | 0 | -0.06 | 28.42 | 0.29 | 28.48 | 0.06 | 28.48 | 0.06 | 28.42 | -0.06
52.576 | 0 | -0.06 | 21.45 | -0.23 | 21.5 | -0.23 | 21.33 | 0.47 | 0 | -9.49 | 0.06 | -32.46 | -0.18 | -27.01 | -0.06 | 21.21 | 0.29 | 21.21 | -3.57 | 0 | -8.67 | 11.72 | -10.96 | 0 | -37.85 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.48 | 0.12 | 28.3 | 0 | 28.59 | 0.12
52.608 | 0 | -0.06 | 21.45 | -0.18 | 21.5 | -0.29 | 21.39 | 0.29 | -0.18 | -7.56 | -0.06 | -32.76 | -0.18 | -31.47 | 0 | 21.45 | 0.29 | 21.27 | -2.87 | 4.34 | -7.44 | 19.22 | -9.55 | -0.06 | -36.62 | 0 | 0 | 0 | 0.29 | 0.12 | 28.59 | 0.18 | 28.42 | 0.18 | 28.48 | 0.12
52.64 | 0 | 0 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 3.75 | -5.33 | -0.06 | -31.11 | 0 | -35.51 | 0.06 | 21.21 | 0.23 | 21.21 | -2.4 | 12.66 | -6.39 | 26.78 | -8.09 | 0 | -35.22 | 0 | 0 | -0.18 | 0.06 | 0.18 | 28.83 | 0.23 | 28.54 | 0.23 | 28.59 | 0.12
52.672 | 0 | -0.06 | 14.12 | -0.23 | 21.45 | -0.18 | 21.39 | 0.53 | 11.66 | -3.34 | 0 | -29.06 | -0.06 | -33.99 | 0.06 | 21.21 | 0.23 | 21.39 | -1.7 | 20.22 | -5.1 | 28.54 | -7.32 | 0 | -32.76 | -0.06 | 0 | -0.18 | 0 | 0.18 | 28.48 | 0.18 | 28.48 | 0.18 | 28.3 | 0.06
52.704 | -29.83 | -0.06 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.53 | 17.11 | -1.7 | 0.06 | -26.13 | -0.06 | -30.41 | 0 | 17.46 | 0.23 | 21.33 | -1.23 | 29.06 | -4.22 | 28.3 | -5.16 | 6.27 | -30.47 | 0 | 0 | -0.18 | 0.06 | 0.12 | 28.42 | 0.12 | 28.42 | 0.12 | 28.42 | 0.06
52.736 | -62.23 | -0.06 | 0 | -0.23 | 21.45 | -0.18 | 21.39 | 0.47 | 21.68 | -0.94 | -0.06 | -24.08 | -0.06 | -27.6 | -0.06 | 0.12 | 0.18 | 21.39 | -1 | 28.77 | -3.46 | 28.36 | -4.75 | 14.12 | -28.24 | -0.06 | -20.63 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.59 | 0.18 | 28.71 | 0.12
52.768 | -45.18 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.33 | 0.29 | 21.39 | -0.53 | 0 | -22.21 | -0.06 | -27.19 | 0 | 0 | 0.18 | 21.45 | -0.82 | 28.48 | -2.99 | 28.65 | -3.69 | 21.91 | -26.66 | 0 | -90.59 | -0.18 | 0 | 0.06 | 28.36 | 0.06 | 28.24 | 0.18 | 28.59 | 0.29
52.8 | -40.55 | -0.06 | 0 | -0.18 | 21.39 | -0.06 | 21.39 | 0.35 | 21.39 | -0.23 | -0.06 | -20.16 | -0.06 | -26.13 | 0.06 | 0 | 0.23 | 21.39 | -0.76 | 28.48 | -2.17 | 28.42 | -2.75 | 29.06 | -24.84 | -0.06 | -58.36 | 0 | 0 | 0 | 28.65 | 0.12 | 28.54 | 0.29 | 28.71 | 0.12
52.832 | -39.08 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.39 | 0.23 | 21.39 | 0.18 | -0.06 | -18.46 | -0.06 | -24.14 | 0.06 | 0 | 0.23 | 21.39 | -0.53 | 28.71 | -1.52 | 28.89 | -2.23 | 28.54 | -23.67 | 0 | -48.4 | 0 | 0 | -0.06 | 28.54 | 0.06 | 28.65 | 0.23 | 28.48 | 0.18
52.864 | -35.16 | 0 | 0 | -0.06 | 0.06 | -0.18 | 21.5 | 0.18 | 21.39 | 0.18 | -0.06 | -15.94 | 0 | -21.39 | -0.06 | 0 | 0.18 | 21.39 | -0.53 | 28.48 | -1.29 | 28.59 | -1.52 | 28.54 | -20.98 | 0 | -42.66 | -0.06 | 0 | 0 | 28.42 | 0 | 28.36 | 0.12 | 28.48 | 0.18
52.896 | -33.81 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.18 | 21.45 | 0.47 | 0 | -14.12 | -0.06 | -19.69 | 0 | 0 | 0.29 | 21.33 | -0.35 | 28.36 | -0.82 | 28.59 | -0.94 | 28.59 | -18.46 | -0.18 | -36.62 | -0.06 | 0 | 0 | 28.42 | 0 | 28.54 | 0.06 | 28.42 | 0.23
52.928 | -31.58 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | -0.06 | -12.07 | -0.06 | -18.4 | 0 | 0 | 0.18 | 21.15 | -0.47 | 28.77 | -0.76 | 28.59 | -0.7 | 28.36 | -16.7 | -0.06 | -36.97 | -0.06 | 0 | 0 | 0.23 | -0.06 | 28.42 | 0.12 | 28.54 | 0.23
52.96 | -30 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.39 | 0.41 | -0.18 | -10.02 | -0.06 | -16.82 | 0 | 0 | 0.18 | 21.33 | -0.41 | 28.65 | -0.47 | 28.24 | -0.53 | 28.36 | -15.18 | 0 | -37.68 | 0 | 0 | -0.06 | 0.06 | 0 | 28.42 | 0.12 | 13.95 | 0.06
52.992 | -28.01 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.39 | 0.47 | -0.06 | -7.91 | 0 | -15.06 | -0.06 | 0 | 0.18 | 21.33 | -0.35 | 28.48 | -0.47 | 28.65 | -0.23 | 28.36 | -13.71 | -0.06 | -37.91 | -0.06 | 0 | 0.12 | 0 | 0 | 18.34 | 0.12 | 0.12 | 0.18
53.024 | -24.43 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 3.22 | -5.22 | 0 | -12.89 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.65 | -0.35 | 28.65 | -0.35 | 28.36 | -10.78 | -0.06 | -35.68 | 0 | -38.38 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 0 | 0.12
53.056 | -23.03 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 21.39 | 0.41 | 9.55 | -3.11 | 0 | -11.07 | -0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.24 | -0.23 | 28.54 | -0.23 | 28.42 | -8.32 | 0 | -33.87 | -0.18 | -83.26 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12
53.088 | -21.15 | -0.06 | -0.29 | -0.06 | 0 | -0.06 | 0.06 | 0.06 | 21.27 | 0.35 | 14.77 | -1.88 | -0.06 | -9.43 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.95 | -0.29 | 28.54 | -0.18 | 28.3 | -6.21 | -0.23 | -31.17 | 0 | -55.37 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12
53.12 | -19.45 | 0 | -56.13 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 20.39 | -1.17 | -0.06 | -7.56 | -0.06 | 0 | 0.23 | 21.21 | -0.35 | 28.48 | -0.23 | 28.48 | -0.23 | 28.24 | -4.04 | 6.91 | -29.65 | -0.06 | -48.4 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.23
53.152 | -17.05 | 0 | -52.85 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -0.59 | -0.18 | -5.98 | -0.06 | -60 | 0.23 | 21.33 | -0.35 | 28.48 | -0.23 | 28.48 | -0.23 | 28.42 | -2.75 | 14.65 | -27.77 | -0.18 | -44.36 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23
53.184 | -15.23 | -0.06 | -39.85 | 0 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.35 | 21.39 | -0.06 | -0.06 | -3.69 | -0.06 | -59.94 | 0 | 21.33 | -0.35 | 28.65 | -0.35 | 28.54 | -0.29 | 28.42 | -1.46 | 24.84 | -25.55 | 0 | -37.38 | -0.06 | 0 | -0.23 | 0 | -0.06 | 0 | 0.18
53.216 | -13.95 | 0 | -38.91 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.35 | 21.39 | -0.06 | 7.21 | -2.52 | 0 | -37.62 | 0.12 | 21.39 | -0.35 | 29.24 | -0.23 | 28.54 | -0.18 | 28.36 | -0.82 | 28.77 | -23.91 | 0 | -37.44 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
53.248 | -12.13 | 0.06 | -36.86 | -0.06 | -3.52 | -0.06 | 0 | -0.06 | 21.21 | 0.29 | 21.39 | 0.06 | 12.89 | -1.7 | 0 | -27.13 | 0.06 | 0.23 | -0.35 | 28.59 | -0.29 | 28.54 | -0.35 | 28.48 | -0.23 | 28.65 | -21.86 | -0.06 | -38.44 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.06
53.28 | -10.78 | 0 | -35.16 | -0.18 | -62.23 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.35 | 18.34 | -1.17 | -0.06 | -32.17 | 0.12 | 0.06 | -0.29 | 28.65 | -0.29 | 28.48 | -0.23 | 28.54 | 0 | 28.54 | -20.45 | -0.18 | -37.79 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18
53.312 | -9.61 | 0 | -32.81 | -0.06 | -44.36 | -0.06 | -7.15 | -0.06 | 21.39 | 0.23 | 21.33 | 0.29 | 21.56 | -0.7 | -0.06 | -35.57 | 0.12 | 0 | -0.29 | 28.59 | -0.29 | 28.42 | -0.18 | 28.59 | 0 | 28.65 | -18.52 | 0 | -35.74 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12
53.344 | -7.38 | -0.06 | -30.47 | -0.18 | -39.96 | 0 | -71.66 | -0.06 | 21.33 | 0.12 | 21.39 | 0.35 | 21.39 | -0.47 | 0 | -32.99 | 0.06 | 0 | -0.18 | 28.3 | -0.18 | 28.65 | -0.18 | 28.65 | 0.12 | 28.3 | -16.06 | 0 | -32.99 | 0 | -41.54 | -0.06 | 0 | 0 | 0 | 0
53.376 | -6.21 | 0 | -28.48 | -0.06 | -38.79 | -0.06 | -47.99 | -0.06 | 21.39 | 0.06 | 21.39 | 0.53 | 21.33 | -0.06 | -0.06 | -30.35 | -0.06 | 0 | -0.18 | 28.42 | -0.18 | 28.65 | -0.18 | 28.36 | 0.23 | 28.54 | -14.24 | -0.06 | -31.64 | 0 | -83.79 | -0.18 | 0 | 0 | -5.57 | 0.06
53.408 | -4.69 | -0.18 | -26.37 | -0.06 | -35.98 | -0.06 | -42.83 | -0.06 | 21.15 | 0.06 | 21.39 | 0.41 | 21.33 | -0.06 | 0 | -27.95 | 0 | 0 | -0.18 | 28.36 | -0.18 | 28.77 | -0.18 | 28.54 | 0.23 | 28.65 | -12.31 | 0 | -29.24 | 0 | -53.21 | -0.06 | -0.06 | -0.06 | -86.31 | 0
53.44 | -3.34 | -0.06 | -24.08 | -0.18 | -34.98 | -0.18 | -36.62 | -0.06 | 21.39 | -0.06 | 21.27 | 0.47 | 21.39 | 0.06 | -0.06 | -27.36 | 0 | 0 | -0.18 | 28.65 | -0.06 | 4.51 | -0.06 | 28.59 | 0.35 | 28.65 | -10.61 | -0.06 | -27.6 | -0.06 | -32.4 | 0 | -83.73 | 0 | -64.57 | 0
53.472 | -2.23 | 1.17 | -22.38 | -0.06 | -31.52 | 0 | -32.46 | -0.06 | 21.45 | -0.06 | 21.21 | 0.53 | 21.45 | 0.06 | 5.86 | -25.55 | 0 | 0 | -0.18 | 28.48 | -0.18 | 0.12 | -0.06 | 28.42 | 0.35 | 28.48 | -8.32 | 0 | -25.61 | -0.18 | -32.52 | -0.18 | -69.44 | -0.06 | -32.4 | 0.06
53.504 | -1.17 | 11.13 | -19.75 | 0 | -30.47 | -0.06 | -32.99 | -0.06 | 21.09 | 0 | 21.39 | 0.41 | 21.39 | 0.06 | 13.54 | -22.62 | -0.06 | 0 | -0.18 | 0.29 | -0.18 | 0 | -0.18 | 28.54 | 0.35 | 28.71 | -5.51 | 0 | -23.85 | 0 | -42.6 | -0.18 | -34.63 | 0 | -31.11 | -0.18
53.536 | -0.64 | 16.47 | -17.58 | -0.06 | -28.3 | -0.06 | -32.34 | -0.06 | 0.12 | 0.06 | 21.21 | 0.41 | 21.33 | 0.23 | 19.04 | -21.27 | 0 | 0 | -0.18 | 0.06 | -0.18 | 0.06 | -0.06 | 28.36 | 0.29 | 28.65 | -3.57 | 0 | -22.03 | 0 | -40.31 | -0.18 | -31.52 | 0 | -41.02 | 0
53.568 | -0.18 | 21.8 | -15.94 | -0.06 | -26.13 | -0.06 | -31.47 | -0.06 | 0 | 0 | 21.33 | 0.29 | 21.39 | 0.18 | 21.5 | -19.86 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.54 | 0.29 | 28.71 | -2.17 | 3.69 | -20.1 | -0.06 | -36.39 | -0.18 | -40.02 | -0.18 | -42.07 | 0
53.6 | -0.18 | 21.5 | -14.47 | -0.18 | -23.61 | -0.06 | -29.01 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.33 | 0.12 | 21.45 | -18.16 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.06 | 0.23 | 0.35 | 28.65 | -1.17 | 12.07 | -18.16 | -0.18 | -33.69 | -0.06 | -42.48 | -0.18 | -38.61 | -0.06
53.632 | -0.18 | 21.5 | -13.01 | -0.06 | -20.86 | 0 | -26.84 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.29 | 21.33 | -16.47 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | -0.06 | 0.12 | 0.29 | 28.59 | -0.82 | 19.51 | -16.47 | 0 | -32.93 | -0.18 | -38.38 | -0.18 | -33.4 | 0
53.664 | -0.35 | 21.45 | -10.9 | -0.06 | -19.04 | 0 | -24.32 | -0.06 | 0 | 0 | 21.27 | 0.06 | 21.39 | 0.18 | 21.39 | -14.94 | 0.06 | -22.09 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.65 | 0 | 27.36 | -13.59 | -0.06 | -32.23 | -0.23 | -33.99 | -0.06 | -32.46 | -0.06
53.696 | -0.29 | 21.45 | -9.38 | -0.06 | -17.23 | -0.06 | -22.74 | -0.06 | 0 | -0.18 | 21.33 | 0.18 | 21.33 | 0.23 | 21.33 | -12.42 | -0.06 | -74.07 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.29 | 28.59 | -0.18 | 28.48 | -11.66 | -0.06 | -30.53 | -0.06 | -32.99 | -0.06 | -32.64 | 0
53.728 | -0.35 | 21.5 | -8.2 | -0.06 | -15.7 | -0.06 | -20.22 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.18 | 21.39 | -11.07 | 0.06 | -50.04 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.65 | 0.18 | 28.59 | -10.02 | -0.06 | -28.48 | -0.18 | -32.99 | -0.18 | -31.64 | 0
53.76 | -0.35 | 21.45 | -6.68 | -0.06 | -14.59 | -0.06 | -18.16 | -0.06 | 0 | -0.06 | 21.15 | 0.12 | 21.27 | 0.12 | 21.39 | -9.84 | 0 | -32.17 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.59 | 0.12 | 28.65 | -7.68 | -0.06 | -26.9 | -0.18 | -31.47 | -0.06 | -29.53 | -0.18
53.792 | -0.35 | 21.45 | -5.1 | -0.06 | -11.95 | -0.06 | -16.58 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.15 | 0.18 | 21.27 | -7.15 | 0 | -29.94 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0.23 | 28.3 | 0.23 | 28.42 | -5.74 | -0.18 | -25.61 | -0.18 | -29.59 | -0.18 | -27.01 | -0.06
53.824 | -0.23 | 21.45 | -3.75 | -0.06 | -10.66 | 0 | -14 | 0 | 0 | -0.18 | 21.33 | 0.12 | 21.21 | 0.29 | 21.39 | -5.63 | -0.06 | -35.39 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.18 | 28.48 | 0.35 | 28.13 | -3.28 | -0.06 | -24.08 | -0.18 | -26.66 | -0.18 | -26.25 | 0
53.856 | -0.18 | 21.45 | -2.23 | -0.06 | -9.32 | 0 | -11.66 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.45 | 0.23 | 21.39 | -3.75 | 0 | -35.22 | -0.06 | 0 | 0 | -3.28 | -0.18 | 0 | 0.18 | 28.48 | 0.35 | 28.24 | -1.93 | -0.06 | -22.62 | -0.06 | -26.13 | -0.06 | -25.08 | -0.06
53.888 | -0.23 | 21.45 | -1.29 | 5.68 | -7.97 | -0.06 | -9.9 | -0.18 | 0 | -0.18 | 21.27 | 0.06 | 21.39 | 0.18 | 21.21 | -2.7 | -0.06 | -31.82 | -0.18 | 0 | 0 | -77.99 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.59 | -1.17 | 0 | -20.57 | -0.18 | -24.79 | -0.06 | -22.85 | -0.06
53.92 | -0.23 | 21.5 | -0.7 | 11.78 | -6.56 | 4.51 | -7.62 | -0.06 | 0 | -0.06 | 0.88 | 0.06 | 21.33 | 0.12 | 21.39 | -1.82 | 0 | -28.83 | -0.06 | -19.63 | 0 | -58.24 | 0 | 0 | 0.18 | 28.65 | 0.47 | 28.77 | -0.7 | 0 | -19.22 | -0.06 | -22.91 | 0 | -20.98 | 0.06
53.952 | -0.18 | 21.45 | -0.18 | 17.11 | -4.34 | 10.78 | -5.16 | -0.18 | -39.79 | -0.06 | 0 | 0.12 | 21.27 | 0.12 | 21.33 | -1 | -0.06 | -27.13 | -0.06 | -78.34 | -0.18 | -46.06 | -0.06 | 0 | 0 | 28.3 | 0.29 | 28.36 | -0.35 | 0 | -17.87 | -0.06 | -21.39 | -0.18 | -19.45 | 0
53.984 | -0.18 | 21.45 | -0.29 | 21.62 | -3.22 | 18.16 | -3.11 | -0.06 | -70.84 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.45 | -0.53 | -0.06 | -26.49 | -0.06 | -49.34 | -0.06 | -45.12 | -0.06 | 0 | 0.18 | 25.67 | 0.29 | 28.18 | -0.18 | 0 | -15.18 | -0.18 | -19.16 | -0.18 | -17.87 | 0
54.016 | -0.18 | 21.45 | -0.23 | 21.5 | -2.17 | 21.62 | -1.64 | -0.06 | -48.52 | 0 | 0.06 | 0 | 21.39 | 0.12 | 21.39 | -0.35 | 3.22 | -24.43 | -0.18 | -45.88 | -0.06 | -42.48 | -0.18 | -35.04 | 0 | 0.18 | 0.29 | 28.59 | 0.12 | 2.29 | -13.54 | 0 | -17.87 | 0 | -16.17 | 0
54.048 | -0.29 | 21.45 | -0.29 | 21.45 | -1.17 | 21.5 | -1.17 | -0.18 | -42.01 | 0 | 0 | 0.06 | 21.33 | 0.06 | 21.33 | -0.06 | 9.55 | -22.62 | -0.18 | -43.65 | -0.06 | -41.37 | -0.06 | -75.24 | 0.06 | 0 | 0.18 | 28.54 | 0.12 | 9.84 | -12.36 | 0 | -16.41 | -0.18 | -15 | 0
54.08 | -0.18 | 21.39 | -0.29 | 21.45 | -0.59 | 21.45 | -0.47 | -0.06 | -35.74 | -0.06 | 0 | -0.06 | 21.56 | 0 | 21.33 | -0.06 | 14.71 | -20.63 | -0.06 | -41.9 | 0.06 | -39.67 | -0.06 | -51.15 | 0 | 0 | 0.29 | 28.65 | 0.29 | 17.17 | -10.25 | 8.56 | -14.41 | -0.06 | -13.01 | 0
54.112 | -0.18 | 21.45 | -0.41 | 21.45 | -0.18 | 21.45 | -0.41 | -0.18 | -31.93 | -0.18 | 0 | -0.06 | 0.12 | 0 | 21.39 | 0 | 20.57 | -19.63 | -0.06 | -40.37 | -0.06 | -36.45 | -0.06 | -46.53 | 0.06 | 0 | 0.18 | 28.48 | 0.35 | 27.42 | -8.91 | 15.7 | -13.18 | 9.14 | -10.61 | 2.46
54.144 | -0.06 | 21.45 | -0.41 | 21.45 | -0.06 | 21.45 | 0.18 | -0.06 | -32.58 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | -17.7 | 0 | -37.33 | -0.18 | -34.92 | 0 | -44.24 | -0.06 | 0 | 0.18 | 28.54 | 0.29 | 28.59 | -6.62 | 25.84 | -11.13 | 18.87 | -9.32 | 12.72
54.176 | -0.18 | 21.5 | -0.23 | 21.45 | -0.18 | 21.45 | 0.12 | -0.06 | -32.64 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | 0.18 | 21.39 | -15.76 | 0 | -35.74 | -0.06 | -33.22 | -0.06 | -41.19 | 0 | 0 | 0.18 | 28.59 | 0.29 | 28.48 | -5.04 | 28.65 | -9.43 | 26.84 | -7.68 | 20.27
54.208 | -0.18 | 21.39 | -0.23 | 21.45 | -0.35 | 21.45 | 0.47 | -0.06 | -30.53 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -13.71 | -0.06 | -34.51 | -0.06 | -31 | -0.06 | -39.96 | -0.06 | 0 | 0.18 | 28.54 | 0.29 | 28.59 | -3.4 | 28.65 | -7.97 | 28.65 | -6.33 | 28.13
54.24 | -0.06 | 21.39 | -0.29 | 21.45 | -0.29 | 21.5 | 0.41 | -0.06 | -29.06 | 0.12 | 0 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | -12.13 | -0.18 | -32.34 | -0.06 | -29.12 | 0 | -38.5 | 0 | 0 | 0 | 28.42 | 0.29 | 28.54 | -2.23 | 28.65 | -6.62 | 28.42 | -4.63 | 28.71
54.272 | -0.18 | 21.45 | -0.18 | 21.5 | -0.41 | 21.39 | 0.41 | -0.06 | -25.9 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.39 | -10.14 | -0.06 | -30.82 | 0 | -26.31 | -0.06 | -37.15 | 0 | 0 | 0.06 | 28.48 | 0.23 | 28.3 | -1.52 | 28.42 | -4.75 | 28.59 | -2.75 | 28.77
54.304 | -0.06 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | 0.53 | 6.8 | -23.85 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.12 | 21.39 | -8.67 | -0.06 | -28.01 | 0 | -24.84 | -0.06 | -33.75 | -0.06 | 0 | 0.06 | 28.65 | 0.23 | 28.42 | -0.94 | 28.18 | -2.93 | 28.48 | -2.17 | 28.42
54.336 | -0.06 | 21.39 | -0.18 | 21.45 | -0.23 | 21.5 | 0.47 | 12.31 | -21.56 | 0.06 | -6.56 | -0.18 | 0 | -0.18 | 17.46 | 0.18 | 21.39 | -6.68 | -0.18 | -26.02 | 0.06 | -22.79 | 0 | -32.23 | -0.06 | 0 | 0.12 | 28.54 | 0.18 | 28.42 | -0.41 | 28.48 | -1.99 | 28.42 | -1.41 | 28.54
54.368 | -0.18 | 0.12 | -0.23 | 21.45 | -0.23 | 21.45 | 0.47 | 17.99 | -20.39 | -0.18 | -77 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.04 | -5.22 | 4.92 | -24.2 | 0 | -20.98 | 0 | -30.59 | 0 | 0 | 0.12 | 28.48 | 0.12 | 28.42 | -0.18 | 28.54 | -1.7 | 28.65 | -0.64 | 28.65
54.4 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.5 | 0.35 | 21.56 | -17.93 | -0.06 | -54.2 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | -3.63 | 13.07 | -22.33 | 0 | -19.16 | 0 | -28.48 | 0 | 0 | -0.06 | 28.59 | 0 | 28.48 | -0.06 | 28.48 | -0.53 | 28.48 | -0.41 | 28.65
54.432 | -0.18 | 0 | -0.35 | 21.45 | -0.18 | 21.45 | 0.47 | 21.39 | -15.29 | 0 | -41.95 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.39 | -2.23 | 20.63 | -20.57 | -0.06 | -16.29 | 0 | -26.6 | 0 | -49.75 | -0.06 | 28.42 | 0.12 | 28.48 | 0.06 | 28.65 | -0.53 | 28.48 | -0.06 | 28.59
54.464 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.5 | 0.41 | 21.39 | -14.12 | 0 | -35.98 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.45 | -1.46 | 29.12 | -18.16 | 0 | -15.18 | 0.06 | -24.26 | -0.06 | -70.14 | -0.06 | 1.41 | 0 | 28.42 | 0 | 28.48 | -0.35 | 28.48 | 0 | 28.42
54.496 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.29 | 21.39 | -11.6 | -0.18 | -31.93 | 0 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.82 | 28.83 | -16.29 | 0 | -13.48 | 0 | -22.15 | -0.18 | -48.87 | -0.06 | 0.12 | 0 | 28.42 | 0.18 | 28.42 | 0 | 28.77 | 0.06 | 28.59
54.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | -9.84 | -0.06 | -32.76 | -0.18 | -51.45 | -0.06 | 0 | 0.18 | 21.39 | -0.47 | 28.48 | -14.77 | 0.64 | -12.07 | -0.06 | -20.39 | -0.06 | -47.23 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.59 | 0 | 28.42 | 0.06 | 28.54
54.56 | -0.18 | 0 | -0.18 | 21.45 | -0.35 | 21.39 | 0.23 | 21.45 | -7.62 | -0.06 | -32.64 | 0 | -63.99 | -0.12 | 0 | 0.18 | 21.5 | -0.29 | 28.59 | -13.24 | 10.72 | -10.84 | 1.99 | -18.46 | -0.06 | -41.08 | -0.18 | 0.06 | 0 | 28.65 | 0.18 | 28.48 | -0.06 | 28.65 | 0.06 | 28.42
54.592 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.18 | 21.39 | -4.45 | 0.06 | -31.11 | -0.06 | -39.26 | -0.06 | 0 | 0.12 | 21.45 | -0.06 | 28.65 | -11.84 | 21.15 | -9.32 | 8.5 | -16.47 | -0.06 | -36.33 | 0 | 0 | -0.06 | 28.65 | 0.18 | 28.59 | 0.12 | 28.48 | 0.35 | 28.65
54.624 | -0.12 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.12 | 21.39 | -2.7 | -0.06 | -28.95 | 0 | -28.3 | -0.06 | 0 | 0.18 | 21.27 | -0.06 | 28.89 | -10.37 | 28.54 | -7.85 | 18.28 | -14.65 | -0.06 | -38.32 | 0 | 0 | -0.06 | 28.59 | 0.18 | 28.48 | 0 | 28.48 | 0.29 | 28.48
54.656 | -0.06 | 0 | -0.06 | 21.5 | -0.29 | 21.5 | 0.06 | 21.27 | -1.99 | -0.06 | -27.48 | -0.06 | -33.52 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.42 | -8.85 | 28.54 | -6.56 | 25.96 | -13.24 | -0.06 | -38.09 | 0 | 0 | -0.06 | 28.59 | 0.12 | 28.42 | 0.12 | 28.48 | 0.29 | 28.59
54.688 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.12 | 21.39 | -0.35 | 3.75 | -24.96 | 0 | -35.68 | -0.18 | 0 | 0.12 | 21.74 | 0.12 | 28.65 | -7.27 | 28.59 | -5.39 | 28.77 | -11.78 | 0 | -36.68 | 0 | 0 | -0.06 | 28.59 | 0.12 | 28.71 | 0.23 | 28.36 | 0.23 | 28.42
54.72 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | 0.06 | 21.39 | -0.64 | 9.96 | -23.15 | -0.18 | -33.28 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 28.83 | -6.45 | 28.59 | -4.51 | 28.65 | -10.14 | -0.06 | -35.1 | -0.18 | 0 | -0.06 | 28.59 | 0.12 | 28.71 | 0.06 | 28.65 | 0.18 | 28.42
54.752 | -0.06 | 0 | -0.06 | 21.5 | -0.18 | 21.5 | -0.06 | 21.33 | -0.06 | 15.18 | -20.16 | -0.06 | -30.35 | -0.06 | 0.06 | 0.12 | 21.39 | 0.12 | 28.48 | -5.16 | 28.48 | -3.22 | 28.71 | -9.02 | -0.06 | -33.17 | 0 | 0 | -0.06 | 28.77 | 0 | 28.48 | 0.23 | 28.48 | 0.18 | 28.13
54.784 | -0.06 | -23.85 | -0.18 | 0.12 | -0.18 | 21.45 | 0 | 21.39 | 0.18 | 21.68 | -18.16 | -0.06 | -27.19 | -0.06 | -63.28 | 0.18 | 21.39 | 0.06 | 28.65 | -4.1 | 28.59 | -2.52 | 28.65 | -7.32 | 0 | -30.94 | -0.18 | 0 | -0.06 | 28.48 | 0 | 28.18 | 0.29 | 28.83 | 0.35 | 28.42
54.816 | -0.06 | -62.35 | -0.18 | 0 | -0.18 | 0.12 | 0.06 | 21.45 | 0.18 | 21.45 | -16.99 | -0.18 | -26.72 | -0.06 | -60.82 | 0.12 | 21.39 | 0.12 | 28.42 | -3.46 | 28.42 | -1.99 | 28.83 | -6.27 | 0 | -29.24 | 0.06 | 0 | 0 | 28.48 | 0 | 28.48 | 0.18 | 28.42 | 0.18 | 28.48
54.848 | -0.06 | -46.53 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.45 | -14.47 | -0.06 | -25.55 | -0.06 | -35.8 | 0.12 | 21.39 | 0.23 | 28.48 | -2.64 | 28.3 | -1.52 | 28.3 | -4.86 | -0.06 | -27.42 | 0 | 0 | -0.06 | 28.54 | 0.06 | 28.48 | 0.18 | 28.48 | 0.29 | 28.54
54.88 | -0.18 | -40.55 | -0.18 | 0.06 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.39 | -12.77 | -0.06 | -23.44 | -0.06 | -27.6 | 0.18 | 21.33 | 0.12 | 28.95 | -2.23 | 28.48 | -1.41 | 28.18 | -4.39 | -0.06 | -25.43 | -0.06 | -15.94 | -0.06 | 28.13 | 0 | 28.48 | 0.06 | 28.48 | 0.18 | 28.48
54.912 | -0.06 | -39.44 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.45 | -10.2 | -0.06 | -21.56 | -0.06 | -32.58 | 0.12 | 1.11 | 0.18 | 28.42 | -1.7 | 29.12 | -0.64 | 28.48 | -3.28 | 0 | -23.5 | 0 | -92.58 | -0.06 | 0.76 | -0.06 | 28.48 | 0.12 | 28.48 | 0.06 | 28.42
54.944 | 0 | -34.92 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.45 | -7.56 | -0.06 | -19.75 | 0 | -35.74 | 0.06 | 0 | 0.18 | 28.48 | -1.17 | 28.59 | -0.59 | 29.12 | -2.46 | 0 | -21.15 | 0 | -63.52 | -0.06 | 0.12 | 0 | 17.87 | 0.12 | 28.48 | 0.18 | 28.48
54.976 | -0.06 | -34.69 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.27 | -5.8 | 0 | -18.22 | -0.18 | -33.46 | 0.18 | 0 | 0.18 | 28.54 | -0.94 | 28.3 | -0.53 | 28.54 | -1.7 | -0.06 | -19.63 | 0 | -47.99 | -0.06 | 0 | -0.18 | 0.18 | 0.12 | 13.71 | 0.23 | 28.48
55.008 | -0.06 | -32.17 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -3.57 | 0.06 | -16.29 | 0.06 | -29.53 | 0.12 | 0.06 | 0.23 | 28.48 | -0.7 | 28.65 | -0.35 | 28.42 | -1.17 | -0.18 | -17.4 | 0 | -43.48 | -0.06 | 0 | 0 | 0 | 0.12 | 0.12 | 0.12 | 0.29
55.04 | -0.06 | -29.77 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.35 | 21.21 | -2.17 | -0.06 | -14.53 | -0.06 | -27.72 | 0.12 | 0 | 0.18 | 28.36 | -0.53 | 28.48 | -0.35 | 28.48 | -0.94 | 7.32 | -15.82 | -0.06 | -37.91 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0.12
55.072 | -0.06 | -28.18 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.05 | 2.99 | -12.77 | -0.06 | -27.42 | 0.06 | 0 | 0.18 | 28.54 | -0.35 | 28.54 | -0.29 | 28.36 | -0.53 | 14.82 | -13.71 | 0 | -36.97 | 0 | 0 | -0.18 | 0 | 0.12 | 0 | -0.06 | 0
55.104 | -0.06 | -25.31 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -0.7 | 11.07 | -10.72 | -0.06 | -25.25 | 0.12 | 0 | 0.23 | 28.42 | -0.29 | 28.36 | -0.29 | 28.48 | -0.47 | 25.31 | -10.78 | 0 | -38.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0.06
55.136 | 0 | -23.67 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.35 | 21.39 | -0.06 | 16.47 | -9.08 | -0.06 | -23.09 | 0 | 0 | 0.18 | 28.83 | -0.35 | 28.59 | -0.29 | 28.48 | -0.35 | 28.59 | -9.32 | 0 | -38.44 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0
55.168 | -0.06 | -21.21 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.29 | 21.27 | 0 | 21.97 | -7.38 | 5.33 | -20.98 | -0.06 | 0 | 0.12 | 28.42 | -0.35 | 28.65 | -0.23 | 28.36 | -0.23 | 28.48 | -6.8 | 0 | -35.86 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.12 | 0
55.2 | -0.06 | -18.99 | 0 | -30.29 | -0.18 | 0 | -0.18 | 0.12 | 0.29 | 21.39 | 0.23 | 21.5 | -6.15 | 11.13 | -20.57 | -0.06 | 0 | 0 | 28.65 | -0.35 | 28.42 | -0.29 | 28.36 | -0.29 | 28.59 | -4.63 | 0 | -34.04 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
55.232 | -0.06 | -17.29 | -0.06 | -61.12 | -0.06 | -19.16 | -0.18 | 0 | 0.23 | 21.45 | 0.18 | 21.45 | -4.04 | 16.64 | -18.28 | 0.06 | 0 | 0.18 | 28.54 | -0.41 | 28.48 | -0.23 | 28.65 | -0.35 | 28.48 | -2.75 | 0 | -31.93 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
55.264 | -0.06 | -15.06 | -0.06 | -41.02 | -0.06 | -58.13 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -2.29 | 21.5 | -16.11 | 0 | 0 | 0.23 | 0.23 | -0.29 | 28.77 | -0.23 | 28.48 | -0.29 | 28.48 | -1.64 | -0.06 | -30.47 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0
55.296 | -0.06 | -13.89 | -0.06 | -39.08 | -0.18 | -42.31 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.45 | -1.52 | 21.39 | -14.47 | 0.06 | 0 | 0.06 | 0.06 | -0.29 | 28.65 | -0.29 | 28.54 | -0.23 | 28.59 | -0.82 | 0 | -28.65 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
55.328 | 0 | -12.48 | -0.06 | -37.85 | -0.06 | -39.61 | -0.18 | 0 | 0.06 | 21.27 | 0.53 | 21.39 | -1.11 | 21.39 | -12.83 | 0 | -4.28 | 0.23 | 0 | -0.29 | 28.65 | -0.18 | 28.83 | -0.23 | 28.42 | -0.41 | -0.06 | -26.13 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
55.36 | 0 | -10.61 | -0.06 | -35.57 | -0.06 | -38.61 | -0.18 | 0 | 0.06 | 21.39 | 0.41 | 21.21 | -0.59 | 21.45 | -11.37 | 0 | -72.83 | 0.18 | 0.06 | -0.29 | 28.48 | -0.18 | 28.65 | -0.18 | 28.95 | -0.23 | 0 | -24.14 | -0.06 | -79.46 | -0.06 | -0.06 | -0.06 | 0 | 0 | 0
55.392 | 1.23 | -9.08 | 0 | -34.16 | -0.06 | -35.86 | -0.18 | 0 | 0.06 | 21.45 | 0.53 | 21.39 | -0.18 | 21.39 | -9.55 | 0.06 | -55.08 | 0.23 | 0 | -0.29 | 28.36 | -0.23 | 28.48 | -0.18 | 28.42 | 0 | 0 | -22.21 | -0.06 | -59.71 | -0.06 | -54.67 | 0 | 0.06 | -0.06 | 0
55.424 | 11.07 | -7.79 | -0.06 | -30.7 | -0.06 | -33.52 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.39 | -0.18 | 21.27 | -6.97 | 0 | -29.47 | 0.12 | 0 | -0.29 | 0.29 | -0.18 | 28.42 | -0.29 | 28.42 | 0.23 | 3.34 | -20.74 | 0 | -48.11 | -0.06 | -69.49 | 0 | -87.72 | -0.06 | -18.93
55.456 | 16.41 | -6.33 | -0.06 | -29.12 | -0.06 | -31.17 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | 21.21 | 0 | 21.21 | -5.68 | 0 | -29.24 | 0.12 | 0 | -0.18 | 0.12 | -0.18 | 17.34 | -0.18 | 28.65 | 0.29 | 11.6 | -18.69 | -0.06 | -45.35 | -0.18 | -41.66 | -0.06 | -66.86 | -0.06 | -89.59
55.488 | 21.68 | -4.51 | 0 | -27.36 | -0.06 | -29.71 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | 0.12 | 21.33 | -3.75 | 0 | -34.4 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0.18 | -0.23 | 28.59 | 0.29 | 19.22 | -16.06 | 0 | -40.02 | -0.18 | -29.06 | -0.06 | -38.44 | -0.06 | -61.99
55.52 | 21.39 | -3.46 | -0.06 | -25.37 | 0 | -27.89 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0.06 | 21.39 | -2.7 | -0.06 | -34.98 | 0.06 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 27.01 | -14.59 | -0.06 | -36.39 | -0.06 | -37.62 | -0.06 | -29.47 | -0.18 | -34.63
55.552 | 21.39 | -2.23 | -0.06 | -22.62 | 0 | -25.55 | -0.18 | 0 | 0 | 21.39 | 0.29 | 21.45 | 0.29 | 21.33 | -1.7 | -0.06 | -32.93 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.48 | -12.36 | 0 | -37.5 | -0.06 | -43.3 | -0.18 | -38.91 | -0.18 | -29.77
55.584 | 21.5 | -1.11 | 0 | -20.39 | -0.06 | -22.91 | -0.18 | 0 | -0.06 | 0.12 | 0.23 | 21.45 | 0.18 | 21.39 | -0.94 | -0.06 | -28.13 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.29 | 28.59 | -10.37 | -0.06 | -37.5 | -0.18 | -38.15 | 0 | -42.42 | -0.06 | -41.66
55.616 | 21.45 | -0.47 | 0 | -18.52 | -0.06 | -20.57 | -0.18 | -51.15 | -0.06 | 0 | 0.12 | 21.5 | 0.29 | 21.15 | -0.64 | -0.06 | -27.25 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.23 | 28.36 | -8.14 | 0 | -36.15 | -0.18 | -34.98 | -0.06 | -38.67 | 0 | -41.84
55.648 | 21.45 | -0.06 | -0.06 | -16.76 | -0.18 | -18.93 | -0.06 | -64.51 | -0.06 | 0 | 0.18 | 21.39 | 0.23 | 21.33 | -0.18 | -0.06 | -26.54 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.42 | -5.39 | -0.06 | -34.04 | 0 | -32.93 | -0.06 | -34.04 | 0 | -37.68
55.68 | 21.45 | -0.23 | 0 | -15.59 | 0.06 | -17.05 | -0.06 | -45.65 | -0.18 | 0 | 0.18 | 21.27 | 0.23 | 21.27 | -0.18 | -0.06 | -25.31 | 0.12 | -17.7 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.23 | 28.42 | -3.63 | 0 | -32.17 | -0.18 | -33.17 | -0.18 | -32.93 | -0.06 | -34.04
55.712 | 21.39 | -0.06 | -0.06 | -13.89 | -0.06 | -15.29 | -0.06 | -40.02 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | -0.06 | -0.18 | -22.97 | 0.23 | -77.29 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.35 | 28.54 | -1.82 | 2.64 | -29.83 | -0.06 | -32.7 | -0.06 | -32.93 | -0.06 | -32.7
55.744 | 21.45 | -0.29 | -0.06 | -11.13 | -0.06 | -13.71 | -0.18 | -33.69 | -0.18 | 0 | 0.06 | 21.39 | 0.12 | 21.5 | -0.18 | -0.06 | -20.92 | 0.12 | -47.99 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.42 | -1.46 | 13.24 | -28.36 | -0.06 | -29.42 | -0.18 | -31.76 | -0.06 | -32.46
55.776 | 21.39 | -0.35 | 0.23 | -10.14 | -0.06 | -11.84 | -0.18 | -32.81 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.33 | 0.23 | -0.18 | -19.69 | 0.12 | -44.59 | 0 | 0 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.3 | -0.7 | 20.74 | -27.01 | 0 | -27.42 | -0.06 | -29.65 | 0 | -31.64
55.808 | 21.5 | -0.41 | 7.44 | -9.02 | -0.06 | -10.25 | -0.06 | -33.34 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | 0.23 | -0.06 | -17.99 | 0.23 | -44.24 | -0.06 | 0 | 0 | 0 | -0.18 | 28.48 | 0.23 | 28.65 | -0.35 | 29.01 | -24.9 | -0.18 | -26.43 | -0.06 | -28.3 | 0 | -29.06
55.84 | 21.5 | -0.41 | 13.13 | -7.27 | 0 | -9.14 | 0 | -31.76 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.39 | 0.23 | 4.92 | -15.94 | 0.18 | -42.13 | 0 | -0.76 | -0.12 | 0 | -0.18 | 28.48 | 0.18 | 28.65 | -0.18 | 28.65 | -22.68 | 0 | -25.67 | -0.06 | -26.31 | -0.06 | -27.07
55.872 | 21.5 | -0.41 | 18.63 | -6.15 | 0 | -7.38 | -0.18 | -30.59 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.5 | 0.29 | 12.89 | -14.24 | 0.12 | -39.85 | 0 | -78.23 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.65 | 0.06 | 28.65 | -20.86 | -0.06 | -23.85 | -0.06 | -24.96 | 0 | -26.31
55.904 | 21.45 | -0.18 | 21.5 | -4.1 | -0.06 | -6.09 | -0.06 | -27.77 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.35 | 18.34 | -11.89 | 0.06 | -38.09 | -0.06 | -54.79 | -0.06 | -69.67 | -0.18 | 23.85 | 0.12 | 28.48 | 0 | 28.42 | -18.87 | 0 | -21.5 | -0.06 | -22.91 | 0 | -24.55
55.936 | 21.45 | -0.18 | 21.5 | -3.11 | 0.82 | -4.45 | -0.18 | -25.84 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.33 | 0.29 | 21.56 | -10.08 | 0.12 | -35.57 | -0.06 | -45.59 | -0.06 | -60.65 | -0.18 | 0.18 | 0.06 | 28.01 | 0.35 | 28.65 | -16.99 | 0 | -19.81 | -0.06 | -21.21 | -0.06 | -22.56
55.968 | 21.45 | -0.23 | 21.45 | -1.99 | 8.73 | -3.22 | -0.06 | -23.67 | -0.06 | 0 | 0 | 0.12 | 0.18 | 21.21 | 0.18 | 21.39 | -9.08 | 0.12 | -34.4 | -0.06 | -45.29 | 0 | -45.35 | 0 | 0 | -0.06 | 28.59 | 0.29 | 28.42 | -15.06 | -0.18 | -18.52 | -0.18 | -19.69 | 0 | -21.27
56 | 21.45 | -0.18 | 21.5 | -1.17 | 14.36 | -1.99 | 1.46 | -21.74 | 0 | -18.46 | -0.06 | 0.06 | 0.06 | 21.33 | 0.23 | 21.27 | -7.38 | 0.06 | -32.93 | -0.06 | -42.78 | -0.06 | -45.47 | -0.18 | 0 | 0.06 | 28.48 | 0.29 | 28.54 | -13.24 | 0.06 | -17.17 | -0.06 | -18.34 | 0 | -19.81
56.032 | 21.45 | -0.23 | 21.45 | -0.53 | 21.68 | -1.05 | 8.79 | -19.98 | -0.06 | -77.05 | -0.06 | 0 | 0.06 | 21.21 | 0.18 | 21.39 | -5.45 | 0 | -30.18 | -0.06 | -40.55 | -0.06 | -44.01 | -0.06 | 0 | 0.06 | 28.83 | 0.29 | 28.36 | -10.66 | 1.88 | -15.23 | -0.06 | -16.47 | 0 | -18.52
56.064 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -0.53 | 15.94 | -17.23 | -0.06 | -47.64 | -0.06 | 0 | 0.06 | 0.12 | 0.35 | 21.39 | -3.69 | 0 | -27.95 | -0.06 | -38.38 | -0.06 | -41.08 | 0 | 0 | 0.06 | 28.59 | 0.29 | 28.3 | -8.32 | 13.77 | -12.89 | -0.18 | -14.3 | 2.46 | -16.06
56.096 | 21.5 | -0.29 | 21.45 | -0.23 | 21.5 | -0.18 | 21.5 | -15.35 | -0.06 | -41.02 | -0.06 | 0 | 0 | 0 | 0.29 | 21.45 | -2.29 | 0.23 | -25.67 | -0.06 | -36.45 | -0.06 | -39.08 | -0.06 | 0 | 0 | 28.54 | 0.18 | 28.48 | -6.09 | 21.39 | -11.66 | 0.35 | -12.89 | 10.2 | -14.24
56.128 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -13.36 | 0 | -36.04 | 0 | 0 | -0.06 | 0 | 0.18 | 21.33 | -1.52 | 0.12 | -23.91 | -0.18 | -35.63 | 0 | -37.56 | 0 | 0 | 0.06 | 28.54 | 0.23 | 28.54 | -4.28 | 29.01 | -9.84 | 10.14 | -11.48 | 17.7 | -12.72
56.16 | 21.5 | -0.23 | 21.39 | -0.35 | 21.45 | -0.06 | 21.45 | -11.37 | -0.06 | -31.99 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.39 | -0.94 | 0.06 | -22.21 | -0.18 | -33.81 | -0.06 | -35.22 | -0.06 | 0 | 0.06 | 28.48 | 0.18 | 28.59 | -2.7 | 28.48 | -8.56 | 17.4 | -9.9 | 25.43 | -11.02
56.192 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | -9.26 | 0 | -32.58 | -0.06 | 0 | 0.12 | 0 | 0.29 | 21.5 | -0.53 | 0.12 | -20.22 | 0 | -30.7 | -0.06 | -34.69 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.59 | -1.46 | 28.42 | -6.62 | 25.31 | -7.79 | 28.42 | -9.49
56.224 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -0.47 | 21.39 | -6.45 | 0.06 | -32.34 | 0.06 | 0 | 0 | 0 | 0.23 | 21.33 | -0.29 | 0.12 | -17.75 | 0.06 | -29.06 | -0.06 | -31.64 | -0.18 | 0 | 0 | 28.83 | 0.12 | 28.48 | -0.82 | 28.48 | -4.45 | 28.54 | -5.86 | 28.54 | -7.79
56.256 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.41 | 21.39 | -4.57 | -0.06 | -30.94 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.06 | 0.12 | -16.41 | -0.06 | -26.31 | 0 | -29.47 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.65 | -0.29 | 28.54 | -3.16 | 28.48 | -4.63 | 28.54 | -6.21
56.288 | 0.12 | -0.06 | 21.45 | -0.47 | 21.45 | -0.35 | 21.39 | -2.52 | -0.06 | -29.24 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.21 | 0 | 0.12 | -14.24 | 0.06 | -24.84 | 0 | -27.83 | -0.06 | 0 | 0 | 28.36 | 0.12 | 28.59 | -0.23 | 28.65 | -2.17 | 28.54 | -3.11 | 28.48 | -4.22
56.32 | 0.06 | -0.18 | 21.45 | -0.23 | 21.45 | -0.41 | 21.45 | -1.76 | 0.06 | -26.78 | 0 | 0 | -0.06 | 0 | 0.12 | 21.45 | -0.06 | 0.12 | -13.48 | 0 | -23.32 | -0.18 | -25.37 | -0.06 | 0 | 0 | 4.45 | 0 | 28.65 | 0.12 | 28.54 | -1.46 | 28.36 | -2.05 | 28.36 | -2.75
56.352 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.41 | 21.39 | -0.76 | 0.06 | -24.9 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 3.75 | -11.89 | -0.06 | -19.86 | 0 | -23.85 | 0 | -47.46 | 0 | 0.12 | 0.06 | 28.42 | 0.12 | 28.48 | -0.94 | 28.3 | -0.94 | 28.36 | -2.23
56.384 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.39 | -0.35 | -0.06 | -22.38 | -0.18 | -33.52 | 0 | 0 | 0.18 | 21.27 | 0.18 | 9.84 | -10.08 | 0 | -18.46 | 0 | -20.92 | 0 | -71.78 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.48 | -0.41 | 28.65 | -0.94 | 28.65 | -1.17
56.416 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.18 | 21.39 | 0.06 | 0.06 | -20.51 | -0.06 | -72.54 | -0.06 | 0 | 0.06 | 21.27 | 0.12 | 15.18 | -9.02 | 0 | -16.52 | 0 | -19.04 | -0.06 | -48.63 | 0 | 0.06 | 0 | 28.48 | 0.35 | 28.36 | -0.23 | 28.59 | -0.53 | 28.48 | -0.7
56.448 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.18 | 21.33 | 0.12 | -0.06 | -18.69 | -0.06 | -49.63 | 0.06 | 0 | 0.06 | 21.21 | 0.12 | 21.04 | -7.38 | -0.06 | -15.35 | -0.06 | -17.29 | 0 | -47.87 | -0.06 | 0 | -0.06 | 28.59 | 0.29 | 28.42 | -0.29 | 28.42 | -0.41 | 28.42 | -0.53
56.48 | 0.06 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | 0.23 | -0.06 | -16.47 | -0.06 | -41.9 | -0.06 | -36.74 | 0.06 | 21.33 | 0.23 | 21.39 | -6.45 | -0.06 | -13.83 | 0 | -15.65 | 0 | -41.66 | -0.06 | 0 | 0.06 | 28.89 | 0.29 | 28.3 | 0.06 | 28.3 | 0 | 28.36 | -0.35
56.512 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | 0.41 | 0 | -14.41 | 0 | -36.33 | 0 | -70.26 | 0.12 | 21.15 | 0.23 | 21.39 | -5.39 | -0.06 | -11.95 | 0 | -14.3 | -0.18 | -36.27 | -0.06 | 0 | 0 | 28.95 | 0.29 | 28.65 | 0.06 | 28.42 | -0.18 | 28.42 | -0.18
56.544 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | 0.47 | 0 | -12.07 | -0.06 | -32.34 | -0.06 | -39.67 | 0 | 21.39 | 0.23 | 21.15 | -4.28 | 0 | -10.49 | 2.7 | -12.48 | -0.06 | -37.03 | 0 | 0 | 0 | 28.36 | 0.41 | 28.65 | 0.06 | 28.48 | 0 | 28.54 | 0
56.576 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.29 | 0 | -9.55 | 0.12 | -32.81 | -0.18 | -27.01 | 0.12 | 21.15 | 0.29 | 21.33 | -3.52 | 0 | -9.38 | 10.78 | -10.9 | 0 | -37.56 | -0.06 | 0 | -0.06 | 28.42 | 0.23 | 28.65 | 0.12 | 28.59 | 0.12 | 28.36 | 0
56.608 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | 0.47 | 0 | -7.79 | 0.06 | -32.58 | -0.06 | -31.7 | 0.12 | 21.33 | 0.29 | 21.33 | -2.93 | 3.34 | -7.79 | 18.22 | -9.79 | 0.06 | -36.92 | 0 | 0 | -0.06 | 0.29 | 0.35 | 28.65 | 0.12 | 28.54 | 0.18 | 28.59 | 0.12
56.64 | 0 | -0.06 | 21.5 | -0.23 | 21.39 | -0.18 | 21.45 | 0.35 | 3.34 | -5.63 | 0 | -30.94 | -0.06 | -35.98 | 0.06 | 21.39 | 0.18 | 21.39 | -2.52 | 11.66 | -6.68 | 26.02 | -8.09 | -0.06 | -35.33 | 0 | 0 | -0.06 | 0.12 | 0.12 | 28.48 | 0.18 | 28.42 | 0.06 | 28.65 | 0
56.672 | 0 | -0.06 | 20.22 | -0.06 | 21.45 | -0.18 | 21.45 | 0.47 | 11.25 | -3.52 | -0.18 | -28.95 | 0.06 | -34.28 | -0.06 | 21.33 | 0.35 | 21.33 | -1.76 | 19.34 | -5.22 | 28.65 | -7.32 | -0.06 | -33.28 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.59 | 0.18 | 28.48 | 0.23 | 28.59 | 0.12
56.704 | -27.07 | -0.18 | 0.06 | -0.18 | 21.39 | -0.23 | 21.39 | 0.35 | 16.7 | -1.7 | 0.06 | -26.6 | -0.18 | -30 | -0.06 | 21.21 | 0.12 | 21.39 | -1.46 | 29.42 | -4.28 | 28.3 | -5.51 | 5.04 | -30.59 | -0.06 | 0 | -0.06 | 0.06 | 0.06 | 28.24 | 0.18 | 28.59 | 0.18 | 28.54 | 0.06
56.736 | -62.64 | -0.06 | 0 | -0.06 | 21.5 | -0.18 | 21.39 | 0.47 | 21.62 | -0.94 | -0.06 | -24.32 | -0.06 | -27.54 | 0.06 | 0.12 | 0.18 | 21.39 | -1.23 | 28.54 | -3.52 | 28.65 | -4.39 | 13.07 | -28.48 | -0.06 | -5.86 | -0.06 | 0 | 0.06 | 28.48 | 0.06 | 28.59 | 0 | 28.42 | 0.12
56.768 | -46.06 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.33 | 0.29 | 21.45 | -0.7 | 0 | -22.79 | -0.18 | -27.6 | 0.06 | 0 | 0.23 | 21.39 | -0.82 | 28.48 | -2.81 | 28.65 | -3.93 | 20.8 | -27.19 | 0 | -92.76 | 0 | 0 | 0.12 | 28.36 | 0.12 | 28.18 | 0.06 | 28.42 | 0.12
56.8 | -40.02 | -0.06 | 0 | -0.12 | 21.45 | -0.18 | 21.39 | 0.29 | 21.45 | -0.18 | 0 | -20.51 | 0 | -26.13 | -0.06 | 0 | 0.23 | 21.39 | -0.76 | 28.65 | -2.17 | 28.59 | -3.11 | 28.83 | -25.61 | 0 | -61.82 | 0 | 0 | 0.06 | 28.65 | 0.06 | 28.59 | 0.06 | 28.48 | 0.12
56.832 | -38.79 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 21.39 | 0.29 | 21.45 | 0 | -0.06 | -18.69 | -0.18 | -24.32 | 0.06 | 0 | 0.12 | 21.39 | -0.59 | 28.59 | -1.52 | 28.71 | -2.23 | 28.65 | -23.67 | 0 | -47.93 | -0.06 | 0 | 0 | 28.83 | 0 | 28.18 | 0.06 | 28.48 | 0.18
56.864 | -35.74 | -0.18 | 0 | -0.18 | 0.06 | -0.18 | 21.39 | 0.29 | 21.39 | 0.23 | -0.06 | -16.17 | -0.06 | -21.27 | -0.06 | 0 | 0.18 | 21.33 | -0.41 | 28.18 | -1.46 | 28.59 | -1.64 | 28.48 | -21.33 | 0 | -43.89 | -0.06 | 0 | 0.06 | 28.42 | 0 | 28.42 | 0.12 | 28.42 | 0.06
56.896 | -34.51 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.12 | 0.06 | 21.39 | 0.47 | -0.06 | -13.89 | -0.06 | -19.63 | -0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.59 | -1.05 | 28.3 | -1.17 | 28.42 | -19.4 | -0.06 | -38.15 | 0 | 0 | 0 | 28.48 | 0 | 28.42 | 0.12 | 28.42 | 0.18
56.928 | -32.23 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.06 | 0.23 | 21.33 | 0.47 | -0.06 | -11.84 | 0 | -18.69 | -0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.24 | -0.94 | 28.89 | -0.94 | 28.65 | -17.4 | 0 | -37.21 | -0.18 | 0 | -0.06 | 0.23 | 0 | 28.77 | 0.06 | 28.48 | 0.12
56.96 | -30.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.53 | -0.06 | -10.08 | -0.06 | -17.05 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.65 | -0.7 | 28.48 | -0.59 | 28.65 | -15.65 | -0.06 | -38.2 | 0 | 0 | -0.06 | 0.12 | 0 | 28.3 | 0 | 23.26 | 0.23
56.992 | -27.95 | 0 | 0 | -0.18 | 0 | -0.18 | 0.06 | 0.18 | 21.39 | 0.47 | -0.06 | -7.79 | -0.18 | -15.18 | -0.06 | 0 | 0.12 | 21.33 | -0.35 | 28.48 | -0.41 | 28.42 | -0.53 | 28.65 | -13.48 | 0 | -37.85 | 0 | 0 | -0.18 | 0 | -0.06 | 23.85 | 0.12 | 0.12 | 0.06
57.024 | -25.08 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.41 | 2.99 | -5.1 | -0.23 | -12.89 | -0.06 | 0 | 0.18 | 21.33 | -0.47 | 28.59 | -0.35 | 28.54 | -0.23 | 28.65 | -11.02 | -0.06 | -34.92 | 0 | -23.03 | 0.06 | 0 | -0.06 | 0.12 | 0 | 0 | 0.12
57.056 | -23.85 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.35 | 9.32 | -3.11 | -0.18 | -11.31 | -0.18 | 0 | 0.18 | 21.39 | -0.29 | 28.42 | -0.41 | 28.83 | -0.23 | 28.24 | -8.56 | -0.06 | -33.75 | 0 | -90.59 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0.06 | 0.12
57.088 | -21.27 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 14.41 | -1.7 | -0.06 | -9.43 | -0.06 | 0 | 0.23 | 21.39 | -0.41 | 28.48 | -0.35 | 28.3 | -0.35 | 28.59 | -6.86 | -0.06 | -32.23 | 0 | -58.48 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.18
57.12 | -19.45 | -0.06 | -47.23 | -0.18 | 0 | -0.06 | 0 | 0 | 21.45 | 0.41 | 20.39 | -0.94 | -0.06 | -7.79 | -0.18 | 0 | 0.12 | 21.39 | -0.23 | 28.18 | -0.35 | 28.13 | -0.35 | 28.89 | -4.98 | 5.68 | -29.77 | -0.06 | -49.04 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0.12
57.152 | -17.58 | -0.06 | -54.9 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -0.53 | -0.18 | -6.33 | -0.06 | -50.74 | 0.23 | 21.39 | -0.29 | 28.42 | -0.23 | 28.65 | -0.23 | 28.83 | -3.22 | 13.65 | -28.54 | -0.06 | -44.36 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.12
57.184 | -15.23 | -0.06 | -40.02 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -0.18 | -0.06 | -3.75 | -0.06 | -62.35 | 0.12 | 21.39 | -0.23 | 28.71 | -0.35 | 28.48 | -0.18 | 28.3 | -1.99 | 23.85 | -26.13 | -0.06 | -37.38 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0
57.216 | -14 | -0.06 | -39.32 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | 0.06 | 6.27 | -2.64 | -0.06 | -39.73 | 0.12 | 21.39 | -0.29 | 28.54 | -0.23 | 28.65 | -0.35 | 28.3 | -0.94 | 28.71 | -24.32 | -0.18 | -37.15 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12
57.248 | -12.13 | -0.06 | -37.03 | -0.18 | -0.18 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.33 | 0.23 | 11.89 | -1.99 | -0.06 | -27.83 | 0.18 | 0.23 | -0.23 | 28.77 | -0.18 | 28.54 | -0.18 | 28.48 | -0.23 | 28.71 | -22.33 | -0.06 | -37.68 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
57.28 | -11.02 | -0.06 | -35.39 | -0.06 | -56.55 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.35 | 17.52 | -1 | -0.06 | -30.94 | 0.12 | 0.06 | -0.23 | 28.48 | -0.23 | 28.42 | -0.18 | 28.59 | -0.06 | 28.48 | -20.39 | -0.06 | -37.44 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | 0
57.312 | -9.43 | -0.06 | -33.34 | 0 | -46.64 | -0.06 | -1.52 | 0.12 | 21.39 | 0.35 | 21.39 | 0.35 | 21.62 | -0.76 | -0.06 | -35.92 | 0.23 | 0 | -0.23 | 28.54 | -0.23 | 28.65 | -0.18 | 28.42 | 0.12 | 28.65 | -18.4 | -0.06 | -36.21 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
57.344 | -7.27 | 0.06 | -30.53 | -0.06 | -40.14 | -0.06 | -74.18 | 0.06 | 21.21 | 0 | 21.33 | 0.47 | 21.39 | -0.35 | -0.06 | -33.46 | 0.12 | 0.06 | -0.18 | 28.59 | -0.23 | 28.48 | -0.18 | 28.48 | 0.12 | 28.18 | -16.47 | -0.06 | -33.81 | -0.06 | -36.51 | -0.18 | 0 | 0 | 0 | 0.18
57.376 | -6.33 | -0.06 | -28.36 | -0.06 | -39.08 | -0.06 | -49.63 | -0.06 | 21.39 | 0 | 21.33 | 0.41 | 21.39 | -0.18 | -0.06 | -29.83 | 0.12 | 0 | -0.23 | 28.48 | -0.18 | 28.71 | -0.18 | 28.07 | 0.35 | 28.65 | -14.71 | 0 | -31.47 | -0.18 | -86.78 | -0.06 | 0 | 0 | -0.06 | 0.06
57.408 | -4.51 | 0 | -26.13 | -0.18 | -37.21 | -0.06 | -42.83 | -0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.33 | -0.18 | -0.06 | -28.01 | 0.12 | 0 | -0.23 | 28.54 | -0.18 | 28.77 | -0.06 | 28.54 | 0.35 | 29.01 | -12.19 | -0.06 | -29.65 | 0 | -55.43 | -0.18 | -0.06 | -0.18 | -75.24 | 0
57.44 | -3.46 | 0 | -24.55 | -0.06 | -34.98 | -0.06 | -37.21 | -0.06 | 21.39 | -0.06 | 21.45 | 0.53 | 21.39 | 0 | -0.06 | -27.01 | 0.12 | 0 | -0.23 | 28.59 | -0.18 | 18.75 | -0.18 | 28.65 | 0.47 | 28.59 | -10.37 | -0.06 | -28.07 | 0 | -33.34 | -0.06 | -79.28 | -0.06 | -70.96 | 0
57.472 | -2.23 | 0.59 | -22.38 | -0.06 | -32.23 | -0.06 | -32.93 | 0.06 | 21.45 | 0.12 | 21.33 | 0.53 | 21.39 | 0.06 | 5.1 | -25.37 | 0.06 | 0 | -0.18 | 28.59 | -0.06 | 0.12 | -0.06 | 28.24 | 0.47 | 28.65 | -8.14 | 0 | -26.6 | -0.06 | -31.76 | -0.18 | -72.19 | -0.06 | -34.69 | 0
57.504 | -1.11 | 10.43 | -19.86 | -0.06 | -29.77 | -0.06 | -32.58 | -0.06 | 21.33 | -0.06 | 21.33 | 0.41 | 21.39 | 0.06 | 12.95 | -23.44 | -0.06 | 0 | -0.18 | 0.23 | -0.18 | 0 | -0.18 | 28.36 | 0.41 | 28.48 | -5.39 | -0.06 | -24.32 | -0.06 | -43.07 | -0.18 | -35.57 | 0 | -29.83 | 0.12
57.536 | -0.53 | 15.76 | -17.93 | -0.06 | -28.13 | -0.06 | -32.29 | -0.06 | 0.12 | 0.12 | 21.09 | 0.41 | 21.33 | 0.06 | 18.34 | -21.39 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 0 | -0.06 | 28.42 | 0.35 | 28.83 | -3.69 | 0 | -22.74 | -0.06 | -41.84 | -0.06 | -30.24 | -0.06 | -40.02 | 0.06
57.568 | -0.18 | 21.45 | -16.47 | -0.18 | -26.72 | -0.06 | -31.47 | -0.06 | 0.06 | 0 | 21.39 | 0.41 | 21.39 | 0.06 | 21.62 | -20.51 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 27.95 | 0.35 | 28.42 | -1.88 | 3.34 | -20.63 | 0 | -36.68 | -0.18 | -40.14 | -0.06 | -42.89 | -0.06
57.6 | -0.18 | 21.5 | -14.59 | -0.06 | -23.97 | -0.06 | -29.94 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | 0.18 | 21.39 | -18.69 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.29 | 0.35 | 28.42 | -1.64 | 11.78 | -18.4 | 0 | -33.69 | -0.06 | -42.66 | 0 | -38.79 | 0
57.632 | -0.35 | 21.5 | -13.24 | -0.06 | -21.68 | 0 | -27.19 | -0.18 | 0 | 0 | 21.21 | 0.29 | 21.15 | 0.18 | 21.33 | -16.52 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0.12 | 0.35 | 28.71 | -0.59 | 19.22 | -16.7 | -0.06 | -32.93 | -0.06 | -38.5 | 0.06 | -33.63 | 0
57.664 | -0.29 | 21.5 | -11.31 | 0 | -19.28 | -0.06 | -24.26 | 0 | 0 | -0.06 | 21.27 | 0.12 | 21.39 | 0.12 | 21.45 | -14.82 | -0.06 | -17.29 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | 0.35 | 28.42 | -0.18 | 27.01 | -14.36 | -0.06 | -32.46 | -0.18 | -34.22 | -0.06 | -32.7 | -0.18
57.696 | -0.35 | 21.45 | -9.9 | -0.06 | -17.64 | -0.06 | -22.21 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.45 | 0.23 | 21.27 | -13.36 | 0 | -73.25 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.47 | 28.48 | -0.18 | 28.48 | -12.36 | -0.06 | -30.94 | -0.18 | -32.58 | -0.06 | -32.81 | -0.18
57.728 | -0.35 | 21.45 | -8.32 | -0.06 | -15.7 | -0.06 | -20.63 | 0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.15 | 0.23 | 21.21 | -11.31 | 0 | -51.62 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0.29 | 28.54 | 0.06 | 28.36 | -10.61 | -0.18 | -28.83 | -0.18 | -32.93 | -0.06 | -31.7 | 0
57.76 | -0.41 | 21.45 | -7.09 | -0.06 | -14.41 | -0.06 | -18.69 | -0.06 | 0 | -0.06 | 21.09 | 0.06 | 21.15 | 0.29 | 21.39 | -10.02 | 0.06 | -32.34 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.35 | 28.48 | 0.12 | 28.54 | -8.03 | -0.06 | -27.6 | -0.06 | -31.99 | -0.06 | -29.59 | -0.06
57.792 | -0.35 | 21.45 | -5.45 | 0 | -12.42 | 0.06 | -16.52 | -0.18 | 0 | -0.06 | 21.39 | 0.12 | 21.15 | 0.23 | 21.27 | -7.38 | -0.06 | -29.77 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | 0.12 | 28.65 | 0.41 | 28.65 | -5.86 | 0.06 | -26.02 | -0.18 | -30 | -0.06 | -27.36 | 0
57.824 | -0.18 | 21.45 | -3.57 | -0.18 | -10.96 | -0.06 | -14.41 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.21 | 0.12 | 21.39 | -5.98 | 0 | -34.86 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 28.42 | 0.29 | 28.24 | -3.4 | -0.06 | -23.91 | -0.06 | -27.36 | -0.06 | -26.72 | -0.06
57.856 | -0.23 | 21.45 | -2.4 | -0.06 | -9.67 | -0.06 | -11.95 | -0.06 | 0 | 0 | 21.21 | 0.06 | 21.39 | 0.18 | 21.04 | -4.45 | 0 | -35.27 | -0.06 | 0 | -0.06 | -0.12 | 0 | 0 | 0.12 | 28.65 | 0.35 | 28.54 | -2.17 | -0.18 | -22.44 | -0.18 | -26.02 | -0.06 | -25.08 | -0.06
57.888 | -0.18 | 21.39 | -1.41 | 5.45 | -8.14 | -0.06 | -10.2 | -0.18 | 0 | 0 | 21.39 | 0 | 21.27 | 0.23 | 21.27 | -3.11 | -0.06 | -32.05 | -0.06 | 0 | -0.06 | -72.89 | -0.18 | 0 | 0.12 | 28.65 | 0.29 | 28.48 | -1.29 | 0 | -20.92 | 0 | -25.08 | 0 | -23.15 | -0.18
57.92 | -0.23 | 21.45 | -0.7 | 11.48 | -6.91 | 3.93 | -7.97 | -0.06 | 0 | 0 | 3.93 | 0.06 | 21.33 | 0.12 | 21.27 | -2.23 | -0.06 | -28.54 | -0.06 | -7.15 | -0.06 | -62.64 | -0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.54 | -0.76 | 0 | -18.87 | 0 | -23.56 | 0 | -21.45 | -0.06
57.952 | -0.23 | 21.39 | -0.35 | 16.88 | -4.63 | 10.25 | -5.98 | -0.18 | -37.62 | -0.06 | 0.06 | 0.06 | 21.39 | 0.12 | 21.21 | -1.17 | -0.06 | -27.19 | -0.18 | -76 | -0.06 | -45.7 | -0.06 | 0 | 0.06 | 28.42 | 0.23 | 28.42 | -0.18 | 0 | -17.58 | 0 | -21.91 | 0 | -19.69 | -0.18
57.984 | -0.35 | 21.45 | -0.29 | 21.62 | -3.52 | 17.52 | -3.16 | -0.06 | -72.95 | 0 | 0 | 0 | 21.33 | 0.06 | 21.15 | -0.76 | -0.06 | -26.37 | -0.06 | -51.27 | 0 | -45.53 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.48 | -0.06 | 0 | -15.53 | 0 | -19.63 | -0.06 | -18.16 | -0.18
58.016 | -0.23 | 21.45 | -0.18 | 21.5 | -2.29 | 21.86 | -1.64 | -0.06 | -49.04 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -0.35 | 2.99 | -25.2 | 0 | -44.65 | -0.06 | -42.78 | -0.06 | -26.08 | 0.06 | 0.18 | 0.29 | 28.77 | 0.23 | 0 | -14.06 | -0.06 | -18.16 | 0 | -16.29 | -0.06
58.048 | -0.18 | 21.45 | -0.23 | 21.15 | -1.41 | 21.45 | -1.29 | -0.06 | -42.42 | -0.06 | 0 | 0 | 21.21 | 0 | 21.39 | -0.18 | 9.38 | -22.79 | -0.18 | -44.24 | -0.06 | -41.08 | -0.06 | -79.16 | -0.06 | 0 | 0.35 | 28.42 | 0.12 | 9.2 | -12.01 | 0 | -16.76 | -0.06 | -15.35 | 0
58.08 | -0.18 | 21.45 | -0.23 | 21.45 | -0.64 | 21.5 | -0.7 | -0.06 | -36.92 | 0.06 | 0 | 0 | 21.33 | 0.06 | 21.39 | -0.18 | 14.53 | -20.98 | 0 | -42.6 | 0 | -39.38 | -0.06 | -54.2 | -0.06 | 0 | 0.12 | 28.65 | 0.35 | 16.58 | -10.66 | 6.97 | -15.18 | -0.06 | -13.54 | -0.06
58.112 | -0.23 | 21.45 | -0.35 | 21.45 | -0.29 | 21.5 | -0.18 | -0.06 | -32.05 | 0 | 0 | -0.06 | 0.12 | 0 | 21.39 | 0.12 | 20.51 | -19.04 | -0.06 | -41.08 | -0.06 | -36.86 | 0 | -45.35 | -0.06 | 0 | 0.12 | 28.36 | 0.29 | 26.84 | -9.2 | 14.71 | -13.18 | 7.79 | -11.31 | 0.06
58.144 | -0.18 | 21.45 | -0.23 | 21.39 | -0.18 | 21.45 | -0.06 | -0.06 | -32.64 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.62 | -17.64 | -0.18 | -38.32 | -0.06 | -34.98 | 0 | -44.83 | -0.06 | 0 | 0.18 | 28.48 | 0.47 | 28.65 | -7.15 | 24.84 | -11.13 | 17.64 | -9.9 | 11.78
58.176 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.41 | -0.06 | -32.29 | 0 | 0 | 0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | -15.94 | -0.06 | -36.27 | 0.06 | -33.05 | 0.06 | -40.9 | 0 | 0 | 0.23 | 28.71 | 0.23 | 28.48 | -5.68 | 28.71 | -9.55 | 25.43 | -8.44 | 19.22
58.208 | -0.18 | 21.45 | -0.35 | 21.5 | -0.23 | 21.5 | 0.41 | -0.18 | -31.11 | 0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -13.83 | -0.06 | -34.63 | 0 | -31.52 | 0 | -41.25 | 0 | 0 | 0.18 | 28.48 | 0.41 | 28.59 | -3.63 | 28.71 | -8.03 | 28.59 | -6.56 | 27.07
58.24 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.29 | -0.06 | -29.24 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | -12.6 | 0 | -32.4 | -0.06 | -30 | 0 | -38.79 | 0 | 0 | 0.12 | 28.42 | 0.23 | 28.42 | -2.4 | 28.59 | -6.21 | 28.83 | -4.98 | 28.54
58.272 | -0.06 | 21.45 | -0.23 | 21.45 | -0.53 | 21.45 | 0.53 | -0.06 | -26.66 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | -10.37 | -0.06 | -30.53 | 0 | -26.66 | 0 | -37.38 | 0 | 0 | 0.12 | 28.42 | 0.29 | 28.83 | -2.17 | 28.65 | -4.75 | 28.65 | -3.16 | 28.36
58.304 | -0.18 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | 0.35 | 5.74 | -24.61 | 0 | 0 | -0.06 | 0.06 | -0.06 | 21.39 | 0.23 | 21.33 | -7.91 | 0 | -28.77 | 0 | -25.02 | 0 | -34.28 | 0 | 0 | 0 | 28.59 | 0.18 | 28.65 | -1.17 | 28.65 | -2.93 | 28.48 | -2.23 | 28.48
58.336 | -0.06 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.53 | 11.78 | -22.21 | 0.06 | -2.4 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | -6.56 | -0.18 | -26.25 | 0 | -22.97 | 0.06 | -32.93 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 28.48 | -0.35 | 28.65 | -2.17 | 28.48 | -1.41 | 28.42
58.368 | -0.06 | 0.23 | -0.29 | 21.45 | -0.23 | 21.39 | 0.41 | 17.17 | -20.45 | 0.06 | -75.35 | -0.06 | 0 | 0 | 0.12 | 0.23 | 21.33 | -5.39 | 4.34 | -24.32 | 0.06 | -20.92 | 0.06 | -30.35 | -0.06 | 0 | 0 | 28.54 | 0.18 | 28.42 | -0.53 | 28.42 | -1.23 | 28.42 | -1 | 28.42
58.4 | -0.18 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 21.74 | -18.16 | -0.06 | -57.31 | -0.06 | 0 | 0.06 | 0 | 0.35 | 21.39 | -3.75 | 12.89 | -22.56 | 0 | -19.4 | 0 | -29.06 | 0 | 0 | -0.06 | 28.42 | 0.06 | 28.42 | -0.06 | 28.65 | -0.82 | 28.71 | -0.59 | 28.48
58.432 | -0.18 | 0 | -0.23 | 21.45 | -0.29 | 21.39 | 0.47 | 21.5 | -15.94 | -0.06 | -42.25 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.21 | -2.23 | 20.22 | -20.45 | -0.06 | -16.58 | -0.06 | -26.66 | -0.18 | -40.78 | 0 | 28.48 | 0.06 | 28.65 | -0.06 | 28.83 | -0.53 | 28.42 | -0.06 | 28.36
58.464 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.35 | 21.5 | -14.36 | -0.06 | -36.45 | -0.18 | 0 | -0.12 | 0 | 0.23 | 21.45 | -1.7 | 29.12 | -17.75 | -0.06 | -15.18 | 0 | -24.08 | 0 | -73.25 | -0.06 | 3.34 | 0 | 28.65 | 0.06 | 28.42 | -0.35 | 28.42 | 0 | 28.42
58.496 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.35 | 21.39 | -11.78 | -0.06 | -32.64 | -0.06 | 0 | 0 | 0 | 0.18 | 21.45 | -0.7 | 28.18 | -16.23 | 0 | -14.06 | 0.06 | -22.68 | 0 | -51.1 | 0 | 0.12 | 0 | 28.59 | 0.12 | 28.42 | -0.12 | 28.59 | 0 | 28.42
58.528 | -0.06 | 0 | -0.23 | 21.39 | -0.23 | 21.45 | 0.29 | 21.39 | -10.2 | 0 | -32.23 | -0.06 | -42.13 | 0 | 0 | 0.12 | 21.15 | -0.76 | 28.65 | -14.82 | 2.29 | -12.25 | 0.06 | -20.39 | -0.06 | -47.76 | -0.06 | 0 | 0 | 28.59 | 0.23 | 28.42 | 0.06 | 28.83 | 0.23 | 28.48
58.56 | -0.06 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | 0.23 | 21.39 | -8.09 | 0 | -32.29 | -0.06 | -67.03 | 0 | 0 | 0.23 | 21.39 | -0.18 | 28.48 | -13.36 | 9.9 | -11.13 | 0 | -18.46 | 0.06 | -42.66 | 0 | 0.06 | 0.06 | 28.83 | 0 | 28.59 | -0.06 | 28.71 | 0.12 | 28.48
58.592 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.23 | 21.5 | -4.92 | 0.06 | -30.94 | -0.06 | -41.95 | -0.06 | 0 | 0.18 | 21.33 | -0.12 | 28.95 | -12.13 | 19.86 | -9.2 | 6.97 | -16.29 | -0.06 | -37.5 | 0.06 | 0 | 0 | 28.71 | 0.12 | 28.48 | 0 | 28.24 | 0 | 28.3
58.624 | 0 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | 0.18 | 21.39 | -2.99 | 0.06 | -29.01 | 0 | -27.48 | -0.06 | 0 | 0.12 | 21.33 | -0.06 | 28.65 | -10.14 | 27.31 | -7.91 | 17.17 | -15.23 | -0.06 | -37.62 | -0.18 | 0 | -0.06 | 28.13 | 0.18 | 28.48 | 0.23 | 28.42 | 0.29 | 28.3
58.656 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.12 | 21.39 | -1.88 | 0.06 | -27.25 | -0.06 | -32.34 | -0.06 | 0 | 0.12 | 21.33 | 0 | 28.59 | -9.14 | 28.54 | -6.45 | 25.02 | -13.65 | 0.06 | -37.91 | -0.06 | 0 | 0.06 | 28.36 | 0.12 | 28.54 | 0.12 | 28.59 | 0.23 | 28.07
58.688 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -0.76 | 2.99 | -25.02 | -0.06 | -35.63 | -0.06 | 0 | 0.23 | 21.39 | 0 | 28.18 | -7.91 | 28.71 | -5.57 | 28.83 | -11.54 | 0 | -36.51 | -0.18 | 0 | 0 | 28.65 | 0.06 | 28.48 | 0.12 | 28.3 | 0.29 | 28.42
58.72 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.06 | 21.39 | -0.7 | 9.32 | -23.32 | -0.06 | -34.1 | -0.18 | 0 | 0.18 | 21.21 | 0 | 28.48 | -6.45 | 28.59 | -4.51 | 28.83 | -10.61 | 0 | -35.16 | -0.06 | 0 | -0.06 | 28.36 | 0 | 28.59 | 0.18 | 28.59 | 0.23 | 28.36
58.752 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0 | 21.39 | -0.06 | 14.47 | -20.92 | -0.18 | -30.18 | 0 | 0 | 0.29 | 21.39 | 0.18 | 28.54 | -5.33 | 28.42 | -3.57 | 28.65 | -9.14 | 0 | -33.46 | -0.18 | 0 | -0.18 | 28.54 | 0 | 28.42 | 0.12 | 28.54 | 0.29 | 28.71
58.784 | 0 | -13.24 | -0.06 | 0.12 | -0.18 | 21.39 | 0.06 | 21.39 | 0.18 | 21.74 | -18.75 | -0.06 | -27.77 | 0 | -51.92 | 0.18 | 21.39 | 0.06 | 28.36 | -4.22 | 28.3 | -2.99 | 28.71 | -7.32 | 0.06 | -30.59 | 0 | 0 | 0 | 28.48 | 0 | 28.59 | 0.12 | 28.18 | 0.35 | 28.01
58.816 | -0.06 | -64.28 | -0.06 | 0 | -0.18 | 0.12 | 0 | 21.39 | 0.47 | 21.45 | -16.99 | -0.06 | -26.84 | 0.12 | -63.11 | 0.12 | 21.39 | 0.18 | 28.3 | -3.4 | 28.65 | -2.17 | 28.3 | -6.39 | -0.06 | -29.01 | 0 | 0 | -0.06 | 28.42 | 0.06 | 28.42 | 0.12 | 28.59 | 0.18 | 28.42
58.848 | -0.18 | -47.64 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.27 | 0.29 | 21.39 | -14.59 | -0.18 | -25.9 | 0 | -39.02 | 0.12 | 21.33 | 0.23 | 28.48 | -2.87 | 28.65 | -1.46 | 28.54 | -5.16 | 0 | -27.31 | 0 | 0 | 0.06 | 28.89 | 0 | 28.48 | 0.06 | 28.18 | 0.23 | 28.42
58.88 | -0.06 | -39.2 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.27 | 0.35 | 21.39 | -12.6 | -0.06 | -23.91 | -0.06 | -28.13 | 0.06 | 21.21 | 0.12 | 28.54 | -2.23 | 28.48 | -1.29 | 28.59 | -4.04 | 0 | -25.78 | 0 | -8.97 | -0.06 | 28.59 | -0.06 | 28.42 | 0.12 | 28.48 | 0.18 | 28.24
58.912 | -0.06 | -39.55 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.47 | 21.39 | -9.9 | 0 | -21.68 | 0 | -31.52 | 0.06 | 2.46 | 0.18 | 28.65 | -1.82 | 28.59 | -0.76 | 28.48 | -3.22 | 0 | -23.79 | -0.06 | -93.34 | 0 | 8.97 | -0.06 | 28.48 | 0.06 | 28.3 | 0.18 | 28.59
58.944 | -0.06 | -35.33 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.41 | 21.39 | -7.56 | -0.06 | -19.51 | -0.06 | -35.74 | 0.06 | 0.06 | 0.12 | 28.65 | -1.41 | 28.71 | -0.53 | 28.3 | -2.46 | -0.06 | -21.62 | -0.06 | -64.4 | -0.06 | 0.12 | -0.06 | 26.25 | 0.06 | 28.48 | 0.12 | 28.48
58.976 | -0.06 | -35.16 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -5.63 | 0.06 | -17.99 | 0 | -33.75 | 0.06 | 0 | 0.18 | 28.3 | -1 | 28.77 | -0.41 | 28.48 | -1.93 | 0 | -19.51 | -0.06 | -47.99 | 0 | 0 | 0 | 0.18 | 0.18 | 23.2 | 0 | 28.48
59.008 | -0.18 | -31.58 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.45 | -3.75 | 0.06 | -16.47 | -0.06 | -30.06 | 0.12 | 0 | 0.23 | 28.54 | -0.76 | 28.65 | -0.35 | 28.42 | -1.41 | -0.06 | -17.4 | -0.06 | -44.12 | 0.06 | 0 | -0.23 | 0 | 0.12 | 0.18 | 0 | 3.16
59.04 | -0.06 | -30 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.45 | -1.93 | 0.06 | -14.82 | 0 | -28.36 | 0.23 | 0 | 0.12 | 28.42 | -0.64 | 28.42 | -0.29 | 28.48 | -0.94 | 6.97 | -15.94 | -0.06 | -38.85 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0 | 0.12
59.072 | -0.06 | -28.48 | -0.06 | 0 | 0 | 0 | -0.06 | 21.45 | 0.29 | 21.45 | -1.05 | 0.7 | -12.83 | -0.06 | -27.19 | -0.06 | 0 | 0.12 | 28.48 | -0.47 | 28.89 | -0.29 | 28.48 | -0.76 | 14.71 | -13.77 | 0 | -37.15 | 0 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06 | 0
59.104 | -0.06 | -26.54 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -0.47 | 10.31 | -10.66 | -0.06 | -25.08 | -0.06 | 0 | 0.18 | 28.65 | -0.47 | 28.59 | -0.23 | 28.54 | -0.53 | 24.9 | -11.37 | 0.06 | -38.15 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0.06
59.136 | -0.06 | -23.44 | -0.06 | 0 | 0 | 0 | -0.06 | 21.5 | 0.29 | 21.39 | -0.18 | 15.7 | -9.32 | -0.06 | -23.5 | 0 | 0 | 0.18 | 28.24 | -0.47 | 28.54 | -0.29 | 28.3 | -0.35 | 28.71 | -9.43 | 0 | -37.85 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
59.168 | -0.06 | -21.39 | -0.06 | 0 | 0 | 0 | -0.18 | 21.39 | 0.35 | 21.33 | 0.06 | 21.56 | -7.44 | 4.98 | -21.62 | 0.06 | 0 | 0.23 | 28.59 | -0.35 | 28.42 | -0.29 | 28.48 | -0.35 | 28.54 | -6.68 | 0.06 | -35.92 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0
59.2 | -0.06 | -19.28 | -0.18 | -30.7 | 0 | 0 | -0.18 | 0.12 | 0.23 | 21.27 | 0.23 | 21.45 | -5.51 | 10.96 | -20.16 | 0 | 0 | 0.23 | 28.71 | -0.35 | 28.48 | -0.23 | 28.59 | -0.23 | 28.48 | -5.1 | 0 | -33.93 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
59.232 | -0.06 | -17.7 | -0.06 | -61.76 | -0.18 | -13.83 | -0.06 | 0 | 0.23 | 21.39 | 0.35 | 21.39 | -3.93 | 16.47 | -18.75 | 0.06 | 0 | 0 | 28.18 | -0.18 | 28.36 | -0.23 | 28.24 | -0.29 | 28.65 | -2.93 | 0 | -31.76 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
59.264 | -0.06 | -15.7 | -0.06 | -40.78 | -0.06 | -59.71 | -0.18 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -2.46 | 21.62 | -16.41 | 0.06 | 0 | 0.18 | 0.23 | -0.35 | 28.48 | -0.23 | 28.42 | -0.29 | 28.07 | -1.76 | 0 | -30 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
59.296 | -0.06 | -13.95 | -0.06 | -39.85 | -0.06 | -43.19 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | -1.46 | 21.39 | -14.82 | 0 | 0 | 0.18 | 0.06 | -0.35 | 28.48 | -0.18 | 28.42 | -0.23 | 28.48 | -0.76 | -0.06 | -28.13 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | -0.06 | 0
59.328 | -0.06 | -12.77 | -0.06 | -37.97 | -0.06 | -40.72 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -1 | 21.39 | -12.6 | -0.06 | -2.23 | 0.18 | 0 | -0.29 | 28.48 | -0.18 | 28.54 | -0.29 | 28.59 | -0.53 | 0 | -26.13 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0 | 0
59.36 | 0 | -11.07 | -0.06 | -36.33 | -0.06 | -38.67 | -0.06 | 0 | 0 | 21.45 | 0.53 | 21.39 | -0.41 | 21.39 | -11.66 | -0.06 | -71.43 | 0.18 | 0 | -0.23 | 28.48 | -0.18 | 28.48 | -0.18 | 28.42 | -0.18 | 0 | -24.26 | -0.06 | -75.3 | -0.18 | -0.06 | 0 | 0 | 0 | 0
59.392 | 0.53 | -9.84 | -0.06 | -34.69 | 0.06 | -35.63 | -0.18 | 0.06 | -0.06 | 21.5 | 0.41 | 21.33 | -0.41 | 21.39 | -9.84 | -0.06 | -57.48 | 0.12 | 0 | -0.29 | 28.42 | -0.18 | 28.42 | -0.18 | 28.95 | 0 | -0.06 | -22.44 | 0 | -61.47 | -0.06 | -45.7 | 0 | 0 | -0.18 | 0
59.424 | 10.2 | -7.68 | -0.06 | -30.59 | -0.06 | -34.51 | 0 | 0 | 0.12 | 21.39 | 0.41 | 21.33 | 0 | 21.33 | -7.79 | 0 | -30.18 | 0.23 | 0 | -0.23 | 0.76 | -0.18 | 28.42 | -0.18 | 28.65 | 0.12 | 1.7 | -20.92 | -0.18 | -47.76 | -0.18 | -72.54 | 0 | -80.57 | 0 | -2.64
59.456 | 15.82 | -6.33 | -0.06 | -29.53 | -0.06 | -31.17 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.39 | 0 | 21.33 | -5.57 | -0.06 | -28.95 | 0.23 | 0 | -0.23 | 0.12 | -0.18 | 25.96 | -0.18 | 28.59 | 0.23 | 11.13 | -19.16 | -0.18 | -45.76 | -0.18 | -43.71 | 0 | -70.73 | -0.18 | -88.54
59.488 | 21.5 | -4.98 | 0 | -27.42 | 0 | -29.83 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | 0.06 | 21.39 | -3.93 | -0.06 | -34.51 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.18 | -0.18 | 28.48 | 0.29 | 18.52 | -16.29 | 0 | -40.49 | -0.18 | -28.54 | 0 | -40.61 | -0.18 | -65.45
59.52 | 21.5 | -3.46 | -0.06 | -25.37 | -0.18 | -27.95 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0 | 21.33 | -2.81 | -0.18 | -35.33 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 26.19 | -14.53 | -0.06 | -36.86 | -0.06 | -35.98 | -0.06 | -28.59 | -0.06 | -37.5
59.552 | 21.5 | -2.46 | -0.06 | -22.85 | -0.18 | -25.72 | -0.18 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.23 | 21.33 | -1.99 | -0.06 | -32.76 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.71 | 0.29 | 28.83 | -12.36 | 0 | -37.79 | -0.06 | -43.24 | 0 | -36.97 | -0.06 | -29.06
59.584 | 21.45 | -1.05 | -0.06 | -20.22 | -0.06 | -23.09 | -0.06 | 0.12 | 0 | 0.12 | 0.23 | 21.15 | 0.23 | 21.39 | -1.35 | -0.06 | -28.54 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.35 | 28.65 | -10.72 | 0 | -38.09 | -0.06 | -39.08 | -0.06 | -42.25 | 0 | -40.9
59.616 | 21.45 | -0.64 | -0.06 | -18.4 | -0.18 | -21.09 | -0.18 | -41.37 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 21.39 | -0.35 | 0 | -27.31 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.23 | 28.42 | -8.32 | -0.18 | -36.27 | -0.06 | -35.22 | -0.06 | -39.14 | -0.18 | -41.78
59.648 | 21.45 | -0.23 | -0.06 | -16.88 | -0.06 | -19.1 | -0.06 | -68.79 | 0 | 0 | 0.12 | 21.39 | 0.12 | 21.45 | -0.18 | -0.06 | -27.01 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.35 | 28.42 | -5.57 | -0.06 | -34.28 | -0.06 | -32.93 | -0.06 | -34.75 | -0.06 | -38.67
59.68 | 21.5 | -0.23 | 0 | -15 | -0.06 | -16.99 | -0.18 | -47.93 | 0 | 0 | 0.06 | 21.27 | 0.12 | 21.33 | -0.18 | -0.06 | -25.43 | 0.18 | -11.07 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.29 | 28.48 | -3.57 | -0.06 | -32.29 | -0.06 | -32.7 | -0.06 | -32.76 | -0.06 | -34.22
59.712 | 21.45 | -0.23 | -0.06 | -13.71 | -0.06 | -15.41 | -0.18 | -41.37 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | 0 | -0.06 | -23.15 | 0.06 | -77.23 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.29 | 28.65 | -2.05 | 0.29 | -30.12 | -0.06 | -32.46 | -0.06 | -32.76 | -0.06 | -32.64
59.744 | 21.45 | -0.29 | 0 | -11.66 | -0.06 | -13.36 | -0.06 | -34.4 | -0.18 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | 0.06 | -0.06 | -20.86 | 0.12 | -49.34 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.65 | 0.35 | 28.48 | -1.64 | 12.13 | -28.48 | 0 | -30.24 | -0.06 | -32.05 | -0.06 | -32.7
59.776 | 21.45 | -0.35 | -0.06 | -10.14 | -0.06 | -11.72 | -0.06 | -31.82 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | 0.18 | -0.06 | -19.22 | 0.23 | -45.35 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.18 | 28.48 | -0.64 | 19.81 | -26.72 | -0.06 | -27.77 | 0 | -29.71 | -0.06 | -31.41
59.808 | 21.5 | -0.29 | 6.97 | -8.85 | -0.06 | -10.37 | -0.18 | -32.58 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | 0.06 | 0 | -17.99 | 0.06 | -43.89 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.65 | 0.18 | 28.42 | -0.18 | 27.6 | -25.31 | -0.06 | -26.66 | 0.12 | -27.6 | 0 | -29.12
59.84 | 21.39 | -0.35 | 12.48 | -7.56 | -0.06 | -9.08 | -0.06 | -32.4 | 0.06 | 0 | -0.06 | 21.21 | 0.12 | 21.39 | 0.29 | 4.45 | -15.94 | 0.06 | -41.78 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.65 | 0.12 | 28.48 | -0.06 | 28.77 | -23.5 | -0.06 | -25.61 | 0 | -26.84 | -0.06 | -27.36
59.872 | 21.5 | -0.29 | 18.22 | -5.92 | -0.06 | -7.79 | -0.06 | -31.23 | 0 | 0 | -0.18 | 21.39 | 0.12 | 21.21 | 0.23 | 12.25 | -13.95 | 0.23 | -40.78 | -0.06 | -78.99 | -0.18 | 0 | -0.18 | 28.71 | 0.18 | 28.59 | 0 | 28.71 | -21.09 | -0.18 | -23.61 | -0.06 | -25.14 | 0 | -26.49
59.904 | 21.39 | -0.23 | 21.5 | -3.87 | -0.06 | -5.63 | -0.18 | -28.48 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.18 | 17.7 | -12.01 | 0.06 | -37.44 | -0.06 | -56.9 | -0.06 | -58.95 | -0.18 | 25.49 | 0.18 | 28.42 | 0.12 | 28.48 | -19.16 | 0 | -21.91 | 0 | -23.32 | -0.18 | -24.55
59.936 | 21.45 | -0.29 | 21.45 | -3.11 | 0.41 | -3.98 | -0.18 | -25.9 | -0.06 | 0 | -0.18 | 21.39 | 0.12 | 21.33 | 0.23 | 21.5 | -10.31 | 0.12 | -35.63 | -0.06 | -45 | -0.06 | -63.52 | -0.18 | 0.18 | 0.06 | 28.54 | 0.12 | 28.65 | -17.23 | 0 | -20.39 | -0.06 | -21.86 | 0 | -22.91
59.968 | 21.45 | -0.18 | 21.39 | -1.93 | 8.14 | -2.99 | -0.18 | -23.91 | -0.06 | 0 | -0.18 | 0.23 | 0.06 | 21.39 | 0.18 | 21.45 | -8.91 | 0.12 | -34.69 | 0 | -45.53 | 0 | -45.59 | 0 | 0 | 0 | 28.42 | 0.35 | 28.65 | -14.82 | 0.06 | -18.63 | -0.18 | -20.16 | -0.06 | -21.97
60 | 21.45 | -0.23 | 21.5 | -1.17 | 13.71 | -2.05 | 0.41 | -22.21 | -0.06 | -6.86 | -0.06 | 0.06 | 0 | 21.33 | 0.29 | 21.39 | -6.68 | 0.06 | -32.76 | -0.06 | -43.13 | -0.06 | -45.7 | -0.06 | 0 | -0.18 | 28.42 | 0.23 | 28.59 | -13.36 | -0.06 | -17.4 | 0 | -18.69 | 0 | -19.98
60.032 | 21.45 | -0.23 | 21.45 | -0.59 | 21.39 | -1.17 | 7.97 | -20.27 | -0.06 | -74.65 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.27 | -5.45 | 0.18 | -29.94 | -0.06 | -40.78 | -0.18 | -44.24 | -0.18 | 0 | -0.06 | 28.48 | 0.29 | 28.54 | -10.66 | 2.4 | -15.59 | 0 | -17.23 | -0.06 | -18.4
60.064 | 21.45 | -0.18 | 21.45 | -0.18 | 21.5 | -0.41 | 15.23 | -17.75 | -0.06 | -50.39 | -0.06 | 0 | 0.06 | 0.12 | 0.23 | 21.45 | -3.22 | 0.12 | -27.6 | 0 | -38.56 | 0 | -40.9 | 0 | 0 | 0.06 | 28.54 | 0.23 | 28.48 | -8.85 | 12.6 | -13.24 | -0.06 | -14.71 | 0.23 | -16.41
60.096 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | -0.06 | 20.98 | -15.94 | -0.18 | -41.95 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -2.17 | 0.06 | -26.37 | 0 | -37.33 | -0.06 | -39.02 | -0.06 | 0 | 0.06 | 28.48 | 0.23 | 28.48 | -6.39 | 20.45 | -11.48 | 0.23 | -13.48 | 9.61 | -14.82
60.128 | 21.45 | -0.23 | 21.5 | -0.23 | 21.45 | -0.06 | 21.39 | -13.71 | -0.06 | -37.33 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.45 | -1.46 | 0.06 | -23.91 | -0.06 | -34.98 | 0 | -37.56 | -0.06 | 0 | -0.06 | 28.65 | 0.23 | 28.3 | -4.45 | 28.24 | -10.14 | 9.96 | -11.84 | 16.99 | -13.18
60.16 | 21.45 | -0.23 | 21.5 | -0.29 | 21.45 | -0.18 | 21.39 | -11.66 | 0.06 | -32.05 | -0.06 | 0 | 0.12 | 0.06 | 0.29 | 21.39 | -0.7 | 0.12 | -22.56 | 0 | -33.75 | 0 | -36.68 | 0 | 0 | -0.06 | 28.36 | 0.12 | 28.42 | -2.64 | 28.83 | -8.91 | 16.99 | -10.25 | 24.67 | -11.37
60.192 | 21.5 | -0.18 | 21.39 | -0.29 | 21.5 | -0.23 | 21.39 | -9.79 | -0.18 | -32.34 | 0 | 0 | -0.06 | 0 | 0.29 | 21.39 | -0.7 | 0.06 | -20.86 | 0.06 | -30.53 | 0 | -33.69 | -0.18 | 0 | -0.06 | 28.42 | 0.18 | 28.54 | -1.29 | 28.48 | -7.03 | 25.02 | -8.32 | 28.83 | -9.96
60.224 | 21.45 | -0.23 | 21.39 | -0.29 | 21.45 | -0.35 | 21.33 | -6.86 | 0 | -32.17 | -0.06 | 0 | 0 | 0 | 0.18 | 21.45 | -0.06 | 0.12 | -17.7 | -0.06 | -29.06 | 0 | -31.47 | -0.18 | 0 | 0 | 28.42 | 0.12 | 28.48 | -0.94 | 28.42 | -5.1 | 28.48 | -5.92 | 28.83 | -7.56
60.256 | 21.45 | -0.18 | 21.45 | -0.29 | 21.5 | -0.35 | 21.39 | -4.92 | 0.06 | -30.53 | -0.06 | 0 | 0.06 | 0 | 0.18 | 21.5 | -0.18 | 0.06 | -15.94 | 0 | -27.42 | 0 | -29.83 | -0.06 | 0 | 0 | 28.24 | 0.12 | 28.48 | -0.35 | 28.36 | -3.22 | 28.48 | -4.86 | 28.18 | -5.98
60.288 | 0.23 | -0.18 | 21.5 | -0.41 | 21.5 | -0.29 | 21.39 | -2.75 | 0 | -29.12 | -0.06 | 0 | -0.18 | 0 | 0.23 | 21.33 | 0.06 | -0.06 | -14.77 | 0 | -25.14 | 0 | -28.18 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.18 | -0.41 | 28.65 | -2.17 | 28.59 | -3.46 | 28.65 | -4.28
60.32 | 0.06 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.33 | -1.93 | 0 | -27.01 | -0.06 | 0 | 0 | 0 | 0.12 | 21.62 | 0.06 | 0 | -13.3 | 0 | -23.32 | 0 | -25.37 | -0.06 | 0 | -0.06 | 14.65 | 0.06 | 28.42 | 0.18 | 28.54 | -1.46 | 28.65 | -2.29 | 28.48 | -2.99
60.352 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -0.76 | -0.06 | -25.2 | 0 | 0 | 0 | 0 | 0.18 | 21.45 | 0.06 | 2.99 | -12.13 | 0.06 | -20.57 | -0.18 | -23.85 | -0.06 | -43.65 | 0 | 0.12 | 0.12 | 28.59 | 0.12 | 28.89 | -1 | 28.54 | -1.52 | 28.77 | -2.23
60.384 | 0 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.39 | -0.35 | 0.06 | -22.15 | 0 | -23.44 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 9.43 | -10.37 | -0.06 | -18.87 | 0 | -21.5 | -0.06 | -71.66 | -0.06 | 0 | -0.06 | 28.3 | 0.12 | 28.59 | -0.47 | 28.77 | -0.82 | 28.42 | -1.05
60.416 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.33 | -0.18 | -0.06 | -20.45 | -0.06 | -76.35 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 14.47 | -9.14 | -0.06 | -16.7 | 0 | -19.45 | -0.06 | -48.81 | -0.06 | 0.06 | 0 | 28.54 | 0.23 | 28.42 | -0.23 | 28.3 | -0.59 | 28.42 | -0.7
60.448 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.23 | 21.39 | 0.18 | 0 | -18.69 | -0.06 | -52.15 | 0 | 0 | 0.06 | 21.39 | 0.29 | 20.27 | -7.97 | -0.06 | -15.18 | -0.06 | -17.4 | -0.18 | -47.93 | 0 | 0 | -0.06 | 28.42 | 0.41 | 28.71 | -0.18 | 28.48 | -0.29 | 28.42 | -0.53
60.48 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.23 | 21.33 | 0.29 | -0.06 | -16.7 | -0.06 | -42.6 | -0.06 | -34.28 | 0.06 | 21.39 | 0.18 | 21.5 | -6.8 | 0 | -13.77 | -0.18 | -15.7 | 0 | -42.36 | 0 | 0 | -0.06 | 28.36 | 0.29 | 28.65 | 0 | 28.48 | -0.18 | 28.36 | -0.18
60.512 | 0 | -0.06 | 21.39 | -0.35 | 21.39 | -0.18 | 21.39 | 0.23 | 0 | -14.82 | -0.18 | -37.33 | -0.06 | -70.84 | 0.12 | 21.39 | 0.23 | 21.33 | -5.86 | -0.06 | -11.95 | 0.06 | -14.59 | 0 | -37.21 | 0 | 0 | -0.06 | 28.95 | 0.35 | 28.48 | 0.12 | 28.48 | -0.06 | 28.83 | -0.18
60.544 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.18 | 21.33 | 0.35 | -0.06 | -12.07 | -0.06 | -31.82 | 0.06 | -40.61 | 0.06 | 21.33 | 0.29 | 21.39 | -4.34 | -0.06 | -11.02 | 2.29 | -12.36 | -0.06 | -37.21 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.54 | 0.12 | 28.65 | 0 | 28.59 | 0.06
60.576 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.47 | 0.06 | -10.14 | 0 | -32.29 | -0.06 | -27.72 | 0.06 | 21.39 | 0.29 | 21.39 | -3.69 | -0.06 | -9.26 | 9.79 | -10.84 | -0.06 | -37.68 | 0 | 0 | -0.06 | 28.65 | 0.18 | 28.48 | 0.23 | 28.59 | 0.18 | 28.48 | 0
60.608 | 0 | -0.06 | 21.45 | -0.23 | 21.39 | -0.23 | 21.39 | 0.47 | -0.06 | -8.03 | -0.06 | -32.23 | -0.06 | -31.47 | 0.12 | 21.39 | 0.29 | 21.39 | -2.81 | 2.64 | -8.14 | 17.23 | -9.43 | 0 | -36.86 | -0.06 | 0 | -0.18 | 0.88 | 0.23 | 28.42 | 0.06 | 28.54 | 0.12 | 28.65 | 0.18
60.64 | 0 | -0.06 | 21.5 | -0.23 | 21.45 | -0.18 | 21.33 | 0.47 | 2.05 | -5.63 | 0.06 | -31 | -0.06 | -35.74 | 0.12 | 21.33 | 0.35 | 21.27 | -2.4 | 10.84 | -7.03 | 25.14 | -8.32 | -0.06 | -35.45 | -0.06 | 0 | 0 | 0.12 | 0.18 | 28.54 | 0.12 | 28.42 | 0 | 28.48 | 0.18
60.672 | 0 | -0.06 | 21.45 | -0.29 | 21.39 | -0.18 | 21.39 | 0.47 | 10.96 | -3.87 | -0.06 | -29.42 | -0.06 | -34.98 | -0.06 | 21.39 | 0.18 | 21.39 | -1.88 | 18.28 | -5.1 | 28.89 | -7.15 | -0.06 | -33.52 | -0.06 | 0 | 0 | 0 | 0.12 | 28.48 | 0.29 | 28.71 | 0.12 | 28.65 | 0.06
60.704 | -17.05 | -0.06 | 0.06 | -0.23 | 21.45 | -0.18 | 21.39 | 0.29 | 16.35 | -1.88 | -0.06 | -26.31 | -0.06 | -30.18 | 0.06 | 21.33 | 0.29 | 21.33 | -1.35 | 28.89 | -4.69 | 27.95 | -5.74 | 4.28 | -30.76 | -0.06 | 0 | 0 | 0.06 | 0.12 | 28.48 | 0.12 | 28.42 | 0.18 | 28.42 | 0.12
60.736 | -63.4 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.41 | 21.56 | -1.41 | -0.06 | -24.38 | -0.18 | -27.77 | 0 | 0.12 | 0.18 | 21.33 | -1.05 | 28.71 | -3.52 | 28.54 | -4.63 | 12.48 | -28.54 | -0.06 | -1.46 | -0.18 | 0 | 0 | 28.59 | 0.18 | 28.95 | 0.12 | 28.36 | 0.23
60.768 | -47.64 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 21.5 | -0.53 | -0.06 | -22.74 | -0.06 | -26.78 | -0.06 | 0 | 0.23 | 21.33 | -0.82 | 28.65 | -2.81 | 28.65 | -3.75 | 20.22 | -27.13 | 0 | -91.41 | 0 | 0 | 0.12 | 28.48 | 0.12 | 28.59 | 0.18 | 28.59 | 0.18
60.8 | -39.85 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 21.45 | -0.18 | -0.06 | -20.63 | -0.06 | -26.08 | -0.06 | 0 | 0.23 | 21.39 | -0.7 | 28.65 | -2.17 | 28.48 | -2.99 | 28.18 | -25.37 | -0.06 | -65.45 | 0 | 0 | 0 | 28.54 | 0.12 | 28.42 | 0.18 | 28.54 | 0.18
60.832 | -39.44 | -0.06 | 0 | -0.18 | 0.23 | -0.18 | 21.39 | 0.29 | 21.33 | 0 | -0.06 | -18.75 | -0.06 | -25.08 | 0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.48 | -1.7 | 28.59 | -2.46 | 28.83 | -23.67 | 0 | -48.4 | -0.06 | 0 | 0.06 | 28.59 | 0.06 | 28.89 | 0.06 | 28.42 | 0.23
60.864 | -35.39 | -0.06 | 0 | -0.06 | 0.06 | 0 | 21.39 | 0.23 | 21.39 | 0.18 | -0.06 | -16.29 | -0.06 | -21.33 | -0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.59 | -1.29 | 28.71 | -1.82 | 28.48 | -21.33 | -0.06 | -43.48 | 0 | 0 | 0 | 28.54 | 0 | 28.48 | 0.23 | 28.83 | 0.18
60.896 | -34.1 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.29 | 0.23 | 21.39 | 0.29 | -0.06 | -14.36 | -0.06 | -19.92 | -0.06 | 0 | 0.18 | 21.33 | -0.53 | 28.59 | -0.94 | 28.54 | -1.05 | 28.65 | -18.93 | -0.18 | -38.32 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.48 | 0.18 | 28.48 | 0.18
60.928 | -32.29 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.06 | 0.12 | 21.39 | 0.41 | -0.06 | -12.48 | -0.06 | -18.69 | -0.06 | 0 | 0.18 | 21.21 | -0.41 | 28.59 | -0.76 | 28.3 | -0.7 | 28.65 | -17.52 | -0.18 | -36.8 | -0.18 | 0 | 0 | 0.29 | -0.06 | 28.48 | 0.12 | 28.48 | 0.23
60.96 | -30.59 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.45 | 0.47 | -0.06 | -10.02 | -0.06 | -17.17 | -0.18 | 0 | 0.18 | 21.33 | -0.41 | 28.54 | -0.59 | 28.3 | -0.59 | 28.95 | -15.65 | -0.18 | -37.21 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 28.48 | 0.12 | 27.48 | 0.29
60.992 | -28.36 | 0 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.18 | 21.39 | 0.41 | -0.06 | -8.32 | 0.06 | -15.65 | -0.06 | 0 | 0.23 | 21.33 | -0.41 | 28.48 | -0.41 | 28.77 | -0.41 | 28.77 | -13.89 | -0.06 | -38.15 | 0 | 0 | -0.06 | 0 | -0.06 | 25.84 | 0.06 | 0.12 | 0.18
61.024 | -25.43 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 1.52 | -5.04 | 0 | -12.77 | -0.06 | 0 | 0.12 | 21.33 | -0.35 | 28.65 | -0.29 | 28.65 | -0.35 | 28.59 | -11.13 | -0.18 | -36.51 | -0.06 | -6.21 | -0.06 | 0.06 | -0.06 | 0.12 | 0.12 | 0 | 0.18
61.056 | -23.79 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | 9.02 | -3.34 | 0 | -11.07 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.54 | -0.41 | 28.42 | -0.23 | 28.71 | -8.5 | -0.06 | -34.51 | -0.06 | -93.28 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.23
61.088 | -21.45 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 14 | -1.99 | 0 | -9.84 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.95 | -0.29 | 28.65 | -0.29 | 28.59 | -6.91 | -0.06 | -32.23 | 0 | -61.64 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.18
61.12 | -19.63 | -0.06 | -41.54 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | 19.75 | -1.05 | -0.06 | -7.91 | 0 | 0 | 0.23 | 21.39 | -0.29 | 28.59 | -0.23 | 28.89 | -0.29 | 28.65 | -4.63 | 4.39 | -30.12 | -0.06 | -48.52 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0.12
61.152 | -17.46 | -0.06 | -56.55 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -0.53 | -0.18 | -5.92 | -0.18 | -44.47 | 0.18 | 21.39 | -0.23 | 28.36 | -0.29 | 28.42 | -0.35 | 28.59 | -2.87 | 12.66 | -27.89 | 0 | -45.29 | -0.06 | 0 | -0.12 | 0 | 0.06 | 0 | 0.06
61.184 | -15.47 | -0.06 | -40.61 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.21 | 0.29 | 21.45 | 0 | -0.06 | -3.93 | -0.18 | -65.1 | 0.06 | 21.39 | -0.29 | 28.18 | -0.18 | 28.89 | -0.23 | 28.54 | -1.46 | 23.03 | -26.31 | 0.06 | -37.44 | 0 | 0 | -0.18 | 0 | 0.06 | 0 | 0.12
61.216 | -14.06 | 0 | -39.44 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0 | 5.63 | -2.52 | -0.18 | -41.54 | 0.18 | 21.39 | -0.18 | 28.54 | -0.29 | 28.48 | -0.23 | 28.48 | -0.7 | 28.83 | -24.79 | 0 | -36.8 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12
61.248 | -12.31 | -0.06 | -37.09 | -0.06 | -0.18 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.23 | 11.48 | -2.05 | -0.18 | -28.18 | 0.18 | 3.34 | -0.23 | 28.65 | -0.18 | 28.42 | -0.18 | 28.36 | -0.53 | 28.65 | -22.38 | 0 | -37.91 | 0 | 0 | -0.18 | 0 | 0.06 | 0 | -0.06
61.28 | -10.96 | -0.06 | -35.63 | -0.06 | -52.62 | -0.06 | 0 | -0.06 | 21.27 | 0.23 | 21.33 | 0.23 | 16.99 | -1.17 | -0.06 | -30.53 | 0.06 | 0.12 | -0.23 | 28.65 | -0.18 | 28.42 | -0.29 | 28.65 | -0.06 | 28.54 | -20.51 | -0.06 | -37.38 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06
61.312 | -9.49 | -0.06 | -33.75 | -0.06 | -47.23 | -0.18 | -0.06 | 0 | 21.33 | 0.06 | 21.39 | 0.41 | 21.68 | -0.53 | -0.18 | -35.74 | 0.06 | 0 | -0.23 | 28.59 | -0.18 | 28.48 | -0.18 | 28.65 | 0 | 28.65 | -18.75 | 0 | -36.74 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0
61.344 | -7.62 | 0 | -30.7 | -0.06 | -40.08 | -0.06 | -77.35 | -0.06 | 21.39 | 0.06 | 21.33 | 0.53 | 21.33 | -0.35 | -0.06 | -34.1 | 0.12 | 0 | -0.29 | 28.42 | -0.18 | 28.36 | -0.18 | 28.77 | 0.06 | 28.42 | -16.58 | -0.06 | -33.99 | 0 | -23.09 | 0 | 0 | 0.06 | 0 | 0.06
61.376 | -6.15 | 0 | -28.89 | -0.18 | -39.38 | -0.18 | -52.44 | 0 | 21.39 | 0 | 21.39 | 0.53 | 21.39 | -0.12 | -0.06 | -30.24 | 0.06 | 0 | -0.23 | 28.42 | -0.18 | 28.54 | -0.06 | 28.48 | 0.41 | 28.65 | -14.71 | 0 | -31.76 | 0 | -90.53 | -0.18 | 0 | 0 | 0 | 0.06
61.408 | -4.75 | 0 | -27.25 | -0.06 | -36.86 | 0 | -42.89 | -0.06 | 21.39 | 0 | 21.39 | 0.35 | 21.33 | 0 | -0.06 | -28.01 | -0.06 | 0 | -0.18 | 28.54 | -0.18 | 28.54 | -0.18 | 28.54 | 0.29 | 28.65 | -12.77 | -0.06 | -30.18 | -0.06 | -59.06 | -0.06 | 0.12 | 0 | -70.78 | 0
61.44 | -3.52 | 0.06 | -24.43 | -0.06 | -35.45 | -0.06 | -37.38 | 0 | 21.5 | 0 | 21.39 | 0.47 | 21.33 | 0.18 | 0 | -27.31 | 0 | 0 | -0.23 | 28.89 | -0.18 | 25.61 | -0.06 | 28.48 | 0.23 | 28.24 | -10.49 | -0.06 | -28.01 | -0.06 | -34.86 | -0.18 | -74.83 | 0 | -71.55 | -0.06
61.472 | -2.46 | 2.11 | -23.32 | 0 | -32.23 | -0.06 | -32.99 | 0 | 21.15 | 0 | 21.39 | 0.41 | 21.39 | -0.06 | 4.16 | -25.72 | 0.06 | 0 | -0.06 | 28.65 | -0.18 | 0.12 | -0.23 | 28.48 | 0.29 | 28.54 | -8.56 | 0 | -26.72 | -0.06 | -30.18 | -0.18 | -72.83 | -0.06 | -36.1 | 0
61.504 | -1.17 | 9.49 | -19.81 | -0.06 | -30 | -0.06 | -32.23 | -0.06 | 21.27 | 0.06 | 21.45 | 0.41 | 21.39 | 0.18 | 11.95 | -23.79 | 0.06 | 0 | -0.18 | 0.29 | -0.23 | 0 | -0.06 | 28.48 | 0.41 | 28.65 | -5.57 | -0.06 | -24.14 | 0 | -42.6 | -0.18 | -36.86 | -0.18 | -30.06 | 0
61.536 | -0.53 | 15.18 | -18.05 | -0.06 | -28.54 | -0.06 | -31.76 | -0.06 | 0.12 | 0.06 | 21.39 | 0.35 | 21.39 | 0.23 | 17.52 | -21.62 | -0.06 | 0 | -0.23 | 0.12 | -0.18 | 0 | -0.06 | 28.42 | 0.35 | 28.59 | -3.75 | -0.06 | -22.62 | -0.18 | -42.01 | -0.06 | -30.24 | -0.18 | -39.79 | -0.18
61.568 | -0.18 | 20.74 | -16.47 | -0.18 | -26.66 | -0.06 | -31.23 | -0.18 | 0.06 | -0.06 | 21.45 | 0.29 | 21.39 | 0.23 | 21.56 | -19.92 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.48 | -2.05 | 0.94 | -21.09 | 0 | -37.15 | -0.18 | -39.9 | -0.06 | -42.42 | -0.06
61.6 | -0.18 | 21.5 | -14.65 | -0.06 | -24.08 | -0.06 | -29.94 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.35 | 21.33 | -18.87 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 0 | -0.18 | 0.29 | 0.29 | 28.13 | -1.7 | 10.78 | -18.87 | 0 | -34.22 | -0.06 | -42.42 | -0.06 | -39.08 | 0
61.632 | -0.23 | 21.5 | -13.3 | -0.06 | -21.56 | 0 | -27.95 | -0.06 | 0 | 0 | 21.27 | 0.29 | 21.39 | 0.35 | 21.45 | -16.76 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.35 | 28.48 | -0.7 | 18.22 | -16.99 | -0.18 | -32.93 | -0.06 | -38.85 | -0.06 | -34.63 | 0
61.664 | -0.35 | 21.45 | -10.96 | -0.06 | -19.45 | 0.06 | -24.84 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 21.33 | -14.94 | -0.06 | -6.5 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.54 | -0.23 | 25.96 | -14.36 | 0 | -32.76 | -0.06 | -33.99 | -0.06 | -32.93 | 0
61.696 | -0.35 | 21.45 | -9.67 | -0.06 | -17.58 | -0.06 | -22.62 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.29 | 21.33 | -13.3 | -0.06 | -70.08 | -0.18 | 0 | -0.06 | 0 | 0 | 0.06 | 0.35 | 28.36 | 0 | 28.59 | -12.6 | 0.06 | -31.11 | -0.06 | -33.34 | -0.18 | -32.93 | 0
61.728 | -0.35 | 21.45 | -8.56 | -0.06 | -16.11 | -0.06 | -20.98 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.35 | 21.33 | -11.84 | 0.06 | -54.44 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.29 | 28.65 | 0 | 28.48 | -10.55 | -0.18 | -28.71 | -0.18 | -33.05 | -0.06 | -31.76 | -0.18
61.76 | -0.35 | 21.45 | -6.97 | -0.18 | -14.41 | -0.06 | -18.93 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.56 | 0.35 | 21.33 | -10.08 | -0.06 | -35.45 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.23 | 28.48 | 0.12 | 28.42 | -8.56 | 0 | -27.54 | -0.06 | -31.99 | 0 | -29.36 | -0.06
61.792 | -0.29 | 21.45 | -5.51 | -0.18 | -12.77 | -0.06 | -17.05 | -0.06 | 0 | -0.06 | 21.45 | 0 | 21.39 | 0.18 | 21.33 | -7.97 | 0.06 | -29.06 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 28.48 | 0.23 | 28.71 | -6.62 | 0 | -26.08 | -0.06 | -30.24 | -0.06 | -27.07 | -0.06
61.824 | -0.18 | 21.5 | -3.75 | -0.06 | -11.37 | -0.06 | -14.47 | 0 | 0 | 0 | 21.39 | 0 | 21.45 | 0.35 | 21.09 | -6.21 | -0.06 | -34.22 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 28.48 | 0.29 | 28.3 | -3.87 | 0 | -23.85 | -0.06 | -26.9 | 0 | -26.13 | 0
61.856 | -0.23 | 21.45 | -2.4 | -0.06 | -9.43 | -0.06 | -12.48 | -0.06 | 0 | 0 | 21.39 | 0 | 21.33 | 0.06 | 21.09 | -4.45 | -0.06 | -35.22 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 0 | 0.06 | 28.48 | 0.23 | 28.42 | -1.93 | 0 | -22.85 | -0.18 | -25.78 | 0 | -25.31 | -0.06
61.888 | -0.18 | 21.45 | -1.41 | 4.63 | -8.32 | 0 | -10.2 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.39 | 0.23 | 21.21 | -3.22 | -0.06 | -33.28 | -0.18 | 0 | -0.06 | -63.75 | -0.06 | 0 | 0.12 | 28.65 | 0.35 | 28.36 | -1.46 | -0.06 | -21.09 | -0.06 | -24.84 | -0.06 | -23.5 | -0.06
61.92 | -0.18 | 21.39 | -0.82 | 10.9 | -6.86 | 3.75 | -8.14 | -0.06 | 0 | -0.18 | 11.37 | 0.12 | 21.39 | 0.06 | 21.39 | -1.76 | -0.06 | -29.47 | -0.06 | -1.23 | -0.06 | -65.39 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.42 | -0.94 | -0.06 | -18.87 | -0.06 | -23.32 | 0 | -21.68 | 0
61.952 | -0.29 | 21.45 | -0.64 | 16.06 | -4.92 | 10.02 | -5.98 | 0 | -35.63 | 0 | 0.06 | 0.12 | 21.39 | 0.12 | 21.39 | -1.23 | -0.06 | -27.36 | -0.18 | -69.91 | -0.06 | -44.3 | -0.06 | 0 | 0 | 28.65 | 0.35 | 28.42 | -0.18 | 0 | -17.7 | -0.18 | -21.86 | 0 | -19.75 | 0
61.984 | -0.29 | 21.45 | -0.23 | 21.62 | -3.87 | 17.23 | -3.22 | -0.06 | -73.48 | 0 | 0 | -0.06 | 21.33 | 0.06 | 21.33 | -0.76 | -0.06 | -26.9 | -0.06 | -52.97 | -0.06 | -45.47 | -0.06 | 0 | 0.06 | 28.48 | 0.18 | 28.54 | -0.18 | 0 | -15.41 | 0 | -19.69 | 0 | -18.16 | 0
62.016 | -0.23 | 21.39 | -0.29 | 21.39 | -2.52 | 21.74 | -1.88 | -0.06 | -49.22 | 0 | 0 | 0 | 21.39 | 0 | 21.33 | -0.35 | 2.46 | -25.37 | -0.18 | -44.53 | -0.06 | -42.78 | 0 | -15.53 | 0 | 0.29 | 0.29 | 28.65 | 0.18 | -0.06 | -13.77 | 0 | -18.52 | -0.06 | -16.23 | -0.06
62.048 | -0.18 | 21.45 | -0.35 | 21.45 | -1.64 | 21.62 | -1.23 | -0.18 | -41.78 | -0.18 | 0 | 0.12 | 21.15 | 0.06 | 21.33 | -0.23 | 8.56 | -22.97 | -0.18 | -45.53 | 0 | -41.13 | 0 | -79.05 | 0 | 0.06 | 0.23 | 28.59 | 0.18 | 9.2 | -12.42 | -0.06 | -16.7 | 0 | -15 | -0.06
62.08 | -0.18 | 21.45 | -0.41 | 21.45 | -0.82 | 21.45 | -0.64 | -0.06 | -36.45 | 0.06 | 0 | 0.06 | 21.39 | -0.06 | 21.33 | 0 | 13.83 | -20.92 | -0.06 | -42.95 | 0 | -38.67 | -0.06 | -56.9 | 0 | 0 | 0.23 | 28.71 | 0.23 | 16.11 | -10.84 | 6.09 | -14.77 | -0.06 | -13.48 | 0
62.112 | -0.18 | 21.45 | -0.47 | 21.45 | -0.29 | 21.45 | -0.18 | -0.06 | -32.4 | 0 | 0 | -0.06 | 0.12 | 0.06 | 21.33 | 0.12 | 19.4 | -19.28 | -0.06 | -40.37 | 0 | -37.33 | -0.06 | -45.35 | 0 | 0.06 | 0.12 | 28.42 | 0.47 | 26.43 | -9.38 | 13.95 | -13.24 | 6.33 | -11.02 | 0
62.144 | -0.29 | 21.39 | -0.41 | 21.45 | -0.18 | 21.5 | 0.12 | 0.12 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.06 | 21.45 | -17.64 | -0.06 | -38.56 | -0.06 | -35.57 | -0.06 | -44.42 | -0.06 | 0 | 0.18 | 28.42 | 0.35 | 28.77 | -6.91 | 24.2 | -11.13 | 16.64 | -9.84 | 10.72
62.176 | -0.23 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | 0.18 | 0 | -32.58 | -0.06 | 0 | -0.06 | 0 | 0 | 21.27 | 0.12 | 21.39 | -15.94 | -0.06 | -36.33 | -0.06 | -33.52 | -0.06 | -42.19 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.42 | -5.16 | 28.77 | -9.43 | 24.2 | -8.14 | 18.16
62.208 | -0.06 | 21.45 | -0.29 | 21.45 | -0.35 | 21.45 | 0.35 | 0 | -31.23 | 0 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.33 | -14.36 | -0.06 | -34.86 | 0 | -31.64 | 0 | -40.84 | 0 | 0 | 0.06 | 28.48 | 0.29 | 28.54 | -3.87 | 28.65 | -8.32 | 28.83 | -6.33 | 25.96
62.24 | -0.18 | 21.45 | -0.29 | 21.39 | -0.35 | 21.45 | 0.41 | 0.06 | -29.12 | 0.06 | 0 | 0 | 0 | 0 | 21.39 | 0.23 | 21.33 | -12.54 | 0 | -33.05 | -0.06 | -29.24 | -0.06 | -38.91 | 0 | 0 | 0.06 | 28.59 | 0.41 | 28.42 | -2.46 | 28.42 | -6.09 | 28.48 | -4.8 | 28.54
62.272 | -0.23 | 21.45 | -0.29 | 21.45 | -0.41 | 21.5 | 0.41 | 0.06 | -26.08 | 0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | -10.61 | -0.06 | -30.82 | -0.06 | -27.01 | 0 | -36.68 | 0 | 0 | 0 | 28.48 | 0.29 | 28.65 | -1.93 | 28.24 | -4.86 | 28.54 | -2.99 | 28.54
62.304 | -0.18 | 21.39 | -0.29 | 21.39 | -0.47 | 21.45 | 0.41 | 5.22 | -24.08 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.33 | -8.61 | -0.06 | -28.54 | 0 | -25.37 | 0 | -35.33 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.42 | -1 | 28.48 | -2.75 | 28.48 | -2.23 | 28.48
62.336 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | 0.47 | 11.25 | -22.09 | 0 | -0.82 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | -6.91 | -0.06 | -26.6 | -0.06 | -23.56 | 0.06 | -32.99 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.65 | -0.64 | 28.42 | -2.17 | 28.42 | -1.46 | 28.48
62.368 | -0.18 | 0.64 | -0.29 | 21.45 | -0.35 | 21.45 | 0.29 | 16.7 | -20.63 | -0.06 | -73.66 | -0.18 | 0 | -0.06 | 0.12 | 0.29 | 21.39 | -5.39 | 3.98 | -25.02 | 0.06 | -20.86 | -0.06 | -31.06 | -0.18 | 0 | 0.12 | 28.77 | 0.18 | 28.36 | -0.47 | 28.48 | -1.17 | 28.42 | -0.82 | 28.36
62.4 | -0.18 | 0.06 | -0.29 | 21.5 | -0.18 | 21.5 | 0.47 | 21.62 | -18.46 | -0.06 | -58.3 | -0.18 | 0 | 0 | 0 | 0.18 | 21.39 | -3.46 | 12.19 | -23.15 | -0.06 | -19.45 | -0.06 | -28.89 | 0 | 0 | -0.18 | 28.42 | 0.06 | 28.65 | -0.06 | 28.48 | -0.7 | 28.48 | -0.47 | 28.3
62.432 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.47 | 21.39 | -15.76 | 0.06 | -43.48 | -0.18 | 0 | 0 | 0 | 0.23 | 21.45 | -2.23 | 19.75 | -20.86 | 0 | -16.99 | 0 | -27.6 | 0 | -26.31 | 0.06 | 28.48 | 0.12 | 28.18 | -0.06 | 28.54 | -0.64 | 28.48 | -0.18 | 28.65
62.464 | -0.18 | 0 | -0.29 | 21.45 | -0.29 | 21.45 | 0.47 | 21.39 | -14 | -0.06 | -36.39 | -0.06 | 0 | 0 | 0 | 0.18 | 21.21 | -1.7 | 29.24 | -18.63 | 0.06 | -15.29 | 0.06 | -24.38 | -0.06 | -78.46 | 0 | 17.64 | -0.06 | 28.65 | -0.06 | 28.65 | -0.29 | 28.48 | -0.23 | 28.42
62.496 | -0.06 | 0 | -0.23 | 21.5 | -0.23 | 21.45 | 0.35 | 21.39 | -12.31 | -0.06 | -32.52 | -0.06 | -0.06 | -0.06 | 0 | 0.18 | 21.33 | -0.76 | 28.65 | -16.7 | -0.06 | -14.36 | -0.06 | -22.38 | -0.06 | -53.21 | 0 | 0.18 | -0.06 | 28.59 | 0.18 | 28.65 | -0.06 | 28.71 | -0.18 | 28.42
62.528 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.29 | 21.33 | -10.25 | -0.06 | -31.82 | 0.06 | -38.73 | 0.06 | 0 | 0.18 | 21.15 | -0.64 | 28.3 | -15.35 | 1.99 | -12.6 | -0.06 | -20.86 | -0.06 | -48.22 | 0 | 0 | 0.06 | 28.48 | 0.18 | 28.48 | -0.06 | 28.83 | 0.18 | 28.71
62.56 | -0.18 | 0 | -0.18 | 21.45 | -0.35 | 21.45 | 0.29 | 21.33 | -8.14 | 0 | -31.99 | -0.18 | -68.21 | 0.06 | 0 | 0.23 | 21.33 | -0.29 | 28.65 | -13.83 | 9.32 | -11.13 | 0.06 | -18.93 | -0.06 | -43.89 | 0 | 0.06 | 0.12 | 28.65 | 0.12 | 28.65 | 0.06 | 28.59 | 0.06 | 28.48
62.592 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.23 | 21.39 | -4.92 | 0 | -30.76 | -0.06 | -44.18 | -0.06 | 0 | 0.12 | 21.39 | -0.18 | 28.71 | -12.07 | 19.16 | -9.49 | 5.68 | -17.05 | 0 | -37.79 | -0.06 | 0 | -0.18 | 28.65 | 0.12 | 28.36 | 0.06 | 28.36 | 0 | 28.24
62.624 | -0.06 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | 0.18 | 21.21 | -2.99 | -0.06 | -29.59 | 0 | -27.01 | 0 | 0 | 0.18 | 21.39 | -0.06 | 28.65 | -10.55 | 26.72 | -8.14 | 16.11 | -14.82 | 0 | -37.15 | 0 | 0 | -0.06 | 28.59 | 0.12 | 28.42 | 0 | 28.59 | 0.12 | 28.65
62.656 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.23 | 21.39 | -1.88 | -0.06 | -27.36 | -0.06 | -32.46 | 0 | 0 | 0.12 | 21.45 | 0.12 | 28.59 | -9.14 | 28.83 | -6.91 | 23.79 | -13.59 | -0.06 | -38.44 | 0 | 0 | 0 | 28.42 | 0.18 | 28.59 | 0.12 | 28.48 | 0.18 | 28.59
62.688 | -0.06 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | 0 | 21.39 | -1.23 | 0.88 | -25.55 | -0.06 | -35.98 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 28.54 | -7.97 | 28.65 | -5.63 | 28.71 | -12.36 | -0.06 | -37.79 | -0.06 | 0 | -0.18 | 28.71 | 0.06 | 28.48 | 0.29 | 28.18 | 0.18 | 28.42
62.72 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0 | 21.39 | -0.23 | 8.79 | -23.67 | -0.06 | -33.69 | -0.06 | 0 | 0.29 | 21.39 | 0.06 | 28.65 | -6.68 | 28.65 | -4.8 | 28.65 | -10.84 | 0 | -36.27 | -0.06 | 0 | 0.06 | 28.42 | 0.06 | 28.42 | 0.23 | 28.18 | 0.18 | 28.48
62.752 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -0.18 | 14.12 | -20.74 | 0 | -30.12 | 0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.42 | -5.33 | 28.42 | -3.63 | 28.59 | -9.08 | 0 | -33.17 | -0.06 | 0 | -0.06 | 28.65 | 0 | 28.59 | 0.23 | 28.42 | 0.12 | 28.48
62.784 | -0.06 | -3.52 | -0.06 | 0.12 | -0.23 | 21.39 | -0.06 | 21.39 | 0.35 | 21.56 | -18.75 | 0 | -27.6 | -0.06 | -43.89 | 0.18 | 21.33 | 0.29 | 28.89 | -4.39 | 28.48 | -2.99 | 28.71 | -7.44 | -0.18 | -31.17 | 0 | 0 | 0 | 28.54 | 0 | 28.48 | 0.18 | 28.54 | 0.12 | 28.48
62.816 | -0.06 | -62.11 | -0.06 | 0 | -0.06 | 0.18 | 0 | 21.33 | 0.29 | 21.39 | -17.29 | -0.06 | -27.07 | -0.18 | -65.51 | 0.18 | 21.39 | 0.29 | 28.83 | -3.52 | 28.3 | -2.17 | 28.83 | -6.5 | 0 | -28.83 | -0.06 | 0 | 0 | 28.48 | -0.06 | 28.18 | 0.12 | 28.77 | 0.12 | 28.42
62.848 | -0.06 | -50.74 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.29 | 21.39 | -15.18 | -0.06 | -25.78 | -0.06 | -41.84 | 0.23 | 21.39 | 0.29 | 28.65 | -2.93 | 28.54 | -1.7 | 28.65 | -5.33 | -0.06 | -27.36 | 0 | 0 | 0 | 28.54 | 0.06 | 28.65 | 0.18 | 28.48 | 0.12 | 28.36
62.88 | -0.06 | -40.31 | -0.18 | 0.06 | -0.18 | 0 | 0 | 21.39 | 0.35 | 21.39 | -12.95 | -0.06 | -23.56 | -0.06 | -27.83 | 0.12 | 21.39 | 0.29 | 28.59 | -2.4 | 28.59 | -1.29 | 28.65 | -4.34 | 0 | -25.67 | -0.06 | -1.23 | 0 | 28.48 | 0.06 | 28.13 | 0.12 | 28.48 | 0.12 | 28.36
62.912 | -0.06 | -39.61 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.45 | 0.35 | 21.39 | -10.37 | 0 | -21.68 | -0.06 | -30.94 | 0.18 | 13.54 | 0.29 | 28.18 | -1.82 | 28.71 | -0.76 | 28.77 | -3.28 | -0.06 | -23.91 | -0.06 | -87.84 | -0.18 | 13.95 | 0 | 28.48 | 0.06 | 28.48 | 0.18 | 28.48
62.944 | 0 | -35.74 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.53 | 21.39 | -8.2 | 0.12 | -19.98 | -0.06 | -35.63 | 0.12 | 0.06 | 0.23 | 28.48 | -1.29 | 28.36 | -0.7 | 28.42 | -2.75 | -0.06 | -21.09 | 0 | -68.56 | 0 | 0.12 | 0 | 28.59 | 0.06 | 28.48 | 0.12 | 28.42
62.976 | -0.06 | -34.63 | -0.18 | 0.06 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.39 | -6.15 | -0.06 | -18.46 | -0.06 | -33.93 | 0.06 | 0 | 0.23 | 28.71 | -1.05 | 28.59 | -0.29 | 28.54 | -2.11 | -0.06 | -19.4 | -0.23 | -48.58 | -0.06 | 0 | 0 | 0.23 | 0 | 28.48 | 0.12 | 28.48
63.008 | 0 | -32.93 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.53 | 21.39 | -3.93 | -0.06 | -16.88 | -0.06 | -30.24 | 0.12 | 0 | 0.23 | 29.12 | -0.94 | 28.54 | -0.23 | 28.36 | -1.46 | -0.06 | -17.64 | -0.18 | -44.24 | 0 | 0 | -0.06 | 0 | 0.06 | 0.18 | 0.06 | 17.93
63.04 | -0.06 | -29.77 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.27 | 0.47 | 21.39 | -1.99 | 0.06 | -15 | 0 | -27.95 | 0.06 | 0 | 0.18 | 28.54 | -0.7 | 28.54 | -0.29 | 28.48 | -0.94 | 5.68 | -16.47 | -0.18 | -39.2 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.18 | 0.18
63.072 | -0.06 | -29.01 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.27 | 0.47 | 21.39 | -1.41 | 2.23 | -12.89 | -0.18 | -26.72 | 0.06 | 0 | 0.23 | 28.65 | -0.47 | 28.48 | -0.35 | 28.48 | -0.76 | 13.71 | -14.3 | -0.18 | -36.8 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
63.104 | -0.18 | -25.9 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.41 | 21.39 | -0.7 | 9.55 | -10.9 | 0 | -26.31 | 0.06 | 0 | 0.35 | 28.65 | -0.29 | 28.59 | -0.29 | 28.54 | -0.59 | 23.97 | -11.48 | 0 | -37.79 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0.06
63.136 | -0.18 | -23.85 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | -0.18 | 14.77 | -9.08 | -0.18 | -24.08 | 0.06 | 0 | 0.12 | 28.54 | -0.47 | 28.83 | -0.35 | 28.48 | -0.41 | 28.71 | -9.61 | 0 | -38.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
63.168 | -0.06 | -21.86 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.35 | 21.27 | -0.18 | 20.45 | -7.62 | 4.22 | -21.68 | 0.06 | 0 | 0.18 | 28.65 | -0.41 | 28.59 | -0.29 | 28.18 | -0.41 | 28.48 | -7.21 | -0.06 | -35.92 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
63.2 | -0.06 | -20.1 | -0.18 | -24.14 | -0.18 | 0 | -0.18 | 0.12 | 0.35 | 21.27 | 0.29 | 21.39 | -5.92 | 10.2 | -20.45 | 0.12 | 0 | 0.18 | 28.54 | -0.35 | 28.48 | -0.29 | 28.77 | -0.23 | 28.59 | -5.22 | -0.06 | -33.63 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0 | 0
63.232 | 0 | -17.58 | 0 | -63.28 | 0 | -8.61 | -0.18 | 0.06 | 0.12 | 21.39 | 0.41 | 21.39 | -4.16 | 15.65 | -19.4 | 0.06 | 0 | 0.12 | 28.48 | -0.29 | 28.59 | -0.29 | 28.59 | -0.23 | 28.48 | -3.4 | -0.06 | -31.99 | 0.06 | 0 | -0.18 | 0 | 0 | 0 | -0.06 | 0
63.264 | -0.06 | -15.76 | -0.06 | -43.13 | -0.18 | -61.23 | -0.18 | 0 | 0.23 | 21.39 | 0.35 | 21.39 | -2.29 | 21.74 | -16.7 | 0.06 | 0 | 0.23 | 0.23 | -0.29 | 28.59 | -0.23 | 28.18 | -0.23 | 28.42 | -1.99 | -0.06 | -30.53 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
63.296 | -0.06 | -14.06 | -0.06 | -40.08 | -0.06 | -43.65 | -0.18 | 0 | 0.18 | 21.45 | 0.53 | 21.39 | -1.46 | 21.45 | -15 | 0 | 0 | 0.23 | 0.06 | -0.35 | 28.48 | -0.18 | 28.48 | -0.29 | 28.65 | -1.05 | -0.18 | -28.65 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
63.328 | -0.06 | -12.77 | -0.06 | -37.85 | -0.06 | -40.37 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -1.17 | 21.33 | -13.13 | 0.06 | -0.12 | 0.23 | 0 | -0.29 | 28.42 | -0.18 | 28.48 | -0.18 | 28.59 | -0.59 | -0.18 | -26.02 | 0 | 0 | -0.18 | 0 | 0 | 0 | -0.06 | 0
63.36 | -0.06 | -11.07 | -0.06 | -35.63 | -0.06 | -38.85 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -0.41 | 21.39 | -11.48 | -0.06 | -65.74 | 0.23 | 0 | -0.23 | 28.48 | -0.23 | 28.59 | -0.18 | 28.59 | -0.23 | 0 | -24.61 | 0 | -69.67 | -0.18 | 0 | -0.06 | 0 | 0 | 0
63.392 | 2.11 | -9.84 | -0.06 | -34.63 | -0.06 | -35.86 | -0.18 | 0 | 0.06 | 21.39 | 0.41 | 21.39 | -0.35 | 21.33 | -10.08 | 0 | -59.94 | 0.23 | 0 | -0.18 | 28.42 | -0.18 | 28.42 | -0.29 | 28.48 | -0.06 | -0.06 | -22.68 | 0 | -62.99 | -0.18 | -40.08 | 0 | 0 | 0 | 0
63.424 | 9.43 | -7.62 | -0.06 | -31.64 | -0.06 | -33.87 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.27 | -0.06 | 21.39 | -7.44 | 0 | -31.99 | 0.29 | 0 | -0.23 | 9.08 | -0.18 | 28.65 | -0.23 | 28.59 | 0.18 | 2.64 | -20.86 | -0.06 | -48.28 | -0.06 | -74.42 | -0.06 | -67.74 | -0.18 | -0.12
63.456 | 14.88 | -6.45 | 0 | -29.42 | -0.06 | -31.76 | -0.18 | 0 | 0 | 21.27 | 0.41 | 21.33 | 0 | 21.39 | -5.98 | 0 | -28.89 | 0.18 | 0 | -0.18 | 0.12 | -0.18 | 28.65 | -0.18 | 28.48 | 0.06 | 10.9 | -18.69 | 0 | -45.12 | 0 | -47.11 | -0.06 | -75.76 | -0.06 | -82.5
63.488 | 20.63 | -4.98 | 0.06 | -27.31 | -0.06 | -30 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.15 | 0.06 | 21.39 | -4.28 | 0.06 | -34.45 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0.18 | -0.18 | 28.3 | 0.23 | 18.22 | -16.52 | 0 | -40.14 | 0 | -29.01 | -0.18 | -44.71 | 0 | -70.26
63.52 | 21.5 | -3.63 | -0.06 | -25.78 | -0.06 | -27.77 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | 0.12 | 21.33 | -2.87 | 0.06 | -35.39 | 0.23 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.71 | 0.35 | 26.25 | -14.53 | 0 | -36.92 | -0.18 | -35.27 | -0.06 | -29.42 | 0 | -40.37
63.552 | 21.56 | -2.52 | -0.06 | -23.32 | -0.06 | -26.31 | -0.06 | 0 | 0.06 | 21.5 | 0.35 | 21.39 | 0.12 | 21.33 | -2.05 | 0.06 | -33.4 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.53 | 28.65 | -13.18 | 0 | -37.91 | -0.06 | -43.07 | -0.06 | -35.16 | 0 | -28.71
63.584 | 21.45 | -1.29 | -0.06 | -20.92 | 0 | -23.09 | 0 | 0 | 0 | 0.12 | 0.29 | 21.45 | 0.29 | 21.33 | -1.23 | -0.06 | -28.71 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.35 | 28.3 | -10.78 | -0.06 | -37.74 | -0.23 | -38.97 | -0.06 | -42.13 | 0 | -39.61
63.616 | 21.45 | -0.82 | 0 | -18.69 | -0.06 | -21.27 | -0.18 | -32.81 | -0.06 | 0.06 | 0.29 | 21.39 | 0.35 | 21.45 | -0.29 | -0.18 | -27.48 | 0.12 | 0 | -0.18 | 0 | -0.23 | 0 | -0.18 | 28.59 | 0.53 | 28.59 | -8.91 | 0 | -36.51 | 0 | -35.92 | -0.06 | -39.61 | -0.18 | -41.49
63.648 | 21.45 | -0.29 | 0 | -17.05 | -0.06 | -19.16 | -0.18 | -71.49 | 0 | 0 | 0.29 | 21.33 | 0.23 | 21.5 | -0.29 | 0 | -26.84 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.35 | 28.48 | -5.86 | -0.06 | -34.63 | 0 | -33.22 | -0.06 | -34.98 | -0.06 | -38.67
63.68 | 21.45 | -0.18 | 0.06 | -15.18 | -0.06 | -17.17 | -0.06 | -49.69 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -0.23 | -0.06 | -25.02 | 0.18 | -6.45 | -0.06 | 0 | -0.18 | 0 | -0.23 | 28.48 | 0.41 | 28.59 | -4.28 | 0 | -32.17 | -0.06 | -33.11 | 0 | -32.81 | -0.18 | -34.98
63.712 | 21.5 | -0.18 | 0 | -13.77 | -0.06 | -15.7 | -0.18 | -42.31 | -0.18 | 0 | 0.12 | 21.15 | 0.18 | 21.39 | 0 | 0 | -23.38 | 0.18 | -76.94 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.42 | -2.23 | 1.99 | -30.29 | -0.06 | -32.64 | 0.06 | -32.34 | 0 | -33.4
63.744 | 21.45 | -0.47 | 0 | -11.72 | -0.06 | -13.83 | -0.18 | -34.22 | -0.06 | 0 | 0.06 | 21.27 | 0.18 | 21.21 | 0.06 | -0.06 | -20.92 | 0.12 | -49.28 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.23 | 28.77 | -1.41 | 11.48 | -28.59 | -0.06 | -30.18 | -0.06 | -31.93 | -0.18 | -32.7
63.776 | 21.45 | -0.35 | 1.99 | -10.61 | -0.06 | -12.01 | -0.18 | -31.82 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -0.06 | -0.06 | -19.63 | 0.06 | -44.88 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.23 | 28.48 | -0.47 | 19.1 | -26.43 | -0.06 | -27.77 | 0 | -30.47 | 0 | -31.58
63.808 | 21.39 | -0.41 | 6.68 | -9.02 | 0 | -10.78 | -0.06 | -32.81 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | 0.18 | -0.06 | -18.28 | 0.12 | -44.3 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.65 | -0.41 | 26.78 | -25.2 | -0.18 | -26.72 | 0 | -28.3 | 0 | -29.3
63.84 | 21.39 | -0.29 | 12.42 | -8.03 | 0 | -8.73 | -0.18 | -32.64 | -0.18 | 0 | 0 | 21.27 | 0.12 | 21.33 | 0.29 | 4.22 | -16.11 | 0 | -42.78 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.18 | 28.42 | -0.06 | 28.65 | -22.91 | 0 | -25.67 | -0.06 | -27.01 | -0.18 | -27.31
63.872 | 21.45 | -0.47 | 17.93 | -6.15 | -0.06 | -8.09 | -0.06 | -30.94 | -0.06 | 0 | 0.06 | 21.33 | 0.06 | 21.33 | 0.23 | 12.13 | -14.59 | 0.12 | -40.78 | -0.06 | -76.88 | -0.06 | 0 | -0.18 | 28.71 | 0.12 | 28.48 | 0.12 | 28.65 | -20.92 | -0.18 | -24.08 | 0 | -25.67 | -0.06 | -26.43
63.904 | 21.45 | -0.29 | 21.62 | -4.04 | 0 | -5.86 | 0 | -28.36 | 0 | 0 | 0 | 21.45 | 0.12 | 21.27 | 0.18 | 17.58 | -12.36 | 0.12 | -38.15 | -0.06 | -59.59 | -0.06 | -46.7 | -0.06 | 28.65 | 0 | 28.42 | 0.35 | 28.65 | -18.93 | -0.18 | -21.68 | -0.18 | -23.32 | -0.06 | -25.31
63.936 | 21.45 | -0.23 | 21.39 | -2.99 | 0.35 | -4.45 | -0.06 | -26.49 | 0 | 0 | 0 | 21.39 | 0.06 | 21.15 | 0.18 | 21.62 | -10.37 | 0.06 | -35.63 | -0.06 | -45.59 | -0.06 | -67.03 | -0.06 | 0.23 | 0.06 | 28.36 | 0.23 | 28.65 | -17.23 | 0.06 | -20.57 | 0.06 | -21.97 | 0 | -22.97
63.968 | 21.45 | -0.23 | 21.45 | -1.88 | 7.85 | -3.34 | -0.06 | -24.08 | -0.06 | 0 | 0 | 0.59 | 0.06 | 21.39 | 0.29 | 21.39 | -9.38 | 0 | -34.69 | -0.06 | -45.47 | -0.06 | -47.99 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.42 | -15.29 | -0.06 | -18.87 | 0 | -20.27 | 0.06 | -21.45
64 | 21.45 | -0.35 | 21.39 | -1.23 | 13.54 | -2.05 | -0.06 | -22.38 | 0 | -2.64 | 0.12 | 0.06 | 0.06 | 21.39 | 0.23 | 21.33 | -7.27 | 0.06 | -32.23 | -0.06 | -43.65 | -0.18 | -46.47 | -0.06 | 0 | 0 | 28.36 | 0.47 | 28.77 | -12.95 | -0.18 | -17.17 | -0.06 | -19.28 | 0 | -20.1
64.032 | 21.45 | -0.29 | 21.45 | -0.59 | 21.04 | -1.29 | 7.09 | -20.22 | -0.06 | -70.78 | -0.06 | 0 | -0.06 | 21.45 | 0.29 | 21.39 | -5.8 | -0.06 | -30.59 | -0.18 | -40.67 | -0.06 | -43.89 | -0.06 | 0 | 0.12 | 28.65 | 0.29 | 28.42 | -10.96 | 0 | -16 | 0.06 | -17.58 | -0.18 | -18.69
64.064 | 21.45 | -0.18 | 21.39 | -0.29 | 21.5 | -0.47 | 14.41 | -17.58 | 0 | -51.27 | -0.06 | 0 | -0.06 | 0.18 | 0.29 | 21.39 | -3.87 | 0.12 | -27.83 | 0 | -38.44 | 0 | -41.25 | -0.06 | 0 | 0.06 | 28.59 | 0.35 | 28.59 | -8.67 | 11.66 | -13.24 | -0.18 | -15.18 | -0.06 | -16.41
64.096 | 21.45 | -0.23 | 21.45 | -0.23 | 21.5 | -0.06 | 20.45 | -15.7 | 0.06 | -42.07 | -0.18 | 0 | -0.06 | 0 | 0.35 | 21.39 | -2.46 | 0.06 | -26.13 | 0 | -36.27 | -0.06 | -39.85 | -0.18 | 0 | 0 | 28.65 | 0.35 | 28.59 | -6.39 | 19.22 | -11.78 | -0.06 | -13.54 | 9.38 | -14.77
64.128 | 21.39 | -0.23 | 21.39 | -0.23 | 21.45 | -0.06 | 21.39 | -13.83 | 0 | -37.56 | -0.06 | 0 | 0.06 | 0 | 0.23 | 21.39 | -1.99 | 0 | -24.2 | -0.06 | -35.16 | -0.18 | -37.91 | -0.06 | 0 | 0.06 | 28.42 | 0.29 | 28.42 | -4.45 | 26.84 | -10.08 | 8.97 | -11.72 | 16.64 | -13.18
64.16 | 21.5 | -0.23 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | -12.07 | -0.06 | -32.93 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | -1.05 | -0.06 | -22.85 | 0 | -33.28 | 0 | -36.27 | -0.06 | 0 | 0.06 | 28.36 | 0.23 | 28.83 | -2.58 | 28.65 | -8.67 | 16.11 | -10.55 | 24.32 | -11.72
64.192 | 21.5 | -0.18 | 21.45 | -0.47 | 21.5 | -0.29 | 21.39 | -9.67 | -0.06 | -32.64 | 0 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.41 | -0.06 | -20.74 | 0 | -31.64 | 0 | -34.28 | 0 | 0 | -0.06 | 28.65 | 0.12 | 28.42 | -1.41 | 28.59 | -6.86 | 23.67 | -8.38 | 28.48 | -9.79
64.224 | 21.39 | -0.23 | 21.45 | -0.35 | 21.39 | -0.35 | 21.39 | -6.62 | 0.06 | -32.58 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.45 | -0.18 | -0.18 | -18.22 | -0.18 | -29.59 | 0 | -32.34 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.42 | -0.94 | 28.48 | -4.8 | 28.65 | -6.39 | 28.59 | -8.14
64.256 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | -4.86 | -0.06 | -30.94 | 0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.29 | -0.06 | -16.06 | -0.06 | -27.6 | 0 | -29.83 | -0.06 | 0 | 0 | 28.65 | 0.23 | 28.48 | -0.18 | 28.3 | -3.46 | 28.59 | -4.86 | 28.42 | -6.45
64.288 | 0.59 | -0.06 | 21.45 | -0.29 | 21.5 | -0.29 | 21.39 | -2.7 | -0.06 | -29.24 | 0 | 0 | -0.06 | 0 | 0.18 | 21.27 | 0.06 | -0.18 | -14.77 | -0.06 | -25.37 | -0.06 | -27.95 | 0 | 0 | 0 | 28.89 | 0.06 | 28.65 | -0.18 | 28.42 | -2.23 | 28.59 | -3.4 | 28.59 | -4.8
64.32 | 0.06 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | -1.99 | -0.06 | -27.19 | 0 | 0 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 0 | -13.77 | -0.06 | -23.2 | -0.06 | -26.13 | 0 | 0 | -0.06 | 17.58 | 0.06 | 28.54 | 0 | 28.65 | -1.76 | 28.48 | -2.4 | 28.36 | -2.99
64.352 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -0.82 | -0.06 | -25.55 | 0 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 2.64 | -11.66 | -0.06 | -21.27 | 0 | -24.08 | 0 | -33.4 | -0.06 | 0.12 | 0 | 28.65 | 0.12 | 28.3 | -1 | 28.48 | -1.29 | 28.42 | -2.17
64.384 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -0.18 | -0.06 | -22.62 | -0.06 | -8.2 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 8.79 | -10.2 | -0.06 | -19.63 | 0 | -21.74 | -0.06 | -76.35 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.42 | -0.47 | 28.42 | -0.7 | 28.65 | -1.23
64.416 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -0.06 | -0.06 | -20.92 | -0.06 | -75.06 | 0 | 0 | 0.06 | 21.33 | 0.29 | 14.06 | -8.61 | -0.06 | -17.17 | 0 | -19.81 | -0.06 | -49.92 | -0.06 | 0.06 | 0.06 | 28.36 | 0.23 | 28.59 | -0.35 | 28.54 | -0.7 | 28.48 | -0.7
64.448 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.23 | -0.06 | -18.52 | -0.06 | -54.32 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 19.81 | -7.79 | 0 | -15.35 | 0 | -17.93 | -0.06 | -48.05 | -0.18 | 0 | 0 | 28.42 | 0.35 | 28.59 | -0.23 | 28.48 | -0.18 | 28.48 | -0.53
64.48 | 0 | -0.06 | 21.45 | -0.41 | 21.39 | -0.18 | 21.39 | 0.29 | -0.06 | -16.7 | -0.06 | -43.13 | -0.06 | -24.32 | 0.06 | 21.39 | 0.23 | 21.45 | -6.68 | -0.06 | -14.3 | -0.06 | -16 | 0 | -42.78 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.42 | 0 | 28.48 | -0.18 | 28.54 | -0.59
64.512 | 0 | -0.06 | 21.39 | -0.18 | 21.39 | -0.29 | 21.39 | 0.41 | 0 | -14.94 | 0.12 | -38.15 | -0.18 | -74.12 | 0 | 21.21 | 0.18 | 21.39 | -5.33 | 0 | -11.95 | 0 | -14.82 | -0.06 | -37.21 | -0.06 | 0 | 0 | 28.3 | 0.18 | 28.42 | 0 | 28.48 | -0.18 | 28.42 | 0.06
64.544 | 0 | -0.18 | 21.5 | -0.35 | 21.45 | -0.23 | 21.39 | 0.53 | -0.06 | -12.48 | 0 | -32.17 | -0.18 | -43.71 | 0 | 21.39 | 0.18 | 21.39 | -4.34 | 0 | -10.84 | 0.06 | -12.89 | 0 | -37.21 | -0.06 | 0 | 0 | 28.42 | 0.29 | 28.65 | 0.12 | 28.48 | 0 | 28.59 | 0
64.576 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.18 | 21.39 | 0.53 | 0.06 | -10.2 | -0.06 | -32.34 | -0.06 | -28.24 | -0.06 | 21.33 | 0.23 | 21.39 | -3.87 | -0.06 | -9.08 | 9.08 | -11.31 | 0 | -37.5 | -0.18 | 0 | 0 | 28.24 | 0.29 | 28.71 | 0.12 | 28.71 | 0.12 | 28.48 | -0.06
64.608 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.33 | 0.41 | 0.06 | -8.14 | -0.06 | -32.05 | -0.06 | -30.53 | 0.06 | 21.33 | 0.29 | 21.39 | -2.93 | 0.29 | -8.14 | 16.11 | -10.31 | -0.06 | -36.62 | -0.06 | 0 | 0.06 | 8.73 | 0.29 | 28.54 | 0.18 | 28.3 | 0.12 | 28.42 | -0.18
64.64 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 0.88 | -5.8 | -0.06 | -31.41 | 0.06 | -35.39 | 0 | 21.33 | 0.23 | 21.15 | -2.29 | 9.84 | -6.91 | 24.02 | -8.56 | -0.06 | -35.51 | 0 | 0 | -0.06 | 0.12 | 0.06 | 28.59 | 0.29 | 28.59 | 0.18 | 28.48 | 0.06
64.672 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 10.25 | -3.93 | 0.06 | -30.35 | 0 | -34.86 | 0.06 | 21.39 | 0.29 | 21.45 | -1.93 | 17.11 | -5.45 | 28.95 | -7.44 | -0.06 | -33.69 | 0 | 0 | 0 | 0 | 0.06 | 28.48 | 0.23 | 28.36 | 0.18 | 28.54 | 0.12
64.704 | -6.91 | -0.06 | 0.12 | -0.18 | 21.39 | -0.18 | 21.45 | 0.35 | 15.65 | -1.99 | -0.06 | -27.54 | -0.06 | -30.59 | -0.06 | 21.21 | 0.23 | 21.33 | -1.46 | 27.48 | -4.51 | 28.83 | -5.98 | 3.93 | -30.94 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 28.48 | 0.18 | 28.54 | 0.06 | 28.77 | 0.12
64.736 | -60.24 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 21.39 | -1.17 | -0.06 | -25.31 | 0.06 | -27.77 | -0.06 | 0.12 | 0.18 | 21.39 | -1.17 | 28.65 | -3.57 | 28.77 | -4.63 | 12.19 | -29.06 | -0.06 | -0.53 | -0.06 | 0 | 0.12 | 28.59 | 0.06 | 28.59 | 0.06 | 29.06 | 0.12
64.768 | -48.75 | -0.18 | 0 | -0.18 | 21.39 | -0.06 | 21.39 | 0.35 | 21.45 | -0.53 | -0.06 | -23.09 | -0.06 | -27.31 | -0.06 | 0 | 0.23 | 21.33 | -1 | 28.65 | -2.81 | 28.42 | -4.1 | 19.98 | -27.48 | -0.06 | -89.3 | 0 | 0 | 0 | 28.48 | 0.12 | 28.3 | 0.06 | 28.48 | 0.12
64.8 | -40.72 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.33 | 0.35 | 21.45 | -0.18 | 0 | -21.21 | -0.06 | -26.66 | -0.06 | 0 | 0.23 | 21.15 | -0.64 | 28.48 | -2.29 | 28.65 | -3.16 | 27.6 | -25.67 | -0.06 | -66.39 | -0.18 | 0 | 0.06 | 28.18 | 0.12 | 28.48 | 0.06 | 28.48 | 0
64.832 | -38.91 | 0 | 0 | -0.23 | 0.23 | -0.18 | 21.39 | 0.29 | 21.39 | -0.06 | -0.06 | -18.87 | -0.06 | -24.61 | -0.06 | 0 | 0.18 | 21.39 | -0.64 | 28.83 | -1.64 | 28.42 | -2.46 | 28.65 | -24.08 | 0 | -48.99 | -0.18 | 0 | 0.12 | 28.48 | 0.12 | 28.36 | 0.18 | 28.42 | 0.06
64.864 | -36.39 | 0 | 0 | -0.18 | 0.06 | -0.18 | 21.39 | 0.29 | 21.39 | 0.18 | -0.06 | -16.47 | -0.18 | -22.38 | -0.06 | 0 | 0.29 | 21.39 | -0.47 | 28.36 | -1.11 | 28.54 | -1.7 | 28.48 | -21.39 | -0.06 | -43.83 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.48 | 0.18
64.896 | -34.45 | -0.06 | 0 | -0.18 | 0 | -0.06 | 4.16 | 0.18 | 21.27 | 0.35 | 0 | -14.65 | -0.06 | -20.39 | 0 | 0 | 0.23 | 21.39 | -0.41 | 28.42 | -0.94 | 28.59 | -1.23 | 28.48 | -19.45 | 0 | -38.67 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.42 | 0.06
64.928 | -32.23 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.12 | 21.27 | 0.35 | -0.06 | -12.72 | -0.06 | -19.04 | 0.12 | 0 | 0.29 | 21.33 | -0.35 | 28.36 | -0.64 | 28.65 | -0.76 | 28.24 | -17.75 | 0 | -37.15 | -0.23 | 0 | 0 | 4.57 | 0.12 | 28.48 | 0.18 | 28.48 | 0.18
64.96 | -31.11 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.41 | -0.06 | -10.61 | 0 | -16.99 | 0 | 0 | 0.29 | 21.33 | -0.35 | 28.18 | -0.53 | 28.42 | -0.7 | 28.48 | -16.29 | 0 | -37.91 | -0.06 | 0 | 0 | 0.12 | -0.06 | 28.48 | 0.06 | 28.48 | 0.12
64.992 | -29.24 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 0 | -8.56 | -0.06 | -15.65 | 0 | 0 | 0.12 | 21.33 | -0.35 | 28.59 | -0.29 | 29.01 | -0.53 | 29.01 | -14.24 | -0.06 | -37.85 | -0.18 | 0 | -0.06 | 0 | 0 | 28.48 | 0.06 | 0.12 | 0.18
65.024 | -26.49 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.53 | 0.94 | -5.63 | 0 | -13.01 | 0 | 0 | 0.23 | 21.39 | -0.29 | 28.59 | -0.35 | 28.42 | -0.23 | 28.65 | -11.78 | -0.06 | -35.74 | -0.06 | -3.11 | 0 | 0.06 | 0.12 | 0.12 | 0.06 | 0 | 0.06
65.056 | -24.26 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 8.61 | -3.75 | 0 | -11.31 | 0 | 0 | 0.23 | 21.39 | -0.23 | 28.71 | -0.23 | 29.3 | -0.41 | 28.48 | -9.26 | -0.06 | -34.51 | -0.18 | -93.11 | 0 | 0 | 0 | 0 | 0.12 | 0 | 0.12
65.088 | -21.09 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.47 | 14 | -1.82 | -0.06 | -10.08 | 0 | 0 | 0.23 | 21.39 | -0.29 | 28.42 | -0.29 | 28.36 | -0.29 | 28.77 | -7.32 | -0.06 | -32.05 | -0.18 | -63.64 | 0 | 0 | 0 | 0 | 0.12 | 0 | 0.12
65.12 | -19.92 | -0.18 | -38.2 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.35 | 19.4 | -1.64 | -0.06 | -8.2 | -0.06 | 0 | 0.23 | 21.33 | -0.29 | 28.42 | -0.41 | 28.59 | -0.35 | 28.59 | -5.16 | 3.69 | -30.47 | 0 | -47.81 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.12
65.152 | -17.93 | -0.06 | -58.3 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.5 | -0.64 | 0 | -6.8 | 0 | -41.84 | 0.18 | 21.33 | -0.18 | 28.24 | -0.29 | 28.42 | -0.29 | 28.42 | -3.34 | 12.13 | -28.65 | -0.06 | -45.35 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0
65.184 | -15.59 | -0.06 | -40.61 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.23 | 21.45 | -0.18 | 0 | -4.45 | -0.06 | -66.39 | 0.12 | 21.39 | -0.23 | 28.65 | -0.23 | 28.59 | -0.18 | 28.18 | -1.88 | 22.27 | -26.08 | -0.06 | -37.85 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0
65.216 | -14 | -0.06 | -40.14 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.15 | 0.18 | 21.45 | 0 | 5.45 | -2.81 | -0.06 | -43.42 | 0.12 | 21.33 | -0.41 | 28.65 | -0.18 | 28.54 | -0.18 | 28.48 | -1 | 29.12 | -24.55 | 0 | -36.92 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
65.248 | -12.77 | 0 | -38.2 | -0.18 | -0.06 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.23 | 11.37 | -2.05 | -0.06 | -28.54 | 0.12 | 11.37 | -0.29 | 28.54 | -0.18 | 28.13 | -0.18 | 28.3 | -0.53 | 28.48 | -22.33 | -0.06 | -37.68 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.23
65.28 | -11.25 | -0.06 | -35.16 | -0.06 | -50.04 | -0.06 | 0 | -0.06 | 21.27 | 0.23 | 21.39 | 0.41 | 16.82 | -1 | -0.06 | -30.7 | 0.06 | 0.12 | -0.29 | 28.48 | -0.18 | 28.36 | -0.23 | 28.48 | -0.29 | 28.59 | -20.57 | -0.06 | -37.62 | -0.18 | 0 | 0 | 0 | 0.06 | 0 | 0.06
65.312 | -9.67 | -0.06 | -33.52 | -0.18 | -47.7 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.41 | 21.39 | -1 | 0 | -35.63 | 0.06 | 0 | -0.18 | 28.42 | -0.29 | 28.59 | -0.23 | 28.42 | 0 | 28.65 | -19.1 | -0.18 | -36.15 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0
65.344 | -7.79 | -0.06 | -30.94 | -0.06 | -40.37 | -0.06 | -75.24 | -0.06 | 21.21 | 0.06 | 21.5 | 0.47 | 21.39 | -0.35 | -0.06 | -33.93 | 0.12 | 0.06 | -0.18 | 28.48 | -0.18 | 28.59 | -0.18 | 28.48 | 0.18 | 28.65 | -16.41 | 0 | -33.99 | -0.06 | -10.08 | -0.18 | 0 | 0 | 0 | 0.06
65.376 | -6.39 | -0.06 | -28.54 | -0.06 | -39.08 | -0.06 | -55.14 | 0 | 21.21 | 0.12 | 21.39 | 0.53 | 21.5 | -0.35 | -0.06 | -31.47 | 0.06 | 0 | -0.18 | 28.36 | -0.23 | 28.3 | -0.18 | 28.65 | 0.23 | 28.24 | -14.82 | 0 | -32.17 | -0.06 | -89.53 | -0.06 | 0 | 0 | 0 | 0
65.408 | -4.86 | -0.06 | -26.78 | -0.06 | -37.33 | -0.18 | -43.54 | -0.06 | 21.27 | 0.18 | 21.39 | 0.41 | 21.33 | 0 | -0.06 | -28.48 | 0.12 | 0 | -0.18 | 28.89 | -0.18 | 28.42 | -0.18 | 27.95 | 0.35 | 28.65 | -12.77 | 0 | -30.41 | -0.06 | -62.7 | -0.06 | 0.06 | 0 | -66.15 | 0
65.44 | -3.63 | -0.06 | -25.2 | -0.06 | -35.16 | 0 | -38.91 | -0.06 | 21.21 | 0 | 21.39 | 0.53 | 21.5 | -0.06 | -0.06 | -26.72 | 0.12 | 0 | -0.29 | 28.65 | -0.06 | 28.48 | -0.06 | 28.3 | 0.29 | 28.54 | -10.9 | -0.06 | -28.13 | -0.18 | -36.92 | -0.18 | -63.93 | 0 | -73.42 | 0
65.472 | -2.29 | 0 | -22.85 | -0.06 | -32.34 | -0.06 | -33.87 | 0 | 21.39 | -0.06 | 21.21 | 0.47 | 21.33 | 0.12 | 3.57 | -25.61 | 0 | 0 | -0.18 | 28.54 | -0.18 | 0.12 | -0.06 | 28.59 | 0.41 | 28.3 | -8.44 | -0.06 | -26.78 | 0 | -29.94 | -0.18 | -76.94 | 0.06 | -37.33 | 0
65.504 | -1.23 | 8.91 | -20.22 | -0.06 | -30.06 | 0 | -32.46 | -0.06 | 21.39 | 0.06 | 21.21 | 0.53 | 21.39 | 0.18 | 11.37 | -23.91 | 0 | 0 | -0.18 | 0.76 | -0.18 | 0 | -0.23 | 28.54 | 0.35 | 28.54 | -5.51 | -0.06 | -23.97 | -0.06 | -41.66 | -0.18 | -38.61 | 0 | -29.53 | 0.06
65.536 | -0.41 | 21.45 | -11.66 | -0.06 | -20.39 | -0.18 | -25.14 | 0.06 | 0 | -0.06 | 21.27 | 0.18 | 21.45 | 0.18 | 21.39 | -14.94 | -0.06 | -2.46 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.48 | -0.41 | 24.84 | -14.41 | 0.06 | -32.34 | -0.18 | -34.75 | -0.06 | -33.17 | 0
65.568 | -0.18 | 21.45 | -10.02 | -0.18 | -17.64 | 0 | -22.97 | 0.06 | 0 | 0 | 21.33 | 0.18 | 21.21 | 0.29 | 21.27 | -13.54 | -0.06 | -66.15 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.35 | 28.54 | -0.06 | 28.83 | -12.13 | -0.06 | -31.23 | 0 | -32.81 | 0 | -32.76 | -0.18
65.6 | -0.23 | 21.39 | -8.73 | -0.06 | -15.47 | -0.06 | -21.09 | 0 | 0 | -0.06 | 21.27 | 0.06 | 21.39 | 0.18 | 21.39 | -11.48 | -0.06 | -56.6 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.29 | 28.42 | 0.06 | 28.54 | -10.61 | -0.06 | -29.59 | 0 | -32.4 | 0 | -32.23 | -0.23
65.632 | -0.47 | 21.45 | -6.68 | -0.06 | -13.83 | 0 | -18.4 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.23 | 21.39 | -9.67 | 0 | -30.47 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.29 | 28.42 | 0.12 | 28.42 | -7.79 | -0.06 | -26.84 | 0 | -31.64 | 0 | -29.42 | -0.18
65.664 | -0.47 | 21.45 | -5.16 | -0.06 | -12.25 | 0 | -16.7 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.56 | 0.18 | 21.39 | -7.79 | -0.06 | -28.83 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0.18 | 28.48 | 0.29 | 28.48 | -5.22 | 0 | -25.84 | -0.06 | -29.42 | 0 | -27.77 | -0.18
65.696 | -0.23 | 21.45 | -4.1 | -0.06 | -11.48 | -0.06 | -14.71 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.39 | 0.23 | 21.45 | -6.09 | 0.06 | -34.63 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.48 | 0.23 | 28.65 | -3.75 | 0 | -24.43 | -0.06 | -28.07 | 0 | -25.9 | -0.18
65.728 | -0.29 | 21.45 | -2.7 | -0.18 | -9.79 | 0.06 | -12.77 | 0 | 0 | 0.06 | 21.33 | 0.06 | 21.39 | 0.18 | 21.5 | -4.22 | 0 | -35.16 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 28.59 | 0.29 | 28.48 | -2.29 | 0 | -23.03 | -0.06 | -26.54 | -0.06 | -25.08 | 0
65.76 | -0.18 | 21.45 | -1.46 | 3.75 | -8.03 | -0.06 | -10.66 | -0.06 | 0 | -0.06 | 21.15 | 0.12 | 21.33 | 0.18 | 21.39 | -2.99 | -0.06 | -32.7 | -0.06 | 0 | -0.06 | -51.97 | -0.18 | 0 | 0.18 | 28.65 | 0.29 | 28.65 | -1.23 | 0 | -21.68 | -0.18 | -25.2 | 0 | -23.5 | -0.18
65.792 | -0.29 | 21.45 | -0.76 | 11.78 | -6.5 | 4.92 | -7.44 | -0.18 | 0 | -0.06 | 0.23 | -0.06 | 21.09 | 0.12 | 21.09 | -1.88 | -0.06 | -28.48 | -0.06 | -9.9 | 0 | -58.71 | -0.06 | 0 | 0.06 | 28.59 | 0.29 | 28.65 | -0.59 | -0.06 | -18.93 | -0.06 | -22.97 | -0.06 | -21.21 | -0.18
65.824 | -0.23 | 21.45 | -0.18 | 17.17 | -4.98 | 11.13 | -5.68 | 0 | -25.2 | -0.06 | 0.06 | -0.06 | 21.33 | 0.18 | 21.33 | -1.29 | -0.18 | -27.42 | -0.18 | -76.7 | -0.06 | -45 | -0.06 | 0 | 0.12 | 28.71 | 0.35 | 28.71 | -0.29 | -0.06 | -17.23 | -0.06 | -21.5 | 0 | -19.63 | -0.18
65.856 | -0.29 | 21.45 | -0.18 | 21.8 | -3.93 | 16.47 | -3.57 | -0.06 | -77.35 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.45 | -0.41 | 0.12 | -26.78 | -0.06 | -55.96 | 0 | -45.82 | 0 | 0 | 0.12 | 28.65 | 0.29 | 28.36 | -0.06 | -0.06 | -15.82 | -0.06 | -19.98 | -0.06 | -18.22 | -0.18
65.888 | -0.18 | 21.39 | -0.18 | 21.5 | -2.46 | 21.74 | -2.05 | -0.18 | -51.51 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.21 | -0.41 | 0.23 | -25.37 | -0.18 | -44.42 | -0.06 | -43.48 | 0 | -9.79 | 0 | 0.23 | 0.18 | 28.59 | 0 | 1.99 | -14.24 | -0.06 | -18.34 | -0.06 | -16.76 | -0.06
65.92 | -0.23 | 21.45 | -0.23 | 21.5 | -1.17 | 21.5 | -1.52 | -0.06 | -42.6 | 0 | 0 | -0.18 | 21.27 | 0.18 | 21.27 | -0.18 | 7.5 | -23.15 | -0.06 | -44.59 | -0.06 | -41.84 | 0 | -78.99 | 0 | 0.06 | 0.23 | 28.65 | 0 | 7.68 | -12.77 | -0.06 | -16.7 | 0 | -15.35 | -0.06
65.952 | -0.29 | 21.45 | -0.41 | 21.39 | -0.76 | 21.45 | -0.64 | -0.06 | -35.74 | 0 | 0 | 0.12 | 19.28 | 0.06 | 21.33 | -0.06 | 14.82 | -20.74 | -0.18 | -42.25 | 0 | -39.2 | -0.06 | -49.51 | 0 | 0 | 0.12 | 28.71 | 0.29 | 17.75 | -10.49 | 9.2 | -14.36 | -0.06 | -13.07 | -0.18
65.984 | -0.18 | 21.45 | -0.47 | 21.45 | -0.29 | 21.45 | -0.18 | -0.06 | -31.93 | 0.06 | 0 | 0 | 0.12 | 0.12 | 21.39 | -0.06 | 20.57 | -19.51 | -0.18 | -40.78 | -0.06 | -37.03 | 0 | -47.05 | 0 | 0 | 0.18 | 28.54 | 0.35 | 25.37 | -9.2 | 16.11 | -12.72 | 8.79 | -11.13 | 0.12
66.016 | -0.18 | 21.45 | -0.35 | 21.45 | -0.06 | 21.45 | -0.18 | 0 | -32.58 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.12 | 21.39 | -18.05 | -0.06 | -39.08 | 0.06 | -36.21 | 0 | -45.06 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.65 | -7.44 | 23.85 | -11.54 | 16.11 | -9.84 | 9.84
66.048 | -0.18 | 21.45 | -0.29 | 21.45 | -0.06 | 21.45 | 0.23 | -0.06 | -32.4 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.23 | 21.39 | -16.29 | -0.06 | -36.86 | 0 | -34.04 | 0 | -41.9 | 0.06 | 0 | 0 | 28.42 | 0.29 | 28.71 | -5.8 | 28.65 | -9.84 | 23.67 | -8.32 | 16.99
66.08 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.29 | 0 | -31.29 | 0.06 | 0 | 0 | 0 | 0 | 21.39 | 0.06 | 21.39 | -14.53 | -0.06 | -34.75 | -0.06 | -31.99 | -0.06 | -40.9 | 0 | 0.06 | 0.06 | 28.48 | 0.29 | 28.71 | -4.16 | 28.48 | -8.14 | 28.77 | -6.56 | 24.9
66.112 | 0 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | 0.29 | -0.06 | -29.24 | -0.06 | 0 | 0.06 | 0 | 0.06 | 21.21 | 0.23 | 21.21 | -11.95 | -0.06 | -32.7 | -0.06 | -29.24 | -0.06 | -37.74 | -0.06 | 0 | 0.06 | 28.48 | 0.41 | 28.48 | -2.46 | 28.42 | -6.27 | 28.48 | -4.45 | 28.65
66.144 | -0.23 | 21.39 | -0.29 | 21.45 | -0.47 | 21.45 | 0.41 | -0.06 | -26.84 | 0 | 0 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -10.78 | -0.06 | -31.17 | -0.06 | -27.07 | -0.06 | -36.33 | -0.06 | 0 | -0.06 | 28.24 | 0.41 | 28.65 | -1.7 | 28.42 | -4.16 | 28.65 | -3.16 | 28.54
66.176 | -0.18 | 21.39 | -0.29 | 21.45 | -0.35 | 21.45 | 0.41 | 4.98 | -24.55 | 0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | -8.91 | -0.18 | -28.89 | -0.06 | -25.55 | 0.06 | -34.98 | 0.06 | 0 | 0 | 28.54 | 0.23 | 28.48 | -1.23 | 28.48 | -2.99 | 28.42 | -2.11 | 28.59
66.208 | -0.18 | 21.39 | -0.18 | 21.45 | -0.29 | 21.45 | 0.47 | 11.07 | -22.62 | 0 | -0.18 | 0 | 0 | -0.18 | 21.45 | 0.18 | 21.62 | -7.09 | -0.06 | -26.84 | 0 | -23.61 | -0.06 | -33.4 | 0.06 | 0 | 0 | 28.48 | 0.18 | 28.42 | -0.23 | 28.48 | -2.23 | 28.48 | -1.35 | 28.65
66.24 | -0.18 | 5.27 | -0.23 | 21.45 | -0.18 | 21.45 | 0.53 | 16.41 | -21.09 | -0.06 | -71.19 | 0.12 | 0 | 0 | 0.12 | 0.23 | 21.33 | -5.39 | 1.7 | -24.9 | 0 | -21.5 | -0.06 | -30.76 | 0 | 0 | 0 | 28.48 | 0.18 | 28.65 | -0.47 | 28.42 | -1.46 | 28.48 | -0.82 | 28.59
66.272 | -0.18 | 0.06 | -0.29 | 21.45 | -0.23 | 21.39 | 0.53 | 21.56 | -18.22 | 0 | -51.56 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.33 | -3.57 | 14.06 | -22.09 | -0.06 | -18.69 | -0.06 | -28.3 | -0.18 | 0 | 0 | 28.48 | 0.06 | 28.42 | -0.18 | 28.59 | -0.7 | 28.48 | -0.35 | 28.71
66.304 | -0.18 | 0 | -0.23 | 21.39 | -0.23 | 21.45 | 0.53 | 21.45 | -15.94 | -0.18 | -42.78 | 0 | 0 | -0.18 | 0 | 0.12 | 21.33 | -2.23 | 21.86 | -20.51 | -0.06 | -16.76 | 0 | -26.49 | 0 | -44.47 | -0.06 | 28.59 | 0.18 | 28.71 | -0.18 | 28.48 | -0.59 | 28.54 | -0.41 | 28.42
66.336 | -0.18 | 0.06 | -0.29 | 21.39 | -0.18 | 21.39 | 0.35 | 21.39 | -14.59 | -0.06 | -36.86 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -1.46 | 29.24 | -18.4 | 0 | -15.76 | 0 | -25.14 | 0 | -81.39 | -0.06 | 25.84 | 0.12 | 28.59 | 0.06 | 28.65 | -0.18 | 28.54 | -0.18 | 28.59
66.368 | -0.18 | 0 | -0.23 | 21.45 | -0.18 | 21.45 | 0.29 | 21.45 | -12.07 | -0.06 | -32.7 | 0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | -1.17 | 28.65 | -16.93 | 0 | -14.24 | -0.06 | -22.68 | 0 | -54.32 | 0 | 0.18 | -0.06 | 28.36 | 0 | 28.36 | -0.18 | 28.83 | -0.06 | 28.48
66.4 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | 0.29 | 21.39 | -10.37 | 0 | -32.93 | -0.06 | -37.09 | -0.06 | 0 | 0.23 | 21.33 | -0.35 | 28.65 | -15.82 | 1.99 | -12.77 | 0 | -20.86 | -0.06 | -48.4 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.59 | 0 | 28.48 | 0 | 28.42
66.432 | -0.06 | 0 | -0.18 | 21.5 | -0.29 | 21.45 | 0.29 | 21.45 | -7.62 | 0 | -32.76 | 0 | -60.88 | -0.18 | 0 | 0.23 | 21.39 | -0.47 | 28.48 | -13.48 | 10.61 | -10.72 | 1.99 | -18.16 | 0 | -40.9 | 0 | 0 | 0.06 | 28.59 | 0.18 | 28.65 | 0 | 28.48 | 0.06 | 28.59
66.464 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.29 | 21.39 | -5.22 | 0 | -31.29 | -0.06 | -39.44 | 0 | 0 | 0.18 | 21.39 | -0.18 | 28.59 | -11.95 | 18.11 | -9.61 | 8.32 | -16.11 | -0.06 | -36.97 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.36 | 0.12 | 28.36 | 0.06 | 28.48
66.496 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.12 | 21.39 | -3.16 | -0.18 | -29.36 | -0.06 | -27.42 | 0 | 0 | 0.12 | 21.33 | 0.06 | 28.3 | -10.72 | 25.61 | -8.32 | 15.29 | -14.94 | -0.06 | -37.5 | -0.18 | 0 | -0.06 | 28.54 | 0.12 | 28.48 | 0.12 | 28.59 | 0.29 | 28.65
66.528 | -0.18 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | 0 | 21.45 | -2.11 | 0 | -27.31 | 0 | -31.88 | -0.06 | 0 | 0.12 | 21.15 | 0.06 | 28.83 | -9.61 | 28.71 | -6.97 | 23.73 | -13.89 | -0.18 | -38.09 | -0.06 | 0 | 0.06 | 28.48 | 0.12 | 28.48 | 0.18 | 28.42 | 0.29 | 28.71
66.56 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.06 | 21.39 | -1.17 | 0.59 | -25.31 | 0.06 | -35.92 | -0.06 | 0 | 0.12 | 21.45 | 0.18 | 28.48 | -8.14 | 28.65 | -5.63 | 28.83 | -12.07 | 0 | -37.97 | 0 | 0 | 0 | 28.65 | 0.18 | 28.3 | 0.12 | 28.59 | 0.23 | 28.54
66.592 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0 | 21.39 | -0.29 | 10.14 | -22.56 | -0.06 | -33.22 | -0.06 | 0 | 0.18 | 21.27 | 0.12 | 28.42 | -6.33 | 28.59 | -4.63 | 28.65 | -10.37 | 0 | -34.98 | 0 | 0 | -0.06 | 28.65 | 0 | 28.48 | 0.23 | 28.54 | 0.12 | 28.48
66.624 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.06 | 21.39 | -0.06 | 15.59 | -20.92 | -0.06 | -30.18 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.65 | -5.39 | 28.3 | -3.63 | 28.77 | -9.2 | -0.06 | -32.81 | -0.18 | 0 | -0.18 | 28.83 | 0.06 | 28.71 | 0.35 | 28.65 | 0.18 | 28.48
66.656 | -0.06 | -1.76 | -0.18 | 0.12 | -0.18 | 21.39 | 0 | 21.39 | 0 | 21.33 | -18.99 | -0.06 | -28.07 | 0 | -38.97 | 0.18 | 21.21 | 0.23 | 28.42 | -4.45 | 28.3 | -2.75 | 28.65 | -7.27 | -0.06 | -31.06 | -0.06 | 0 | -0.06 | 28.65 | 0 | 28.48 | 0.12 | 28.36 | 0.23 | 28.48
66.688 | -0.18 | -60.41 | -0.06 | 0 | -0.18 | 0.12 | 0 | 21.39 | 0.23 | 21.39 | -16.76 | -0.06 | -27.48 | -0.06 | -67.09 | 0.18 | 21.39 | 0.29 | 28.54 | -3.57 | 28.36 | -2.17 | 28.65 | -6.33 | -0.06 | -29.65 | -0.06 | 0 | -0.06 | 28.65 | 0 | 28.65 | 0.12 | 28.48 | 0.18 | 28.48
66.72 | -0.06 | -50.98 | -0.18 | 0 | -0.23 | 0 | 0.06 | 21.45 | 0.23 | 21.39 | -15 | 0 | -25.9 | -0.06 | -44.01 | 0.18 | 21.33 | 0.23 | 28.36 | -2.81 | 28.3 | -1.7 | 28.54 | -5.22 | 0 | -27.36 | -0.06 | 0 | 0 | 28.59 | 0 | 28.42 | 0.12 | 28.48 | 0.18 | 28.42
66.752 | -0.06 | -40.08 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -12.13 | 0.06 | -22.97 | 0.06 | -27.25 | 0.12 | 21.39 | 0.18 | 28.65 | -2.23 | 28.48 | -1.11 | 28.89 | -4.16 | 0 | -25.14 | -0.06 | -0.06 | 0 | 28.48 | 0 | 28.48 | 0.18 | 28.54 | 0.23 | 28.48
66.784 | -0.06 | -38.67 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -10.37 | -0.06 | -21.33 | -0.06 | -32.4 | 0.18 | 0.59 | 0.23 | 28.42 | -1.7 | 28.59 | -1.05 | 28.59 | -3.28 | -0.18 | -23.85 | -0.06 | -80.39 | 0 | 23.03 | -0.06 | 28.48 | 0 | 28.48 | 0.18 | 28.48
66.816 | -0.06 | -36.86 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | -8.14 | -0.06 | -20.22 | 0 | -35.27 | 0.18 | 0.06 | 0.23 | 28.48 | -1.46 | 28.3 | -0.64 | 28.48 | -2.58 | -0.06 | -21.56 | -0.18 | -72.48 | -0.18 | 0.18 | -0.06 | 28.48 | 0.18 | 28.3 | 0.12 | 28.48
66.848 | -0.06 | -34.75 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.45 | 0.47 | 21.39 | -6.15 | -0.06 | -18.69 | -0.06 | -34.16 | -0.06 | 0 | 0.29 | 28.83 | -1.11 | 28.65 | -0.59 | 28.59 | -2.17 | -0.06 | -19.98 | -0.06 | -49.75 | -0.18 | 0 | 0 | 0.23 | 0.12 | 28.48 | 0.18 | 28.42
66.88 | 0 | -32.34 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -4.1 | 0.06 | -16.88 | -0.06 | -30.94 | 0 | 0 | 0.23 | 28.48 | -0.94 | 28.48 | -0.29 | 28.48 | -1.46 | 0 | -17.58 | 0 | -45 | 0 | 0 | 0 | 0 | -0.06 | 0.23 | 0.18 | 26.25
66.912 | -0.06 | -30.06 | -0.06 | 0 | -0.12 | 0 | -0.06 | 21.33 | 0.47 | 21.45 | -2.05 | -0.18 | -14.24 | -0.06 | -27.6 | 0 | 0 | 0.23 | 28.54 | -0.59 | 28.24 | -0.23 | 28.3 | -1 | 7.62 | -15.82 | 0 | -39.67 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.23 | 0.12
66.944 | -0.06 | -28.36 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -1.17 | 2.7 | -12.72 | 0 | -27.01 | 0 | 0 | 0.06 | 28.48 | -0.41 | 28.42 | -0.23 | 28.42 | -0.76 | 14.94 | -13.89 | -0.06 | -36.68 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
66.976 | -0.06 | -26.13 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.5 | 0.35 | 21.39 | -0.82 | 8.85 | -10.96 | -0.06 | -25.55 | 0.06 | 0 | 0.23 | 28.65 | -0.41 | 28.48 | -0.23 | 28.83 | -0.7 | 22.91 | -11.84 | -0.06 | -37.91 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 0
67.008 | -0.06 | -23.61 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.5 | 0.41 | 21.39 | -0.23 | 14.18 | -9.32 | 0 | -23.79 | 0.06 | 0 | 0.23 | 28.54 | -0.41 | 28.48 | -0.41 | 28.36 | -0.35 | 29.06 | -9.55 | 0 | -38.09 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 0
67.04 | 0 | -22.62 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | -0.06 | 19.98 | -7.38 | 3.4 | -21.62 | 0.06 | 0 | 0.23 | 28.54 | -0.41 | 28.42 | -0.29 | 28.59 | -0.18 | 28.54 | -7.38 | 0 | -36.33 | -0.18 | 0 | -0.12 | 0 | -0.06 | 0 | 0.06 | 0
67.072 | -0.06 | -19.45 | -0.18 | -31.52 | -0.06 | 0 | -0.18 | 0.12 | 0.29 | 21.39 | 0.29 | 21.45 | -5.68 | 11.37 | -20.16 | 0.06 | 0 | 0.18 | 28.54 | -0.47 | 28.59 | -0.29 | 28.42 | -0.35 | 28.65 | -4.45 | -0.06 | -34.45 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
67.104 | 0 | -17.17 | -0.06 | -61.35 | -0.06 | -27.89 | -0.06 | 0 | 0.29 | 21.39 | 0.29 | 21.39 | -3.57 | 16.76 | -18.52 | 0.06 | 0 | 0.18 | 28.48 | -0.35 | 28.65 | -0.18 | 28.71 | -0.29 | 28.65 | -2.93 | 0 | -31.99 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0 | 0
67.136 | -0.06 | -15.47 | -0.06 | -44.36 | -0.06 | -63.11 | -0.06 | 0 | 0.18 | 21.21 | 0.35 | 21.33 | -2.17 | 21.74 | -16.99 | -0.06 | 0 | 0.12 | 0.18 | -0.23 | 28.59 | -0.18 | 28.59 | -0.29 | 28.48 | -2.05 | -0.06 | -30.94 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
67.168 | -0.06 | -14.36 | -0.06 | -41.08 | -0.06 | -45.12 | -0.18 | 0 | 0.12 | 21.39 | 0.29 | 21.39 | -1.7 | 21.39 | -15.18 | -0.06 | 0 | 0.12 | 0.12 | -0.35 | 28.65 | -0.18 | 28.83 | -0.29 | 28.48 | -1.05 | 0 | -28.65 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
67.2 | -0.06 | -12.72 | -0.06 | -38.56 | -0.06 | -40.37 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -1.05 | 21.39 | -13.24 | 0 | 0 | 0.18 | 0 | -0.41 | 28.54 | -0.18 | 28.83 | -0.23 | 28.65 | -0.59 | -0.06 | -26.25 | 0 | 0.06 | -0.06 | 0 | 0 | 0 | 0.06 | 0
67.232 | -0.06 | -10.96 | -0.18 | -35.92 | -0.06 | -37.85 | -0.18 | 0 | 0.06 | 21.21 | 0.53 | 21.39 | -0.35 | 21.33 | -11.13 | -0.06 | -70.84 | 0.06 | 0 | -0.18 | 28.42 | -0.18 | 28.48 | -0.18 | 28.48 | -0.35 | -0.06 | -24.55 | -0.06 | -83.85 | -0.18 | 0 | -0.06 | 0 | 0 | 0
67.264 | 0.82 | -9.32 | -0.06 | -33.99 | -0.18 | -36.15 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -0.35 | 21.39 | -9.55 | -0.18 | -54.14 | 0.23 | 0 | -0.23 | 28.48 | -0.18 | 28.48 | -0.23 | 28.54 | 0 | -0.06 | -22.91 | 0 | -67.39 | -0.18 | -55.43 | -0.18 | 0 | -0.06 | 0
67.296 | 8.67 | -7.79 | -0.06 | -31.7 | -0.18 | -34.04 | -0.06 | 0 | 0 | 21.21 | 0.53 | 21.33 | -0.18 | 21.21 | -7.85 | 0 | -34.22 | 0.23 | 0 | -0.29 | 19.22 | -0.18 | 28.54 | -0.18 | 28.18 | 0.06 | 2.23 | -21.45 | -0.18 | -49.04 | -0.06 | -75.18 | -0.06 | -62.76 | -0.06 | 0
67.328 | 14.12 | -6.86 | -0.06 | -29.77 | -0.06 | -31.7 | -0.06 | 0 | 0.06 | 21.27 | 0.41 | 21.39 | -0.06 | 21.39 | -6.45 | 0 | -27.89 | 0.18 | 0 | -0.18 | 0.18 | -0.18 | 28.54 | -0.18 | 28.95 | 0.18 | 9.96 | -19.45 | 0 | -46.35 | -0.06 | -48.11 | 0 | -77.35 | -0.06 | -71.49
67.36 | 19.81 | -5.39 | 0 | -27.83 | -0.06 | -29.77 | -0.18 | 0 | 0.12 | 21.39 | 0.41 | 21.39 | -0.06 | 21.33 | -4.45 | 0 | -33.52 | 0.12 | 0 | -0.23 | 0 | -0.06 | 0.23 | -0.18 | 28.65 | 0.18 | 17.17 | -16.76 | -0.18 | -41.49 | -0.06 | -29.18 | -0.06 | -46.41 | -0.18 | -74.3
67.392 | 21.62 | -3.52 | -0.06 | -25.43 | 0 | -27.48 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.33 | 0.12 | 21.33 | -2.64 | -0.06 | -35.68 | 0.12 | 0 | -0.23 | 0.06 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 27.42 | -14.82 | 0 | -37.38 | -0.06 | -38.32 | 0 | -28.59 | 0 | -37.62
67.424 | 21.5 | -2.23 | -0.06 | -22.91 | -0.18 | -25.14 | -0.06 | 0 | 0 | 21.45 | 0.12 | 21.39 | 0.23 | 21.39 | -1.99 | -0.06 | -32.29 | 0.12 | 0 | -0.23 | 0 | -0.23 | 0 | -0.18 | 28.3 | 0.35 | 28.77 | -13.01 | 0 | -37.85 | -0.06 | -42.36 | 0 | -37.21 | -0.06 | -29.94
67.456 | 21.45 | -1.23 | -0.06 | -21.56 | -0.06 | -23.32 | -0.06 | 0 | -0.06 | 0.23 | 0.12 | 21.33 | 0.29 | 21.39 | -1.23 | -0.06 | -29.01 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.36 | 0.29 | 28.54 | -11.13 | -0.18 | -37.97 | -0.18 | -39.2 | -0.06 | -43.3 | -0.06 | -38.38
67.488 | 21.45 | -0.76 | -0.06 | -19.4 | -0.06 | -21.86 | 0 | -33.58 | 0 | 0.06 | 0.23 | 21.15 | 0.12 | 21.33 | -0.7 | -0.06 | -27.19 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.54 | -8.67 | -0.06 | -36.86 | 0 | -35.74 | -0.06 | -40.2 | -0.06 | -42.48
67.52 | 21.45 | -0.35 | -0.18 | -17.23 | -0.06 | -19.34 | -0.06 | -73.01 | -0.06 | 0 | 0.23 | 21.33 | 0.12 | 21.39 | -0.35 | -0.06 | -26.72 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.35 | 28.48 | -5.57 | 0.06 | -35.33 | -0.06 | -33.99 | 0.06 | -35.74 | 0 | -39.79
67.552 | 21.45 | -0.18 | 0 | -15.06 | -0.06 | -16.88 | -0.18 | -45.12 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.15 | -0.18 | 0 | -25.14 | 0.18 | -36.56 | 0 | 0 | -0.18 | 0 | -0.23 | 28.59 | 0.29 | 28.54 | -3.87 | -0.06 | -32.4 | 0 | -32.81 | 0 | -33.05 | -0.06 | -33.69
67.584 | 21.45 | -0.18 | -0.06 | -13.54 | -0.06 | -15.06 | -0.06 | -41.08 | -0.06 | 0 | 0 | 21.45 | 0.18 | 21.33 | -0.06 | -0.18 | -23.15 | 0.18 | -74.01 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.77 | -2.23 | 3.05 | -30.35 | 0 | -32.4 | 0 | -32.93 | -0.06 | -33.11
67.616 | 21.45 | -0.23 | -0.06 | -12.07 | -0.06 | -13.71 | -0.06 | -34.98 | -0.18 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -0.06 | -0.06 | -21.39 | 0.12 | -50.16 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.23 | 28.65 | -1.88 | 11.25 | -28.65 | -0.06 | -29.71 | 0 | -32.17 | -0.18 | -32.81
67.648 | 21.45 | -0.35 | -0.06 | -10.49 | -0.06 | -12.25 | -0.18 | -31.99 | 0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.33 | 0.18 | -0.06 | -20.1 | 0.06 | -44.94 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 28.48 | -0.76 | 18.87 | -26.84 | -0.06 | -28.24 | 0 | -30.12 | -0.06 | -31.82
67.68 | 21.45 | -0.35 | 5.86 | -9.55 | -0.06 | -10.96 | -0.06 | -32.7 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.33 | 0.12 | -0.06 | -18.46 | 0.12 | -44.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.83 | -0.47 | 26.37 | -24.84 | -0.06 | -26.84 | 0 | -27.54 | 0 | -29.77
67.712 | 21.45 | -0.29 | 13.42 | -7.62 | 0 | -9.02 | -0.06 | -31.99 | -0.06 | 0 | 0.12 | 21.27 | 0.23 | 21.33 | 0.23 | 5.45 | -15.7 | 0.12 | -41.43 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.12 | 28.65 | -0.18 | 28.59 | -23.09 | -0.06 | -25.25 | -0.06 | -26.37 | 0 | -28.24
67.744 | 21.45 | -0.18 | 18.99 | -5.86 | 0 | -7.68 | -0.06 | -30.7 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | 0.29 | 11.31 | -14 | 0.12 | -40.08 | -0.18 | -70.96 | -0.06 | 0 | -0.18 | 28.65 | 0.18 | 28.42 | 0 | 28.59 | -20.92 | 0 | -23.85 | -0.06 | -25.37 | -0.06 | -26.31
67.776 | 21.45 | -0.23 | 21.62 | -4.63 | -0.06 | -5.92 | 0 | -28.59 | -0.06 | 0 | 0 | 21.45 | 0.06 | 21.04 | 0.23 | 16.88 | -12.72 | 0.12 | -37.79 | -0.06 | -63.34 | 0 | -42.19 | -0.06 | 28.59 | 0.18 | 28.65 | 0.12 | 28.77 | -19.1 | -0.06 | -21.68 | 0 | -23.67 | 0 | -25.43
67.808 | 21.39 | -0.29 | 21.5 | -3.34 | 1.99 | -4.98 | -0.18 | -26.43 | 0 | 0 | 0 | 21.39 | -0.06 | 21.15 | 0.29 | 21.62 | -11.25 | 0 | -35.92 | -0.06 | -45.18 | 0 | -69.08 | -0.06 | 0.23 | 0.06 | 28.48 | 0.18 | 28.48 | -17.23 | -0.06 | -20.57 | -0.06 | -21.97 | 0 | -23.15
67.84 | 21.39 | -0.18 | 21.5 | -2.11 | 9.2 | -3.52 | -0.18 | -24.84 | -0.06 | 0 | 0 | 10.9 | 0.06 | 21.33 | 0.23 | 21.39 | -9.49 | 0 | -34.1 | -0.06 | -45.18 | -0.06 | -48.05 | -0.06 | 0.06 | 0.18 | 28.54 | 0.41 | 28.42 | -14.94 | -0.06 | -18.87 | 0 | -20.16 | 0 | -21.74
67.872 | 21.5 | -0.23 | 21.45 | -0.94 | 14.65 | -1.88 | 1.17 | -22.09 | 0.06 | -19.34 | -0.06 | 0.06 | 0.06 | 21.33 | 0.29 | 21.39 | -7.32 | 0 | -31.99 | -0.06 | -43.13 | -0.06 | -45.7 | -0.18 | 0 | 0.06 | 28.48 | 0.29 | 28.59 | -12.6 | -0.06 | -16.99 | -0.06 | -18.4 | 0 | -20.1
67.904 | 21.45 | -0.18 | 21.45 | -0.64 | 20.16 | -0.94 | 8.67 | -19.63 | -0.06 | -76.29 | 0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | -5.22 | 0.12 | -29.77 | -0.06 | -41.13 | -0.18 | -43.13 | -0.18 | 0 | 0 | 28.36 | 0.35 | 28.48 | -10.84 | 2.93 | -15.7 | 0 | -17.23 | 0 | -18.34
67.936 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.47 | 14.12 | -17.75 | -0.06 | -52.62 | -0.06 | 0.06 | -0.06 | 0.18 | 0.29 | 21.21 | -3.75 | 0 | -28.07 | 0 | -39.38 | 0 | -41.25 | 0 | 0 | 0.06 | 28.59 | 0.35 | 28.36 | -8.67 | 11.07 | -13.36 | -0.06 | -15 | -0.06 | -17.23
67.968 | 21.39 | -0.23 | 21.45 | -0.18 | 21.45 | -0.18 | 19.51 | -16.06 | -0.23 | -42.54 | -0.06 | 0 | 0 | 0.06 | 0.23 | 21.39 | -2.46 | 0 | -26.02 | -0.06 | -36.51 | -0.18 | -40.08 | -0.06 | 0 | 0.06 | 28.59 | 0.29 | 28.65 | -6.5 | 18.4 | -11.78 | 0 | -13.59 | 8.67 | -15
68 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.56 | -14.24 | -0.06 | -37.38 | -0.06 | 0 | -0.06 | 0 | 0.35 | 21.39 | -1.88 | 0 | -23.91 | -0.06 | -35.74 | 0 | -38.15 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.65 | -4.75 | 26.25 | -10.55 | 7.68 | -11.95 | 15.65 | -13.48
68.032 | 21.39 | -0.23 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -11.6 | -0.18 | -32.17 | -0.06 | 0 | 0 | 0 | 0.23 | 21.21 | -0.94 | 0.06 | -21.91 | -0.06 | -33.11 | 0 | -35.74 | 0 | 0 | 0 | 28.48 | 0.18 | 28.71 | -2.23 | 29.01 | -8.14 | 17.64 | -9.96 | 25.9 | -11.48
68.064 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.39 | -9.67 | -0.06 | -32.4 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.64 | -0.06 | -20.1 | 0 | -30.7 | -0.06 | -34.22 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.24 | -1.7 | 28.36 | -6.39 | 25.37 | -8.44 | 28.83 | -9.55
68.096 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.29 | 21.39 | -7.44 | 0 | -32.34 | -0.18 | 0 | 0.06 | 0 | 0.18 | 21.39 | -0.35 | 0.06 | -17.87 | 0 | -29.47 | -0.06 | -31.99 | -0.18 | 0 | 0 | 28.42 | 0.12 | 28.3 | -0.64 | 28.42 | -4.92 | 28.71 | -6.68 | 28.65 | -8.14
68.128 | 21.39 | -0.18 | 21.45 | -0.47 | 21.45 | -0.35 | 21.33 | -5.22 | -0.18 | -31.29 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | -0.18 | 0 | -16.29 | 0 | -27.48 | 0 | -29.94 | 0.06 | 0 | 0 | 28.42 | 0 | 28.42 | -0.47 | 28.42 | -3.46 | 28.42 | -5.16 | 28.48 | -6.39
68.16 | 11.02 | -0.18 | 21.45 | -0.18 | 21.39 | -0.35 | 21.27 | -2.99 | -0.06 | -29.83 | 0.06 | 0 | 0.06 | 0 | 0.12 | 21.33 | -0.18 | 0.06 | -15 | 0 | -25.72 | 0 | -28.48 | 0 | 0 | 0.06 | 28.71 | 0.06 | 28.54 | -0.18 | 28.48 | -2.29 | 28.71 | -3.57 | 28.36 | -4.98
68.192 | 0.06 | -0.18 | 21.45 | -0.41 | 21.45 | -0.29 | 21.39 | -1.7 | -0.18 | -27.01 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.12 | 0.06 | -13.13 | -0.06 | -23.32 | -0.06 | -25.55 | -0.06 | 0 | 0 | 0.76 | 0.12 | 28.42 | 0.06 | 28.48 | -1.35 | 28.42 | -2.29 | 28.3 | -2.87
68.224 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.29 | 21.39 | -0.76 | 0 | -25.02 | 0 | 0 | 0 | 0 | 0.12 | 21.33 | 0.18 | 0.47 | -11.66 | -0.06 | -21.09 | -0.06 | -23.44 | 0 | -55.43 | -0.18 | 0.12 | 0.12 | 28.54 | 0.18 | 28.48 | -0.76 | 28.59 | -1.23 | 28.42 | -2.23
68.256 | 0.06 | -0.18 | 21.39 | -0.35 | 21.45 | -0.23 | 21.39 | -0.35 | 0 | -22.79 | -0.06 | -2.05 | -0.06 | 0 | 0.23 | 21.45 | 0.06 | 8.61 | -10.37 | -0.06 | -19.4 | 0.06 | -22.15 | -0.06 | -80.86 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.65 | -0.47 | 28.54 | -0.82 | 28.42 | -1.23
68.288 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.23 | 21.39 | 0 | -0.06 | -20.68 | -0.06 | -70.08 | -0.06 | 0 | 0.12 | 21.45 | 0.18 | 13.83 | -9.2 | -0.06 | -16.7 | -0.06 | -19.75 | 0 | -52.91 | 0 | 0 | 0 | 28.42 | 0.23 | 28.42 | -0.29 | 28.48 | -0.47 | 28.65 | -0.7
68.32 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | -0.23 | 21.39 | 0.06 | 0 | -18.87 | -0.06 | -57.01 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 19.57 | -7.97 | 0 | -15.94 | 0 | -17.75 | 0 | -49.4 | 0 | 0 | 0 | 28.54 | 0.23 | 28.36 | -0.23 | 28.42 | -0.41 | 28.54 | -0.59
68.352 | 0 | -0.06 | 21.45 | -0.29 | 21.39 | -0.18 | 21.45 | 0.35 | -0.06 | -16.29 | -0.06 | -41.95 | -0.18 | -34.86 | 0.06 | 21.39 | 0.18 | 21.27 | -6.09 | 0.12 | -14.06 | 0 | -15.35 | -0.06 | -41.6 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.59 | -0.18 | 28.48 | -0.06 | 28.65 | -0.06
68.384 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.18 | 21.39 | 0.35 | -0.06 | -14.41 | -0.06 | -37.21 | -0.06 | -67.74 | 0.06 | 21.39 | 0.23 | 21.39 | -5.51 | -0.06 | -12.36 | 0 | -14.41 | -0.06 | -36.27 | 0 | 0 | -0.06 | 28.3 | 0.47 | 28.42 | 0 | 28.36 | 0 | 28.65 | -0.06
68.416 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | -0.29 | 21.45 | 0.29 | -0.06 | -12.6 | -0.06 | -32.58 | 0 | -46 | 0.06 | 21.39 | 0.23 | 21.21 | -4.22 | -0.06 | -11.31 | 1.99 | -12.77 | -0.06 | -36.92 | -0.18 | 0 | 0 | 28.59 | 0.23 | 28.65 | 0 | 28.54 | 0.06 | 28.59 | -0.18
68.448 | 0 | -0.06 | 21.45 | -0.29 | 21.39 | -0.18 | 21.39 | 0.41 | -0.06 | -10.61 | -0.06 | -32.81 | 0 | -29.12 | 0 | 21.33 | 0.18 | 21.39 | -3.75 | 0 | -9.61 | 7.62 | -11.54 | -0.06 | -37.33 | 0 | 0 | -0.06 | 28.65 | 0.18 | 28.48 | 0.06 | 28.48 | 0.12 | 28.36 | 0.06
68.48 | 0 | 0 | 21.45 | -0.18 | 21.5 | -0.23 | 21.39 | 0.53 | -0.06 | -8.14 | -0.06 | -32.64 | -0.18 | -29.36 | 0.06 | 21.39 | 0.18 | 21.27 | -2.87 | 1.99 | -8.73 | 15.06 | -10.37 | -0.06 | -36.86 | -0.06 | 0 | 0 | 13.48 | 0.23 | 28.36 | 0.06 | 28.42 | 0.18 | 28.42 | 0.06
68.512 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.35 | 3.57 | -5.22 | -0.06 | -31.64 | -0.18 | -35.86 | 0 | 21.39 | 0.18 | 21.39 | -2.17 | 11.54 | -7.09 | 25.43 | -8.14 | 0 | -35.45 | -0.06 | 0 | 0 | 0.12 | 0.12 | 28.36 | 0.12 | 28.59 | 0.23 | 29.12 | 0.06
68.544 | 0 | -0.18 | 17.58 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 9.96 | -3.28 | -0.06 | -29.36 | 0 | -34.51 | -0.06 | 21.56 | 0.23 | 21.33 | -1.93 | 19.4 | -5.63 | 28.65 | -6.86 | 0 | -32.7 | -0.18 | 0 | -0.06 | 0 | 0.18 | 28.59 | 0.06 | 28.42 | 0.06 | 28.3 | 0.12
68.576 | -1.05 | -0.06 | 0.12 | -0.23 | 21.39 | -0.06 | 21.39 | 0.41 | 15.18 | -2.05 | -0.06 | -26.9 | 0 | -30.94 | 0 | 21.39 | 0.18 | 21.39 | -1.35 | 26.66 | -4.63 | 28.71 | -5.74 | 1.11 | -31.11 | 0 | 0 | 0 | 0.06 | 0.06 | 28.42 | 0.06 | 28.54 | 0.23 | 28.48 | 0.18
68.608 | -52.85 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 20.92 | -1.05 | 0 | -25.14 | -0.18 | -28.24 | -0.06 | 0.12 | 0.06 | 21.39 | -1.17 | 28.77 | -3.63 | 28.48 | -4.98 | 11.25 | -29.77 | -0.06 | 0 | 0 | 0 | 0.12 | 28.42 | 0.06 | 28.48 | 0.06 | 28.48 | 0.18
68.64 | -51.45 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 21.45 | -0.82 | 0.06 | -23.32 | -0.06 | -27.54 | 0 | 0.06 | 0.18 | 21.33 | -0.76 | 28.59 | -3.16 | 28.54 | -4.1 | 18.87 | -27.6 | -0.06 | -81.51 | -0.06 | 0 | 0 | 28.59 | 0.12 | 28.36 | 0.12 | 28.42 | 0.18
68.672 | -40.14 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.29 | 21.45 | -0.06 | -0.06 | -20.98 | -0.06 | -26.31 | -0.06 | 0 | 0.18 | 21.39 | -0.7 | 28.42 | -2.4 | 28.54 | -2.93 | 29.01 | -24.9 | -0.06 | -59.83 | 0 | 0 | 0.06 | 28.42 | 0.12 | 28.54 | 0.18 | 28.48 | 0.18
68.704 | -39.26 | -0.06 | 0 | -0.18 | 2.99 | -0.06 | 21.39 | 0.23 | 21.39 | 0.06 | -0.06 | -19.4 | -0.06 | -23.79 | -0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.59 | -1.7 | 28.71 | -2.4 | 28.77 | -23.32 | -0.06 | -48.46 | -0.18 | 0 | -0.06 | 28.48 | 0.12 | 28.36 | 0.18 | 28.48 | 0.18
68.736 | -37.15 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.39 | 0.41 | 21.39 | 0.18 | -0.06 | -17.23 | -0.06 | -22.09 | -0.06 | 0 | 0.23 | 21.39 | -0.53 | 28.48 | -1.41 | 28.36 | -1.76 | 28.54 | -21.33 | 0 | -45.12 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.83 | 0.18 | 28.42 | 0.18
68.768 | -34.69 | -0.18 | 0 | -0.18 | 0 | -0.06 | 11.48 | 0.18 | 21.5 | 0.41 | -0.06 | -15.18 | -0.06 | -20.39 | 0 | 0 | 0.18 | 21.39 | -0.47 | 28.48 | -1.23 | 28.65 | -1.29 | 28.71 | -19.86 | -0.06 | -39.08 | 0 | 0 | -0.06 | 28.42 | 0 | 28.48 | 0.12 | 28.48 | 0.18
68.8 | -32.81 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.39 | 0.47 | -0.06 | -12.95 | 0.06 | -18.93 | -0.06 | 0 | 0.18 | 21.39 | -0.41 | 28.42 | -0.76 | 28.59 | -1 | 28.59 | -17.93 | -0.18 | -36.33 | -0.06 | 0 | 0.06 | 15 | -0.06 | 28.65 | 0.18 | 28.48 | 0.18
68.832 | -29.53 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | -0.06 | -9.9 | 0 | -16.93 | -0.18 | 0 | 0.12 | 21.21 | -0.47 | 28.65 | -0.53 | 28.65 | -0.47 | 28.48 | -15.35 | 0 | -37.91 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 28.48 | 0.18 | 17.93 | 0.18
68.864 | -28.48 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.06 | 21.39 | 0.35 | 0.06 | -8.32 | 0 | -15.23 | -0.18 | 0 | 0.23 | 21.45 | -0.35 | 28.48 | -0.53 | 28.36 | -0.18 | 28.65 | -13.48 | -0.06 | -37.5 | -0.06 | 0 | 0 | 0 | -0.06 | 23.32 | 0 | 0.18 | 0.18
68.896 | -26.25 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.33 | 0.47 | 2.11 | -5.98 | -0.06 | -13.36 | 0.06 | 0 | 0.18 | 21.39 | -0.35 | 28.54 | -0.41 | 28.48 | -0.41 | 28.89 | -11.48 | 0.06 | -36.27 | 0 | -1.52 | -0.06 | 0.06 | -0.18 | 0.18 | 0.12 | 0 | 0.23
68.928 | -24.26 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 7.62 | -3.93 | 0 | -11.6 | 0 | 0 | 0.12 | 21.33 | -0.35 | 28.42 | -0.23 | 28.89 | -0.41 | 28.36 | -9.43 | -0.18 | -34.22 | 0 | -92.64 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.12
68.96 | -22.38 | -0.06 | -0.06 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 13.24 | -2.05 | -0.06 | -10.08 | -0.06 | 0 | 0.06 | 21.27 | -0.23 | 28.42 | -0.23 | 28.48 | -0.23 | 28.48 | -7.56 | -0.06 | -32.58 | 0 | -65.22 | -0.06 | 0 | 0 | 0 | 0.12 | 0 | 0.18
68.992 | -19.51 | -0.06 | -48.46 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 20.45 | -1 | 0 | -7.79 | -0.06 | 0.06 | 0.12 | 21.33 | -0.29 | 28.54 | -0.23 | 28.48 | -0.41 | 28.3 | -4.92 | 6.39 | -30.18 | 0 | -48.75 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.12
69.024 | -17.93 | -0.06 | -52.85 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -0.76 | -0.06 | -6.33 | 0 | -35.98 | 0.06 | 21.33 | -0.41 | 28.48 | -0.35 | 28.54 | -0.23 | 28.48 | -2.87 | 14.12 | -28.59 | -0.06 | -44.01 | 0 | 0 | 0 | 0 | 0.12 | 0 | 0.06
69.056 | -15.76 | -0.06 | -41.78 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -0.06 | -0.06 | -4.28 | -0.06 | -69.96 | 0.12 | 21.33 | -0.29 | 28.77 | -0.29 | 28.42 | -0.18 | 28.65 | -1.99 | 22.03 | -26.54 | -0.18 | -38.56 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12
69.088 | -14.59 | 0 | -39.61 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.18 | 21.21 | 0.06 | 4.63 | -2.87 | -0.06 | -45.7 | 0.18 | 21.33 | -0.35 | 28.48 | -0.29 | 28.59 | -0.35 | 28.65 | -0.94 | 29.01 | -24.96 | -0.06 | -37.33 | 0 | 0 | 0 | 0 | 0 | 0 | 0.12
69.12 | -12.95 | -0.06 | -38.67 | -0.06 | -3.4 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.33 | 0.23 | 10.66 | -2.17 | 0.06 | -29.94 | 0.18 | 13.3 | -0.35 | 28.42 | -0.23 | 28.48 | -0.18 | 28.24 | -0.64 | 28.42 | -23.15 | 0 | -37.68 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06
69.152 | -10.66 | -0.06 | -35.27 | -0.06 | -57.07 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | 0.41 | 18.05 | -1.11 | -0.18 | -30.94 | 0.23 | 0.06 | -0.23 | 28.65 | -0.23 | 28.77 | -0.18 | 29.12 | -0.18 | 28.48 | -20.39 | -0.06 | -37.38 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
69.184 | -9.38 | 0 | -33.05 | -0.18 | -50.57 | -0.06 | -3.22 | -0.06 | 21.27 | 0.12 | 21.39 | 0.35 | 21.62 | -0.59 | -0.06 | -35.98 | 0.12 | 0 | -0.18 | 28.42 | -0.18 | 28.59 | -0.18 | 28.42 | 0 | 28.48 | -18.69 | -0.06 | -36.04 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.06
69.216 | -7.91 | 0 | -31 | -0.06 | -39.61 | -0.06 | -72.07 | 0 | 21.5 | 0 | 21.39 | 0.47 | 21.39 | -0.47 | -0.06 | -34.69 | 0.18 | 0 | -0.23 | 28.42 | -0.18 | 28.65 | -0.18 | 28.65 | 0.12 | 28.65 | -16.82 | -0.06 | -33.81 | -0.06 | -4.16 | -0.06 | 0 | 0 | 0 | 0.06
69.248 | -6.68 | -0.06 | -30 | -0.18 | -38.85 | -0.06 | -56.84 | -0.06 | 21.39 | 0.06 | 21.45 | 0.47 | 21.39 | -0.18 | 0.06 | -31.47 | 0.18 | 0 | -0.18 | 28.42 | -0.18 | 28.83 | -0.18 | 28.24 | 0.23 | 28.65 | -15.18 | 0 | -31.7 | -0.06 | -86.08 | -0.18 | 0 | -0.06 | 0 | 0
69.28 | -5.39 | -0.06 | -27.01 | 0 | -36.62 | -0.06 | -43.77 | -0.18 | 21.39 | 0.06 | 21.21 | 0.53 | 21.33 | -0.18 | 0 | -28.48 | 0.06 | 0 | -0.06 | 28.59 | -0.18 | 28.42 | -0.18 | 28.42 | 0.18 | 28.48 | -12.89 | 0 | -30.18 | -0.06 | -65.39 | -0.18 | 0 | 0.06 | -53.62 | 0
69.312 | -3.28 | -0.06 | -24.08 | -0.06 | -35.33 | -0.06 | -37.56 | 0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.21 | 0.06 | 0 | -27.01 | 0.18 | 0 | -0.18 | 28.48 | -0.06 | 23.26 | -0.18 | 28.48 | 0.41 | 28.59 | -10.49 | -0.06 | -28.24 | -0.06 | -32.29 | -0.06 | -75.82 | 0 | -68.44 | 0
69.344 | -2.29 | 0.94 | -22.09 | -0.06 | -32.93 | -0.06 | -32.64 | -0.06 | 21.45 | 0.06 | 21.39 | 0.47 | 21.39 | 0.18 | 5.45 | -25.78 | 0.06 | 0 | -0.18 | 28.48 | -0.06 | 0.18 | -0.18 | 28.59 | 0.35 | 28.42 | -8.38 | -0.06 | -25.9 | 0 | -31.29 | -0.18 | -68.97 | 0 | -40.67 | 0.06
69.376 | -1.41 | 8.67 | -20.45 | -0.18 | -30.94 | -0.06 | -32.64 | -0.18 | 21.39 | 0 | 21.45 | 0.47 | 21.39 | 0.06 | 11.37 | -24.08 | 0.12 | 0 | -0.06 | 9.14 | -0.18 | 0 | -0.18 | 28.36 | 0.47 | 28.48 | -6.21 | 0 | -24.38 | 0 | -41.49 | -0.06 | -43.13 | 0 | -29.01 | -0.06
69.408 | -0.64 | 14.3 | -18.16 | -0.06 | -28.77 | 0 | -32.23 | 0.06 | 0.29 | 0.06 | 21.39 | 0.41 | 21.33 | 0.12 | 16.7 | -22.33 | 0.12 | 0 | -0.23 | 0.12 | -0.18 | 0 | -0.18 | 28.48 | 0.53 | 28.42 | -4.45 | -0.18 | -22.44 | -0.18 | -41.9 | -0.06 | -31.47 | -0.06 | -36.62 | 0
69.44 | -0.18 | 19.86 | -16.76 | 0.06 | -26.66 | -0.06 | -31.41 | -0.06 | 0.06 | 0.06 | 21.39 | 0.35 | 21.39 | 0.06 | 21.5 | -20.92 | 0.12 | 0 | -0.18 | 0 | -0.23 | 0 | -0.18 | 28.59 | 0.41 | 28.3 | -2.4 | 1.99 | -20.74 | -0.06 | -37.62 | -0.06 | -39.08 | 0 | -42.6 | -0.06
69.472 | 0 | 21.45 | -14.24 | -0.06 | -24.2 | -0.06 | -29.71 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.12 | 21.68 | -18.63 | 0 | 0 | -0.18 | 0.06 | -0.12 | 0 | -0.18 | 0.29 | 0.47 | 28.48 | -1.46 | 9.32 | -18.05 | -0.06 | -33.4 | -0.18 | -44.12 | -0.06 | -38.67 | -0.06
69.504 | -0.18 | 21.5 | -12.72 | 0 | -22.15 | -0.06 | -27.48 | -0.06 | 0 | -0.06 | 21.27 | 0.23 | 21.33 | 0.12 | 21.39 | -17.23 | 0.12 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.35 | 28.42 | -0.76 | 16.41 | -16.23 | 0.06 | -32.93 | 0 | -39.9 | 0.06 | -34.98 | -0.06
69.536 | -0.23 | 21.5 | -11.6 | -0.06 | -19.98 | -0.06 | -25.08 | 0 | 0 | 0.06 | 21.39 | 0.29 | 21.15 | 0.18 | 21.27 | -15.47 | 0.18 | -0.88 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.35 | 28.89 | -0.29 | 24.26 | -14.3 | -0.06 | -32.7 | -0.06 | -36.1 | -0.06 | -32.7 | 0
69.568 | -0.47 | 21.45 | -10.02 | -0.18 | -17.93 | -0.06 | -23.15 | -0.06 | 0 | -0.18 | 21.33 | 0.18 | 21.33 | 0.23 | 21.15 | -13.54 | -0.06 | -62.99 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.29 | 28.59 | -0.29 | 28.71 | -12.6 | 0 | -31.17 | 0 | -33.87 | -0.06 | -32.93 | 0
69.6 | -0.41 | 21.45 | -8.85 | -0.06 | -15.76 | -0.06 | -21.33 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.27 | 0.23 | 21.39 | -12.07 | 0 | -57.66 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.42 | 0.12 | 28.42 | -10.66 | -0.06 | -29.3 | -0.23 | -34.1 | 0 | -32.4 | -0.06
69.632 | -0.23 | 21.45 | -6.86 | -0.06 | -14.24 | 0 | -18.75 | 0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | 0.29 | 21.39 | -10.02 | 0.06 | -31.06 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0.23 | 28.71 | 0.18 | 28.83 | -7.79 | -0.18 | -27.25 | 0 | -32.76 | 0 | -29.77 | 0
69.664 | -0.23 | 21.45 | -5.33 | 0 | -12.72 | -0.06 | -16.76 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | 0.18 | 21.39 | -7.97 | 0 | -28.48 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.89 | 0.29 | 28.48 | -5.8 | 0 | -25.84 | -0.18 | -30 | -0.06 | -27.42 | 0
69.696 | -0.35 | 21.45 | -3.93 | -0.06 | -11.48 | -0.06 | -14.82 | 0 | 0 | 0.06 | 21.39 | 0 | 21.33 | 0.29 | 21.39 | -6.33 | -0.06 | -34.51 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.59 | 0.23 | 28.42 | -3.75 | 0 | -24.2 | -0.18 | -28.48 | -0.18 | -26.25 | 0
69.728 | -0.23 | 21.45 | -2.99 | 0 | -9.61 | -0.06 | -12.95 | -0.06 | 0 | -0.18 | 21.39 | 0.06 | 21.39 | 0.18 | 21.39 | -4.69 | 0 | -35.27 | -0.18 | 0 | -0.06 | 0.06 | -0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.65 | -1.99 | -0.06 | -22.91 | -0.18 | -25.9 | -0.06 | -25.31 | -0.06
69.76 | -0.23 | 21.45 | -1.82 | 3.16 | -8.14 | 0 | -10.61 | 0 | 0 | 0 | 21.45 | -0.06 | 21.33 | 0.23 | 21.33 | -3.28 | 0 | -33.22 | -0.18 | 0 | 0 | -41.31 | 0 | 0 | 0.12 | 28.42 | 0.29 | 28.48 | -1.7 | -0.06 | -21.45 | 0 | -24.32 | 0 | -23.85 | 0
69.792 | -0.23 | 21.45 | -0.53 | 11.25 | -6.68 | 4.16 | -8.09 | -0.06 | 0 | -0.06 | 0.59 | 0 | 21.39 | 0.12 | 21.45 | -1.88 | -0.06 | -29.01 | -0.06 | -4.28 | -0.12 | -63.11 | -0.18 | 0 | 0.06 | 28.48 | 0.35 | 28.3 | -0.64 | -0.06 | -19.16 | -0.18 | -22.15 | -0.18 | -21.86 | 0
69.824 | -0.18 | 21.45 | -0.29 | 16.64 | -5.22 | 10.31 | -5.92 | -0.06 | -16.76 | -0.06 | 0.06 | -0.06 | 21.21 | 0.12 | 21.21 | -1.41 | -0.06 | -27.25 | -0.18 | -73.19 | -0.06 | -45.41 | -0.18 | 0 | 0.12 | 28.42 | 0.29 | 28.42 | -0.29 | 0 | -17.75 | -0.06 | -20.63 | -0.06 | -19.98 | 0
69.856 | -0.29 | 21.45 | -0.18 | 21.74 | -3.98 | 15.59 | -3.75 | -0.06 | -76.82 | -0.06 | 0 | -0.18 | 21.39 | 0.12 | 21.39 | -0.64 | 0 | -26.66 | -0.06 | -57.48 | 0.06 | -45.29 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.59 | -0.06 | -0.18 | -15.82 | -0.06 | -18.99 | -0.06 | -18.4 | 0.06
69.888 | -0.29 | 21.45 | -0.18 | 21.5 | -2.7 | 21.45 | -2.11 | -0.06 | -53.03 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.45 | -0.64 | -0.06 | -25.31 | -0.06 | -44.36 | -0.06 | -44.3 | 0 | -1.52 | -0.06 | 0.23 | 0.18 | 28.48 | 0.06 | 0 | -14.65 | -0.06 | -17.87 | -0.06 | -16.99 | 0
69.92 | -0.23 | 21.39 | -0.23 | 21.45 | -1.29 | 21.5 | -1.52 | 0 | -42.48 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | -0.18 | 6.62 | -22.97 | -0.06 | -44.94 | 0 | -41.25 | 0 | -74.83 | 0 | 0.12 | 0.29 | 28.54 | 0.12 | 6.33 | -13.01 | -0.18 | -16.47 | -0.06 | -15.29 | 0
69.952 | -0.23 | 21.5 | -0.41 | 21.45 | -0.64 | 21.5 | -0.7 | 0 | -36.15 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 21.33 | -0.18 | 14.18 | -20.74 | -0.06 | -41.49 | 0 | -39.38 | 0 | -52.44 | 0 | 0 | 0.06 | 28.42 | 0.29 | 16.7 | -10.66 | 7.79 | -13.89 | -0.06 | -12.89 | 0
69.984 | -0.18 | 21.39 | -0.41 | 21.5 | -0.29 | 21.45 | -0.06 | -0.18 | -31.88 | -0.06 | 0 | -0.18 | 0.12 | 0.12 | 21.39 | 0 | 19.81 | -19.45 | -0.18 | -41.31 | -0.06 | -37.21 | 0 | -46.64 | -0.18 | 0 | 0.12 | 28.59 | 0.29 | 24.38 | -9.02 | 15 | -12.48 | 8.67 | -11.84 | -0.06
70.016 | -0.18 | 21.45 | -0.41 | 21.45 | 0 | 21.45 | -0.06 | 0.06 | -32.64 | -0.06 | 0 | 0 | 0 | -0.06 | 21.21 | 0.06 | 21.39 | -17.93 | 0 | -38.32 | -0.06 | -35.74 | -0.06 | -45.29 | 0 | 0 | 0 | 28.48 | 0.35 | 28.77 | -7.44 | 22.91 | -10.37 | 15.7 | -10.14 | 9.26
70.048 | -0.23 | 21.45 | -0.35 | 21.5 | -0.18 | 21.45 | 0.29 | 0.12 | -32.46 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0 | 21.39 | -16.29 | -0.06 | -37.21 | -0.06 | -34.28 | 0 | -43.13 | 0 | 0 | 0.18 | 28.65 | 0.29 | 28.59 | -5.8 | 29.18 | -9.02 | 23.38 | -8.44 | 16.47
70.08 | -0.18 | 21.39 | -0.35 | 21.45 | -0.41 | 21.45 | 0.29 | 0 | -31.47 | -0.06 | 0 | 0 | 0 | -0.06 | 21.21 | 0.18 | 21.15 | -14.24 | 0 | -35.68 | 0 | -32.29 | -0.18 | -41.49 | -0.06 | 0 | 0.12 | 28.54 | 0.29 | 28.42 | -4.28 | 28.48 | -7.44 | 28.83 | -7.32 | 24.14
70.112 | -0.06 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | 0.35 | 0 | -29.12 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.23 | 21.15 | -12.31 | -0.06 | -32.76 | -0.06 | -29.12 | 0.06 | -38.2 | -0.06 | 0 | 0 | 28.42 | 0.29 | 28.42 | -2.46 | 28.65 | -5.8 | 28.54 | -4.69 | 28.48
70.144 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | 0 | -26.66 | -0.06 | 0 | 0 | 0 | -0.06 | 21.33 | 0.29 | 21.15 | -10.96 | -0.06 | -31 | 0 | -27.89 | 0 | -37.5 | 0 | 0 | 0 | 28.36 | 0.29 | 28.59 | -1.46 | 28.83 | -3.69 | 28.54 | -3.22 | 28.42
70.176 | -0.06 | 21.45 | -0.35 | 21.45 | -0.35 | 21.39 | 0.47 | 4.22 | -24.43 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.33 | -8.79 | -0.06 | -28.95 | -0.06 | -25.61 | -0.06 | -34.98 | 0.06 | 0 | 0 | 28.42 | 0.29 | 28.65 | -1.29 | 28.48 | -2.4 | 28.59 | -2.17 | 28.48
70.208 | -0.06 | 21.39 | -0.23 | 21.45 | -0.35 | 21.45 | 0.41 | 10.43 | -22.38 | 0.06 | 0.06 | -0.18 | 0 | 0 | 21.39 | 0.12 | 21.27 | -6.97 | 0 | -27.77 | -0.06 | -24.14 | -0.06 | -32.99 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.59 | -0.59 | 28.65 | -1.76 | 28.59 | -1.64 | 28.42
70.24 | -0.06 | 10.55 | -0.29 | 21.45 | -0.29 | 21.45 | 0.47 | 15.53 | -21.21 | -0.06 | -63.58 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.33 | -5.8 | 0.7 | -25.14 | -0.06 | -21.68 | 0 | -31.52 | -0.06 | 0 | 0 | 28.54 | 0.18 | 28.48 | -0.29 | 28.71 | -1 | 28.59 | -0.94 | 28.71
70.272 | -0.18 | 0.06 | -0.23 | 21.39 | -0.23 | 21.45 | 0.47 | 21.62 | -18.69 | -0.06 | -54.03 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | -3.4 | 13.18 | -22.38 | -0.06 | -19.28 | 0 | -28.48 | 0 | 0 | -0.06 | 28.36 | 0.12 | 28.59 | -0.18 | 28.42 | -0.76 | 28.42 | -0.53 | 28.42
70.304 | -0.18 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | 0.47 | 21.39 | -16.17 | -0.06 | -42.48 | -0.06 | 0 | 0 | 0 | 0.18 | 21.27 | -2.52 | 20.8 | -20.86 | 0 | -17.64 | 0.06 | -26.84 | 0 | -40.43 | -0.06 | 28.65 | 0 | 28.36 | 0.06 | 28.54 | -0.29 | 28.48 | -0.53 | 28.48
70.336 | -0.06 | 0 | -0.29 | 21.39 | -0.18 | 21.45 | 0.41 | 21.39 | -14.36 | 0 | -37.38 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | -1.7 | 28.59 | -18.52 | 0 | -15.7 | 0 | -24.96 | 0 | -83.85 | -0.06 | 28.48 | 0.18 | 28.42 | 0.06 | 28.59 | -0.23 | 28.65 | -0.18 | 28.42
70.368 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.39 | 0.35 | 21.39 | -12.6 | -0.06 | -32.93 | 0 | 0 | -0.06 | 0 | 0.12 | 21.33 | -0.94 | 28.48 | -16.7 | -0.06 | -14.65 | 0.06 | -23.32 | -0.06 | -54.61 | 0 | 0.23 | -0.06 | 28.48 | 0.18 | 28.54 | -0.06 | 28.48 | -0.18 | 28.42
70.4 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | -10.49 | 0.06 | -32.81 | -0.06 | -30.47 | 0 | 0 | 0.12 | 21.27 | -0.7 | 28.71 | -15.47 | 1.99 | -12.95 | 0 | -20.92 | 0 | -48.22 | 0 | 0 | -0.06 | 28.36 | 0.18 | 28.3 | 0 | 28.54 | 0 | 28.36
70.432 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.18 | 21.39 | -7.68 | -0.06 | -32.29 | -0.06 | -63.75 | -0.06 | 0 | 0.18 | 21.27 | -0.35 | 28.36 | -13.48 | 10.31 | -11.02 | 1.99 | -17.99 | 0 | -41.25 | -0.18 | 0 | -0.06 | 28.48 | 0.23 | 28.42 | 0.18 | 28.71 | 0 | 28.48
70.464 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.18 | 21.39 | -5.39 | 0.06 | -31.64 | -0.06 | -41.78 | -0.06 | 0 | 0.23 | 21.33 | -0.06 | 28.42 | -11.95 | 17.87 | -10.02 | 7.91 | -16.52 | 0 | -36.68 | -0.06 | 0 | -0.06 | 28.36 | 0.23 | 28.48 | 0.06 | 28.36 | 0.06 | 28.36
70.496 | -0.06 | 0 | -0.18 | 21.5 | -0.23 | 21.45 | 0.18 | 21.39 | -3.46 | -0.06 | -30.18 | 0 | -27.36 | -0.06 | 0 | 0.12 | 21.33 | -0.06 | 28.54 | -11.02 | 25.43 | -8.32 | 14.94 | -15.18 | 0.12 | -37.09 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.36 | 0.06 | 28.42 | 0.06 | 28.42
70.528 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.23 | 21.39 | -1.76 | 0 | -28.36 | 0 | -31 | -0.06 | 0 | 0.12 | 21.27 | -0.06 | 28.54 | -9.38 | 28.48 | -7.09 | 23.2 | -13.83 | -0.18 | -37.79 | -0.18 | 0 | -0.06 | 28.42 | 0.18 | 28.65 | 0.12 | 28.48 | 0.12 | 28.42
70.56 | -0.18 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | 0.06 | 21.39 | -1.41 | 0.18 | -26.08 | -0.06 | -35.74 | -0.06 | 0 | 0.12 | 21.27 | 0.12 | 28.59 | -8.32 | 28.59 | -5.92 | 28.89 | -12.13 | -0.06 | -37.09 | -0.06 | 0 | -0.18 | 28.65 | 0.18 | 28.24 | 0.06 | 28.59 | 0.12 | 28.13
70.592 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0 | 21.45 | -0.53 | 9.49 | -23.15 | -0.06 | -33.52 | -0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.42 | -6.62 | 28.65 | -4.45 | 28.71 | -10.25 | -0.06 | -35.1 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.48 | 0.06 | 28.48 | 0.12 | 28.3
70.624 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.06 | 21.39 | -0.06 | 14.94 | -21.09 | -0.06 | -30.06 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 28.71 | -5.51 | 28.65 | -3.69 | 28.71 | -8.91 | -0.06 | -33.34 | -0.06 | 0 | -0.06 | 28.42 | 0.06 | 28.42 | 0.18 | 28.24 | 0.12 | 28.42
70.656 | 0 | -0.64 | -0.18 | 0.18 | -0.18 | 21.45 | 0 | 21.39 | 0.12 | 20.74 | -18.87 | 0 | -27.89 | -0.06 | -36.51 | 0.12 | 21.39 | 0.06 | 28.48 | -4.8 | 28.83 | -2.87 | 28.59 | -7.56 | 0 | -31.29 | -0.18 | 0 | -0.18 | 28.42 | -0.06 | 28.36 | 0.06 | 28.65 | 0.12 | 28.42
70.688 | -0.06 | -58.54 | -0.18 | 0.06 | -0.18 | 0.23 | -0.06 | 21.39 | 0.23 | 21.39 | -16.99 | -0.06 | -26.84 | -0.06 | -68.73 | 0.18 | 21.39 | 0.23 | 28.48 | -3.87 | 28.54 | -2.23 | 28.65 | -6.56 | 0 | -29.36 | -0.06 | 0 | 0 | 28.59 | -0.06 | 28.48 | 0.12 | 28.13 | 0.06 | 28.48
70.72 | -0.06 | -52.33 | 0 | 0 | -0.23 | 0.06 | -0.06 | 21.45 | 0.47 | 21.39 | -15.18 | 0 | -26.02 | -0.06 | -45.06 | 0 | 21.39 | 0.23 | 28.36 | -2.87 | 28.54 | -1.82 | 28.71 | -5.22 | 0 | -27.6 | 0 | 0 | -0.06 | 28.42 | 0 | 28.48 | 0.18 | 28.59 | 0.06 | 28.42
70.752 | -0.06 | -39.61 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -12.6 | -0.18 | -23.15 | -0.06 | -27.31 | 0.06 | 21.15 | 0.12 | 28.42 | -2.23 | 28.3 | -1.17 | 28.77 | -3.93 | -0.18 | -25.37 | -0.06 | 0.06 | 0 | 28.42 | -0.06 | 28.48 | 0.12 | 28.48 | 0.06 | 28.42
70.784 | -0.18 | -39.44 | -0.18 | 0 | -0.18 | 0.06 | 0.06 | 21.39 | 0.41 | 21.39 | -10.49 | -0.06 | -21.62 | -0.06 | -31.64 | 0.12 | 10.78 | 0.18 | 28.48 | -1.64 | 28.77 | -0.82 | 28.42 | -3.28 | -0.06 | -23.56 | -0.18 | -72.6 | -0.06 | 28.48 | -0.06 | 28.89 | 0.06 | 28.54 | 0.18 | 28.48
70.816 | -0.18 | -37.33 | -0.06 | 0 | -0.18 | 0 | 0 | 21.21 | 0.47 | 21.21 | -8.79 | -0.18 | -20.16 | 0.06 | -35.74 | 0.12 | 0.06 | 0.06 | 28.89 | -1.29 | 28.71 | -0.64 | 28.36 | -2.7 | -0.06 | -21.86 | 0 | -74.65 | -0.06 | 0.18 | 0 | 28.65 | 0.12 | 28.42 | 0.06 | 28.48
70.848 | -0.06 | -35.1 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.33 | -6.56 | -0.18 | -18.87 | -0.06 | -34.28 | 0.12 | 0 | 0.35 | 28.54 | -1.11 | 28.48 | -0.47 | 28.59 | -1.99 | -0.06 | -20.33 | 0 | -49.69 | 0 | 0 | -0.18 | 0.23 | 0.06 | 28.42 | 0.18 | 28.54
70.88 | -0.06 | -33.4 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.53 | 21.39 | -4.39 | -0.06 | -17.17 | -0.06 | -30.76 | 0.18 | 0 | 0.12 | 28.71 | -0.82 | 28.65 | -0.23 | 28.48 | -1.41 | 0 | -17.93 | -0.06 | -45.29 | -0.06 | 0 | 0 | 0.06 | 0 | 0.18 | 0.12 | 28.42
70.912 | 0 | -30.7 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.27 | -2.05 | -0.06 | -15.18 | -0.06 | -27.36 | 0.06 | 0 | 0.06 | 28.77 | -0.41 | 28.54 | -0.23 | 28.89 | -0.94 | 6.68 | -15.7 | -0.06 | -39.32 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.18 | 0.12
70.944 | -0.06 | -27.89 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.33 | -1.35 | 0.64 | -13.36 | -0.06 | -27.01 | 0.06 | 0 | 0.12 | 28.83 | -0.47 | 28.65 | -0.23 | 28.59 | -0.76 | 14.41 | -14 | -0.18 | -36.74 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 0
70.976 | -0.06 | -26.13 | -0.18 | 0 | -0.12 | 0 | 0 | 21.39 | 0.35 | 21.33 | -0.64 | 8.44 | -11.48 | -0.06 | -26.31 | 0.18 | 0 | 0.23 | 28.54 | -0.47 | 28.59 | -0.23 | 28.65 | -0.35 | 22.38 | -11.95 | -0.06 | -37.21 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
71.008 | 0 | -24.43 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.29 | 21.5 | -0.35 | 13.83 | -9.61 | -0.06 | -24.14 | -0.06 | 0.06 | 0.18 | 28.59 | -0.35 | 28.65 | -0.29 | 28.48 | -0.29 | 29.06 | -9.9 | -0.18 | -38.61 | -0.18 | 0 | 0.06 | 0 | 0 | 0 | 0.12 | 0
71.04 | 0 | -21.74 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | -0.06 | 19.34 | -7.79 | 2.7 | -21.86 | 0.06 | 0 | 0.18 | 28.65 | -0.29 | 28.48 | -0.29 | 28.65 | -0.29 | 28.77 | -7.85 | -0.06 | -36.62 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 0
71.072 | -0.06 | -19.22 | -0.18 | -29.3 | -0.18 | 0 | -0.18 | 0.12 | 0.23 | 21.39 | 0.06 | 21.5 | -5.57 | 10.66 | -20.1 | 0.06 | 0 | 0.06 | 28.18 | -0.41 | 28.71 | -0.18 | 28.24 | -0.29 | 28.48 | -5.04 | 0 | -34.28 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
71.104 | -0.06 | -17.4 | -0.06 | -62.93 | -0.06 | -16.82 | -0.18 | 0 | 0.18 | 21.33 | 0.29 | 21.39 | -4.04 | 15.94 | -18.87 | 0.06 | 0 | 0.18 | 28.54 | -0.35 | 28.59 | -0.18 | 28.48 | -0.18 | 28.48 | -3.22 | 0 | -32.81 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0
71.136 | -0.06 | -16.41 | -0.18 | -44.83 | -0.18 | -63.58 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 21.39 | -2.52 | 21.68 | -16.7 | 0 | 0 | 0.06 | 0.29 | -0.35 | 28.77 | -0.23 | 28.3 | -0.29 | 28.65 | -2.11 | -0.18 | -30.7 | 0 | 0 | 0 | 0 | -0.18 | 0 | 0.12 | 0
71.168 | -0.06 | -14.36 | -0.18 | -40.37 | -0.06 | -47.7 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -1.82 | 21.27 | -15.23 | 0 | 0 | 0.18 | 0.12 | -0.35 | 28.65 | -0.23 | 28.42 | -0.23 | 28.48 | -1.17 | 0 | -29.24 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
71.2 | -0.06 | -12.77 | -0.18 | -38.5 | -0.18 | -40.14 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | 21.39 | -1.05 | 21.5 | -13.24 | 0 | 0.06 | 0.12 | 0 | -0.18 | 28.59 | -0.18 | 28.48 | -0.29 | 28.3 | -0.59 | 0 | -26.66 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
71.232 | -0.06 | -11.02 | -0.18 | -35.39 | -0.06 | -38.32 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -0.53 | 21.33 | -11.13 | 0.12 | -66.39 | 0.12 | 0.06 | -0.23 | 28.54 | -0.18 | 28.54 | -0.23 | 28.48 | -0.23 | 0 | -24.84 | 0 | -67.68 | -0.18 | -0.06 | -0.06 | 0 | 0 | 0
71.264 | 2.46 | -9.49 | 0 | -34.51 | -0.18 | -35.63 | -0.18 | 0.06 | 0.06 | 21.39 | 0.41 | 21.39 | -0.47 | 21.39 | -9.61 | -0.06 | -56.31 | 0.06 | 0 | -0.23 | 28.54 | -0.18 | 28.48 | -0.18 | 28.59 | -0.18 | 0 | -23.09 | -0.18 | -71.19 | -0.06 | -45.24 | -0.06 | 0 | 0 | 0
71.296 | 8.03 | -8.03 | -0.06 | -31.7 | -0.06 | -34.63 | -0.06 | 0 | -0.06 | 21.27 | 0.29 | 21.39 | -0.06 | 21.39 | -8.03 | 0.06 | -35.33 | 0.18 | 0 | -0.29 | 23.09 | -0.18 | 28.36 | -0.18 | 28.42 | -0.06 | 1.99 | -21.33 | -0.18 | -49.81 | -0.06 | -79.34 | -0.06 | -56.19 | 0.06 | 0
71.328 | 13.59 | -6.68 | -0.06 | -30.35 | -0.06 | -31.82 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.33 | -0.06 | 21.39 | -6.15 | 0.06 | -27.25 | 0.06 | 0 | -0.23 | 0.18 | -0.18 | 28.65 | -0.18 | 28.42 | 0.18 | 9.02 | -19.63 | 0 | -46.99 | -0.18 | -53.09 | -0.18 | -77.87 | 0 | -66.51
71.36 | 19.4 | -5.22 | -0.06 | -28.36 | -0.06 | -30.29 | 0 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.06 | 21.39 | -4.34 | 0.06 | -32.23 | 0.06 | 0 | -0.23 | 0 | -0.06 | 0.29 | -0.18 | 28.83 | 0.23 | 16.11 | -16.99 | 0 | -42.19 | -0.18 | -30.53 | -0.18 | -47.99 | 0 | -75.06
71.392 | 21.5 | -3.28 | -0.06 | -24.84 | -0.06 | -27.54 | -0.06 | 0 | 0 | 21.45 | 0.35 | 21.39 | 0.29 | 21.39 | -2.64 | 0.12 | -35.45 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.65 | 0.29 | 26.37 | -14.94 | -0.06 | -36.97 | -0.06 | -37.44 | -0.06 | -28.83 | 0 | -38.85
71.424 | 21.45 | -2.11 | -0.06 | -23.44 | -0.06 | -25.61 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.33 | 0.23 | 21.33 | -1.88 | 0.06 | -32.7 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.36 | 0.41 | 28.54 | -13.13 | 0 | -37.85 | -0.06 | -42.72 | -0.06 | -36.86 | 0 | -29.94
71.456 | 21.5 | -1.17 | -0.06 | -20.98 | -0.06 | -23.91 | -0.06 | 0 | 0.06 | 0.53 | 0.23 | 21.27 | 0.18 | 21.39 | -0.82 | -0.06 | -29.83 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.24 | 0.35 | 28.42 | -11.31 | 0 | -37.85 | -0.18 | -40.2 | -0.06 | -43.01 | 0 | -37.91
71.488 | 21.45 | -0.82 | -0.06 | -19.51 | -0.06 | -21.68 | -0.18 | -22.62 | -0.06 | 0.06 | 0.23 | 21.27 | 0.18 | 21.39 | -0.76 | 0.12 | -27.77 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.41 | 28.42 | -9.08 | -0.06 | -36.68 | -0.06 | -36.27 | -0.06 | -39.96 | 0 | -42.36
71.52 | 21.45 | -0.29 | -0.06 | -17.23 | -0.06 | -19.75 | -0.18 | -76.94 | -0.18 | 0 | 0.12 | 21.21 | 0.23 | 21.15 | -0.23 | 0 | -27.31 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.47 | 28.42 | -6.45 | 0 | -34.98 | -0.06 | -33.17 | -0.18 | -35.63 | 0 | -39.85
71.552 | 21.45 | -0.06 | -0.06 | -15.35 | -0.06 | -17.46 | -0.06 | -46.7 | -0.06 | 0.06 | 0.23 | 21.33 | 0.18 | 21.15 | -0.06 | -0.06 | -25.67 | 0.06 | -27.42 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.65 | 0.35 | 28.42 | -4.45 | -0.06 | -32.7 | 0 | -32.81 | -0.18 | -33.28 | -0.06 | -34.04
71.584 | 21.45 | -0.18 | -0.06 | -13.83 | -0.06 | -15.41 | -0.06 | -42.19 | -0.06 | 0 | 0.12 | 21.27 | 0.35 | 21.39 | 0 | -0.06 | -23.15 | 0.18 | -77.58 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 28.65 | -2.46 | 0.76 | -30.53 | -0.18 | -32.34 | 0 | -32.52 | -0.06 | -32.7
71.616 | 21.45 | -0.23 | 0 | -12.07 | 0 | -14.12 | -0.18 | -35.63 | -0.06 | 0 | 0.12 | 21.33 | 0.29 | 21.39 | -0.06 | -0.06 | -21.15 | 0.12 | -52.56 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.29 | 28.59 | -1.99 | 10.25 | -29.06 | -0.18 | -30.35 | -0.06 | -32.23 | 0 | -32.58
71.648 | 21.39 | -0.41 | -0.06 | -10.96 | -0.06 | -12.54 | -0.18 | -32.29 | -0.06 | 0 | 0.06 | 21.15 | 0.12 | 21.21 | 0.12 | -0.06 | -19.75 | 0.12 | -44.83 | -0.06 | 0 | 0 | 0 | -0.06 | 28.48 | 0.29 | 28.65 | -0.41 | 17.81 | -27.19 | -0.06 | -28.13 | -0.06 | -31.11 | -0.06 | -32.29
71.68 | 21.45 | -0.35 | 4.98 | -9.55 | -0.06 | -11.07 | -0.18 | -32.4 | 0 | 0 | 0.06 | 21.33 | 0.06 | 21.39 | 0 | 0 | -18.69 | 0.18 | -44.59 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.18 | 28.59 | -0.18 | 25.43 | -25.37 | -0.06 | -26.84 | -0.18 | -28.18 | -0.18 | -29.94
71.712 | 21.45 | -0.18 | 12.72 | -7.79 | -0.06 | -9.2 | -0.06 | -32.4 | 0 | 0 | 0.06 | 21.33 | 0.23 | 21.33 | 0.23 | 4.51 | -16.29 | 0.12 | -40.9 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | -0.06 | 28.65 | -23.15 | -0.06 | -25.55 | -0.06 | -26.13 | 0 | -27.19
71.744 | 21.45 | -0.23 | 18.28 | -6.33 | -0.06 | -7.91 | -0.18 | -31.41 | 0 | 0 | 0.06 | 21.27 | 0.12 | 21.39 | 0.23 | 10.55 | -13.95 | 0.12 | -40.61 | -0.18 | -67.62 | -0.18 | 0.06 | -0.18 | 28.48 | 0.18 | 28.54 | 0.06 | 28.54 | -21.21 | -0.06 | -23.91 | -0.18 | -25.31 | -0.18 | -26.02
71.776 | 21.39 | -0.23 | 21.62 | -4.63 | 0 | -6.62 | -0.06 | -28.95 | 0 | 0 | 0.06 | 21.45 | 0.12 | 21.39 | 0.23 | 15.94 | -12.6 | 0.18 | -38.32 | -0.06 | -64.69 | 0 | -40.31 | 0 | 28.65 | 0.18 | 28.65 | 0.12 | 28.54 | -19.63 | -0.06 | -21.91 | -0.06 | -23.44 | -0.06 | -24.96
71.808 | 21.45 | -0.41 | 21.33 | -3.46 | -0.18 | -4.75 | -0.06 | -26.66 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | 0.23 | 21.45 | -11.13 | 0.18 | -36.15 | -0.06 | -45.06 | -0.06 | -70.84 | -0.06 | 0.23 | 0.18 | 28.42 | 0.23 | 28.59 | -17.7 | -0.06 | -20.92 | -0.06 | -22.03 | 0 | -23.67
71.84 | 21.5 | -0.18 | 21.45 | -1.88 | 8.38 | -3.75 | -0.06 | -24.43 | -0.06 | 0 | -0.06 | 14.71 | 0.06 | 21.09 | 0.29 | 21.39 | -9.38 | 0.06 | -35.16 | -0.06 | -46.23 | 0 | -49.28 | -0.06 | 0.06 | 0.12 | 28.59 | 0.29 | 28.24 | -16.06 | 0 | -19.1 | -0.06 | -20.22 | -0.18 | -21.86
71.872 | 21.39 | -0.29 | 21.68 | -1.11 | 14 | -1.99 | 0.59 | -21.86 | 0.06 | -6.5 | -0.06 | 0.06 | 0.12 | 21.33 | 0.18 | 21.39 | -7.15 | 0.12 | -31.76 | -0.06 | -43.24 | -0.06 | -45.59 | 0 | 0 | 0.06 | 28.42 | 0.35 | 28.48 | -13.3 | -0.06 | -16.7 | -0.06 | -18.4 | 0 | -20.1
71.904 | 21.39 | -0.23 | 21.56 | -0.59 | 19.45 | -1.11 | 8.44 | -20.16 | 0.06 | -77.23 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | -5.39 | 0 | -30.06 | -0.06 | -41.49 | -0.06 | -44.24 | -0.06 | 0 | -0.06 | 28.36 | 0.41 | 28.48 | -10.84 | 0.94 | -15.23 | -0.18 | -16.64 | -0.18 | -18.34
71.936 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -0.64 | 14.12 | -17.93 | 0 | -54.67 | -0.06 | 0.06 | 0.06 | 0.18 | 0.23 | 21.39 | -4.28 | 0.12 | -28.3 | -0.06 | -39.44 | 0 | -42.01 | -0.06 | 0 | 0 | 28.59 | 0.23 | 28.3 | -9.02 | 10.78 | -14.18 | -0.06 | -15.18 | -0.18 | -16.93
71.968 | 21.5 | -0.23 | 21.5 | -0.18 | 21.5 | -0.35 | 19.45 | -16.52 | -0.06 | -42.72 | -0.18 | 0 | -0.06 | 0.06 | 0.23 | 21.39 | -2.64 | 0 | -26.54 | -0.06 | -37.03 | -0.06 | -39.32 | -0.06 | 0 | 0 | 28.59 | 0.35 | 28.48 | -6.56 | 18.16 | -12.01 | 0 | -13.59 | 7.03 | -14.94
72 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | -0.06 | 21.45 | -14.47 | -0.06 | -38.5 | -0.06 | 0 | 0 | 0 | 0.29 | 21.27 | -1.93 | 0.12 | -24.84 | -0.06 | -36.33 | -0.18 | -38.15 | -0.06 | 0 | 0 | 28.42 | 0.29 | 28.59 | -4.69 | 26.13 | -10.61 | 6.8 | -12.07 | 14.65 | -13.83
72.032 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.18 | 21.39 | -11.66 | -0.06 | -31.99 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -1.17 | 0.06 | -22.21 | -0.06 | -33.87 | 0 | -36.45 | -0.18 | 0 | 0 | 28.65 | 0.23 | 28.48 | -2.46 | 28.54 | -8.32 | 16.76 | -10.02 | 24.96 | -11.31
72.064 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | -0.41 | 21.39 | -9.84 | 0 | -32.29 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -0.29 | 0 | -20.39 | -0.18 | -31.58 | -0.06 | -34.45 | 0 | 0 | 0 | 28.01 | 0.12 | 28.65 | -1.88 | 28.54 | -6.86 | 24.61 | -7.97 | 28.83 | -10.08
72.096 | 21.45 | -0.18 | 21.39 | -0.35 | 21.45 | -0.35 | 21.39 | -7.68 | 0.06 | -32.34 | -0.18 | 0 | 0 | 0 | 0.23 | 21.39 | -0.18 | 0.06 | -18.52 | -0.18 | -29.59 | 0 | -32.23 | 0 | 0 | -0.18 | 28.54 | 0.06 | 28.65 | -0.94 | 28.65 | -4.86 | 28.83 | -6.68 | 28.71 | -8.38
72.128 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.39 | -5.39 | -0.06 | -31.52 | -0.06 | 0 | 0 | 0 | 0.29 | 21.39 | -0.23 | 0.06 | -16.76 | -0.06 | -27.13 | 0 | -29.59 | 0 | 0 | -0.06 | 28.65 | 0.18 | 28.42 | -0.59 | 28.48 | -3.75 | 28.48 | -5.33 | 28.65 | -6.45
72.16 | 14.88 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | -3.22 | 0.06 | -30 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.27 | 0.06 | 0.06 | -15.18 | -0.06 | -25.84 | 0 | -28.3 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 28.36 | -0.23 | 28.54 | -2.23 | 28.59 | -3.69 | 28.42 | -5.1
72.192 | 0.06 | -0.06 | 21.5 | -0.29 | 21.45 | -0.23 | 21.39 | -1.7 | 0 | -27.54 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0 | 0.06 | -13.07 | -0.06 | -23.15 | 0 | -25.37 | -0.06 | 0.06 | 0 | 13.77 | 0.12 | 28.42 | 0 | 28.48 | -1.17 | 28.65 | -2.23 | 28.48 | -3.28
72.224 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | -1.17 | 0.06 | -24.9 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.21 | 0.18 | 2.11 | -12.07 | -0.06 | -21.33 | 0 | -23.44 | 0 | -48.05 | 0 | 0.12 | -0.06 | 28.48 | 0.06 | 28.65 | -0.94 | 28.48 | -1.7 | 28.3 | -2.23
72.256 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.18 | 21.33 | -0.18 | 0.06 | -22.85 | -0.06 | -0.29 | 0 | 0 | 0.06 | 21.39 | 0.23 | 7.62 | -10.37 | -0.06 | -19.63 | 0.06 | -22.33 | -0.06 | -84.44 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.48 | -0.35 | 28.65 | -0.76 | 28.42 | -1.41
72.288 | 0 | -0.18 | 21.5 | -0.29 | 21.45 | -0.18 | 21.33 | -0.18 | 0.06 | -20.86 | -0.06 | -64.22 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 13.18 | -9.08 | 0 | -17.29 | 0 | -20.1 | -0.06 | -53.38 | 0 | 0 | 0 | 28.65 | 0.23 | 28.48 | -0.23 | 28.65 | -0.64 | 28.65 | -0.94
72.32 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.27 | -0.06 | 0.06 | -18.93 | 0 | -59.47 | -0.06 | 0 | 0 | 21.33 | 0.18 | 18.81 | -7.91 | -0.06 | -15.7 | 0 | -17.81 | 0 | -48.69 | -0.18 | 0 | 0.06 | 28.48 | 0.18 | 28.54 | -0.23 | 28.3 | -0.35 | 28.48 | -0.59
72.352 | 0 | -0.06 | 21.45 | -0.18 | 21.5 | -0.23 | 21.39 | 0.23 | 0.06 | -16.76 | 0 | -42.42 | 0 | -25.84 | 0.06 | 21.33 | 0.18 | 21.39 | -6.33 | 0 | -14.06 | 0 | -15.76 | 0.06 | -42.36 | 0 | 0 | -0.06 | 28.59 | 0.29 | 28.48 | -0.06 | 28.54 | 0 | 28.18 | -0.35
72.384 | 0 | -0.18 | 21.39 | -0.29 | 21.39 | -0.18 | 21.21 | 0.29 | -0.06 | -14.77 | -0.06 | -38.09 | -0.06 | -70.32 | 0.06 | 21.21 | 0.29 | 21.39 | -5.39 | 0 | -12.54 | 0 | -14.71 | 0 | -37.21 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.48 | -0.06 | 28.3 | -0.06 | 28.48 | -0.06
72.416 | 0 | -0.06 | 21.45 | -0.29 | 21.5 | -0.18 | 21.33 | 0.41 | 0 | -12.54 | -0.06 | -32.76 | 0 | -49.75 | 0.06 | 21.39 | 0.35 | 21.15 | -4.45 | -0.06 | -11.13 | 0 | -12.89 | 0 | -37.5 | -0.18 | 0 | 0 | 28.07 | 0.35 | 28.59 | 0.06 | 28.54 | 0 | 28.59 | -0.18
72.448 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | -0.23 | 21.39 | 0.35 | 0.06 | -10.37 | -0.06 | -32.29 | -0.06 | -29.77 | 0.06 | 21.39 | 0.29 | 21.45 | -3.75 | 0 | -9.79 | 6.86 | -10.96 | -0.06 | -37.68 | -0.18 | 0 | 0 | 28.65 | 0.23 | 28.48 | 0 | 28.59 | 0.12 | 28.3 | 0.06
72.48 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.47 | -0.06 | -8.61 | 0 | -32.64 | 0.12 | -28.59 | 0 | 21.39 | 0.29 | 21.39 | -2.99 | 1.11 | -8.56 | 14.41 | -9.9 | -0.06 | -37.21 | -0.06 | 0 | -0.06 | 22.91 | 0.23 | 28.48 | 0.18 | 28.48 | 0.12 | 28.18 | 0.18
72.512 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.33 | 0.35 | 3.34 | -5.74 | -0.06 | -30.82 | 0.06 | -35.45 | 0 | 21.15 | 0.23 | 21.39 | -2.23 | 11.19 | -6.8 | 24.79 | -8.38 | -0.06 | -34.98 | -0.06 | 0 | -0.18 | 0.12 | 0.18 | 28.48 | 0.12 | 28.54 | 0.12 | 28.59 | 0.12
72.544 | 0 | -0.06 | 21.5 | -0.18 | 21.39 | -0.18 | 21.39 | 0.29 | 9.55 | -3.46 | -0.06 | -30.24 | -0.06 | -34.75 | 0.12 | 21.39 | 0.18 | 21.33 | -1.82 | 18.87 | -5.68 | 28.59 | -7.15 | -0.06 | -32.93 | 0 | 0 | 0 | 0 | 0.18 | 28.48 | 0.29 | 28.65 | 0.12 | 28.65 | 0.12
72.576 | -0.29 | 0 | 0.12 | -0.18 | 21.45 | -0.18 | 21.33 | 0.41 | 14.77 | -1.93 | 0 | -27.54 | -0.06 | -32.17 | 0 | 21.39 | 0.18 | 21.39 | -1.46 | 26.43 | -4.69 | 28.54 | -5.68 | 2.46 | -30.94 | 0 | 0 | 0 | 0 | 0.06 | 28.42 | 0.12 | 28.42 | 0.29 | 28.59 | 0.29
72.608 | -49.1 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.27 | 0.47 | 20.51 | -1.52 | 0 | -25.37 | -0.18 | -28.59 | -0.06 | 0.12 | 0.23 | 21.39 | -1.17 | 28.65 | -3.87 | 28.36 | -4.92 | 10.37 | -29.06 | -0.18 | -0.06 | 0 | 0 | -0.06 | 28.48 | 0.23 | 28.65 | 0.18 | 28.42 | 0.12
72.64 | -52.5 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.53 | 21.45 | -0.7 | -0.06 | -23.79 | -0.18 | -27.01 | -0.06 | 0.06 | 0.18 | 21.33 | -0.94 | 28.89 | -2.87 | 28.54 | -3.63 | 17.81 | -27.72 | -0.06 | -66.8 | 0 | 0 | 0.12 | 27.95 | 0.18 | 28.65 | 0.23 | 29.12 | 0.23
72.672 | -39.61 | -0.06 | 0 | -0.18 | 21.39 | -0.06 | 21.39 | 0.35 | 21.39 | -0.18 | -0.18 | -20.63 | -0.06 | -26.6 | 0 | 0 | 0.23 | 21.33 | -0.59 | 28.48 | -2.23 | 28.65 | -2.93 | 28.3 | -25.55 | -0.06 | -63.75 | 0 | 0 | -0.06 | 28.65 | 0 | 28.77 | 0.18 | 28.48 | 0.18
72.704 | -39.67 | -0.06 | 0 | -0.18 | 14.06 | -0.18 | 21.39 | 0.35 | 21.39 | 0.06 | 0 | -18.87 | -0.06 | -24.55 | -0.06 | 0 | 0.18 | 21.39 | -0.59 | 28.48 | -1.82 | 28.3 | -2.4 | 28.83 | -23.61 | -0.18 | -48.4 | -0.06 | 0 | 0.12 | 28.48 | 0.06 | 28.48 | 0.12 | 28.54 | 0.18
72.736 | -36.45 | -0.06 | 0 | -0.06 | 0.12 | -0.18 | 21.21 | 0.29 | 21.39 | 0.23 | -0.06 | -17.05 | -0.06 | -22.85 | -0.06 | 0 | 0.29 | 21.39 | -0.47 | 28.48 | -1.46 | 28.59 | -1.64 | 28.59 | -21.62 | -0.06 | -45.24 | -0.18 | 0 | -0.06 | 28.3 | 0.12 | 28.42 | 0.18 | 28.48 | 0.23
72.768 | -34.98 | -0.06 | 0 | -0.06 | 0 | -0.18 | 13.54 | 0.29 | 21.39 | 0.35 | -0.06 | -15.18 | -0.06 | -20.74 | -0.06 | 0 | 0.29 | 21.39 | -0.53 | 28.48 | -1.17 | 29.01 | -1.17 | 28.59 | -19.86 | 0 | -40.14 | -0.18 | 0 | -0.18 | 28.48 | 0.06 | 28.54 | 0.12 | 28.48 | 0.12
72.8 | -33.4 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.18 | 21.39 | 0.41 | -0.06 | -13.18 | 0 | -19.04 | -0.06 | 0 | 0.29 | 21.27 | -0.35 | 28.42 | -0.94 | 28.65 | -0.76 | 28.59 | -18.28 | -0.06 | -35.92 | -0.06 | 0 | 0 | 17.52 | -0.06 | 28.42 | 0.18 | 28.54 | 0.18
72.832 | -29.59 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | -0.06 | -10.08 | -0.06 | -16.88 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.3 | -0.59 | 28.83 | -0.59 | 28.48 | -15.7 | -0.06 | -37.21 | -0.18 | 0 | 0 | 0.12 | 0.06 | 28.42 | 0.18 | 26.25 | 0.29
72.864 | -28.48 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | -0.06 | -8.32 | 0 | -15.41 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.48 | -0.41 | 28.89 | -0.47 | 28.54 | -13.54 | 0 | -38.09 | -0.06 | 0 | -0.18 | 0 | 0 | 27.54 | 0 | 0.18 | 0.18
72.896 | -26.13 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.21 | 0.35 | 1.99 | -5.92 | 0 | -13.89 | -0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.83 | -0.41 | 28.71 | -0.29 | 28.42 | -11.84 | -0.06 | -36.62 | 0 | 0.06 | -0.06 | 0.06 | 0 | 0.18 | 0.06 | 0 | 0.12
72.928 | -23.91 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.41 | 7.03 | -4.04 | -0.06 | -11.66 | -0.06 | 0 | 0.23 | 21.39 | -0.41 | 28.95 | -0.47 | 28.36 | -0.29 | 28.54 | -9.49 | -0.06 | -34.98 | -0.18 | -85.9 | -0.06 | 0 | 0.12 | 0 | 0.06 | 0 | 0.18
72.96 | -23.03 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.47 | 12.66 | -2.23 | -0.06 | -10.14 | -0.06 | 0 | 0.23 | 21.39 | -0.41 | 28.42 | -0.35 | 28.48 | -0.23 | 28.59 | -7.62 | 0.06 | -32.29 | -0.18 | -69.55 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.06
72.992 | -19.16 | -0.06 | -39.44 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 20.16 | -0.94 | -0.18 | -8.2 | 0 | 0 | 0.23 | 21.45 | -0.29 | 28.48 | -0.29 | 28.42 | -0.29 | 28.59 | -4.98 | 5.16 | -29.77 | 0 | -48.69 | 0.12 | 0 | 0 | 0 | 0.06 | 0 | 0.18
73.024 | -17.75 | 0 | -55.49 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.21 | 0.41 | 21.39 | -1 | 0 | -6.33 | -0.18 | -26.19 | 0.18 | 21.39 | -0.29 | 28.95 | -0.29 | 28.48 | -0.29 | 28.65 | -3.16 | 13.18 | -28.07 | -0.06 | -44.77 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.12
73.056 | -15.94 | 0 | -41.72 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -0.06 | -0.06 | -4.63 | -0.06 | -71.66 | 0.06 | 21.33 | -0.29 | 28.65 | -0.23 | 28.83 | -0.23 | 28.42 | -1.88 | 20.86 | -26.78 | -0.06 | -38.85 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0
73.088 | -14.71 | -0.06 | -40.14 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.33 | -0.06 | 3.75 | -3.16 | 0 | -49.87 | 0.06 | 21.39 | -0.29 | 28.59 | -0.23 | 28.54 | -0.23 | 28.65 | -0.94 | 28.65 | -25.25 | 0 | -36.51 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
73.12 | -12.89 | -0.06 | -37.91 | -0.06 | -0.53 | 0 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.23 | 9.9 | -2.17 | -0.06 | -30.94 | 0.18 | 19.4 | -0.18 | 28.54 | -0.18 | 28.48 | -0.23 | 29.06 | -0.64 | 28.71 | -23.15 | -0.18 | -38.32 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.06
73.152 | -11.02 | 0 | -36.04 | -0.06 | -50.98 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | 0.23 | 17.17 | -1.41 | -0.06 | -30.7 | 0.18 | 0.06 | -0.18 | 28.65 | -0.29 | 28.48 | -0.23 | 28.24 | -0.23 | 28.65 | -20.57 | -0.06 | -37.85 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | -0.06
73.184 | -9.79 | -0.06 | -33.05 | -0.18 | -51.92 | -0.06 | -1.23 | -0.06 | 21.39 | 0.06 | 21.33 | 0.41 | 21.8 | -0.59 | -0.06 | -35.27 | 0.06 | 0 | -0.23 | 28.48 | -0.18 | 28.54 | -0.23 | 28.42 | 0 | 28.65 | -18.63 | -0.06 | -35.92 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0
73.216 | -8.2 | -0.06 | -31.7 | -0.06 | -40.67 | 0 | -68.91 | -0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.45 | -0.35 | 0 | -34.45 | 0.23 | 0 | -0.23 | 28.3 | -0.18 | 28.36 | -0.18 | 28.54 | 0.06 | 28.59 | -17.46 | -0.18 | -34.28 | -0.23 | -1.76 | -0.06 | 0 | 0 | 0 | -0.06
73.248 | -6.97 | -0.06 | -28.77 | -0.06 | -38.97 | 0 | -58.65 | -0.06 | 21.39 | 0.06 | 21.39 | 0.53 | 21.39 | -0.18 | 0 | -31.7 | 0 | 0 | -0.18 | 28.59 | -0.18 | 28.77 | -0.23 | 28.42 | 0.12 | 28.42 | -15.18 | -0.06 | -32.7 | -0.06 | -83.21 | -0.06 | 0 | -0.06 | 0 | -0.06
73.28 | -5.39 | 0 | -27.19 | -0.06 | -36.97 | -0.06 | -43.89 | -0.06 | 21.45 | 0.12 | 21.45 | 0.47 | 21.33 | -0.18 | -0.06 | -28.65 | 0.12 | 0 | -0.06 | 28.48 | -0.18 | 28.3 | -0.18 | 28.48 | 0.23 | 28.65 | -13.36 | -0.06 | -30.35 | 0 | -69.26 | -0.18 | 0 | 0 | -44.94 | 0
73.312 | -3.4 | 0 | -24.84 | -0.06 | -34.92 | -0.06 | -37.68 | -0.06 | 21.39 | 0.06 | 21.39 | 0.47 | 21.39 | -0.06 | -0.06 | -27.6 | 0.12 | 0 | -0.23 | 28.48 | -0.18 | 27.77 | -0.06 | 28.48 | 0.29 | 28.59 | -10.31 | -0.18 | -27.77 | 0 | -33.87 | -0.18 | -68.56 | 0 | -71.37 | 0
73.344 | -2.46 | 2.29 | -22.44 | -0.06 | -33.17 | -0.06 | -33.46 | -0.06 | 21.39 | 0.12 | 21.45 | 0.41 | 21.39 | 0.06 | 4.57 | -26.37 | 0.06 | 0 | -0.18 | 28.42 | -0.18 | 0.18 | -0.06 | 28.59 | 0.47 | 28.59 | -8.14 | -0.06 | -26.13 | -0.18 | -30.7 | -0.18 | -71.84 | 0 | -44.65 | -0.06
73.376 | -1.46 | 7.68 | -20.57 | -0.06 | -30.76 | -0.06 | -32.76 | -0.06 | 21.39 | 0.06 | 21.33 | 0.41 | 21.39 | 0.12 | 10.55 | -25.43 | 0.06 | 0 | -0.18 | 13.65 | -0.18 | 0 | -0.18 | 28.42 | 0.47 | 28.59 | -6.21 | -0.18 | -24.61 | -0.18 | -41.02 | -0.23 | -45.06 | 0.06 | -29.3 | -0.06
73.408 | -0.82 | 13.65 | -18.87 | -0.06 | -28.95 | 0 | -32.64 | -0.18 | 0.53 | -0.06 | 21.39 | 0.35 | 21.33 | 0 | 16.11 | -22.68 | 0.06 | 0 | -0.18 | 0.12 | -0.18 | 0 | -0.18 | 28.54 | 0.47 | 28.54 | -4.34 | 0.06 | -22.79 | -0.18 | -42.54 | -0.06 | -28.95 | 0 | -35.98 | -0.06
73.44 | -0.18 | 18.99 | -16.76 | -0.06 | -25.72 | 0 | -31.29 | -0.06 | 0.06 | -0.06 | 21.39 | 0.35 | 21.39 | 0.41 | 21.45 | -21.09 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.36 | 0.47 | 28.59 | -2.64 | -0.18 | -21.33 | 0 | -37.91 | -0.18 | -35.57 | 0 | -43.01 | -0.06
73.472 | -0.29 | 21.5 | -14.71 | 0 | -24.08 | 0 | -29.77 | 0 | 0 | -0.06 | 21.27 | 0.23 | 21.33 | 0.29 | 21.39 | -19.4 | 0.06 | 0 | -0.06 | 0.06 | -0.06 | 0 | -0.18 | 0.29 | 0.41 | 28.65 | -1.29 | 9.08 | -18.87 | -0.18 | -33.4 | -0.18 | -43.24 | -0.06 | -38.97 | 0
73.504 | -0.23 | 21.45 | -13.07 | -0.06 | -22.21 | -0.06 | -26.9 | 0 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.35 | 21.5 | -17.4 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0.12 | 0.35 | 28.65 | -0.76 | 16 | -16.58 | 0 | -32.76 | -0.06 | -39.38 | -0.06 | -34.92 | 0
73.536 | -0.23 | 21.45 | -11.78 | -0.06 | -19.98 | -0.06 | -25.08 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.15 | 0.23 | 21.39 | -15.35 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.35 | 28.54 | -0.47 | 23.91 | -14.41 | -0.18 | -32.81 | -0.18 | -35.1 | -0.06 | -33.11 | 0
73.568 | -0.23 | 21.5 | -10.2 | -0.06 | -17.99 | -0.06 | -23.56 | 0 | 0 | 0.06 | 21.39 | 0.12 | 21.33 | 0.23 | 21.39 | -13.77 | 0 | -53.03 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0.23 | 28.89 | 0 | 28.89 | -12.89 | 0 | -31.47 | 0 | -32.81 | 0 | -33.17 | -0.06
73.6 | -0.35 | 21.45 | -8.91 | -0.06 | -15.94 | -0.18 | -21.56 | 0 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.35 | 21.33 | -12.42 | 0.06 | -60.71 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.65 | -0.06 | 28.48 | -10.96 | -0.06 | -29.77 | -0.18 | -32.76 | -0.18 | -32.76 | -0.06
73.632 | -0.35 | 21.45 | -6.97 | -0.06 | -14.53 | -0.06 | -18.52 | 0 | 0 | -0.06 | 21.45 | 0.18 | 21.27 | 0.29 | 21.39 | -10.14 | 0 | -33.28 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.65 | 0.23 | 28.65 | -8.56 | 0 | -27.48 | -0.18 | -31.99 | -0.18 | -29.77 | 0
73.664 | -0.23 | 21.45 | -5.57 | -0.06 | -13.13 | -0.06 | -16.7 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.27 | 0.35 | 21.39 | -8.32 | -0.06 | -28.24 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.42 | 0.29 | 28.59 | -6.33 | -0.18 | -25.43 | -0.18 | -29.65 | -0.18 | -28.24 | -0.18
73.696 | -0.35 | 21.45 | -4.04 | -0.06 | -11.37 | -0.06 | -14.53 | -0.06 | 0 | 0 | 21.21 | 0.12 | 21.39 | 0.29 | 21.45 | -6.62 | -0.06 | -33.63 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 28.65 | 0.23 | 28.42 | -4.39 | -0.06 | -24.61 | -0.06 | -27.83 | -0.18 | -27.01 | -0.18
73.728 | -0.23 | 21.39 | -2.81 | -0.06 | -9.9 | 0 | -12.89 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.21 | 0.18 | 21.39 | -4.86 | 0.06 | -35.27 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.3 | 0.23 | 28.65 | -2.29 | 0 | -23.03 | -0.18 | -26.72 | -0.06 | -25.78 | -0.06
73.76 | -0.23 | 21.39 | -1.76 | 1.23 | -8.2 | -0.06 | -11.13 | 0 | 0 | -0.06 | 21.39 | -0.06 | 21.39 | 0.06 | 21.27 | -3.28 | -0.06 | -33.28 | -0.18 | 0 | 0 | -38.61 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.71 | -1.76 | -0.06 | -21.62 | -0.06 | -25.78 | -0.06 | -23.85 | 0
73.792 | -0.18 | 21.39 | -0.94 | 11.13 | -7.03 | 3.57 | -8.32 | -0.06 | 0 | -0.06 | 6.91 | -0.06 | 21.33 | 0.18 | 21.45 | -2.17 | 0 | -29.24 | -0.06 | -1.99 | -0.06 | -64.34 | -0.06 | 0 | 0 | 28.54 | 0.29 | 28.42 | -0.64 | -0.06 | -19.34 | -0.06 | -23.2 | -0.06 | -21.74 | 0
73.824 | -0.29 | 21.45 | -0.41 | 16.41 | -5.39 | 9.9 | -5.92 | 0 | -11.48 | -0.06 | 0.12 | -0.06 | 21.33 | 0.12 | 21.39 | -1.35 | -0.06 | -27.72 | -0.06 | -70.67 | 0 | -46.17 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.48 | -0.35 | 0 | -17.75 | -0.18 | -21.33 | -0.18 | -20.27 | 0.12
73.856 | -0.23 | 21.45 | -0.18 | 21.86 | -3.98 | 15.18 | -3.98 | -0.06 | -76 | -0.06 | 0 | -0.06 | 21.21 | 0.23 | 21.15 | -0.64 | -0.06 | -26.78 | -0.18 | -59.12 | -0.06 | -45.7 | 0 | 0 | 0.06 | 28.65 | 0.18 | 28.36 | 0 | 0.06 | -16.29 | -0.06 | -20.1 | -0.18 | -19.04 | -0.18
73.888 | -0.23 | 21.45 | -0.18 | 21.45 | -2.64 | 20.8 | -2.17 | 0 | -54.03 | -0.06 | 0 | -0.18 | 21.33 | -0.06 | 21.45 | -0.53 | 1.99 | -25.61 | -0.18 | -44.53 | 0 | -44.06 | -0.18 | 0 | 0 | 0.29 | 0.35 | 28.59 | 0 | 0 | -14.53 | -0.18 | -18.87 | -0.18 | -16.99 | -0.06
73.92 | -0.18 | 21.45 | -0.35 | 21.74 | -1.41 | 21.5 | -1.7 | -0.06 | -43.3 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.39 | -0.29 | 6.15 | -23.2 | -0.06 | -45.24 | 0 | -41.54 | 0 | -67.74 | 0 | 0.12 | 0.23 | 28.65 | 0.18 | 5.45 | -12.77 | 0 | -17.46 | -0.18 | -15.53 | -0.06
73.952 | -0.23 | 21.45 | -0.35 | 21.5 | -0.76 | 21.45 | -0.41 | 0.06 | -36.68 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | -0.06 | 13.83 | -20.98 | -0.06 | -41.9 | -0.06 | -39.61 | -0.06 | -54.67 | -0.06 | 0 | 0.29 | 28.3 | 0.23 | 15.94 | -10.9 | 6.33 | -14.94 | -0.06 | -13.48 | 0
73.984 | -0.23 | 21.45 | -0.53 | 21.45 | -0.23 | 21.45 | -0.23 | 0 | -32.52 | 0 | 0 | -0.06 | 0.12 | 0 | 21.33 | 0.06 | 19.28 | -19.63 | -0.06 | -41.02 | 0.06 | -37.68 | -0.18 | -45.24 | 0 | 0 | 0.12 | 28.59 | 0.23 | 23.67 | -9.2 | 14.18 | -12.95 | 7.03 | -11.84 | -0.06
74.016 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.18 | -0.06 | -33.05 | -0.06 | 0 | 0 | 0 | 0 | 21.09 | 0.06 | 21.45 | -18.16 | -0.06 | -38.79 | -0.06 | -35.74 | -0.06 | -45.76 | -0.06 | 0 | 0 | 28.48 | 0.41 | 28.95 | -7.44 | 22.03 | -11.37 | 14.65 | -10.08 | 9.08
74.048 | -0.18 | 21.39 | -0.29 | 21.45 | -0.18 | 21.45 | 0.06 | 0 | -32.93 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.15 | -16.82 | -0.06 | -37.38 | -0.06 | -33.63 | -0.06 | -43.77 | -0.06 | 0 | 0.29 | 28.42 | 0.35 | 28.71 | -5.92 | 28.95 | -10.14 | 22.33 | -9.02 | 16.23
74.08 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.29 | -0.06 | -31.47 | -0.06 | 0 | -0.06 | 0.06 | 0.06 | 21.39 | 0.12 | 21.33 | -14.94 | -0.18 | -35.63 | -0.06 | -32.4 | 0 | -41.54 | -0.06 | 0 | 0 | 28.48 | 0.41 | 28.42 | -4.45 | 28.48 | -8.56 | 29.06 | -7.44 | 23.91
74.112 | -0.18 | 21.39 | -0.29 | 21.45 | -0.23 | 21.45 | 0.29 | 0 | -29.42 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | -12.77 | -0.06 | -32.4 | -0.06 | -28.89 | -0.06 | -38.91 | 0 | 0 | -0.06 | 28.36 | 0.41 | 28.42 | -2.52 | 28.01 | -6.5 | 28.71 | -5.1 | 28.71
74.144 | -0.06 | 21.45 | -0.29 | 21.45 | -0.41 | 21.45 | 0.41 | -0.06 | -27.01 | 0 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.12 | 21.33 | -11.31 | -0.18 | -31.29 | 0 | -28.13 | 0 | -37.21 | -0.06 | 0 | 0 | 28.42 | 0.47 | 28.36 | -1.99 | 28.83 | -4.45 | 28.54 | -3.69 | 28.42
74.176 | -0.06 | 21.45 | -0.23 | 21.39 | -0.35 | 21.45 | 0.29 | 3.34 | -25.14 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -9.2 | -0.06 | -29.3 | 0 | -25.78 | 0 | -35.1 | 0 | 0 | 0 | 28.48 | 0.12 | 28.54 | -0.94 | 28.42 | -3.11 | 28.42 | -2.29 | 28.48
74.208 | -0.18 | 21.45 | -0.23 | 21.45 | -0.41 | 21.5 | 0.35 | 9.61 | -22.79 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.29 | 21.09 | -7.38 | -0.06 | -27.31 | -0.18 | -23.91 | -0.06 | -33.46 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.42 | -0.64 | 28.65 | -2.29 | 28.83 | -1.7 | 28.42
74.24 | -0.18 | 17.7 | -0.23 | 21.39 | -0.35 | 21.45 | 0.35 | 14.94 | -21.09 | 0 | -56.31 | -0.06 | 0 | -0.06 | 0.12 | 0.29 | 21.33 | -5.8 | 0.41 | -25.08 | -0.18 | -21.86 | 0 | -31.47 | 0 | 0 | 0 | 28.36 | 0.12 | 28.48 | -0.35 | 28.48 | -1.52 | 28.54 | -1 | 28.42
74.272 | -0.18 | 0.06 | -0.29 | 21.39 | -0.23 | 21.39 | 0.47 | 21.74 | -18.46 | 0 | -54.9 | 0.06 | 0 | -0.06 | 0 | 0.29 | 21.39 | -3.93 | 12.54 | -23.15 | 0 | -19.16 | 0 | -28.95 | 0 | 0 | 0 | 28.65 | 0.06 | 28.59 | -0.06 | 28.59 | -0.59 | 28.42 | -0.53 | 28.42
74.304 | -0.06 | 0 | -0.23 | 21.45 | -0.18 | 21.45 | 0.41 | 21.39 | -16.7 | -0.06 | -43.24 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | -2.4 | 20.22 | -20.92 | -0.18 | -17.7 | -0.06 | -26.66 | 0 | -25.84 | -0.06 | 28.18 | 0.06 | 28.54 | 0 | 28.59 | -0.53 | 28.48 | -0.7 | 28.48
74.336 | -0.06 | 0 | -0.23 | 21.5 | -0.18 | 21.45 | 0.41 | 21.39 | -14.47 | -0.06 | -38.15 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -2.17 | 28.18 | -18.99 | 0 | -15.47 | -0.06 | -25.25 | 0 | -91.12 | -0.06 | 28.48 | 0.06 | 28.42 | 0 | 28.36 | -0.29 | 28.48 | 0 | 28.42
74.368 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.29 | 21.39 | -12.48 | -0.06 | -33.05 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.21 | -0.64 | 28.65 | -17.23 | 0.12 | -14.41 | -0.06 | -23.67 | 0 | -57.13 | 0 | 0.18 | 0.06 | 28.59 | 0.12 | 28.48 | -0.18 | 28.83 | 0 | 28.71
74.4 | -0.06 | 0 | -0.18 | 21.39 | -0.41 | 21.39 | 0.35 | 21.39 | -10.61 | -0.06 | -32.17 | -0.06 | -18.16 | 0 | 0 | 0.12 | 21.15 | -0.76 | 28.65 | -15.29 | -0.06 | -12.83 | 0.06 | -21.27 | -0.06 | -48.87 | -0.06 | 0 | 0.12 | 28.59 | 0.18 | 28.54 | 0 | 28.48 | 0.06 | 28.3
74.432 | -0.06 | 0 | -0.18 | 21.45 | -0.35 | 21.39 | 0.35 | 21.39 | -7.56 | -0.18 | -32.99 | -0.06 | -66.39 | 0 | 0.06 | 0.18 | 21.21 | -0.53 | 28.71 | -13.54 | 9.32 | -11.31 | 0 | -18.46 | 0 | -43.13 | -0.06 | 0 | 0 | 28.83 | 0.12 | 28.59 | 0 | 28.3 | 0.06 | 28.36
74.464 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.23 | 21.39 | -5.39 | -0.06 | -31.99 | -0.06 | -44.77 | -0.06 | 0 | 0.18 | 21.39 | -0.18 | 28.42 | -12.31 | 16.58 | -9.84 | 6.27 | -17.23 | -0.06 | -38.56 | -0.06 | 0 | 0 | 29.06 | 0.23 | 28.48 | 0 | 28.54 | 0.12 | 28.13
74.496 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.12 | 21.39 | -3.69 | -0.06 | -29.77 | -0.06 | -28.18 | -0.06 | 0 | 0.18 | 21.39 | 0 | 28.42 | -10.66 | 24.32 | -8.56 | 14.12 | -15.53 | -0.18 | -37.38 | -0.06 | 0 | 0.06 | 28.77 | 0.23 | 28.36 | 0.18 | 28.65 | 0.06 | 28.18
74.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.06 | 21.33 | -1.82 | 0 | -27.48 | 0.06 | -30.47 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 28.59 | -9.84 | 28.89 | -7.56 | 21.97 | -13.89 | 0 | -38.44 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.54 | 0.29 | 28.42 | 0.06 | 28.65
74.56 | 0 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.06 | 21.21 | -1.41 | 1.99 | -25.43 | -0.06 | -35.27 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.59 | -8.2 | 28.3 | -6.21 | 29.12 | -12.25 | -0.06 | -37.68 | -0.18 | 0 | 0.06 | 28.65 | 0.18 | 28.59 | 0.12 | 28.71 | 0.06 | 28.65
74.592 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.06 | 21.33 | -0.47 | 8.85 | -23.09 | 0.06 | -33.93 | -0.06 | 0 | 0.23 | 21.39 | 0 | 28.54 | -6.8 | 28.59 | -4.75 | 28.71 | -10.66 | -0.06 | -35.98 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.48 | 0.23 | 28.48 | 0.12 | 28.65
74.624 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.06 | 21.27 | -0.35 | 14.3 | -20.92 | 0 | -30.59 | 0 | 0 | 0.18 | 21.39 | 0.12 | 29.06 | -5.68 | 28.71 | -3.87 | 28.65 | -9.32 | -0.06 | -33.69 | 0 | 0 | 0 | 28.65 | 0 | 28.48 | 0.18 | 28.13 | 0.18 | 28.3
74.656 | -0.06 | -0.12 | -0.18 | 0.12 | -0.18 | 21.45 | 0 | 21.33 | 0.12 | 19.75 | -19.16 | -0.18 | -27.6 | -0.06 | -30.41 | 0.12 | 21.39 | 0.06 | 28.36 | -4.51 | 28.65 | -3.16 | 28.48 | -7.79 | -0.18 | -31.64 | -0.06 | 0 | -0.06 | 28.65 | 0.06 | 28.42 | 0.18 | 28.65 | 0.12 | 28.42
74.688 | -0.06 | -52.03 | -0.18 | 0 | -0.18 | 0.29 | 0.12 | 21.33 | 0.18 | 21.45 | -17.46 | -0.06 | -27.54 | 0 | -70.96 | 0.23 | 21.39 | 0.06 | 28.71 | -3.57 | 28.65 | -2.58 | 28.83 | -6.5 | -0.18 | -29.71 | 0 | 0 | -0.06 | 28.65 | 0.06 | 28.65 | 0.29 | 28.42 | 0.18 | 28.71
74.72 | -0.06 | -54.14 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.33 | 0.35 | 21.45 | -15.7 | -0.18 | -26.25 | -0.06 | -48.17 | 0.12 | 21.15 | 0.12 | 28.48 | -2.99 | 28.48 | -1.99 | 28.71 | -5.74 | 0 | -27.6 | 0 | 0 | 0.06 | 28.48 | 0 | 28.65 | 0.12 | 28.54 | 0.12 | 28.42
74.752 | -0.06 | -39.32 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.41 | 21.39 | -12.83 | -0.06 | -23.67 | 0 | -27.77 | 0.18 | 21.39 | 0.23 | 28.59 | -2.29 | 28.65 | -1.23 | 28.42 | -4.28 | -0.06 | -25.61 | -0.18 | 0 | -0.06 | 28.48 | -0.06 | 28.48 | 0.12 | 28.48 | 0.23 | 28.48
74.784 | -0.06 | -38.91 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.33 | 0.35 | 21.39 | -10.84 | -0.06 | -22.09 | -0.06 | -30.35 | 0.18 | 14.77 | 0.23 | 28.42 | -1.76 | 28.48 | -1 | 28.65 | -3.4 | -0.18 | -23.97 | -0.06 | -67.27 | 0 | 28.54 | 0.06 | 28.48 | 0.06 | 28.54 | 0.12 | 28.48
74.816 | -0.06 | -36.92 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -8.79 | 0 | -19.86 | -0.06 | -35.68 | 0.18 | 0.12 | 0.06 | 28.36 | -1.35 | 28.89 | -0.76 | 28.83 | -2.58 | -0.06 | -22.21 | -0.06 | -76.35 | -0.06 | 0.23 | 0 | 28.48 | 0.06 | 28.59 | 0.06 | 28.48
74.848 | -0.06 | -35.22 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.33 | 0.29 | 21.27 | -6.68 | -0.06 | -18.75 | -0.06 | -34.63 | 0.12 | 0 | 0.06 | 28.77 | -1 | 28.48 | -0.53 | 28.42 | -2.17 | 0 | -20.27 | -0.06 | -50.63 | -0.18 | 0.06 | -0.06 | 0.29 | 0 | 28.48 | 0.12 | 28.48
74.88 | -0.06 | -33.4 | -0.18 | 0 | -0.18 | 0 | 0 | 21.27 | 0.47 | 21.39 | -4.22 | -0.18 | -17.4 | 0 | -31.23 | 0.23 | 0 | 0.18 | 28.42 | -0.82 | 28.48 | -0.41 | 28.59 | -1.29 | -0.06 | -18.16 | 0 | -45.35 | 0.06 | 0 | 0 | 0.06 | 0.06 | 0.18 | 0.06 | 28.48
74.912 | 0 | -30.12 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.27 | 0.41 | 21.39 | -1.93 | -0.06 | -14.82 | -0.06 | -28.13 | 0.18 | 0 | 0.18 | 28.48 | -0.59 | 28.77 | -0.18 | 28.36 | -0.94 | 6.33 | -15.29 | 0 | -40.2 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0.12
74.944 | -0.06 | -29.01 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.41 | 21.39 | -1.41 | 0.41 | -13.07 | 0 | -27.01 | 0.06 | 0 | 0.23 | 28.36 | -0.41 | 28.59 | -0.29 | 28.42 | -0.7 | 14.18 | -14.24 | -0.18 | -36.45 | -0.06 | 0 | 0 | 0.06 | -0.18 | 0 | 0.06 | 0
74.976 | -0.06 | -26.25 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.33 | -0.7 | 8.03 | -11.48 | 0 | -25.9 | 0.06 | 0 | 0.12 | 28.54 | -0.41 | 28.65 | -0.29 | 28.42 | -0.47 | 22.03 | -11.84 | -0.18 | -37.38 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
75.008 | -0.06 | -24.84 | 0 | 0 | -0.18 | 0 | -0.06 | 21.21 | 0.35 | 21.39 | -0.41 | 13.48 | -9.9 | 0 | -24.38 | 0.06 | 0 | 0.12 | 28.48 | -0.41 | 28.42 | -0.23 | 28.42 | -0.23 | 29.12 | -10.14 | -0.06 | -38.15 | 0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12 | 0
75.04 | -0.06 | -22.85 | -0.06 | 0 | -0.18 | 0 | 0 | 21.27 | 0.29 | 21.27 | 0 | 19.04 | -7.79 | 2.4 | -22.44 | 0.06 | 0 | 0.12 | 28.65 | -0.35 | 28.54 | -0.41 | 28.24 | -0.35 | 28.48 | -7.85 | -0.18 | -36.86 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
75.072 | 0 | -19.63 | -0.06 | -26.54 | -0.18 | 0 | -0.06 | 0.12 | 0.35 | 21.39 | 0 | 21.62 | -5.86 | 10.14 | -19.98 | 0.12 | 0 | 0.23 | 28.54 | -0.23 | 28.59 | -0.29 | 28.48 | -0.35 | 28.65 | -5.22 | 0 | -35.16 | -0.06 | 0 | -0.23 | 0 | 0 | 0 | 0.06 | 0
75.104 | 0 | -17.81 | -0.06 | -62.7 | -0.06 | -13.24 | 0 | 0.06 | 0.23 | 21.39 | 0.18 | 21.39 | -4.22 | 15.41 | -18.52 | 0 | 0 | 0.12 | 28.59 | -0.23 | 28.48 | -0.23 | 28.83 | -0.35 | 28.65 | -3.16 | -0.06 | -32.52 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
75.136 | 0 | -16.06 | -0.06 | -45.94 | -0.06 | -62.7 | -0.18 | 0 | 0.18 | 21.39 | 0.41 | 21.39 | -2.81 | 21.33 | -17.05 | -0.06 | 0 | 0.12 | 0.29 | -0.18 | 28.83 | -0.23 | 28.54 | -0.29 | 28.59 | -2.23 | -0.06 | -30.94 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
75.168 | 0 | -14.82 | -0.06 | -39.9 | -0.18 | -48.28 | -0.06 | 0.06 | 0.23 | 21.45 | 0.41 | 21.39 | -2.46 | 21.39 | -15.18 | 0 | 0 | 0.12 | 0.12 | -0.29 | 28.54 | -0.18 | 28.59 | -0.35 | 28.65 | -0.94 | 0 | -29.06 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
75.2 | 0 | -13.18 | -0.06 | -38.85 | -0.06 | -40.08 | -0.18 | 0 | 0 | 21.21 | 0.41 | 21.39 | -1.11 | 21.39 | -13.95 | 0 | -0.06 | 0.18 | 0 | -0.29 | 28.83 | -0.29 | 28.59 | -0.41 | 28.48 | -0.94 | -0.06 | -26.72 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
75.232 | 0 | -11.13 | -0.06 | -35.86 | -0.06 | -38.38 | -0.18 | 0 | 0.06 | 21.39 | 0.41 | 21.39 | -0.47 | 21.39 | -11.31 | -0.06 | -61.29 | 0.12 | 0 | -0.18 | 28.59 | -0.18 | 28.95 | -0.18 | 28.18 | -0.18 | 0 | -24.9 | -0.06 | -62.58 | -0.18 | 0 | 0 | 0 | 0 | 0
75.264 | 2.34 | -9.67 | -0.06 | -34.1 | -0.06 | -35.86 | -0.18 | 0.06 | 0 | 21.33 | 0.47 | 21.21 | -0.41 | 21.33 | -10.14 | 0.06 | -58.71 | 0.18 | 0 | -0.23 | 28.54 | -0.18 | 28.42 | -0.18 | 28.83 | -0.18 | 0 | -22.91 | -0.06 | -74.42 | -0.18 | -37.21 | -0.06 | 0.06 | 0 | 0
75.296 | 7.73 | -8.32 | -0.18 | -32.05 | -0.06 | -35.22 | -0.23 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.18 | 21.27 | -8.32 | 0.06 | -37.97 | 0.12 | 0 | -0.18 | 28.59 | -0.23 | 28.42 | -0.18 | 28.42 | 0.12 | -0.06 | -21.97 | 0 | -51.45 | -0.06 | -83.56 | -0.06 | -47.93 | 0 | 0.06
75.328 | 13.54 | -6.91 | -0.06 | -30.53 | 0 | -32.58 | -0.06 | 0 | 0.06 | 21.33 | 0.41 | 21.39 | 0.06 | 21.39 | -6.86 | 0 | -27.77 | 0.06 | 0 | -0.18 | 0.18 | -0.23 | 28.54 | -0.18 | 28.83 | 0.06 | 8.32 | -19.69 | -0.06 | -47.99 | 0 | -54.85 | 0 | -81.21 | -0.06 | -61.35
75.36 | 19.04 | -5.68 | -0.06 | -28.36 | -0.06 | -29.94 | -0.23 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | 0 | 21.15 | -4.92 | 0 | -32.46 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0.23 | -0.18 | 28.3 | 0.23 | 15.35 | -17.46 | 0 | -42.25 | -0.06 | -31.88 | 0 | -50.22 | 0 | -76.76
75.392 | 21.62 | -3.63 | -0.06 | -25.37 | -0.06 | -28.3 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0 | 21.33 | -2.87 | -0.06 | -35.57 | 0.06 | 0 | -0.23 | 0 | -0.23 | 0 | -0.18 | 28.65 | 0.18 | 25.61 | -15.47 | -0.06 | -36.92 | 0 | -35.92 | -0.06 | -28.36 | 0 | -39.79
75.424 | 21.45 | -2.29 | -0.06 | -23.79 | -0.06 | -25.61 | -0.06 | 0 | -0.18 | 21.21 | 0.29 | 21.33 | 0.29 | 21.39 | -2.23 | 0.12 | -32.93 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.65 | 0.35 | 28.71 | -13.48 | -0.06 | -37.68 | -0.06 | -42.13 | -0.18 | -36.39 | -0.06 | -29.42
75.456 | 21.45 | -1.64 | -0.06 | -21.15 | -0.06 | -23.67 | -0.06 | 0 | -0.12 | 10.78 | 0.18 | 21.33 | 0.18 | 21.39 | -1.23 | -0.06 | -29.71 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 29.18 | 0.35 | 28.48 | -11.31 | -0.06 | -38.09 | -0.06 | -40.61 | 0 | -42.78 | 0 | -37.79
75.488 | 21.39 | -0.94 | -0.06 | -19.22 | -0.06 | -21.86 | -0.18 | -3.52 | 0 | 0.06 | 0.18 | 21.39 | 0.29 | 21.15 | -0.64 | -0.06 | -28.13 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.35 | 28.42 | -9.61 | 0 | -36.74 | -0.06 | -36.21 | -0.06 | -40.55 | 0 | -42.72
75.52 | 21.5 | -0.29 | 0 | -17.46 | -0.06 | -19.86 | -0.06 | -76.47 | -0.06 | 0 | 0.23 | 21.15 | 0.23 | 21.39 | -0.35 | -0.06 | -27.01 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.35 | 28.65 | -6.33 | -0.06 | -34.98 | -0.06 | -33.69 | -0.06 | -36.39 | 0 | -39.96
75.552 | 21.39 | -0.18 | -0.18 | -15.29 | 0 | -17.81 | -0.06 | -49.28 | 0 | 0.06 | 0.12 | 21.39 | 0.35 | 21.27 | -0.18 | 0.06 | -25.37 | 0.12 | -18.46 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.89 | 0.35 | 28.36 | -4.45 | 0 | -33.22 | 0 | -33.17 | -0.06 | -33.52 | -0.06 | -34.63
75.584 | 21.45 | -0.18 | -0.06 | -13.95 | 0 | -15.53 | -0.06 | -42.6 | -0.06 | 0 | 0.18 | 21.33 | 0.29 | 21.33 | 0 | -0.06 | -22.85 | 0.12 | -77.58 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.29 | 28.59 | -2.58 | 2.05 | -30.41 | -0.06 | -32.93 | 0 | -32.93 | 0 | -33.17
75.616 | 21.45 | -0.29 | -0.18 | -12.25 | 0 | -14 | -0.18 | -36.04 | -0.18 | 0 | 0.12 | 21.21 | 0.29 | 21.39 | -0.06 | -0.06 | -22.09 | 0.12 | -54.38 | -0.18 | 0 | 0 | 0 | -0.18 | 28.65 | 0.29 | 28.59 | -1.76 | 9.32 | -28.59 | -0.18 | -31.29 | -0.06 | -32.23 | 0 | -33.28
75.648 | 21.5 | -0.35 | -0.06 | -11.07 | 0 | -12.48 | 0 | -31.99 | -0.06 | 0 | 0.06 | 21.15 | 0.29 | 21.33 | 0.12 | -0.18 | -20.39 | 0.18 | -44.3 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.65 | -1.23 | 16.64 | -27.19 | -0.06 | -28.48 | -0.18 | -30.53 | 0 | -32.64
75.68 | 21.5 | -0.35 | 4.45 | -9.38 | -0.06 | -11.02 | -0.06 | -32.46 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.18 | -0.06 | -18.87 | 0.12 | -44.24 | 0 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.36 | -0.41 | 24.32 | -25.2 | -0.06 | -27.01 | -0.18 | -27.95 | 0 | -30.24
75.712 | 21.45 | -0.35 | 12.42 | -7.56 | -0.06 | -9.43 | -0.06 | -32.93 | -0.06 | 0 | 0.06 | 21.45 | 0.12 | 21.27 | 0.23 | 3.98 | -16.29 | 0.18 | -41.54 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.54 | 0.18 | 28.48 | -0.29 | 28.65 | -23.38 | -0.06 | -25.37 | -0.18 | -26.66 | -0.06 | -27.48
75.744 | 21.45 | -0.41 | 17.81 | -6.09 | -0.06 | -8.14 | -0.06 | -31.23 | -0.06 | 0 | 0 | 21.33 | 0.06 | 21.27 | 0.18 | 10.14 | -14.36 | 0.18 | -40.02 | -0.06 | -55.02 | -0.06 | 0 | -0.18 | 28.54 | 0.12 | 28.48 | 0.06 | 28.48 | -21.68 | -0.06 | -23.67 | -0.18 | -25.43 | 0 | -26.37
75.776 | 21.45 | -0.23 | 21.74 | -4.45 | -0.06 | -6.45 | -0.06 | -29.24 | 0 | 0 | 0.06 | 21.27 | 0.06 | 21.33 | 0.23 | 15.41 | -13.18 | 0.12 | -37.97 | -0.18 | -67.85 | -0.18 | -33.05 | -0.06 | 28.54 | 0.06 | 28.77 | 0.12 | 28.48 | -19.75 | -0.06 | -22.62 | -0.06 | -24.08 | 0 | -25.67
75.808 | 21.45 | -0.23 | 21.5 | -3.4 | -0.06 | -4.75 | -0.18 | -27.13 | 0 | 0 | 0 | 21.27 | 0 | 21.5 | 0.35 | 21.15 | -11.54 | 0.12 | -36.04 | -0.06 | -45.82 | -0.06 | -75.47 | -0.06 | 0.23 | 0.12 | 28.42 | 0.06 | 28.71 | -17.4 | -0.18 | -21.09 | -0.06 | -22.38 | -0.18 | -23.61
75.84 | 21.45 | -0.18 | 21.45 | -2.23 | 7.44 | -3.57 | -0.12 | -24.79 | -0.06 | 0 | -0.06 | 17.52 | 0.12 | 21.33 | 0.23 | 21.56 | -9.67 | 0.12 | -35.1 | -0.06 | -45.76 | -0.06 | -51.74 | -0.18 | 0.12 | 0 | 28.48 | 0.41 | 28.54 | -15.47 | -0.06 | -19.16 | -0.18 | -20.63 | -0.18 | -22.15
75.872 | 21.45 | -0.23 | 21.45 | -1.05 | 13.24 | -1.88 | 0.18 | -22.33 | -0.06 | -0.82 | 0 | 0.06 | 0 | 21.39 | 0.18 | 21.39 | -7.44 | 0.06 | -32.52 | -0.06 | -43.48 | 0 | -46.17 | -0.06 | 0 | 0.12 | 28.48 | 0.29 | 28.42 | -12.83 | -0.06 | -17.52 | -0.06 | -18.87 | 0 | -19.98
75.904 | 21.39 | -0.18 | 21.5 | -0.53 | 18.63 | -1.05 | 7.56 | -20.63 | 0 | -74.07 | 0.12 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -5.22 | 0.12 | -30.41 | -0.06 | -41.49 | -0.06 | -44.59 | -0.18 | 0 | 0.12 | 28.18 | 0.29 | 28.65 | -11.13 | 0.18 | -16 | 0 | -17.29 | -0.06 | -18.87
75.936 | 21.39 | -0.18 | 21.39 | -0.29 | 21.56 | -0.53 | 13.18 | -18.16 | -0.18 | -58.6 | 0 | 0 | 0.06 | 0.18 | 0.18 | 21.39 | -3.93 | 0.06 | -28.59 | 0 | -39.67 | 0 | -41.31 | 0 | 0 | -0.06 | 28.71 | 0.29 | 28.3 | -9.32 | 9.96 | -14 | 0 | -15.53 | 0 | -16.99
75.968 | 21.39 | -0.23 | 21.45 | -0.35 | 21.45 | -0.47 | 18.81 | -16.41 | -0.18 | -42.78 | -0.06 | 0 | 0.12 | 0.06 | 0.23 | 21.39 | -2.64 | 0.12 | -26.84 | -0.06 | -37.97 | -0.06 | -40.37 | -0.06 | 0 | -0.06 | 28.59 | 0.23 | 28.48 | -7.03 | 16.99 | -12.54 | -0.06 | -13.89 | 6.09 | -15.18
76 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.23 | 21.62 | -14.24 | -0.06 | -38.15 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | -1.88 | 0.12 | -24.84 | -0.06 | -35.68 | -0.06 | -38.2 | 0 | 0 | -0.06 | 28.71 | 0.29 | 28.54 | -4.69 | 24.9 | -11.02 | 6.33 | -12.36 | 14 | -13.54
76.032 | 21.45 | -0.18 | 21.5 | -0.35 | 21.45 | -0.23 | 21.45 | -11.89 | -0.06 | -32.46 | -0.06 | 0 | 0 | 0.06 | 0.18 | 21.39 | -1.11 | 0 | -22.38 | -0.06 | -33.87 | -0.06 | -35.92 | 0 | 0 | -0.06 | 28.48 | 0.23 | 28.42 | -2.64 | 28.42 | -8.73 | 16.7 | -10.14 | 24.08 | -11.48
76.064 | 21.5 | -0.18 | 21.45 | -0.23 | 21.5 | -0.35 | 21.39 | -10.02 | -0.06 | -32.64 | -0.18 | 0 | -0.06 | 0 | 0.12 | 21.45 | -0.29 | 0.06 | -20.98 | 0 | -31.7 | -0.06 | -34.22 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 28.59 | -1.7 | 28.65 | -6.8 | 24.32 | -8.44 | 28.83 | -9.67
76.096 | 21.39 | -0.18 | 21.5 | -0.41 | 21.45 | -0.41 | 21.39 | -7.62 | -0.06 | -32.76 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | -0.06 | -18.69 | 0.06 | -29.77 | -0.06 | -32.23 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | -0.94 | 28.36 | -5.22 | 28.71 | -6.86 | 28.59 | -8.14
76.128 | 21.45 | -0.18 | 21.5 | -0.35 | 21.45 | -0.41 | 21.39 | -5.92 | 0.12 | -31.99 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.18 | 0 | -16.88 | 0 | -28.13 | -0.06 | -30.24 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.42 | -0.7 | 28.59 | -3.63 | 28.54 | -5.1 | 28.65 | -6.45
76.16 | 17.58 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | -3.34 | 0 | -29.94 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.5 | -0.18 | 0.12 | -15 | 0 | -25.78 | 0 | -28.54 | -0.06 | 0 | -0.06 | 28.3 | 0.06 | 28.42 | -0.06 | 28.36 | -2.4 | 28.54 | -3.69 | 28.42 | -5.1
76.192 | 0.06 | -0.18 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | -1.88 | -0.06 | -27.31 | -0.06 | 0 | 0 | 0 | 0.18 | 21.21 | 0.06 | -0.06 | -13.36 | 0 | -24.08 | -0.06 | -25.84 | 0 | 0 | -0.06 | 20.51 | 0.18 | 28.48 | 0 | 28.48 | -1 | 28.42 | -2.4 | 28.65 | -3.34
76.224 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | -1 | -0.18 | -24.9 | 0 | 0 | -0.06 | 0 | 0.18 | 21.21 | 0 | 1.99 | -12.01 | 0 | -21.68 | 0 | -24.43 | 0 | -43.24 | 0 | 0.18 | 0.06 | 28.48 | 0.29 | 28.3 | -1 | 28.48 | -1.46 | 28.48 | -2.46
76.256 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -0.53 | -0.06 | -23.2 | 0 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 6.68 | -10.61 | -0.06 | -19.28 | -0.06 | -22.74 | -0.06 | -87.07 | -0.06 | 0 | 0.06 | 28.36 | 0.23 | 28.54 | -0.53 | 28.71 | -0.64 | 28.48 | -1.11
76.288 | 0 | -0.06 | 21.5 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | -0.06 | -21.15 | -0.18 | -60.35 | -0.06 | 0 | 0.06 | 21.33 | 0.06 | 12.48 | -9.55 | 0.06 | -17.75 | -0.06 | -19.81 | 0 | -54.38 | 0 | 0 | 0.12 | 28.48 | 0.35 | 28.65 | -0.41 | 28.54 | -0.53 | 28.77 | -0.94
76.32 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | -0.06 | 0.06 | -19.04 | -0.18 | -60.88 | 0.06 | 0 | 0.06 | 21.27 | 0.06 | 17.99 | -8.03 | 0.06 | -15.82 | 0.06 | -17.99 | -0.06 | -49.16 | -0.06 | 0 | -0.06 | 28.71 | 0.29 | 28.48 | -0.18 | 28.3 | -0.29 | 28.42 | -0.7
76.352 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | -0.06 | -16.7 | -0.06 | -42.48 | -0.06 | -17.46 | 0 | 21.39 | 0.23 | 21.39 | -6.33 | 0 | -14.24 | -0.06 | -16.29 | 0 | -42.13 | 0 | 0 | 0 | 28.48 | 0.35 | 28.54 | 0 | 28.59 | -0.18 | 28.48 | -0.18
76.384 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | -0.23 | 21.39 | 0.23 | -0.18 | -14.82 | -0.06 | -37.56 | -0.06 | -74.07 | 0.18 | 21.39 | 0.18 | 21.39 | -5.45 | -0.06 | -12.54 | 0 | -14.36 | 0 | -36.15 | 0 | 0 | -0.06 | 28.59 | 0.23 | 28.42 | 0 | 28.59 | 0 | 28.54 | -0.18
76.416 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.35 | 21.39 | 0.29 | -0.06 | -13.24 | -0.06 | -32.93 | -0.06 | -51.97 | 0.06 | 21.33 | 0.18 | 21.39 | -4.45 | 0 | -11.48 | 0 | -12.89 | 0.06 | -36.51 | -0.18 | 0 | 0 | 28.36 | 0.35 | 28.48 | -0.06 | 28.18 | 0.06 | 28.54 | 0
76.448 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | -0.35 | 21.39 | 0.35 | 0.06 | -10.61 | -0.06 | -32.29 | -0.06 | -30.76 | 0.06 | 21.39 | 0.29 | 21.39 | -3.57 | -0.06 | -9.96 | 6.45 | -11.66 | 0 | -38.15 | -0.06 | 0 | -0.06 | 28.65 | 0.35 | 28.65 | 0.12 | 28.48 | 0 | 28.48 | 0.06
76.48 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.29 | 21.39 | 0.41 | -0.18 | -8.73 | 0 | -32.76 | -0.18 | -29.12 | -0.06 | 21.39 | 0.12 | 21.21 | -2.99 | 1.99 | -8.44 | 14.12 | -10.55 | 0 | -37.44 | 0 | 0 | -0.06 | 28.65 | 0.23 | 28.42 | 0.29 | 28.71 | 0.12 | 28.42 | 0
76.512 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.41 | 2.7 | -5.63 | -0.06 | -31.64 | -0.06 | -35.74 | 0 | 21.39 | 0.12 | 21.39 | -2.29 | 10.25 | -6.91 | 24.43 | -8.56 | 0 | -35.51 | 0 | 0 | 0 | 0.12 | 0.18 | 28.48 | 0.18 | 28.65 | 0.06 | 28.48 | 0.06
76.544 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.06 | 21.39 | 0.47 | 8.79 | -3.75 | -0.18 | -30 | -0.18 | -34.75 | -0.06 | 21.21 | 0.23 | 21.45 | -1.82 | 17.87 | -5.74 | 28.89 | -7.56 | -0.06 | -33.52 | -0.06 | 0 | -0.18 | 0 | 0.12 | 28.48 | 0.18 | 28.59 | 0.12 | 28.36 | 0.18
76.576 | -0.12 | -0.06 | 0.12 | -0.06 | 21.45 | -0.18 | 21.39 | 0.41 | 14.24 | -2.05 | 0 | -27.25 | -0.06 | -31.41 | 0 | 21.33 | 0.23 | 21.39 | -1.46 | 25.43 | -4.8 | 28.89 | -6.09 | 0.23 | -31.29 | -0.06 | 0 | 0 | 0.06 | 0.06 | 28.3 | 0.18 | 28.36 | 0.18 | 28.18 | 0.12
76.608 | -46.76 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 19.86 | -1.7 | 0.06 | -25.43 | -0.06 | -28.36 | -0.06 | 0.23 | 0.18 | 21.39 | -1.17 | 28.65 | -4.04 | 28.48 | -5.33 | 9.79 | -29.77 | -0.06 | 0 | 0.12 | 0 | 0.12 | 28.89 | 0.12 | 28.59 | 0.06 | 28.48 | 0.18
76.64 | -53.67 | 0 | 0 | -0.23 | 21.45 | -0.18 | 21.39 | 0.41 | 21.56 | -0.59 | -0.06 | -23.15 | -0.06 | -28.01 | -0.06 | 0.06 | 0.12 | 21.27 | -0.94 | 28.59 | -3.16 | 28.54 | -3.98 | 16.99 | -27.95 | -0.06 | -58.01 | -0.06 | 0 | 0.12 | 28.59 | 0.12 | 28.42 | 0.18 | 28.42 | 0.18
76.672 | -39.96 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.35 | 21.39 | -0.29 | -0.06 | -21.33 | 0 | -26.43 | 0 | 0 | 0.18 | 21.39 | -0.7 | 28.48 | -2.29 | 28.36 | -3.16 | 27.36 | -25.55 | -0.06 | -67.74 | -0.06 | 0 | 0.06 | 28.65 | 0.06 | 28.42 | 0.18 | 28.59 | 0.12
76.704 | -38.97 | -0.06 | 0 | -0.06 | 20.04 | -0.06 | 21.45 | 0.29 | 21.45 | 0.12 | -0.06 | -19.04 | -0.18 | -24.55 | 0.06 | 0 | 0.12 | 21.33 | -0.53 | 28.65 | -2.11 | 28.42 | -2.4 | 28.71 | -24.38 | -0.18 | -48.63 | -0.18 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | 0.12 | 28.42 | 0.12
76.736 | -36.92 | -0.06 | 0 | -0.06 | 0.12 | -0.18 | 21.39 | 0.29 | 21.39 | 0.06 | -0.06 | -16.52 | -0.18 | -22.21 | -0.06 | 0 | 0.12 | 21.39 | -0.41 | 28.54 | -1.41 | 28.48 | -1.88 | 28.65 | -22.74 | -0.18 | -45.88 | 0 | 0 | -0.06 | 28.42 | 0 | 28.95 | 0.18 | 28.48 | 0.18
76.768 | -35.27 | 0 | 0 | -0.18 | 0 | -0.06 | 19.81 | 0.29 | 21.39 | 0.23 | -0.06 | -15 | 0 | -20.57 | 0 | 0 | 0.06 | 21.33 | -0.35 | 28.65 | -1.05 | 28.42 | -1.29 | 28.71 | -20.39 | -0.18 | -41.02 | 0 | 0 | -0.18 | 28.48 | 0 | 28.42 | 0.06 | 28.48 | 0.18
76.8 | -33.17 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.06 | 0.29 | 21.39 | 0.23 | -0.06 | -13.3 | -0.18 | -19.28 | 0 | 0 | 0.18 | 21.39 | -0.53 | 28.48 | -0.7 | 28.48 | -1 | 28.48 | -18.05 | -0.12 | -37.21 | -0.18 | 0 | -0.06 | 26.08 | 0 | 28.48 | 0.12 | 28.48 | 0.06
76.832 | -30.47 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | -0.06 | -10.37 | -0.06 | -17.75 | 0.06 | 0 | 0.06 | 21.39 | -0.35 | 28.54 | -0.59 | 28.59 | -0.59 | 28.71 | -15.7 | -0.06 | -37.74 | 0 | 0 | 0 | 0.12 | 0 | 28.48 | 0.06 | 28.42 | 0.12
76.864 | -28.77 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.45 | 0.47 | -0.06 | -8.67 | -0.18 | -15.7 | -0.06 | 0 | 0.06 | 21.39 | -0.35 | 28.65 | -0.47 | 28.48 | -0.41 | 28.18 | -13.95 | -0.18 | -37.68 | -0.06 | 0.06 | 0 | 0 | -0.06 | 28.48 | 0.12 | 0.23 | 0.23
76.896 | -26.13 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.47 | 0 | -6.21 | 0 | -13.89 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.3 | -0.29 | 28.48 | -0.35 | 28.71 | -12.31 | 0 | -36.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.18 | 0.18 | 0 | 0.12
76.928 | -24.67 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 21.45 | 0.53 | 6.74 | -3.93 | -0.06 | -12.07 | 0.06 | 0 | 0.06 | 21.33 | -0.29 | 28.48 | -0.29 | 28.13 | -0.29 | 28.48 | -9.9 | -0.18 | -34.98 | 0 | -73.25 | 0.12 | 0 | 0 | 0 | 0.12 | 0 | 0.18
76.96 | -22.38 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.53 | 12.36 | -2.4 | -0.06 | -10.61 | -0.06 | 0 | 0.23 | 21.39 | -0.41 | 28.59 | -0.35 | 28.42 | -0.29 | 28.48 | -7.91 | -0.06 | -32.64 | 0 | -74.59 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18
76.992 | -19.4 | -0.06 | -35.98 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 19.57 | -1.05 | -0.06 | -8.14 | 0 | 0 | 0.18 | 21.45 | -0.29 | 28.83 | -0.29 | 28.71 | -0.23 | 28.36 | -5.22 | 3.81 | -30.29 | 0 | -47.99 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.23
77.024 | -18.16 | 0 | -57.54 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.82 | -0.06 | -6.91 | 0 | -20.63 | 0.06 | 21.39 | -0.35 | 28.65 | -0.23 | 28.59 | -0.29 | 28.59 | -3.16 | 12.13 | -28.48 | -0.06 | -45.82 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
77.056 | -16.23 | 0 | -42.66 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | -0.23 | -0.06 | -5.04 | -0.06 | -73.66 | 0.12 | 21.39 | -0.23 | 28.65 | -0.35 | 28.59 | -0.23 | 28.59 | -1.99 | 19.92 | -26.72 | -0.06 | -40.72 | 0 | 0 | 0 | 0 | 0 | 0 | 0
77.088 | -14.53 | 0 | -39.96 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0 | 2.4 | -3.69 | 0 | -50.86 | 0.18 | 21.33 | -0.18 | 28.36 | -0.35 | 28.71 | -0.18 | 28.24 | -1.41 | 27.36 | -25.31 | -0.06 | -37.21 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0
77.12 | -13.18 | -0.06 | -37.97 | -0.18 | -0.06 | -0.18 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.06 | 9.49 | -2.23 | -0.06 | -32.23 | 0.18 | 21.39 | -0.29 | 28.42 | -0.18 | 28.83 | -0.23 | 28.48 | -0.47 | 28.59 | -23.56 | -0.06 | -37.56 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18
77.152 | -11.07 | 0 | -35.27 | -0.06 | -41.9 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.39 | 0.18 | 16.7 | -1.29 | 0.06 | -29.77 | 0.06 | 0.06 | -0.35 | 28.59 | -0.23 | 28.65 | -0.35 | 28.65 | -0.35 | 28.59 | -21.09 | -0.06 | -37.44 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.06
77.184 | -9.61 | 0 | -33.69 | -0.18 | -54.44 | -0.06 | 0 | -0.06 | 21.45 | 0.12 | 21.39 | 0.35 | 21.86 | -0.82 | 0.06 | -35.51 | 0.12 | 0 | -0.23 | 28.42 | -0.18 | 28.54 | -0.18 | 28.42 | 0.06 | 28.48 | -19.34 | -0.06 | -35.86 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0
77.216 | -8.32 | 0 | -31 | -0.06 | -41.66 | -0.18 | -60.53 | -0.06 | 21.39 | 0.06 | 21.39 | 0.41 | 21.39 | -0.47 | 0 | -34.86 | 0.12 | 0 | -0.23 | 28.54 | -0.06 | 29.06 | -0.23 | 28.59 | 0.06 | 28.42 | -16.88 | 0 | -34.04 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
77.248 | -6.68 | -0.06 | -29.47 | -0.18 | -39.44 | -0.06 | -62.05 | -0.06 | 21.39 | 0 | 21.5 | 0.41 | 21.39 | -0.47 | -0.06 | -31.99 | 0.12 | 0 | -0.23 | 28.42 | -0.23 | 28.89 | -0.18 | 28.42 | 0.29 | 28.3 | -15.47 | -0.06 | -32.23 | -0.06 | -74.65 | -0.18 | 0.06 | -0.06 | 0 | 0.12
77.28 | -5.1 | 0 | -27.89 | -0.06 | -37.03 | 0.06 | -45.65 | -0.06 | 21.39 | 0.12 | 21.39 | 0.41 | 21.45 | 0 | 0 | -29.01 | 0.06 | 0 | -0.18 | 28.48 | -0.06 | 28.65 | -0.18 | 28.65 | 0.47 | 28.42 | -13.36 | -0.06 | -30.24 | 0 | -73.36 | 0 | 0 | -0.06 | -39.38 | 0
77.312 | -3.69 | 0 | -24.79 | -0.06 | -35.92 | -0.06 | -38.97 | -0.06 | 21.39 | 0 | 21.39 | 0.29 | 21.39 | 0.06 | 0 | -27.36 | 0.18 | 0 | -0.18 | 28.3 | -0.06 | 28.42 | -0.18 | 28.59 | 0.29 | 28.3 | -10.9 | 0 | -28.48 | 0 | -36.74 | 0 | -63.58 | 0 | -73.66 | 0
77.344 | -2.46 | -0.06 | -22.91 | -0.06 | -32.76 | 0 | -33.69 | 0 | 21.39 | 0.12 | 21.39 | 0.35 | 21.39 | 0.06 | 3.75 | -26.31 | 0 | 0 | -0.18 | 28.48 | -0.18 | 0.29 | -0.18 | 28.36 | 0.41 | 28.65 | -8.79 | -0.06 | -26.54 | 0 | -29.77 | -0.18 | -72.48 | 0 | -46.17 | 0
77.376 | -1.41 | 7.09 | -20.86 | -0.06 | -31.17 | 0 | -32.05 | 0 | 21.39 | 0.12 | 21.45 | 0.41 | 21.33 | 0.12 | 9.96 | -24.55 | 0.06 | 0 | -0.18 | 23.15 | -0.18 | 0 | -0.06 | 28.59 | 0.35 | 28.65 | -6.86 | 0 | -24.84 | -0.18 | -40.55 | -0.18 | -45.59 | 0 | -29.42 | 0
77.408 | -0.76 | 13.01 | -18.52 | -0.18 | -29.12 | -0.06 | -32.46 | -0.06 | 10.55 | -0.06 | 21.39 | 0.29 | 21.45 | 0.12 | 15.18 | -22.38 | 0.06 | 0 | -0.18 | 0.18 | -0.18 | 0 | -0.06 | 28.65 | 0.47 | 28.3 | -4.16 | 0 | -23.61 | -0.06 | -42.36 | -0.18 | -28.95 | 0 | -34.63 | 0
77.44 | -0.29 | 18.28 | -17.52 | -0.06 | -26.84 | 0 | -31.58 | 0 | 0.12 | 0.06 | 21.45 | 0.23 | 21.33 | 0.29 | 20.74 | -20.86 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.47 | 28.42 | -2.75 | 1.99 | -21.39 | -0.18 | -38.85 | -0.18 | -35.39 | 0.06 | -42.78 | 0
77.472 | -0.18 | 21.5 | -14.53 | -0.06 | -24.38 | -0.18 | -29.77 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.39 | 0.29 | 21.39 | -19.1 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 1.35 | 0.35 | 28.42 | -1.7 | 7.68 | -18.69 | -0.06 | -35.33 | -0.18 | -42.36 | 0 | -39.85 | 0
77.504 | -0.18 | 21.5 | -13.59 | -0.06 | -21.97 | -0.06 | -27.83 | -0.06 | 0 | 0 | 21.45 | 0.18 | 21.33 | 0.12 | 21.33 | -17.4 | 0 | 0 | -0.18 | 0 | 0 | 0 | -0.06 | 0.12 | 0.35 | 28.42 | -0.94 | 15.12 | -16.82 | 0 | -33.34 | -0.06 | -39.44 | -0.06 | -35.16 | 0.06
77.536 | -0.18 | 21.45 | -11.78 | -0.06 | -20.92 | -0.06 | -25.72 | -0.18 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.18 | 21.27 | -15.65 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.35 | 28.42 | -0.35 | 22.91 | -15.29 | -0.06 | -32.34 | -0.18 | -35.86 | 0 | -32.99 | -0.06
77.568 | -0.41 | 21.45 | -10.78 | -0.06 | -18.34 | -0.06 | -23.38 | 0 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | 0.29 | 21.39 | -14.06 | -0.06 | -44.12 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.29 | 28.48 | -0.23 | 29.01 | -12.89 | 0 | -31.99 | -0.18 | -33.05 | 0 | -32.4 | 0
77.6 | -0.47 | 21.45 | -9.38 | -0.06 | -16.29 | -0.06 | -21.39 | -0.06 | 0 | 0 | 21.21 | 0.06 | 21.27 | 0.29 | 21.33 | -12.6 | -0.06 | -63.46 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.41 | 28.65 | 0 | 28.65 | -10.96 | -0.06 | -29.94 | -0.18 | -32.58 | 0.06 | -32.29 | -0.06
77.632 | -0.41 | 21.45 | -7.15 | -0.18 | -14.82 | -0.06 | -18.87 | 0 | 0 | 0.06 | 21.45 | 0.12 | 21.27 | 0.23 | 21.45 | -10.2 | -0.18 | -34.28 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 28.77 | 0.12 | 28.71 | -8.38 | 0 | -27.25 | 0 | -32.34 | 0.06 | -30.35 | 0.06
77.664 | -0.35 | 21.45 | -5.92 | -0.06 | -13.07 | 0 | -16.93 | -0.18 | 0 | -0.06 | 21.39 | 0.12 | 21.27 | 0.29 | 21.39 | -8.32 | 0 | -27.48 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.23 | 28.59 | 0.23 | 28.48 | -6.21 | -0.18 | -25.72 | -0.06 | -30 | 0 | -28.3 | 0
77.696 | -0.29 | 21.39 | -4.22 | -0.06 | -11.78 | -0.06 | -15.29 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.21 | 0.23 | 21.45 | -6.97 | -0.06 | -32.7 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 28.65 | 0.23 | 28.42 | -4.45 | -0.06 | -24.55 | -0.18 | -28.3 | 0 | -27.01 | 0
77.728 | -0.23 | 21.5 | -3.28 | -0.06 | -10.37 | -0.06 | -13.24 | -0.18 | 0 | -0.18 | 21.39 | -0.06 | 21.39 | 0.18 | 21.27 | -4.69 | 0 | -35.92 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 28.65 | 0.29 | 28.3 | -2.46 | -0.06 | -23.5 | -0.06 | -26.6 | 0 | -26.08 | 0
77.76 | -0.23 | 21.5 | -2.11 | 0.59 | -8.38 | -0.06 | -11.07 | -0.18 | 0 | 0 | 21.39 | 0 | 21.39 | 0.18 | 21.45 | -3.63 | 0 | -34.04 | -0.06 | 0 | 0 | -36.45 | 0 | 0 | 0.06 | 28.59 | 0.29 | 28.65 | -1.64 | -0.06 | -21.91 | -0.06 | -25.37 | 0 | -24.08 | 0.06
77.792 | -0.23 | 21.45 | -1 | 10.31 | -7.09 | 3.34 | -8.32 | -0.06 | 0 | -0.06 | 10.43 | -0.06 | 21.33 | 0.12 | 21.21 | -2.17 | -0.06 | -29.53 | -0.06 | -0.12 | 0 | -65.04 | -0.06 | 0 | 0.12 | 28.48 | 0.29 | 28.48 | -0.7 | 0 | -20.16 | -0.06 | -23.56 | -0.18 | -21.68 | -0.06
77.824 | -0.18 | 21.45 | -0.59 | 15.7 | -5.39 | 9.73 | -5.86 | -0.18 | -2.52 | -0.18 | 0.06 | -0.06 | 21.27 | 0.12 | 21.39 | -1.46 | -0.06 | -27.72 | 0 | -60.82 | -0.06 | -46.12 | -0.06 | 0 | 0.06 | 28.59 | 0.29 | 28.48 | -0.59 | -0.18 | -17.75 | -0.06 | -21.91 | 0.06 | -20.39 | 0
77.856 | -0.23 | 21.45 | -0.18 | 21.45 | -4.45 | 14.94 | -4.1 | -0.06 | -70.61 | 0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | -0.94 | -0.06 | -26.72 | -0.18 | -62.7 | 0 | -46.82 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.48 | -0.23 | 0 | -16.29 | 0 | -20.27 | 0 | -19.4 | 0
77.888 | -0.23 | 21.45 | -0.18 | 21.5 | -2.87 | 20.68 | -2.23 | -0.06 | -57.72 | 0 | 0 | -0.18 | 21.21 | 0.12 | 21.39 | -0.64 | -0.06 | -25.61 | -0.06 | -46 | -0.06 | -44.59 | 0 | 0 | -0.06 | 0.76 | 0.35 | 28.54 | 0 | -0.06 | -14.82 | -0.06 | -18.52 | -0.06 | -17.64 | 0
77.92 | -0.29 | 21.45 | -0.29 | 21.45 | -1.52 | 21.5 | -1.76 | -0.06 | -43.71 | -0.18 | 0 | 0 | 21.33 | 0 | 21.39 | -0.35 | 5.8 | -24.08 | -0.18 | -44.83 | -0.06 | -41.66 | -0.06 | -61.29 | -0.06 | 0.12 | 0.12 | 28.65 | 0 | 5.1 | -13.01 | -0.06 | -17.4 | 0 | -15.82 | 0.06
77.952 | -0.29 | 21.45 | -0.29 | 21.45 | -0.94 | 21.5 | -0.59 | -0.06 | -36.92 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -0.18 | 13.54 | -21.62 | -0.06 | -42.89 | -0.06 | -39.02 | -0.06 | -56.02 | 0 | 0 | 0.12 | 28.42 | 0.18 | 15.41 | -11.25 | 5.45 | -15 | 0 | -13.71 | 0
77.984 | -0.23 | 21.39 | -0.41 | 21.45 | -0.53 | 21.5 | -0.29 | -0.06 | -32.76 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.33 | -0.18 | 19.04 | -19.92 | -0.06 | -40.84 | -0.06 | -37.33 | 0 | -45.7 | 0 | 0.06 | 0.18 | 28.48 | 0.29 | 23.44 | -9.08 | 13.59 | -13.54 | 5.57 | -12.36 | 0
78.016 | -0.23 | 21.45 | -0.47 | 21.5 | -0.18 | 21.5 | 0.06 | 0.06 | -32.4 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.39 | 0.06 | 21.39 | -18.52 | -0.18 | -38.85 | -0.06 | -35.86 | 0 | -46.23 | 0 | 0 | 0.12 | 28.48 | 0.23 | 28.83 | -7.21 | 21.09 | -11.84 | 13.65 | -10.55 | 7.73
78.048 | -0.23 | 21.45 | -0.35 | 21.45 | -0.18 | 21.39 | 0.06 | -0.18 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.21 | -0.06 | 21.39 | -16.88 | -0.18 | -37.33 | -0.06 | -34.69 | 0 | -43.77 | 0 | 0 | 0.06 | 28.59 | 0.23 | 28.83 | -5.92 | 29.01 | -10.02 | 21.39 | -8.67 | 15.06
78.08 | -0.18 | 21.45 | -0.35 | 21.45 | -0.29 | 21.45 | 0.35 | -0.06 | -31.52 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.15 | 0.06 | 21.45 | -14.94 | 0 | -35.33 | -0.06 | -32.29 | -0.06 | -41.02 | 0 | 0 | 0 | 28.42 | 0.29 | 28.65 | -4.63 | 28.54 | -8.32 | 29.18 | -7.27 | 23.2
78.112 | -0.06 | 21.45 | -0.23 | 21.45 | -0.41 | 21.39 | 0.29 | -0.18 | -30 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.45 | -12.83 | 0 | -33.05 | 0 | -29.47 | 0 | -38.85 | 0 | 0 | 0.06 | 28.42 | 0.35 | 28.54 | -2.75 | 28.42 | -6.45 | 28.42 | -5.16 | 28.54
78.144 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | 0.41 | -0.06 | -27.6 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.27 | -11.13 | 0 | -31.99 | 0 | -28.01 | -0.18 | -37.56 | 0.12 | 0 | 0 | 28.65 | 0.35 | 28.42 | -1.88 | 28.48 | -4.63 | 28.54 | -3.57 | 28.42
78.176 | -0.06 | 21.45 | -0.23 | 21.45 | -0.41 | 21.39 | 0.47 | 1.88 | -25.37 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.27 | -9.55 | -0.06 | -29.24 | 0 | -26.31 | -0.06 | -35.51 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.24 | -1.23 | 28.42 | -3.34 | 28.48 | -2.23 | 28.24
78.208 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.41 | 9.08 | -23.03 | -0.06 | -0.06 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.21 | -7.62 | 0 | -27.54 | -0.06 | -24.08 | -0.06 | -33.87 | 0 | 0 | 0 | 28.65 | 0.18 | 28.54 | -0.64 | 28.54 | -2.4 | 28.48 | -1.76 | 28.48
78.24 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.53 | 14.36 | -21.39 | -0.18 | -51.68 | -0.06 | 0 | -0.06 | 0.23 | 0.23 | 21.21 | -5.8 | 0.23 | -25.55 | -0.06 | -22.38 | 0 | -31.99 | -0.06 | 0 | 0 | 28.59 | 0.12 | 28.65 | -0.59 | 28.54 | -1.41 | 28.77 | -0.76 | 28.48
78.272 | -0.06 | 0.06 | -0.18 | 21.5 | -0.18 | 21.45 | 0.41 | 21.86 | -18.52 | -0.06 | -56.78 | -0.06 | 0 | 0.12 | 0 | 0.18 | 21.45 | -3.87 | 12.25 | -22.91 | 0 | -19.22 | -0.06 | -29.06 | 0 | 0 | 0.06 | 28.42 | 0.06 | 28.59 | -0.18 | 28.42 | -0.94 | 28.24 | -0.47 | 28.42
78.304 | -0.18 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | 0.47 | 21.39 | -16.7 | -0.06 | -43.07 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.27 | -2.75 | 19.98 | -21.45 | -0.06 | -17.23 | 0 | -27.6 | 0 | -8.03 | 0 | 28.13 | 0.06 | 28.48 | -0.06 | 28.3 | -0.64 | 28.42 | -0.29 | 28.54
78.336 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.47 | 21.39 | -15.06 | -0.06 | -38.38 | 0 | 0 | -0.06 | 0.06 | 0.23 | 21.27 | -1.93 | 27.13 | -19.22 | 0 | -15.7 | -0.18 | -25.25 | 0 | -92.93 | -0.06 | 28.48 | 0.06 | 28.65 | 0 | 28.42 | -0.23 | 28.36 | -0.18 | 28.48
78.368 | -0.06 | 0 | -0.29 | 21.39 | -0.18 | 21.45 | 0.41 | 21.45 | -13.13 | 0 | -33.4 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.45 | -1.11 | 28.59 | -17.64 | 0 | -14.71 | 0.06 | -23.56 | 0 | -61.64 | 0 | 0.29 | 0.06 | 28.42 | -0.18 | 28.48 | -0.18 | 28.42 | 0.06 | 28.65
78.4 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.5 | 0.23 | 21.39 | -11.02 | 0 | -32.58 | -0.06 | -9.79 | 0 | 0 | 0.23 | 21.5 | -0.76 | 28.65 | -15.65 | 0 | -12.95 | 0 | -21.86 | 0 | -48.69 | 0 | 0 | 0 | 28.3 | 0.18 | 28.42 | -0.18 | 28.48 | -0.06 | 28.59
78.432 | -0.06 | 0 | -0.23 | 21.39 | -0.23 | 21.45 | 0.23 | 21.39 | -7.62 | 0.06 | -32.81 | -0.18 | -68.85 | -0.06 | 0 | 0.18 | 20.98 | -0.23 | 28.59 | -13.83 | 8.73 | -11.13 | 0 | -19.22 | 0 | -43.3 | 0 | 0 | 0.12 | 28.48 | 0.12 | 28.3 | 0 | 28.42 | 0.12 | 28.36
78.464 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.18 | 21.45 | -5.86 | -0.06 | -32.05 | -0.06 | -46.94 | 0.06 | 0 | 0.18 | 21.15 | -0.35 | 28.59 | -12.13 | 16 | -10.2 | 5.51 | -17.23 | 0 | -38.61 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.54 | -0.06 | 28.36 | 0.18 | 28.36
78.496 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.18 | 21.39 | -3.57 | 0 | -30.12 | -0.06 | -28.42 | -0.06 | 0 | 0.18 | 21.21 | 0.06 | 28.42 | -11.02 | 23.56 | -8.79 | 13.48 | -15.23 | -0.06 | -37.79 | 0 | 0 | 0 | 28.59 | 0.12 | 28.42 | 0.18 | 28.3 | 0.23 | 28.36
78.528 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.06 | 21.39 | -2.64 | 0 | -27.89 | -0.18 | -28.77 | 0 | 0 | 0.18 | 21.33 | 0 | 28.59 | -9.84 | 28.95 | -7.38 | 21.21 | -14.24 | 0 | -37.97 | -0.18 | 0 | -0.06 | 28.65 | 0.12 | 28.71 | 0.29 | 28.54 | 0.18 | 28.65
78.56 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.45 | 0.06 | 21.39 | -1 | 1.99 | -25.31 | 0 | -35.16 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.65 | -8.44 | 28.59 | -6.21 | 28.89 | -12.72 | -0.06 | -37.33 | 0 | 0 | -0.06 | 28.42 | 0.18 | 28.42 | 0.18 | 28.59 | 0.12 | 28.42
78.592 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.06 | 21.39 | -0.7 | 7.79 | -23.09 | -0.06 | -33.99 | -0.06 | 0 | 0.23 | 21.39 | 0.12 | 28.48 | -6.68 | 28.59 | -4.69 | 28.3 | -10.55 | -0.06 | -35.45 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.59 | 0.23 | 28.07 | 0.18 | 28.3
78.624 | -0.06 | 0 | -0.06 | 21.5 | -0.23 | 21.39 | 0.06 | 21.27 | -0.18 | 13.59 | -21.39 | -0.06 | -31.35 | -0.06 | 0 | 0.12 | 21.33 | 0.18 | 28.42 | -5.74 | 28.83 | -3.93 | 28.42 | -9.26 | -0.06 | -33.63 | 0 | 0 | -0.06 | 28.42 | 0.18 | 28.83 | 0.23 | 28.13 | 0.06 | 28.83
78.656 | -0.06 | 0 | -0.18 | 0.23 | -0.18 | 21.45 | 0 | 21.39 | -0.06 | 19.04 | -19.22 | 0 | -27.95 | -0.06 | -16.64 | 0.18 | 21.33 | 0.18 | 28.48 | -4.34 | 28.77 | -3.16 | 28.59 | -8.14 | -0.18 | -31.76 | -0.06 | 0 | -0.06 | 28.54 | 0.12 | 28.42 | 0.12 | 28.42 | 0.18 | 28.13
78.688 | -0.06 | -40.43 | -0.06 | 0.06 | -0.18 | 0.53 | -0.06 | 21.33 | 0.12 | 21.56 | -17.4 | -0.06 | -27.01 | -0.06 | -73.54 | 0.18 | 21.39 | 0.29 | 28.48 | -3.87 | 28.48 | -2.58 | 28.54 | -7.09 | -0.06 | -29.47 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.48 | 0.18 | 28.48 | 0.12 | 28.42
78.72 | -0.06 | -57.31 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.39 | 0.29 | 21.45 | -15.41 | -0.06 | -26.84 | -0.06 | -52.15 | 0.12 | 21.39 | 0.23 | 28.48 | -2.99 | 28.48 | -1.88 | 28.59 | -5.51 | -0.06 | -28.01 | -0.18 | 0 | -0.06 | 28.48 | 0.06 | 28.83 | 0 | 28.48 | 0.18 | 28.42
78.752 | 0 | -40.31 | -0.18 | 0.06 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.39 | -12.83 | 0 | -23.97 | -0.06 | -29.01 | 0.06 | 21.39 | 0.06 | 28.71 | -2.29 | 28.13 | -1.35 | 28.42 | -4.28 | 0 | -25.72 | 0 | 0.06 | 0 | 28.42 | 0 | 28.48 | 0.06 | 28.42 | 0.18 | 28.48
78.784 | -0.06 | -39.32 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.41 | 21.39 | -11.13 | 0.06 | -22.15 | -0.18 | -29.71 | 0.18 | 17.64 | 0.23 | 28.65 | -1.52 | 28.54 | -0.94 | 28.42 | -3.87 | 0 | -24.55 | -0.06 | -55.08 | -0.18 | 28.54 | -0.06 | 28.54 | 0.06 | 28.48 | 0.12 | 28.59
78.816 | -0.06 | -38.15 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.41 | 21.39 | -8.91 | 0 | -20.22 | -0.06 | -35.22 | 0.18 | 0.12 | 0.12 | 28.3 | -1.41 | 28.36 | -0.76 | 28.54 | -2.87 | 0 | -22.21 | -0.06 | -79.92 | -0.18 | 0.18 | -0.06 | 28.48 | 0.06 | 28.48 | 0.23 | 28.48
78.848 | -0.06 | -35.39 | 0 | 0 | -0.18 | 0 | 0 | 21.39 | 0.35 | 21.39 | -6.56 | -0.06 | -19.4 | 0 | -34.63 | 0.12 | 0 | 0.23 | 28.01 | -1.05 | 28.48 | -0.47 | 28.3 | -2.29 | 0 | -20.22 | -0.06 | -51.97 | 0 | 0.06 | -0.18 | 0.29 | 0.12 | 28.48 | 0.12 | 28.48
78.88 | -0.06 | -33.52 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -4.39 | -0.06 | -17.23 | -0.18 | -31.99 | 0.18 | 0 | 0.12 | 28.54 | -0.7 | 28.65 | -0.35 | 28.48 | -1.7 | -0.18 | -18.05 | -0.06 | -47.17 | 0 | 0 | 0 | 0.12 | 0 | 0.23 | 0.12 | 28.42
78.912 | 0 | -30.59 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.27 | -2.23 | 0 | -14.94 | -0.18 | -27.89 | 0.12 | 0 | 0.18 | 28.77 | -0.59 | 28.83 | -0.23 | 28.48 | -0.94 | 5.16 | -16.41 | -0.06 | -41.31 | -0.06 | 0.06 | 0 | 0 | 0 | 0 | 0.12 | 0.18
78.944 | -0.18 | -29.01 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -1.29 | 0 | -13.71 | 0 | -27.25 | 0.12 | 0 | 0.23 | 28.3 | -0.53 | 28.77 | -0.23 | 28.42 | -0.7 | 13.18 | -14.53 | -0.06 | -36.92 | 0 | 0 | -0.06 | 0.06 | 0 | 0 | 0 | 0
78.976 | -0.06 | -26.6 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.35 | 21.39 | -0.94 | 7.09 | -11.89 | -0.06 | -26.08 | 0 | 0 | 0.18 | 28.65 | -0.35 | 28.3 | -0.35 | 28.42 | -0.53 | 20.92 | -12.42 | -0.06 | -37.33 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0 | 0
79.008 | 0 | -24.79 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.23 | 21.39 | -0.06 | 12.89 | -9.9 | -0.06 | -24.79 | 0.06 | 0 | 0.18 | 28.59 | -0.41 | 28.59 | -0.23 | 28.54 | -0.41 | 28.36 | -10.08 | -0.06 | -37.91 | 0 | 0.06 | 0 | 0 | -0.18 | 0 | 0.06 | 0
79.04 | -0.06 | -22.91 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | -0.06 | 18.34 | -8.32 | 0.29 | -22.62 | 0.12 | 0 | 0.12 | 28.59 | -0.41 | 28.54 | -0.23 | 28.59 | -0.35 | 28.65 | -8.2 | -0.06 | -36.33 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
79.072 | -0.06 | -20.1 | -0.06 | -15.23 | -0.06 | 0 | 0 | 0.12 | 0.23 | 21.33 | 0.35 | 21.39 | -5.74 | 10.02 | -20.22 | 0 | 0 | 0.18 | 28.54 | -0.35 | 28.89 | -0.23 | 28.65 | -0.41 | 28.65 | -5.04 | -0.06 | -35.16 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
79.104 | -0.06 | -18.34 | -0.06 | -62.99 | -0.18 | -9.38 | -0.06 | 0.06 | 0.18 | 21.45 | 0.29 | 21.39 | -4.22 | 15.29 | -19.34 | 0.12 | 0 | 0.12 | 28.3 | -0.29 | 28.65 | -0.18 | 28.48 | -0.23 | 28.65 | -3.57 | -0.06 | -32.93 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
79.136 | -0.06 | -16.17 | -0.06 | -48.11 | -0.06 | -61.53 | 0.06 | 0 | 0.12 | 21.27 | 0.35 | 21.39 | -2.7 | 21.04 | -18.05 | -0.06 | 0 | 0.12 | 0.7 | -0.29 | 28.54 | -0.18 | 28.54 | -0.23 | 28.54 | -2.17 | 0 | -31.06 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
79.168 | -0.06 | -14.71 | -0.06 | -40.78 | -0.06 | -48.46 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.33 | -2.29 | 21.39 | -15.59 | 0 | 0 | 0.18 | 0.12 | -0.23 | 28.48 | -0.18 | 28.48 | -0.35 | 28.42 | -1.23 | -0.18 | -29.06 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
79.2 | -0.06 | -13.3 | -0.06 | -38.5 | -0.06 | -39.85 | -0.18 | 0 | -0.12 | 21.33 | 0.59 | 21.39 | -1.11 | 21.45 | -14 | 0 | 0 | 0.18 | 0 | -0.35 | 28.65 | -0.23 | 28.36 | -0.29 | 28.48 | -0.76 | 0 | -27.13 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0
79.232 | -0.06 | -11.37 | -0.06 | -35.86 | -0.06 | -38.79 | -0.06 | 0 | -0.06 | 21.33 | 0.41 | 21.39 | -0.29 | 21.33 | -11.48 | 0 | -58.6 | 0.23 | 0.06 | -0.29 | 28.42 | -0.18 | 28.59 | -0.23 | 28.59 | -0.35 | 0 | -25.08 | -0.06 | -57.25 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
79.264 | 0 | -9.79 | -0.06 | -35.16 | -0.06 | -36.51 | -0.18 | 0 | 0 | 21.45 | 0.47 | 21.39 | -0.23 | 21.33 | -10.25 | 0 | -59.47 | 0.18 | 0 | -0.18 | 28.59 | -0.23 | 28.59 | -0.23 | 28.18 | 0 | 0 | -23.44 | -0.18 | -76.82 | -0.18 | -27.83 | -0.06 | 0 | 0 | 0
79.296 | 6.97 | -8.56 | -0.06 | -31.99 | -0.18 | -34.51 | -0.06 | 0 | 0 | 21.33 | 0.41 | 21.39 | -0.23 | 21.33 | -8.32 | 0 | -38.5 | 0.12 | 0 | -0.29 | 28.65 | -0.18 | 28.48 | -0.18 | 28.48 | 0 | 0 | -21.91 | 0 | -51.92 | 0 | -84.38 | 0.06 | -38.97 | 0.06 | 0
79.328 | 12.54 | -7.21 | -0.06 | -30.06 | -0.18 | -32.34 | -0.12 | 0 | 0 | 21.33 | 0.35 | 21.39 | 0 | 21.39 | -6.62 | 0.06 | -27.48 | 0.06 | 0 | -0.29 | 0.23 | -0.18 | 28.48 | -0.18 | 28.65 | 0.29 | 7.79 | -19.45 | -0.18 | -47.05 | -0.18 | -58.42 | 0 | -84.91 | 0 | -51.04
79.36 | 18.34 | -5.63 | -0.06 | -28.24 | -0.06 | -30.7 | -0.06 | 0 | 0.06 | 21.21 | 0.41 | 21.39 | 0 | 21.39 | -5.04 | 0 | -31.82 | 0.06 | 0 | -0.23 | 0 | -0.18 | 0.29 | -0.29 | 28.65 | 0.23 | 15.06 | -17.23 | -0.06 | -42.42 | -0.06 | -33.11 | -0.18 | -54.73 | 0 | -79.69
79.392 | 21.5 | -3.63 | -0.06 | -26.13 | -0.18 | -28.07 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.33 | 0.12 | 21.39 | -3.11 | -0.06 | -35.74 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 25.2 | -15.53 | 0 | -36.45 | -0.18 | -34.75 | -0.18 | -29.06 | -0.06 | -42.19
79.424 | 21.5 | -2.64 | -0.06 | -23.56 | -0.06 | -25.84 | -0.18 | 0 | -0.06 | 21.33 | 0.35 | 21.39 | 0.06 | 21.39 | -2.29 | 0 | -33.46 | 0.12 | 0 | -0.12 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.41 | 28.71 | -13.59 | -0.06 | -38.09 | -0.18 | -42.19 | -0.06 | -34.45 | -0.06 | -28.83
79.456 | 21.5 | -1.46 | -0.06 | -22.09 | -0.18 | -23.91 | -0.06 | 0 | -0.06 | 17.58 | 0.18 | 21.39 | 0.23 | 21.39 | -1.23 | 0 | -29.59 | 0.18 | 0 | -0.23 | 0 | -0.06 | 0 | -0.23 | 28.65 | 0.23 | 28.54 | -11.84 | 0 | -38.15 | -0.18 | -40.78 | -0.18 | -42.13 | 0 | -36.68
79.488 | 21.45 | -0.7 | -0.06 | -19.63 | -0.18 | -21.68 | -0.06 | -1.64 | -0.06 | 0.12 | 0.18 | 21.27 | 0.18 | 21.39 | -0.82 | 0 | -27.77 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.65 | -9.38 | -0.06 | -37.21 | -0.06 | -37.33 | 0 | -41.02 | 0 | -42.78
79.52 | 21.45 | -0.29 | -0.18 | -17.64 | 0 | -19.92 | -0.06 | -75.35 | -0.06 | 0 | 0.29 | 21.39 | 0.35 | 21.39 | -0.41 | 0 | -26.84 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.65 | -6.56 | 0 | -35.16 | -0.18 | -34.45 | 0 | -36.51 | 0 | -40.37
79.552 | 21.39 | -0.06 | -0.18 | -15.53 | -0.06 | -17.52 | -0.18 | -50.1 | -0.06 | 0 | 0.12 | 21.39 | 0.29 | 21.39 | -0.18 | -0.06 | -25.9 | 0.12 | -13.42 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.35 | 28.59 | -4.45 | 0 | -32.34 | -0.18 | -32.58 | -0.06 | -32.99 | -0.06 | -35.16
79.584 | 21.5 | 0 | -0.06 | -14.24 | 0 | -15.65 | -0.18 | -42.36 | -0.06 | 0 | 0.12 | 21.56 | 0.29 | 21.33 | 0 | 0.12 | -23.32 | 0.12 | -77.35 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.3 | 0.35 | 28.48 | -2.58 | -0.23 | -30.94 | -0.06 | -32.05 | 0 | -32.93 | 0 | -33.69
79.616 | 21.45 | -0.06 | -0.06 | -12.54 | -0.06 | -13.95 | -0.06 | -36.45 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | 0 | 0.06 | -21.68 | 0.12 | -56.02 | -0.18 | 0 | 0 | 0 | -0.06 | 28.54 | 0.23 | 28.71 | -1.93 | 8.73 | -29.06 | -0.18 | -31.23 | -0.06 | -32.4 | -0.06 | -33.11
79.648 | 21.45 | -0.35 | -0.06 | -11.37 | 0 | -12.54 | 0 | -32.4 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.15 | 0.18 | 0.12 | -20.16 | 0.18 | -44.01 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.18 | 28.42 | -0.94 | 15.94 | -27.25 | -0.18 | -28.65 | -0.06 | -30.53 | -0.18 | -31.82
79.68 | 21.39 | -0.41 | 4.34 | -10.02 | -0.06 | -11.02 | -0.06 | -32.76 | -0.06 | 0 | 0.06 | 21.45 | 0.23 | 21.5 | 0.12 | -0.06 | -18.93 | 0.06 | -44.77 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.29 | 28.42 | -0.41 | 23.97 | -25.9 | -0.06 | -26.66 | 0 | -28.3 | 0 | -30.24
79.712 | 21.5 | -0.23 | 12.07 | -7.91 | -0.06 | -9.61 | -0.06 | -32.99 | 0 | 0 | 0 | 21.33 | 0.12 | 21.33 | 0.18 | 3.75 | -16.17 | 0.06 | -42.01 | -0.06 | 0.06 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.42 | -0.29 | 28.65 | -23.5 | -0.18 | -25.61 | -0.06 | -26.25 | -0.06 | -27.89
79.744 | 21.45 | -0.23 | 17.58 | -6.62 | -0.06 | -7.91 | -0.06 | -31.11 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.45 | 0.23 | 10.02 | -14.3 | 0 | -40.9 | -0.18 | -50.45 | -0.06 | 0 | -0.06 | 28.83 | 0.12 | 28.42 | 0 | 28.54 | -21.74 | -0.18 | -23.67 | -0.06 | -25.55 | -0.06 | -26.72
79.776 | 21.45 | -0.23 | 21.74 | -4.86 | -0.06 | -6.56 | -0.06 | -29.3 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | 0.29 | 15.35 | -13.01 | 0.06 | -38.67 | -0.18 | -69.49 | -0.18 | -28.24 | -0.18 | 28.59 | 0.06 | 28.42 | 0 | 28.77 | -19.63 | -0.18 | -22.33 | -0.06 | -24.26 | -0.06 | -25.61
79.808 | 21.45 | -0.35 | 21.45 | -3.69 | -0.06 | -5.1 | -0.06 | -27.25 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.29 | 21.15 | -11.07 | 0 | -36.51 | -0.18 | -47.05 | -0.06 | -78.46 | -0.06 | 0.29 | 0.06 | 28.48 | 0.12 | 28.48 | -17.58 | -0.06 | -20.92 | 0 | -22.74 | -0.06 | -23.85
79.84 | 21.45 | -0.23 | 21.5 | -2.23 | 7.03 | -3.69 | -0.06 | -25.61 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.33 | 0.35 | 21.45 | -9.49 | 0.18 | -34.86 | 0 | -45.94 | 0 | -52.62 | -0.06 | 0.12 | 0 | 28.48 | 0.29 | 28.48 | -15.94 | -0.18 | -19.4 | 0 | -20.74 | -0.06 | -22.15
79.872 | 21.45 | -0.18 | 21.45 | -1.11 | 12.48 | -2.05 | -0.06 | -22.21 | -0.06 | -0.06 | 0 | 0.06 | 0 | 21.45 | 0.23 | 21.39 | -7.09 | 0.06 | -32.99 | -0.18 | -43.77 | 0 | -45.94 | 0 | 0 | 0 | 28.36 | 0.29 | 28.65 | -13.77 | -0.06 | -17.7 | 0 | -18.69 | -0.06 | -19.98
79.904 | 21.5 | -0.18 | 21.45 | -0.59 | 18.22 | -1.23 | 6.68 | -20.74 | -0.06 | -67.91 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -5.51 | 0.06 | -30.53 | -0.06 | -41.25 | 0 | -44.59 | 0 | 0 | 0.12 | 28.65 | 0.29 | 28.36 | -11.37 | -0.18 | -15.82 | 0 | -17.23 | 0 | -18.87
79.936 | 21.45 | -0.18 | 21.45 | -0.29 | 21.68 | -0.76 | 12.36 | -18.05 | -0.18 | -61.41 | -0.06 | 0 | 0.12 | 0.23 | 0.29 | 21.39 | -4.16 | 0.06 | -29.01 | -0.06 | -39.79 | 0 | -41.25 | 0 | 0 | -0.06 | 28.48 | 0.23 | 28.65 | -9.38 | 9.32 | -14.71 | -0.18 | -15.29 | -0.06 | -16.99
79.968 | 21.39 | -0.29 | 21.45 | -0.18 | 21.5 | -0.29 | 17.87 | -16.29 | -0.06 | -42.72 | -0.06 | 0 | 0.06 | 0.06 | 0.18 | 21.39 | -2.93 | 0.12 | -26.72 | -0.06 | -38.85 | -0.06 | -40.55 | -0.06 | 0 | 0 | 28.54 | 0.29 | 28.71 | -7.15 | 16.47 | -12.54 | 0 | -13.95 | 5.68 | -15.23
80 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -0.06 | 21.62 | -14.47 | 0 | -38.67 | -0.06 | 0 | 0.06 | 0 | 0.23 | 21.33 | -2.17 | 0 | -24.9 | -0.06 | -35.92 | -0.06 | -38.5 | 0 | 0 | -0.06 | 28.59 | 0.29 | 28.77 | -4.98 | 24.26 | -10.61 | 5.04 | -12.36 | 13.77 | -13.95
80.032 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | -0.06 | 21.45 | -11.72 | 0.06 | -32.58 | 0 | 0 | 0.18 | 0 | 0.35 | 21.5 | -0.94 | 0 | -22.44 | -0.06 | -33.34 | -0.06 | -36.45 | -0.06 | 0 | -0.18 | 28.54 | 0.23 | 28.48 | -2.99 | 28.65 | -8.91 | 15.7 | -10.37 | 23.85 | -11.72
80.064 | 21.45 | -0.29 | 21.45 | -0.41 | 21.45 | -0.18 | 21.39 | -9.9 | 0 | -32.29 | 0 | 0 | 0.06 | 0 | 0.29 | 21.15 | -0.76 | 0.06 | -20.68 | -0.06 | -32.23 | 0 | -34.1 | -0.06 | 0 | 0 | 28.18 | 0.18 | 28.65 | -1.76 | 28.59 | -7.03 | 23.38 | -8.61 | 28.83 | -10.08
80.096 | 21.45 | -0.23 | 21.45 | -0.53 | 21.45 | -0.18 | 21.39 | -7.62 | -0.18 | -32.81 | 0.06 | 0 | 0 | 0 | 0.29 | 21.33 | -0.35 | 0 | -18.52 | 0 | -30.53 | 0 | -32.52 | 0 | 0 | 0 | 28.36 | 0.06 | 28.42 | -1.29 | 28.48 | -5.39 | 28.95 | -7.15 | 28.59 | -8.56
80.128 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.47 | 21.39 | -5.68 | -0.06 | -31.99 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.35 | 0.06 | -16.88 | -0.06 | -27.77 | 0 | -31.64 | -0.06 | 0 | -0.18 | 28.59 | 0.12 | 28.59 | -0.53 | 28.48 | -4.16 | 28.65 | -5.39 | 28.42 | -6.68
80.16 | 21.45 | -0.06 | 21.45 | -0.35 | 21.45 | -0.41 | 21.39 | -3.57 | 0 | -30.29 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | -0.18 | 0 | -15.29 | -0.06 | -26.31 | -0.06 | -28.65 | -0.18 | 0 | -0.18 | 28.48 | 0.06 | 28.13 | -0.23 | 28.48 | -2.64 | 28.65 | -4.28 | 28.54 | -5.33
80.192 | 0.06 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.39 | -1.64 | 0 | -27.6 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.15 | -0.06 | 0 | -13.59 | -0.06 | -23.56 | -0.06 | -25.43 | -0.06 | 0 | -0.18 | 22.91 | 0 | 28.48 | 0 | 28.65 | -1.46 | 28.65 | -2.4 | 28.42 | -3.46
80.224 | 0 | -0.18 | 21.45 | -0.35 | 21.39 | -0.29 | 21.39 | -1 | 0 | -25.25 | -0.18 | 0 | -0.18 | 0 | 0.23 | 21.39 | 0 | 1.99 | -12.07 | -0.06 | -21.39 | -0.06 | -23.91 | 0 | -32.4 | -0.06 | 0.18 | 0.06 | 28.59 | 0.18 | 28.18 | -0.94 | 28.48 | -1.99 | 28.48 | -2.4
80.256 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -0.41 | -0.06 | -23.5 | 0 | 0 | 0 | 0 | 0.18 | 21.39 | 0 | 6.09 | -10.49 | 0 | -19.75 | 0 | -22.15 | -0.18 | -92.11 | 0 | 0 | 0.12 | 28.48 | 0.23 | 28.54 | -0.76 | 28.48 | -1.23 | 28.48 | -1.46
80.288 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.27 | -0.18 | 0.06 | -21.68 | -0.06 | -50.22 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 11.78 | -9.2 | 0.06 | -17.58 | 0.12 | -20.1 | 0 | -56.19 | -0.06 | 0 | 0 | 28.54 | 0.12 | 28.48 | -0.23 | 28.54 | -0.41 | 28.42 | -0.82
80.32 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.27 | 0.06 | 0 | -19.86 | -0.06 | -63.17 | 0 | 0 | 0 | 21.39 | 0.18 | 17.46 | -8.14 | 0 | -16.29 | 0 | -18.52 | 0 | -48.28 | -0.18 | 0 | 0 | 28.42 | 0.47 | 28.48 | -0.23 | 28.65 | -0.35 | 28.42 | -0.7
80.352 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.35 | -0.18 | -17.29 | -0.06 | -42.89 | -0.06 | -12.48 | 0 | 21.39 | 0.12 | 21.39 | -6.39 | 0 | -14.24 | 0 | -16.06 | 0 | -43.13 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.42 | 0 | 28.48 | -0.18 | 28.48 | -0.41
80.384 | 0 | 0 | 21.39 | -0.23 | 21.45 | -0.23 | 21.39 | 0.29 | 0 | -15.06 | -0.06 | -38.09 | -0.18 | -74.01 | 0.06 | 21.39 | 0.06 | 21.39 | -5.39 | 0.06 | -12.77 | -0.06 | -14.59 | 0 | -37.09 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.42 | 0.06 | 28.36 | 0 | 28.48 | -0.23
80.416 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.33 | 0.53 | 0 | -13.07 | -0.06 | -33.22 | 0 | -53.09 | 0.12 | 21.33 | 0.23 | 21.39 | -4.51 | 0 | -11.48 | -0.06 | -13.07 | -0.06 | -37.15 | 0 | 0 | 0 | 28.3 | 0.29 | 28.59 | 0.06 | 28.48 | 0 | 28.42 | -0.29
80.448 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | 0.47 | 0.12 | -10.66 | -0.06 | -32.23 | -0.06 | -31.23 | -0.06 | 21.33 | 0.18 | 21.39 | -3.87 | 0 | -10.37 | 5.1 | -11.6 | -0.06 | -37.97 | -0.06 | 0 | 0.06 | 28.18 | 0.35 | 28.42 | 0.18 | 28.48 | 0.12 | 28.42 | 0
80.48 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | 0.47 | -0.06 | -9.02 | -0.06 | -32.52 | 0 | -27.72 | 0 | 21.27 | 0.18 | 21.39 | -3.22 | 0 | -8.85 | 13.13 | -10.84 | 0 | -37.85 | -0.06 | 0 | -0.06 | 28.54 | 0.23 | 28.42 | 0.23 | 28.71 | 0.06 | 28.42 | 0
80.512 | 0 | -0.18 | 21.39 | 0 | 21.39 | -0.18 | 21.39 | 0.47 | 2.34 | -6.33 | -0.06 | -31.93 | 0 | -35.51 | 0 | 21.33 | 0.06 | 21.39 | -2.23 | 9.49 | -6.91 | 23.44 | -8.56 | 0 | -35.86 | -0.18 | 0 | -0.06 | 0.18 | 0.12 | 28.48 | 0.18 | 28.54 | 0.18 | 28.42 | 0.18
80.544 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | -0.23 | 21.45 | 0.47 | 8.03 | -4.16 | -0.06 | -30.41 | -0.18 | -34.98 | -0.06 | 21.21 | 0.18 | 21.39 | -1.82 | 16.64 | -6.09 | 28.83 | -7.38 | -0.06 | -33.75 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.48 | 0.23 | 28.48 | 0.12 | 28.48 | 0.06
80.576 | -0.06 | -0.06 | 0.12 | -0.06 | 21.39 | -0.06 | 21.39 | 0.47 | 13.65 | -2.46 | -0.06 | -27.72 | 0 | -31.58 | 0.06 | 21.27 | 0.12 | 21.39 | -1.46 | 24.38 | -4.75 | 28.65 | -6.27 | 1.99 | -31.76 | -0.06 | 0 | -0.06 | 0 | 0 | 28.65 | 0.06 | 28.48 | 0.23 | 28.36 | 0.12
80.608 | -37.44 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.53 | 19.04 | -1.52 | 0 | -25.14 | 0 | -28.65 | 0.06 | 0.23 | 0.12 | 21.39 | -1.23 | 28.71 | -4.22 | 28.65 | -5.1 | 9.43 | -30 | -0.06 | 0.06 | -0.06 | 0 | 0.18 | 28.65 | 0.12 | 28.42 | 0.23 | 28.77 | 0.12
80.64 | -55.9 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.56 | -1 | -0.06 | -23.67 | 0.06 | -27.42 | 0 | 0.06 | 0.12 | 21.27 | -1 | 29.01 | -3.28 | 28.59 | -4.16 | 16.64 | -28.48 | -0.06 | -50.86 | -0.06 | 0 | 0.06 | 28.48 | 0.18 | 28.42 | 0.23 | 28.07 | 0.12
80.672 | -40.14 | -0.18 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.41 | 21.39 | -0.41 | 0 | -21.15 | 0 | -27.01 | -0.06 | 0 | 0.18 | 21.45 | -0.76 | 28.65 | -2.29 | 28.71 | -2.99 | 27.13 | -26.02 | 0 | -68.79 | -0.06 | 0 | 0.06 | 28.48 | 0.12 | 28.65 | 0.23 | 28.42 | 0.06
80.704 | -39.08 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 21.45 | -0.06 | 0 | -19.45 | -0.06 | -24.84 | -0.06 | 0 | 0.12 | 21.39 | -0.53 | 28.65 | -1.76 | 28.3 | -2.64 | 28.59 | -24.14 | 0 | -48.63 | -0.06 | 0 | 0.06 | 28.36 | 0.12 | 28.48 | 0.06 | 28.42 | 0.18
80.736 | -36.86 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.27 | 0.35 | 21.39 | 0.06 | -0.06 | -17.23 | -0.06 | -23.03 | -0.06 | 0 | 0.06 | 21.39 | -0.53 | 28.65 | -1.29 | 28.48 | -1.88 | 28.42 | -22.62 | -0.18 | -46.23 | -0.18 | 0 | 0.06 | 28.18 | 0.06 | 28.42 | 0.23 | 28.42 | 0.06
80.768 | -35.22 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | 0.35 | 0 | -15.35 | -0.06 | -21.33 | 0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.48 | -0.94 | 28.59 | -1.35 | 28.54 | -20.45 | -0.06 | -40.37 | -0.06 | 0 | -0.06 | 28.42 | 0 | 28.83 | 0.12 | 28.42 | 0.06
80.8 | -33.69 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.18 | 21.39 | 0.23 | 0 | -13.18 | 0.06 | -19.75 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | 28.48 | -0.82 | 28.59 | -0.94 | 28.65 | -18.28 | 0 | -35.98 | -0.06 | 0 | 0.12 | 28.42 | 0.06 | 28.89 | 0.18 | 28.42 | 0.06
80.832 | -31.06 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.41 | 0.06 | -10.37 | 0 | -17.05 | 0 | 0 | 0.06 | 21.33 | -0.41 | 28.24 | -0.53 | 28.65 | -0.59 | 28.54 | -15.94 | 0 | -37.38 | 0 | 0 | -0.18 | 0.12 | 0 | 28.07 | 0.23 | 28.65 | 0.12
80.864 | -28.18 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | 0.47 | 0 | -8.56 | -0.06 | -15.7 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.65 | -0.35 | 28.59 | -0.53 | 28.48 | -14.24 | 0 | -37.79 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.48 | 0.18 | 0.23 | 0.06
80.896 | -26.72 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 1.99 | -6.15 | 0 | -13.71 | 0 | 0 | 0.12 | 21.39 | -0.35 | 28.54 | -0.35 | 28.48 | -0.35 | 28.42 | -12.6 | 0 | -36.74 | 0 | 0.06 | 0 | 0 | -0.06 | 0.18 | 0.18 | 0 | 0
80.928 | -24.61 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.33 | 0.53 | 5.86 | -4.28 | 0 | -12.13 | -0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.48 | -0.41 | 28.42 | -0.23 | 28.42 | -10.37 | 0 | -35.68 | -0.06 | -67.62 | -0.06 | 0 | -0.06 | 0.06 | 0 | 0 | 0.06
80.96 | -22.33 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.53 | 11.72 | -2.64 | -0.06 | -10.37 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | 28.65 | -0.41 | 28.48 | -0.18 | 28.54 | -8.44 | -0.06 | -33.11 | -0.18 | -76 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.12
80.992 | -20.22 | -0.06 | -33.81 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.53 | 19.16 | -1.41 | 0 | -8.56 | -0.06 | 0 | 0.12 | 21.33 | -0.29 | 28.54 | -0.41 | 28.48 | -0.29 | 28.48 | -5.45 | 2.7 | -30.35 | -0.06 | -48.75 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
81.024 | -17.93 | -0.06 | -58.65 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.29 | 21.5 | -1.11 | -0.06 | -6.56 | -0.06 | -17.7 | 0.18 | 21.39 | -0.23 | 28.77 | -0.29 | 28.24 | -0.29 | 28.42 | -3.22 | 11.54 | -28.65 | -0.06 | -45.94 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.12
81.056 | -16.06 | 0 | -43.54 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.5 | 0.29 | 21.45 | -0.06 | 0 | -4.86 | 0 | -73.42 | 0.18 | 21.39 | -0.18 | 28.54 | -0.29 | 28.59 | -0.29 | 28.83 | -2.11 | 19.1 | -26.9 | 0 | -41.78 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06
81.088 | -14.71 | 0 | -39.79 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.29 | 21.33 | 0 | 1.7 | -3.4 | -0.06 | -51.74 | 0.18 | 21.33 | -0.29 | 28.42 | -0.18 | 28.3 | -0.35 | 28.48 | -1 | 26.72 | -25.31 | -0.06 | -38.61 | 0 | 0 | -0.18 | 0.06 | 0 | 0 | 0.12
81.12 | -12.95 | -0.06 | -38.32 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0.18 | 9.32 | -2.29 | 0 | -32.99 | 0.06 | 21.33 | -0.29 | 28.65 | -0.29 | 28.3 | -0.23 | 28.83 | -0.94 | 28.65 | -23.61 | -0.06 | -39.2 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.06
81.152 | -11.31 | -0.06 | -35.51 | -0.18 | -37.21 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.33 | 0.41 | 16.47 | -1.29 | -0.06 | -29.59 | 0.23 | 0.12 | -0.18 | 28.48 | -0.23 | 28.42 | -0.23 | 28.36 | -0.18 | 28.59 | -21.27 | -0.06 | -37.91 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12
81.184 | -9.79 | 0.06 | -33.52 | -0.06 | -55.9 | -0.18 | 0 | -0.06 | 21.27 | 0.23 | 21.33 | 0.41 | 21.74 | -0.76 | -0.06 | -35.33 | 0.18 | 0 | -0.29 | 28.48 | -0.18 | 28.3 | -0.18 | 28.36 | -0.06 | 28.65 | -19.34 | 0 | -36.51 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
81.216 | -8.79 | -0.06 | -31.41 | -0.18 | -42.95 | -0.06 | -49.04 | -0.06 | 21.39 | 0.18 | 21.39 | 0.47 | 21.33 | -0.35 | 0 | -34.75 | 0.06 | 0 | -0.18 | 28.59 | -0.18 | 28.71 | -0.18 | 28.42 | 0.06 | 28.59 | -17.58 | 0.06 | -34.22 | -0.18 | 0.06 | -0.06 | 0 | -0.06 | 0 | 0.06
81.248 | -7.09 | 0 | -29.42 | -0.06 | -39.61 | -0.06 | -64.34 | -0.06 | 21.39 | 0 | 21.39 | 0.53 | 21.33 | -0.35 | 0.06 | -32.4 | -0.06 | 0 | -0.23 | 28.13 | -0.06 | 28.42 | -0.18 | 28.54 | 0.12 | 28.48 | -15.29 | -0.06 | -32.17 | -0.06 | -66.8 | -0.06 | 0 | 0 | 0 | 0.06
81.28 | -5.57 | 0.06 | -27.6 | -0.18 | -37.62 | -0.18 | -46.12 | 0 | 21.33 | 0.06 | 21.33 | 0.41 | 21.33 | -0.06 | 0 | -29.06 | 0 | 0 | -0.18 | 28.54 | -0.18 | 28.3 | -0.18 | 28.59 | 0.23 | 28.42 | -13.77 | 0.06 | -30.53 | 0 | -75.59 | -0.18 | 0 | 0 | -35.39 | 0
81.312 | -3.57 | -0.06 | -24.96 | -0.06 | -34.98 | -0.18 | -39.55 | 0 | 21.39 | 0.06 | 21.39 | 0.53 | 21.39 | -0.06 | 0.06 | -27.01 | 0.12 | 0 | -0.18 | 28.54 | -0.18 | 28.48 | -0.06 | 28.54 | 0.29 | 28.48 | -10.78 | -0.06 | -28.24 | -0.18 | -37.62 | -0.18 | -50.98 | 0 | -75.47 | 0.12
81.344 | -2.46 | 1.99 | -23.61 | -0.06 | -32.81 | -0.06 | -33.99 | 0.06 | 21.21 | 0.12 | 21.33 | 0.53 | 21.39 | 0.12 | 3.16 | -26.6 | 0.12 | 0 | -0.06 | 28.71 | -0.18 | 0.23 | -0.18 | 28.59 | 0.29 | 28.48 | -8.73 | 0 | -26.13 | -0.18 | -28.65 | -0.06 | -76.35 | 0 | -48.99 | 0.06
81.376 | -1.46 | 6.21 | -21.21 | -0.06 | -31.64 | -0.06 | -32.23 | -0.06 | 21.27 | 0.06 | 21.39 | 0.35 | 21.39 | 0.35 | 9.55 | -25.14 | 0 | 0 | -0.18 | 28.48 | 0 | 0.06 | -0.18 | 28.36 | 0.35 | 28.36 | -6.68 | 0 | -24.67 | 0 | -39.14 | -0.18 | -49.81 | 0 | -29.24 | 0.06
81.408 | -0.94 | 12.13 | -18.87 | -0.06 | -30.12 | -0.06 | -32.17 | -0.06 | 17.7 | 0.06 | 21.33 | 0.35 | 21.27 | 0.18 | 14.71 | -22.62 | 0.18 | 0 | -0.06 | 0.18 | -0.06 | 0 | -0.18 | 28.54 | 0.47 | 28.48 | -4.63 | 0 | -23.15 | -0.06 | -42.95 | 0 | -29.3 | 0 | -33.52 | 0
81.44 | -0.41 | 17.58 | -17.23 | -0.06 | -26.72 | -0.06 | -32.23 | -0.18 | 0.12 | 0.06 | 21.33 | 0.35 | 21.39 | 0.23 | 20.57 | -21.09 | 0 | 0 | -0.18 | 0 | -0.18 | 0.06 | -0.18 | 28.54 | 0.29 | 28.48 | -2.87 | -0.06 | -21.62 | 0 | -39.2 | -0.18 | -33.87 | -0.18 | -41.9 | 0.12
81.472 | -0.18 | 21.62 | -14.71 | -0.06 | -24.32 | -0.06 | -30.35 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.21 | 0.29 | 21.39 | -19.04 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | -0.06 | 3.28 | 0.29 | 28.42 | -1.7 | 6.33 | -18.4 | -0.06 | -34.45 | -0.06 | -42.89 | -0.18 | -39.79 | 0
81.504 | -0.06 | 21.45 | -13.18 | -0.06 | -22.68 | 0 | -28.07 | -0.18 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | 0.23 | 21.45 | -17.23 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.12 | 0.35 | 28.42 | -0.7 | 14.12 | -16.7 | -0.18 | -32.64 | -0.18 | -40.14 | -0.18 | -35.27 | 0
81.536 | -0.18 | 21.45 | -12.25 | -0.06 | -20.63 | 0 | -25.84 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.39 | 0.18 | 21.39 | -15.94 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.35 | 28.24 | -0.41 | 21.91 | -14.94 | -0.18 | -32.93 | -0.06 | -35.74 | -0.06 | -33.34 | 0
81.568 | -0.35 | 21.45 | -10.31 | -0.06 | -18.4 | -0.06 | -24.14 | 0 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | 0.12 | 21.45 | -14.24 | 0.18 | -39.79 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.29 | 28.71 | -0.23 | 29.06 | -12.6 | 0 | -31.82 | -0.06 | -32.7 | -0.18 | -32.81 | 0
81.6 | -0.41 | 21.5 | -9.08 | -0.06 | -16.7 | -0.18 | -22.15 | -0.06 | 0 | -0.18 | 21.39 | 0.18 | 21.04 | 0.18 | 21.33 | -12.72 | 0.06 | -65.51 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.42 | 0 | 28.59 | -10.78 | 0 | -29.94 | -0.18 | -32.64 | -0.06 | -32.64 | 0.06
81.632 | -0.35 | 21.45 | -7.27 | -0.06 | -14.94 | -0.12 | -19.81 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.27 | 0.12 | 21.39 | -10.37 | 0 | -36.39 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 28.59 | 0.12 | 28.42 | -8.14 | -0.06 | -27.72 | -0.06 | -32.17 | -0.06 | -29.71 | 0
81.664 | -0.35 | 21.39 | -5.68 | -0.06 | -13.18 | -0.06 | -17.4 | -0.06 | 0 | -0.18 | 21.21 | -0.06 | 21.33 | 0.12 | 21.33 | -8.38 | 0 | -27.6 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 28.48 | 0.29 | 28.77 | -5.86 | 0 | -26.25 | -0.18 | -30.35 | -0.06 | -28.01 | 0
81.696 | -0.29 | 21.45 | -4.16 | -0.06 | -11.84 | -0.06 | -15.47 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.18 | 21.39 | -6.68 | 0 | -32.05 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 28.3 | 0.35 | 28.59 | -3.98 | -0.18 | -25.37 | -0.06 | -28.07 | 0 | -26.84 | -0.18
81.728 | -0.35 | 21.45 | -2.99 | -0.06 | -10.37 | 0 | -13.36 | -0.06 | 0 | -0.06 | 21.33 | -0.06 | 21.39 | 0.18 | 21.39 | -4.8 | 0.06 | -35.39 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 28.42 | 0.29 | 28.48 | -2.29 | -0.06 | -23.44 | -0.06 | -26.84 | -0.06 | -25.61 | -0.06
81.76 | -0.18 | 21.45 | -2.11 | 0.18 | -8.56 | 0 | -11.54 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.5 | 0.18 | 21.39 | -3.98 | -0.18 | -33.99 | -0.06 | 0 | -0.06 | -28.07 | -0.06 | 0 | 0.12 | 28.42 | 0.29 | 28.54 | -1.7 | 0 | -22.09 | -0.06 | -25.72 | -0.06 | -24.2 | 0
81.792 | -0.29 | 21.39 | -1.11 | 9.55 | -7.09 | 2.75 | -8.44 | -0.06 | 0 | -0.06 | 17.23 | 0.06 | 21.45 | 0.18 | 21.33 | -2.17 | 0 | -29.83 | -0.06 | 0.06 | -0.12 | -68.32 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.48 | -0.64 | -0.06 | -19.28 | -0.06 | -23.85 | -0.06 | -21.91 | 0
81.824 | -0.23 | 21.45 | -0.41 | 15.12 | -5.22 | 8.79 | -6.68 | -0.06 | -0.06 | -0.06 | 0.12 | 0 | 21.27 | 0.06 | 21.39 | -1.82 | 0 | -27.83 | -0.18 | -50.8 | -0.06 | -48.22 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.65 | -0.23 | 0 | -17.99 | -0.06 | -22.09 | -0.06 | -20.63 | 0
81.856 | -0.23 | 21.39 | -0.18 | 20.57 | -4.1 | 14.18 | -4.34 | -0.06 | -61.58 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -0.76 | -0.06 | -26.54 | 0 | -65.1 | -0.06 | -45.7 | -0.06 | 0 | 0.06 | 28.42 | 0.29 | 28.71 | -0.18 | -0.06 | -16.7 | 0 | -20.68 | -0.06 | -18.99 | -0.06
81.888 | -0.18 | 21.39 | -0.18 | 21.5 | -3.16 | 19.86 | -2.64 | -0.06 | -60.59 | -0.06 | 0 | 0 | 21.15 | 0.23 | 21.33 | -0.7 | -0.06 | -26.08 | -0.06 | -47.64 | -0.06 | -44.24 | -0.06 | 0 | 0 | 9.02 | 0.29 | 28.48 | 0.12 | -0.18 | -14.94 | -0.06 | -18.69 | -0.18 | -17.4 | -0.06
81.92 | -0.18 | 21.45 | -0.23 | 21.5 | -1.52 | 21.45 | -2.05 | 0 | -44.24 | -0.06 | 0.06 | -0.06 | 21.27 | 0.12 | 21.39 | -0.29 | 5.33 | -23.79 | -0.06 | -45.35 | 0.06 | -41.25 | -0.06 | -58.01 | -0.06 | 0.12 | 0.23 | 28.59 | 0.12 | 3.81 | -13.36 | -0.18 | -17.23 | -0.06 | -16 | -0.06
81.952 | -0.18 | 21.39 | -0.35 | 21.5 | -0.82 | 21.45 | -0.7 | -0.06 | -37.68 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -0.18 | 13.18 | -21.27 | -0.18 | -44.24 | -0.06 | -38.85 | -0.06 | -57.66 | -0.18 | 0 | 0.23 | 28.42 | 0.12 | 14.71 | -11.07 | 5.16 | -15.23 | 0 | -13.3 | 0
81.984 | -0.18 | 21.45 | -0.47 | 21.45 | -0.47 | 21.45 | -0.29 | 0.06 | -33.05 | -0.06 | 0 | 0 | 0.12 | 0.06 | 21.45 | 0 | 18.46 | -20.22 | 0 | -40.9 | -0.18 | -37.97 | -0.06 | -45.18 | -0.06 | 0.06 | 0.12 | 28.42 | 0.35 | 22.33 | -9.61 | 13.18 | -13.36 | 4.75 | -12.25 | 0
82.016 | -0.18 | 21.39 | -0.47 | 21.45 | -0.06 | 21.45 | -0.06 | -0.06 | -32.29 | 0.06 | 0 | 0 | 0.06 | -0.06 | 21.39 | 0.18 | 21.5 | -18.87 | 0 | -39.02 | 0 | -36.51 | 0 | -45.53 | 0 | 0 | 0.12 | 28.83 | 0.41 | 29.06 | -8.14 | 20.8 | -11.48 | 13.01 | -10.49 | 6.33
82.048 | -0.18 | 21.39 | -0.35 | 21.5 | -0.18 | 21.39 | 0 | 0 | -32.52 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.39 | -17.46 | -0.06 | -37.03 | -0.06 | -34.28 | -0.06 | -43.77 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.48 | -6.09 | 28.89 | -10.31 | 20.68 | -8.91 | 14.18
82.08 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | 0.35 | -0.06 | -31.99 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.06 | 21.39 | -14.82 | 0 | -35.86 | 0 | -32.81 | 0 | -41.66 | 0 | 0 | 0.12 | 28.48 | 0.29 | 28.54 | -4.45 | 28.65 | -8.38 | 28.3 | -7.27 | 21.68
82.112 | -0.18 | 21.39 | -0.35 | 21.5 | -0.23 | 21.5 | 0.29 | -0.06 | -29.53 | 0 | 0 | 0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.33 | -12.77 | -0.06 | -33.81 | 0 | -29.71 | -0.06 | -38.85 | -0.18 | 0 | 0.06 | 28.42 | 0.23 | 28.59 | -2.75 | 28.54 | -7.03 | 28.65 | -4.86 | 28.83
82.144 | -0.18 | 21.45 | -0.29 | 21.5 | -0.29 | 21.45 | 0.35 | -0.06 | -27.95 | 0 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.23 | 21.27 | -11.07 | -0.06 | -31.29 | 0.06 | -27.77 | -0.06 | -37.44 | 0 | 0 | 0.12 | 28.3 | 0.29 | 28.13 | -1.88 | 28.3 | -4.8 | 28.65 | -3.57 | 28.54
82.176 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | 0.53 | 2.58 | -25.37 | 0 | 0 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | -9.43 | -0.06 | -29.71 | 0 | -25.9 | -0.06 | -35.33 | 0 | 0 | 0 | 28.42 | 0.12 | 28.65 | -1.29 | 28.48 | -3.63 | 28.48 | -2.58 | 28.42
82.208 | -0.06 | 21.45 | -0.23 | 21.39 | -0.35 | 21.45 | 0.59 | 8.79 | -23.32 | 0 | -0.06 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -7.79 | -0.06 | -27.89 | 0 | -24.61 | -0.06 | -33.46 | 0 | 0 | -0.06 | 28.48 | 0.23 | 28.59 | -0.7 | 28.59 | -2.29 | 28.42 | -1.7 | 28.59
82.24 | -0.06 | 21.45 | -0.35 | 21.45 | -0.23 | 21.39 | 0.47 | 14.41 | -21.68 | 0 | -40.31 | 0 | 0 | -0.06 | 1 | 0.29 | 21.39 | -5.98 | 1.99 | -26.02 | 0 | -22.21 | -0.06 | -31.82 | 0.06 | 0 | 0.06 | 28.18 | 0.18 | 28.42 | -0.53 | 28.42 | -1.7 | 28.3 | -1.05 | 28.48
82.272 | -0.06 | 0.12 | -0.29 | 21.5 | -0.23 | 21.39 | 0.53 | 21.68 | -19.04 | -0.06 | -60.35 | -0.06 | 0 | 0 | 0.06 | 0.29 | 21.15 | -4.16 | 11.37 | -23.5 | -0.06 | -19.98 | -0.06 | -29.24 | 0 | 0 | 0 | 28.77 | 0.06 | 28.59 | -0.29 | 28.48 | -1.11 | 28.42 | -0.59 | 28.54
82.304 | -0.18 | 0 | -0.35 | 21.45 | -0.23 | 21.39 | 0.47 | 21.5 | -16.7 | -0.06 | -44.42 | -0.06 | 0 | 0.06 | 0 | 0.23 | 21.27 | -2.75 | 18.87 | -21.97 | -0.06 | -17.75 | 0 | -27.19 | 0 | -4.45 | 0 | 28.48 | 0.12 | 28.42 | -0.06 | 28.13 | -0.47 | 28.54 | -0.29 | 28.36
82.336 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.29 | 21.45 | -14.82 | -0.06 | -39.08 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.33 | -1.88 | 26.43 | -19.34 | -0.06 | -16.29 | 0 | -25.43 | -0.06 | -92.7 | -0.06 | 28.48 | 0 | 28.42 | -0.06 | 28.42 | -0.35 | 28.54 | -0.35 | 28.42
82.368 | -0.06 | 0.06 | -0.23 | 21.39 | -0.29 | 21.45 | 0.35 | 21.39 | -12.89 | 0 | -34.16 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | -1.29 | 28.71 | -17.64 | -0.06 | -14.82 | 0 | -23.44 | -0.06 | -63.4 | -0.06 | 0.29 | 0 | 28.77 | 0.06 | 28.42 | -0.18 | 28.65 | 0 | 28.54
82.4 | -0.06 | 0 | -0.35 | 21.39 | -0.23 | 21.45 | 0.35 | 21.39 | -11.37 | -0.06 | -32.23 | -0.06 | -5.8 | -0.06 | 0 | 0.18 | 21.45 | -0.76 | 28.48 | -15.94 | -0.06 | -13.18 | -0.06 | -21.68 | 0 | -47.93 | -0.06 | 0 | -0.18 | 28.48 | 0.06 | 28.65 | -0.06 | 28.48 | 0.06 | 28.3
82.432 | 0 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.29 | 21.33 | -8.32 | -0.06 | -32.76 | -0.18 | -69.96 | -0.06 | 0 | 0.18 | 21.39 | -0.53 | 28.48 | -14.18 | 8.56 | -11.13 | -0.06 | -18.87 | -0.06 | -43.89 | 0 | 0 | 0 | 28.48 | 0.12 | 28.59 | 0.06 | 28.42 | -0.06 | 27.89
82.464 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.35 | 21.39 | -6.33 | -0.18 | -32.76 | -0.18 | -48.05 | -0.06 | 0 | 0.29 | 21.15 | -0.06 | 28.65 | -12.72 | 15.59 | -9.84 | 4.98 | -17.05 | -0.06 | -38.61 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.48
82.496 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.23 | 21.33 | -4.34 | 0.12 | -30.47 | -0.06 | -29.53 | -0.06 | 0 | 0.18 | 21.39 | -0.06 | 28.77 | -11.25 | 23.38 | -8.56 | 13.18 | -15.29 | -0.06 | -37.44 | 0 | 0 | 0.06 | 28.65 | 0.12 | 28.59 | 0.18 | 28.89 | 0.12 | 28.42
82.528 | 0 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.18 | 21.39 | -2.17 | 0 | -28.18 | -0.18 | -29.77 | -0.06 | 0 | 0.18 | 21.33 | 0.06 | 28.54 | -10.02 | 28.83 | -7.38 | 20.86 | -14.47 | -0.06 | -37.68 | 0 | 0 | -0.06 | 28.59 | 0.29 | 28.01 | 0.12 | 28.42 | 0.35 | 28.42
82.56 | 0 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.12 | 21.39 | -1.93 | -0.06 | -26.02 | -0.06 | -35.63 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 28.65 | -8.67 | 28.48 | -6.5 | 28.89 | -12.42 | 0 | -37.79 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.54 | 0.23 | 28.65 | 0.23 | 28.71
82.592 | 0 | 0 | -0.18 | 21.5 | -0.23 | 21.39 | 0.12 | 21.39 | -0.41 | 7.38 | -23.38 | -0.18 | -34.22 | -0.06 | 0 | 0.12 | 21.21 | 0.18 | 28.48 | -6.86 | 28.48 | -4.8 | 28.3 | -10.78 | 0.06 | -36.21 | 0 | 0 | -0.06 | 28.54 | 0.12 | 28.3 | 0.23 | 29.01 | 0.18 | 28.59
82.624 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.12 | 21.33 | -0.23 | 13.07 | -21.5 | -0.06 | -31.17 | -0.18 | 0 | 0.29 | 21.33 | 0.06 | 28.65 | -5.98 | 28.59 | -3.98 | 28.83 | -9.55 | -0.06 | -34.16 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | 0.12 | 28.01 | 0.18 | 28.42
82.656 | -0.06 | 0 | 0 | 0.64 | -0.23 | 21.45 | 0.12 | 21.33 | -0.06 | 18.57 | -19.63 | -0.06 | -28.13 | -0.06 | -11.66 | 0.18 | 21.39 | 0.23 | 28.48 | -4.45 | 28.36 | -3.34 | 28.65 | -8.32 | 0 | -31.47 | 0 | 0 | 0.06 | 28.48 | 0.12 | 28.54 | 0.23 | 28.83 | 0.12 | 28.42
82.688 | -0.06 | -35.45 | -0.18 | 0.06 | -0.18 | 10.61 | -0.06 | 21.33 | 0.12 | 21.5 | -17.99 | -0.06 | -27.6 | -0.06 | -72.6 | 0.18 | 21.21 | 0.35 | 28.54 | -3.93 | 28.54 | -2.64 | 28.65 | -7.32 | 0.06 | -30 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.48 | 0.23 | 28.42 | 0.23 | 28.65
82.72 | -0.06 | -58.71 | -0.18 | 0 | -0.18 | 0.12 | 0 | 21.33 | 0.23 | 21.39 | -15.94 | -0.18 | -26.6 | 0 | -54.03 | 0.23 | 21.56 | 0.18 | 28.24 | -2.93 | 28.65 | -1.99 | 28.42 | -5.74 | -0.06 | -28.18 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.18 | 0.12 | 28.42 | 0.23 | 28.42
82.752 | -0.06 | -40.2 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.33 | -12.89 | -0.06 | -24.08 | -0.18 | -30.18 | 0.18 | 21.27 | 0.18 | 28.18 | -2.17 | 28.83 | -1.35 | 28.59 | -4.45 | 0.06 | -26.02 | -0.06 | 0 | -0.06 | 28.54 | 0.06 | 28.48 | 0.06 | 28.54 | 0.12 | 28.3
82.784 | 0 | -39.85 | -0.18 | 0 | -0.18 | 0 | 0 | 21.33 | 0.35 | 21.39 | -10.66 | -0.06 | -22.38 | -0.18 | -29.71 | 0.12 | 20.57 | 0.18 | 28.24 | -1.88 | 28.42 | -0.94 | 28.48 | -3.46 | -0.06 | -24.43 | -0.06 | -41.9 | -0.06 | 28.65 | 0 | 28.36 | 0.12 | 28.54 | 0.18 | 28.48
82.816 | -0.06 | -37.91 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.21 | -8.91 | -0.06 | -20.63 | -0.06 | -35.1 | 0.12 | 0.12 | 0.18 | 28.42 | -1.41 | 28.36 | -0.76 | 28.59 | -2.81 | -0.18 | -22.56 | 0 | -87.66 | -0.06 | 0.29 | -0.06 | 28.48 | 0.12 | 28.48 | 0.12 | 28.42
82.848 | -0.06 | -35.1 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.33 | -6.68 | -0.18 | -19.28 | 0 | -35.27 | 0.18 | 0 | 0.23 | 28.54 | -1.17 | 28.3 | -0.64 | 28.42 | -2.23 | 0 | -20.74 | 0 | -54.67 | 0 | 0.06 | -0.18 | 0.29 | 0.06 | 28.48 | 0.12 | 28.42
82.88 | -0.06 | -34.1 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -4.75 | 0 | -17.81 | -0.06 | -32.23 | 0.12 | 0 | 0.23 | 28.59 | -0.94 | 28.48 | -0.41 | 28.54 | -1.64 | -0.06 | -18.63 | 0 | -47.35 | 0 | 0 | 0 | 0.12 | 0.12 | 0.29 | 0.23 | 28.42
82.912 | -0.06 | -31.11 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -2.46 | -0.06 | -15.18 | -0.18 | -28.3 | 0.12 | 0 | 0.23 | 28.89 | -0.64 | 28.36 | -0.18 | 28.65 | -1.05 | 3.81 | -16.23 | -0.18 | -44.53 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.12 | 0.18
82.944 | -0.06 | -28.65 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -1.52 | 1.99 | -13.54 | -0.18 | -27.54 | 0.12 | 0 | 0.18 | 28.42 | -0.59 | 28.36 | -0.29 | 28.48 | -0.64 | 12.19 | -14.53 | 0 | -38.97 | -0.18 | 0 | -0.06 | 0.06 | 0.06 | 0 | 0.06 | 0
82.976 | 0 | -27.42 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | -0.64 | 6.45 | -11.72 | -0.06 | -26.31 | 0.06 | 0 | 0.23 | 28.48 | -0.35 | 28.59 | -0.29 | 28.42 | -0.53 | 19.92 | -12.13 | 0.06 | -38.85 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 0
83.008 | -0.06 | -24.26 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.29 | 21.39 | -0.35 | 12.07 | -10.2 | -0.18 | -24.79 | 0.06 | 0 | 0.29 | 28.54 | -0.35 | 28.59 | -0.29 | 28.24 | -0.47 | 27.25 | -10.55 | -0.18 | -39.32 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06 | 0
83.04 | -0.06 | -23.09 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.33 | -0.18 | 17.87 | -8.38 | 0 | -22.85 | 0.06 | 0 | 0.23 | 28.3 | -0.35 | 28.48 | -0.23 | 28.42 | -0.41 | 28.77 | -8.03 | 0 | -38.09 | 0 | 0 | 0.06 | 0 | 0 | 0 | 0.06 | 0
83.072 | -0.06 | -20.1 | -0.06 | -7.56 | -0.18 | 0 | -0.06 | 0.23 | 0.29 | 21.33 | 0.18 | 21.39 | -5.92 | 9.38 | -20.57 | 0.06 | 0 | 0.23 | 28.59 | -0.47 | 28.65 | -0.18 | 28.48 | -0.23 | 28.71 | -5.39 | -0.06 | -36.45 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 0
83.104 | 0 | -18.16 | -0.06 | -60.76 | -0.06 | -1.76 | -0.06 | 0.06 | 0.18 | 21.33 | 0.12 | 21.39 | -4.39 | 14.53 | -19.1 | 0 | 0 | 0.18 | 28.54 | -0.53 | 28.48 | -0.18 | 28.54 | -0.35 | 28.48 | -3.57 | -0.18 | -33.87 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0
83.136 | -0.06 | -16.76 | -0.18 | -48.93 | -0.06 | -54.67 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | -2.81 | 20.16 | -17.75 | 0 | 0 | 0.23 | 14.06 | -0.29 | 28.36 | -0.23 | 28.54 | -0.23 | 28.42 | -1.99 | 0 | -32.4 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0
83.168 | 0 | -14.77 | -0.06 | -40.37 | -0.06 | -51.51 | 0.06 | 0 | 0.18 | 21.39 | 0.41 | 21.39 | -2.23 | 21.5 | -15.82 | 0.06 | 0 | 0.23 | 0.12 | -0.35 | 28.36 | -0.18 | 28.48 | -0.29 | 28.59 | -1.64 | -0.18 | -30.12 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
83.2 | -0.18 | -13.3 | -0.06 | -39.08 | -0.06 | -39.49 | 0 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -1.46 | 21.27 | -13.71 | 0.06 | 0 | 0.12 | 0 | -0.35 | 28.42 | -0.18 | 28.42 | -0.23 | 28.77 | -0.59 | 0 | -27.6 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0
83.232 | -0.06 | -11.48 | -0.06 | -36.04 | -0.06 | -38.97 | 0 | 0 | 0.06 | 21.33 | 0.53 | 21.15 | -0.29 | 21.33 | -12.13 | -0.06 | -51.68 | 0.12 | 0.06 | -0.29 | 28.48 | -0.18 | 28.42 | -0.35 | 28.89 | -0.18 | 0 | -25.55 | -0.06 | -45.82 | 0 | 0 | 0 | 0 | 0 | 0
83.264 | -0.06 | -10.2 | -0.18 | -34.63 | -0.06 | -37.21 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | 21.39 | -0.29 | 21.39 | -10.14 | 0 | -61.35 | 0.18 | 0 | -0.29 | 28.65 | -0.23 | 28.48 | -0.18 | 28.48 | -0.18 | 0 | -23.56 | 0 | -80.92 | 0 | -22.38 | -0.06 | 0 | 0 | 0
83.296 | 6.27 | -8.73 | -0.06 | -32.34 | -0.06 | -34.98 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.27 | -0.35 | 21.39 | -8.56 | 0.06 | -40.2 | 0.23 | 0 | -0.23 | 28.54 | -0.23 | 28.48 | -0.23 | 28.65 | 0.18 | -0.06 | -22.09 | -0.18 | -53.85 | 0 | -87.89 | -0.06 | -31.41 | 0 | 0
83.328 | 12.19 | -6.97 | -0.18 | -30.82 | -0.06 | -32.4 | 0 | 0 | 0 | 21.15 | 0.41 | 21.39 | 0.12 | 21.39 | -6.5 | 0 | -27.72 | 0.29 | 0 | -0.23 | 0.23 | -0.18 | 28.3 | -0.23 | 28.89 | 0.18 | 6.33 | -19.51 | 0 | -48.17 | -0.18 | -58.95 | 0.06 | -88.89 | 0 | -39.67
83.36 | 17.58 | -5.8 | -0.18 | -28.36 | -0.06 | -30.94 | -0.06 | 0 | -0.06 | 21.33 | 0.47 | 21.39 | 0.29 | 21.15 | -5.1 | 0 | -30.76 | 0.23 | 0 | -0.18 | 0.06 | -0.18 | 0.29 | -0.29 | 28.59 | 0.29 | 14.12 | -16.82 | -0.06 | -42.66 | -0.18 | -33.81 | 0 | -56.9 | 0 | -84.61
83.392 | 21.68 | -4.22 | -0.06 | -26.08 | -0.18 | -28.48 | 0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.33 | 0.23 | 21.39 | -3.22 | 0.06 | -35.57 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 28.65 | 0.29 | 24.43 | -15 | -0.18 | -35.98 | -0.06 | -34.63 | 0 | -29.77 | 0 | -46.58
83.424 | 21.5 | -2.81 | -0.06 | -23.91 | -0.06 | -25.9 | 0 | 0 | 0 | 21.39 | 0.29 | 21.33 | 0.12 | 21.39 | -2.17 | 0 | -33.28 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.29 | 28.65 | -12.6 | 0 | -37.15 | -0.06 | -41.54 | -0.06 | -33.69 | 0 | -29.59
83.456 | 21.45 | -1.76 | -0.06 | -21.97 | 0 | -24.32 | -0.06 | 0 | -0.06 | 20.1 | 0.18 | 21.39 | 0.18 | 21.27 | -1.23 | 0.06 | -30.24 | 0.29 | 0 | -0.18 | 0 | -0.18 | 0.06 | -0.18 | 28.54 | 0.35 | 28.3 | -10.55 | -0.18 | -38.32 | 0 | -40.61 | -0.06 | -42.19 | 0 | -34.98
83.488 | 21.39 | -1 | -0.06 | -19.75 | -0.06 | -22.33 | -0.06 | -0.29 | -0.06 | 0.12 | 0.23 | 21.27 | 0.18 | 21.39 | -0.82 | 0 | -27.77 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.65 | 0.29 | 28.48 | -8.09 | -0.18 | -37.15 | -0.23 | -37.38 | 0 | -40.9 | 0 | -42.89
83.52 | 21.45 | -0.47 | -0.06 | -17.87 | -0.06 | -20.45 | -0.18 | -73.54 | 0.06 | 0 | 0.23 | 21.15 | 0.18 | 21.21 | -0.53 | 0.12 | -27.13 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.3 | -5.39 | -0.18 | -35.27 | -0.06 | -34.22 | 0 | -37.21 | 0 | -41.13
83.552 | 21.45 | -0.18 | -0.06 | -15.7 | -0.06 | -17.81 | -0.18 | -51.51 | 0.06 | 0 | 0.23 | 21.5 | 0.23 | 21.39 | -0.23 | 0.06 | -25.43 | 0.12 | -3.46 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.35 | 28.42 | -3.69 | -0.06 | -32.81 | -0.18 | -32.52 | -0.06 | -33.11 | 0 | -35.86
83.584 | 21.39 | -0.06 | 0 | -14.24 | 0.06 | -15.94 | -0.18 | -42.6 | -0.06 | 0 | 0.12 | 21.27 | 0.29 | 21.21 | 0 | -0.06 | -23.2 | 0.06 | -72.6 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.83 | -2.17 | 0 | -30.59 | -0.18 | -32.7 | -0.06 | -32.93 | -0.18 | -33.75
83.616 | 21.45 | -0.23 | -0.06 | -12.89 | 0.06 | -14.71 | -0.18 | -37.21 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.39 | -0.06 | 0.06 | -21.56 | 0.06 | -59.59 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.29 | 28.71 | -1.46 | 8.44 | -29.77 | -0.06 | -31.23 | -0.06 | -32.58 | -0.06 | -32.93
83.648 | 21.45 | -0.35 | -0.06 | -11.48 | -0.06 | -12.77 | -0.18 | -32.52 | -0.18 | 0 | 0.06 | 21.15 | 0.29 | 21.39 | 0.23 | 0.06 | -20.1 | 0 | -44.3 | -0.06 | 0 | -0.18 | 0 | -0.12 | 28.65 | 0.18 | 28.42 | -0.7 | 15.59 | -27.54 | -0.18 | -28.54 | -0.18 | -31.23 | 0 | -32.17
83.68 | 21.45 | -0.29 | 3.4 | -9.61 | -0.06 | -11.48 | -0.18 | -32.64 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 21.39 | 0.29 | 0.06 | -18.93 | 0.12 | -44.01 | -0.06 | 0 | 0 | 0 | -0.18 | 28.59 | 0.23 | 28.48 | -0.47 | 23.38 | -25.84 | -0.18 | -26.84 | -0.06 | -28.65 | -0.18 | -30.18
83.712 | 21.45 | -0.41 | 11.48 | -7.85 | -0.06 | -9.43 | -0.06 | -32.64 | -0.06 | 0 | 0 | 21.27 | 0.23 | 21.33 | 0.23 | 2.99 | -16.52 | 0.06 | -42.31 | 0 | 0 | -0.06 | 0 | -0.06 | 28.65 | 0.12 | 28.65 | 0.12 | 28.65 | -23.03 | -0.06 | -25.31 | 0 | -27.01 | -0.06 | -28.3
83.744 | 21.45 | -0.29 | 16.76 | -6.62 | -0.06 | -8.14 | -0.18 | -31.52 | -0.06 | 0 | 0.12 | 21.45 | 0.23 | 21.74 | 0.23 | 9.38 | -14.47 | 0 | -41.19 | 0 | -46.35 | -0.18 | 0 | -0.06 | 28.48 | 0.18 | 28.65 | 0.23 | 28.59 | -21.68 | -0.18 | -24.2 | -0.18 | -25.9 | 0 | -27.19
83.776 | 21.45 | -0.29 | 21.8 | -4.8 | -0.06 | -6.86 | -0.18 | -29.77 | -0.18 | 0 | 0 | 21.33 | 0.18 | 21.33 | 0.23 | 14.47 | -13.13 | 0.06 | -38.85 | -0.18 | -70.26 | -0.06 | -22.62 | 0 | 28.65 | -0.06 | 28.24 | 0.23 | 28.83 | -20.39 | -0.18 | -22.97 | -0.06 | -24.84 | 0 | -26.02
83.808 | 21.45 | -0.29 | 21.5 | -3.93 | -0.06 | -5.22 | -0.06 | -27.19 | -0.06 | 0 | 0.06 | 21.33 | 0 | 21.45 | 0.18 | 20.27 | -11.54 | 0 | -36.51 | -0.06 | -48.52 | -0.18 | -79.46 | -0.18 | 0.88 | 0 | 28.71 | 0.29 | 28.48 | -17.99 | -0.18 | -21.27 | -0.18 | -22.62 | -0.06 | -24.14
83.84 | 21.45 | -0.29 | 21.45 | -2.46 | 6.74 | -3.98 | -0.18 | -25.55 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.45 | 0.12 | 21.39 | -10.02 | 0.06 | -35.22 | -0.06 | -46.06 | -0.06 | -53.03 | -0.06 | 0.12 | 0.12 | 28.83 | 0.29 | 28.48 | -16.23 | -0.18 | -19.34 | -0.18 | -21.09 | 0 | -22.38
83.872 | 21.45 | -0.23 | 21.45 | -1.29 | 12.36 | -2.4 | -0.06 | -22.38 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 21.21 | 0.23 | 21.33 | -7.62 | 0.12 | -33.17 | -0.18 | -43.65 | -0.06 | -46.76 | -0.06 | 0 | 0 | 28.65 | 0.53 | 28.65 | -13.54 | -0.06 | -17.29 | -0.06 | -19.1 | 0 | -20.63
83.904 | 21.5 | -0.18 | 21.45 | -0.76 | 17.93 | -1.46 | 6.09 | -20.86 | -0.18 | -63.05 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -5.86 | 0 | -31.11 | -0.06 | -41.84 | -0.06 | -44.71 | 0 | 0.06 | 0 | 28.54 | 0.41 | 28.65 | -11.66 | 0 | -15.94 | 0 | -17.4 | -0.18 | -18.75
83.936 | 21.45 | -0.18 | 21.39 | -0.29 | 21.68 | -0.76 | 11.84 | -18.63 | 0 | -62.35 | -0.06 | 0 | 0.06 | 3.46 | 0.23 | 21.45 | -4.51 | 0.06 | -29.06 | 0 | -39.2 | -0.06 | -41.19 | -0.18 | 0 | -0.06 | 28.48 | 0.41 | 28.48 | -9.61 | 8.91 | -14.24 | 0 | -15.82 | 0 | -17.29
83.968 | 21.45 | -0.29 | 21.45 | -0.23 | 21.56 | -0.53 | 17.46 | -16.93 | -0.06 | -43.95 | -0.06 | 0 | 0.06 | 0.06 | 0.18 | 21.33 | -2.99 | 0 | -26.66 | -0.06 | -37.91 | 0 | -40.78 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.42 | -7.62 | 16.11 | -12.48 | 0 | -13.95 | 4.45 | -15.7
84 | 21.45 | -0.23 | 21.39 | -0.29 | 21.45 | -0.18 | 21.68 | -14.59 | 0 | -39.44 | 0 | 0 | 0 | 0 | 0.12 | 21.33 | -2.17 | 0 | -24.96 | -0.06 | -36.45 | 0 | -38.61 | -0.18 | 0 | 0 | 28.71 | 0.29 | 28.48 | -5.22 | 23.79 | -10.61 | 3.87 | -12.77 | 12.72 | -14.18
84.032 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -11.84 | -0.18 | -33.17 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.39 | -1.11 | 0.06 | -22.38 | -0.06 | -33.99 | -0.18 | -35.98 | -0.06 | 0 | -0.06 | 28.54 | 0.12 | 28.71 | -2.93 | 28.59 | -8.56 | 14.71 | -10.2 | 22.91 | -12.07
84.064 | 21.39 | -0.23 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | -10.37 | -0.06 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.45 | -0.76 | 0 | -20.74 | -0.06 | -32.23 | -0.06 | -34.45 | -0.06 | 0 | -0.06 | 28.54 | 0.23 | 28.65 | -2.11 | 28.59 | -7.03 | 22.5 | -8.91 | 29.06 | -10.49
84.096 | 21.39 | -0.29 | 21.45 | -0.35 | 21.39 | -0.35 | 21.45 | -7.79 | -0.06 | -32.93 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 0.06 | -18.69 | -0.06 | -29.83 | -0.06 | -32.76 | -0.06 | 0 | -0.06 | 28.3 | 0.06 | 28.42 | -1 | 28.42 | -5.16 | 29.01 | -7.09 | 28.65 | -8.61
84.128 | 21.5 | -0.18 | 21.45 | -0.35 | 21.45 | -0.53 | 21.39 | -5.74 | 0 | -32.23 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 0.12 | -16.82 | -0.06 | -28.13 | -0.06 | -30.82 | -0.06 | 0 | 0 | 28.3 | 0.12 | 28.36 | -0.7 | 28.48 | -3.87 | 28.71 | -5.8 | 28.77 | -7.38
84.16 | 21.39 | -0.18 | 21.45 | -0.41 | 21.45 | -0.35 | 21.5 | -3.28 | -0.06 | -30.53 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.27 | -0.06 | 0.06 | -15.76 | -0.06 | -26.43 | 0 | -29.01 | -0.18 | 0 | 0 | 28.42 | 0.12 | 28.13 | -0.18 | 28.36 | -2.75 | 28.54 | -4.04 | 28.42 | -5.57
84.192 | 0.12 | -0.18 | 21.45 | -0.29 | 21.39 | -0.29 | 21.39 | -1.7 | -0.18 | -28.24 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.09 | 0 | 0 | -13.65 | 0 | -23.67 | 0 | -26.31 | 0 | 0 | 0 | 28.71 | 0.23 | 28.48 | 0 | 28.59 | -1.76 | 28.42 | -2.23 | 28.42 | -3.57
84.224 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.23 | 21.33 | -0.82 | -0.06 | -25.31 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | 0.18 | 1.99 | -12.01 | 0.06 | -21.21 | 0 | -24.2 | -0.06 | -13.83 | 0 | 0.18 | 0.06 | 28.54 | 0.12 | 28.65 | -1.11 | 28.42 | -1.7 | 28.48 | -2.52
84.256 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | -0.59 | 0 | -23.91 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.12 | 5.86 | -11.07 | 0 | -19.63 | 0 | -22.38 | 0 | -93.52 | 0 | 0 | 0.12 | 28.42 | 0.29 | 28.42 | -0.7 | 28.42 | -1.11 | 28.3 | -1.64
84.288 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | -0.06 | -0.06 | -21.68 | -0.06 | -44.24 | 0 | 0 | 0.06 | 21.33 | 0.18 | 11.54 | -9.84 | 0 | -17.93 | 0 | -21.27 | -0.18 | -60.65 | 0 | 0 | 0 | 28.42 | 0.23 | 28.59 | -0.41 | 28.42 | -0.47 | 28.36 | -1
84.32 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | 0.12 | -0.06 | -19.81 | -0.06 | -65.16 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 17.23 | -8.32 | -0.06 | -16.17 | -0.06 | -18.16 | -0.06 | -47.76 | -0.06 | 0 | -0.06 | 28.42 | 0.35 | 28.42 | -0.18 | 28.65 | -0.35 | 28.59 | -0.7
84.352 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | 0.29 | 0 | -17.05 | -0.06 | -43.01 | -0.06 | -2.46 | 0 | 21.39 | 0.12 | 21.39 | -6.68 | 0 | -14.06 | 0 | -16.11 | 0 | -44.06 | -0.18 | 0 | -0.06 | 28.3 | 0.29 | 28.48 | 0 | 28.42 | -0.06 | 28.42 | -0.18
84.384 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | 0.41 | 0 | -15.06 | -0.18 | -39.14 | -0.06 | -71.19 | 0 | 21.33 | 0.29 | 21.39 | -5.68 | 0 | -12.77 | -0.06 | -14.65 | -0.06 | -38.15 | -0.06 | 0 | -0.18 | 28.48 | 0.41 | 28.65 | 0 | 28.42 | -0.18 | 28.48 | -0.06
84.416 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | -0.06 | -13.36 | -0.06 | -33.52 | -0.06 | -56.19 | 0 | 21.33 | 0.12 | 21.33 | -4.63 | 0.06 | -11.6 | -0.06 | -13.48 | 0 | -37.09 | -0.06 | 0 | 0 | 28.36 | 0.23 | 28.65 | -0.06 | 28.48 | -0.06 | 28.71 | -0.23
84.448 | 0 | 0 | 21.39 | -0.18 | 21.39 | -0.29 | 21.39 | 0.53 | 0 | -11.37 | 0 | -31.99 | 0.06 | -33.11 | 0.06 | 21.39 | 0.23 | 21.39 | -3.63 | -0.06 | -10.08 | 3.81 | -12.07 | 0 | -37.44 | 0 | 0 | -0.06 | 28.48 | 0.35 | 28.48 | 0.06 | 28.24 | 0.29 | 28.65 | 0
84.48 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 0 | -9.55 | -0.06 | -32.7 | -0.06 | -27.72 | 0.06 | 21.33 | 0.06 | 21.39 | -3.11 | -0.06 | -8.61 | 12.25 | -10.25 | 0 | -38.09 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.59 | 0.06 | 28.42 | 0.29 | 28.71 | 0
84.512 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 0 | -6.15 | -0.06 | -32.17 | -0.18 | -35.16 | -0.06 | 21.39 | 0.29 | 21.27 | -2.29 | 8.73 | -6.91 | 22.38 | -8.91 | 0 | -36.21 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 28.13 | 0.23 | 28.48 | 0.35 | 28.48 | 0.23
84.544 | 0 | -0.06 | 21.5 | -0.18 | 21.45 | -0.18 | 21.45 | 0.41 | 7.09 | -4.28 | 0 | -30.53 | 0 | -34.75 | 0.06 | 21.33 | 0.12 | 21.45 | -1.82 | 15.82 | -5.98 | 29.12 | -7.44 | 0 | -33.87 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.54 | 0.18 | 28.59 | 0.35 | 28.59 | 0.06
84.576 | -0.06 | -0.06 | 0.12 | -0.18 | 21.39 | -0.06 | 21.33 | 0.41 | 12.66 | -2.29 | -0.06 | -27.77 | 0.06 | -32.76 | 0.06 | 21.39 | 0.35 | 21.15 | -1.46 | 23.73 | -4.98 | 28.95 | -6.45 | 1.99 | -32.11 | 0 | 0 | 0 | 0 | 0.12 | 28.42 | 0.18 | 28.42 | 0.35 | 28.18 | 0.12
84.608 | -33.17 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.41 | 18.46 | -1.76 | -0.06 | -25.67 | -0.18 | -29.65 | 0 | 2.87 | 0.18 | 21.39 | -1.29 | 28.83 | -4.04 | 28.59 | -5.33 | 8.32 | -30.12 | 0 | 0 | 0.06 | 0 | 0.12 | 28.48 | 0.18 | 28.42 | 0.35 | 28.42 | 0.12
84.64 | -58.42 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.41 | 21.5 | -0.7 | -0.06 | -23.91 | -0.06 | -27.42 | -0.06 | 0.06 | 0.23 | 21.33 | -0.94 | 28.48 | -3.16 | 28.65 | -3.93 | 15.53 | -28.59 | 0 | -36.51 | 0 | 0 | 0 | 28.48 | 0.12 | 28.48 | 0.41 | 28.59 | 0.12
84.672 | -40.78 | 0 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.35 | 21.33 | -0.06 | -0.06 | -21.45 | -0.06 | -26.37 | -0.06 | 0 | 0.18 | 21.33 | -0.7 | 28.36 | -2.46 | 28.48 | -3.11 | 25.9 | -25.84 | -0.06 | -74.94 | 0.06 | 0 | 0 | 28.48 | 0.06 | 28.54 | 0.29 | 28.36 | 0.12
84.704 | -40.31 | -0.06 | 0 | -0.23 | 21.45 | -0.06 | 21.39 | 0.29 | 21.39 | 0.18 | -0.06 | -19.69 | 0 | -24.9 | -0.06 | 0.06 | 0.12 | 21.39 | -0.64 | 28.65 | -1.88 | 28.65 | -2.46 | 28.71 | -24.32 | -0.06 | -49.4 | 0 | 0 | 0.06 | 28.77 | 0 | 28.3 | 0.29 | 28.48 | 0.23
84.736 | -37.56 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.39 | 0.29 | 21.39 | 0.23 | -0.06 | -17.81 | -0.06 | -22.85 | 0 | 0 | 0.18 | 21.33 | -0.53 | 28.65 | -1.7 | 28.89 | -1.93 | 28.59 | -23.03 | -0.06 | -47.46 | 0 | 0 | 0.06 | 28.54 | 0.06 | 28.36 | 0.18 | 28.48 | 0.06
84.768 | -35.1 | -0.18 | 0.06 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | 0.35 | 0 | -15.59 | 0 | -20.92 | -0.06 | 0 | 0.18 | 21.39 | -0.41 | 28.42 | -1.23 | 28.48 | -1.23 | 28.71 | -21.09 | 0 | -41.66 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.07 | 0.23 | 28.48 | 0.06
84.8 | -33.52 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | 0.41 | -0.06 | -13.71 | -0.06 | -19.81 | -0.18 | 0 | 0.06 | 21.33 | -0.35 | 29.06 | -0.82 | 28.3 | -0.94 | 28.42 | -18.75 | 0 | -36.74 | -0.06 | 0 | -0.06 | 28.48 | -0.06 | 28.48 | 0.18 | 28.48 | 0.18
84.832 | -30.76 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | 0.41 | -0.06 | -10.9 | -0.06 | -17.81 | -0.06 | 0 | 0.12 | 21.33 | -0.35 | 28.54 | -0.59 | 28.59 | -0.64 | 28.59 | -16.23 | -0.18 | -37.38 | 0 | 0 | 0.06 | 0.12 | 0 | 28.48 | 0.06 | 28.48 | 0.18
84.864 | -28.89 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | -0.06 | -8.91 | 0 | -15.94 | -0.06 | 0 | 0.18 | 21.33 | -0.35 | 28.54 | -0.35 | 28.89 | -0.59 | 28.48 | -14.59 | -0.06 | -38.32 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 0.29 | 0.18
84.896 | -27.13 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12 | 21.39 | 0.47 | -0.06 | -6.62 | -0.06 | -14.06 | -0.06 | 0 | 0.12 | 21.39 | -0.41 | 28.36 | -0.41 | 28.77 | -0.35 | 28.54 | -12.6 | 0 | -36.92 | 0 | 0 | -0.18 | 0 | -0.06 | 0.23 | 0.06 | 0.06 | 0.06
84.928 | -24.61 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.47 | 4.92 | -4.63 | -0.06 | -12.48 | 0 | 0 | 0.18 | 21.33 | -0.35 | 28.71 | -0.35 | 28.42 | -0.41 | 28.65 | -10.49 | -0.06 | -35.22 | -0.06 | -62.58 | 0 | 0 | -0.06 | 0.06 | 0 | 0 | 0.23
84.96 | -22.79 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 10.9 | -2.99 | -0.06 | -10.61 | 0.06 | 0 | 0.06 | 21.39 | -0.29 | 27.95 | -0.35 | 28.89 | -0.29 | 28.18 | -8.56 | -0.06 | -33.4 | -0.06 | -78.28 | -0.06 | 0 | 0 | 0 | -0.06 | 0.06 | 0.18
84.992 | -20.16 | 0 | -30.06 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 18.28 | -1.7 | -0.06 | -8.79 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.65 | -0.18 | 28.95 | -0.18 | 28.65 | -5.68 | 2.99 | -30.59 | -0.06 | -48.28 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18
85.024 | -18.05 | -0.18 | -61.88 | -0.12 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.62 | -0.82 | 0 | -6.68 | -0.06 | -4.92 | 0.23 | 21.39 | -0.23 | 28.77 | -0.23 | 28.42 | -0.35 | 28.48 | -3.98 | 11.31 | -28.77 | 0 | -45.35 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
85.056 | -16.7 | 0 | -45 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.47 | 21.39 | -0.47 | -0.06 | -5.22 | -0.06 | -68.97 | 0.18 | 21.39 | -0.35 | 28.48 | -0.23 | 28.3 | -0.29 | 28.42 | -2.23 | 18.87 | -27.19 | -0.06 | -40.31 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.18
85.088 | -14.94 | -0.06 | -40.2 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | -0.06 | 0.59 | -3.87 | 0 | -55.61 | 0.12 | 21.39 | -0.29 | 28.59 | -0.29 | 28.54 | -0.18 | 28.65 | -1.7 | 26.72 | -25.14 | -0.06 | -36.62 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
85.12 | -13.71 | 0 | -38.61 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.06 | 8.5 | -2.52 | -0.06 | -34.86 | 0.23 | 21.39 | -0.23 | 28.48 | -0.23 | 29.3 | -0.18 | 28.3 | -0.76 | 28.48 | -23.97 | 0 | -37.38 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18
85.152 | -11.25 | -0.06 | -34.86 | -0.18 | -35.8 | -0.06 | 0 | 0 | 21.33 | 0.29 | 21.39 | 0.41 | 15.65 | -1.7 | -0.06 | -28.89 | 0.23 | 0.12 | -0.29 | 28.71 | -0.18 | 28.48 | -0.18 | 28.48 | -0.29 | 28.59 | -21.39 | -0.06 | -38.32 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06
85.184 | -10.37 | 0 | -34.22 | -0.06 | -56.66 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | 0.47 | 21.33 | -0.76 | 0 | -34.45 | 0.06 | 0 | -0.23 | 29.01 | -0.18 | 28.54 | -0.29 | 28.48 | -0.06 | 28.42 | -19.28 | 0 | -36.45 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0
85.216 | -8.85 | 0.06 | -31.93 | -0.06 | -43.42 | -0.18 | -43.54 | -0.06 | 21.45 | 0.18 | 21.27 | 0.35 | 21.39 | -0.64 | 0 | -34.63 | 0.23 | 0 | -0.18 | 28.77 | -0.23 | 28.59 | -0.18 | 28.65 | 0 | 28.65 | -17.46 | 0 | -34.92 | 0 | 0.12 | 0 | 0 | 0 | 0 | 0.06
85.248 | -7.32 | -0.06 | -29.83 | -0.18 | -39.9 | -0.06 | -67.74 | 0.06 | 21.39 | 0.06 | 21.39 | 0.53 | 21.39 | -0.23 | 0 | -32.64 | 0.12 | 0 | -0.18 | 28.48 | -0.18 | 28.77 | -0.18 | 28.48 | 0.29 | 28.48 | -16 | 0 | -32.52 | -0.06 | -61.99 | -0.06 | 0 | -0.06 | 0 | 0.06
85.28 | -5.86 | -0.06 | -28.24 | -0.06 | -37.97 | -0.06 | -47.58 | 0.06 | 21.33 | 0.12 | 21.39 | 0.41 | 21.33 | -0.18 | 0 | -29.71 | 0.12 | 0 | -0.18 | 28.48 | -0.18 | 28.65 | -0.18 | 28.36 | 0.23 | 28.65 | -14.06 | 0 | -31.11 | 0 | -77.46 | -0.06 | 0 | 0 | -22.21 | 0.18
85.312 | -3.75 | -0.06 | -25.37 | -0.06 | -34.81 | -0.06 | -38.85 | -0.06 | 21.45 | 0.06 | 21.33 | 0.47 | 21.33 | -0.06 | 0 | -27.19 | 0 | 0 | -0.18 | 28.48 | -0.18 | 28.59 | -0.18 | 28.48 | 0.35 | 28.54 | -11.25 | 0 | -28.36 | -0.06 | -39.38 | -0.06 | -43.36 | 0 | -78.93 | 0
85.344 | -2.75 | -0.06 | -22.97 | -0.06 | -33.05 | -0.18 | -34.63 | -0.06 | 21.33 | 0.12 | 21.39 | 0.35 | 21.33 | 0.12 | 2.99 | -26.66 | 0.06 | 0 | -0.18 | 28.42 | -0.18 | 0.23 | -0.18 | 28.48 | 0.41 | 28.65 | -9.02 | 0 | -26.84 | -0.06 | -28.48 | 0 | -80.28 | -0.06 | -52.21 | 0.18
85.376 | -1.7 | 5.45 | -21.5 | -0.06 | -31.58 | 0 | -31.99 | -0.06 | 21.56 | 0.12 | 21.27 | 0.41 | 21.39 | 0.12 | 9.43 | -25.2 | 0.06 | 0 | -0.18 | 28.48 | -0.06 | 0.06 | -0.18 | 28.48 | 0.41 | 28.42 | -7.03 | 0 | -25.55 | -0.06 | -38.15 | -0.18 | -52.03 | -0.06 | -30.47 | 0.12
85.408 | -0.94 | 11.48 | -19.51 | 0 | -29.65 | -0.06 | -32.4 | -0.06 | 21.39 | 0.06 | 21.33 | 0.41 | 21.27 | 0.06 | 14.53 | -23.09 | -0.06 | 0 | -0.18 | 0.23 | -0.06 | 0 | -0.18 | 28.36 | 0.35 | 28.71 | -4.75 | 0 | -23.56 | -0.06 | -43.07 | -0.06 | -30.94 | 0 | -32.52 | 0
85.44 | -0.41 | 16.82 | -17.4 | -0.06 | -26.54 | -0.06 | -32.52 | 0.06 | 0.12 | -0.06 | 21.39 | 0.35 | 21.27 | 0.12 | 20.39 | -21.21 | 0 | 0 | -0.18 | 0.06 | -0.18 | 0.06 | -0.18 | 28.59 | 0.41 | 28.71 | -3.11 | 0.06 | -21.45 | -0.18 | -39.44 | -0.06 | -32.52 | 0 | -42.01 | 0
85.472 | -0.18 | 21.45 | -15 | -0.18 | -25.02 | -0.06 | -30.82 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.12 | 21.39 | -19.34 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 17.7 | 0.35 | 28.59 | -1.7 | 5.51 | -19.45 | -0.18 | -34.75 | 0 | -42.6 | 0 | -40.02 | -0.06
85.504 | -0.06 | 21.45 | -13.77 | -0.06 | -22.56 | -0.06 | -28.18 | 0 | 0 | -0.06 | 21.33 | 0.23 | 21.56 | 0.18 | 21.45 | -17.81 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.18 | 0.35 | 28.48 | -1 | 13.48 | -17.23 | -0.18 | -32.81 | -0.06 | -40.31 | 0 | -36.1 | -0.06
85.536 | -0.23 | 21.45 | -12.07 | -0.06 | -20.57 | -0.18 | -26.02 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.33 | 0.12 | 21.45 | -16.29 | -0.06 | 0.06 | -0.18 | 0 | -0.18 | 0.06 | -0.06 | 0 | 0.41 | 28.3 | -0.47 | 21.27 | -15.18 | -0.06 | -32.52 | 0 | -35.98 | -0.06 | -33.34 | -0.06
85.568 | -0.29 | 21.45 | -10.84 | -0.06 | -18.63 | -0.06 | -24.43 | 0 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.29 | 21.27 | -14.41 | 0 | -37.85 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.35 | 28.77 | -0.29 | 28.95 | -13.24 | -0.18 | -32.4 | -0.06 | -33.93 | -0.18 | -32.76 | 0.06
85.6 | -0.29 | 21.39 | -9.32 | -0.06 | -16.47 | -0.06 | -22.33 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.18 | 21.33 | -12.48 | -0.06 | -66.21 | -0.12 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 28.65 | 0 | 28.65 | -11.72 | -0.06 | -30.59 | -0.18 | -32.81 | -0.06 | -32.64 | -0.06
85.632 | -0.29 | 21.45 | -7.27 | -0.06 | -15.18 | -0.06 | -19.45 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | 0.06 | 21.33 | -10.37 | -0.18 | -36.8 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.42 | 0.12 | 28.48 | -8.79 | -0.06 | -27.42 | -0.06 | -32.76 | -0.18 | -30.29 | -0.06
85.664 | -0.29 | 21.45 | -5.92 | -0.06 | -13.48 | -0.06 | -17.75 | -0.18 | 0 | 0 | 21.33 | 0.12 | 21.27 | 0.18 | 21.39 | -8.67 | 0 | -27.54 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.29 | 28.42 | 0.12 | 28.42 | -6.8 | -0.18 | -26.37 | -0.18 | -31 | -0.06 | -28.48 | -0.18
85.696 | -0.41 | 21.5 | -4.63 | -0.06 | -11.89 | -0.18 | -15.7 | 0 | 0 | 0 | 21.39 | -0.06 | 21.45 | 0.18 | 21.39 | -7.27 | -0.06 | -32.4 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.29 | 28.48 | 0.29 | 28.42 | -4.69 | -0.18 | -25.43 | -0.06 | -28.3 | 0 | -26.9 | -0.06
85.728 | -0.35 | 21.39 | -3.52 | -0.06 | -10.49 | -0.06 | -13.71 | -0.06 | 0 | 0 | 21.39 | 0 | 21.33 | 0.12 | 21.39 | -5.1 | -0.06 | -35.45 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 28.42 | 0.23 | 28.42 | -2.64 | 0 | -23.67 | -0.06 | -26.54 | 0 | -25.72 | 0
85.76 | -0.23 | 21.45 | -2.05 | 0 | -8.85 | 0 | -11.84 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.06 | 21.5 | -3.69 | 0 | -34.92 | -0.18 | 0 | 0 | -14.82 | -0.06 | 0 | 0.18 | 28.42 | 0.29 | 28.59 | -2.11 | -0.18 | -22.15 | -0.06 | -25.72 | -0.18 | -24.08 | -0.06
85.792 | -0.23 | 21.45 | -1.23 | 9.43 | -7.21 | 0.47 | -9.08 | -0.06 | 0 | -0.06 | 21.33 | 0 | 21.33 | 0.12 | 21.39 | -2.64 | 0 | -30.53 | -0.18 | 0 | 0 | -71.19 | -0.06 | 0 | 0.12 | 28.71 | 0.35 | 28.77 | -0.94 | -0.18 | -19.92 | -0.06 | -23.85 | -0.18 | -22.21 | -0.18
85.824 | -0.23 | 21.45 | -0.53 | 14.94 | -5.98 | 8.14 | -6.91 | 0 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0.12 | 21.33 | -1.46 | -0.06 | -27.83 | -0.06 | -44.24 | -0.06 | -50.16 | -0.06 | 0 | 0.06 | 28.65 | 0.35 | 28.54 | -0.59 | -0.06 | -18.16 | 0 | -21.97 | -0.06 | -20.33 | 0
85.856 | -0.23 | 21.45 | -0.18 | 20.16 | -4.16 | 13.89 | -4.45 | -0.06 | -51.92 | 0 | 0 | -0.06 | 21.33 | 0.06 | 21.39 | -0.94 | -0.06 | -26.84 | -0.06 | -67.8 | 0 | -45.65 | -0.18 | 0 | 0.06 | 28.65 | 0.35 | 28.59 | -0.06 | 0 | -16.88 | -0.18 | -20.68 | 0 | -19.04 | 0
85.888 | -0.23 | 21.45 | -0.23 | 21.5 | -3.16 | 19.22 | -2.81 | -0.06 | -63.05 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.39 | -0.64 | -0.06 | -25.84 | -0.18 | -48.28 | -0.06 | -44.59 | 0 | 0.06 | 0.06 | 13.71 | 0.35 | 28.42 | -0.18 | -0.06 | -15.29 | -0.06 | -18.87 | -0.18 | -17.75 | 0
85.92 | -0.23 | 21.39 | -0.35 | 21.5 | -1.64 | 21.56 | -1.76 | -0.06 | -45.94 | -0.06 | 0 | 0 | 21.33 | 0.06 | 21.39 | -0.18 | 4.98 | -24.08 | -0.18 | -45.18 | -0.06 | -42.48 | -0.06 | -45.65 | 0.06 | 0.12 | 0.23 | 28.42 | 0.23 | 1.52 | -13.36 | 0 | -17.64 | -0.18 | -16.06 | -0.06
85.952 | -0.18 | 21.39 | -0.41 | 21.39 | -1 | 21.5 | -0.82 | -0.06 | -37.97 | -0.06 | 0 | 0 | 21.33 | 0 | 21.39 | -0.06 | 12.95 | -21.5 | -0.06 | -43.71 | -0.06 | -40.31 | -0.06 | -60.65 | 0.12 | 0 | 0.23 | 28.59 | 0.12 | 13.65 | -11.37 | 3.87 | -15.94 | -0.06 | -13.95 | 0
85.984 | -0.29 | 21.45 | -0.41 | 21.45 | -0.47 | 21.45 | -0.18 | 0 | -33.52 | -0.18 | 0 | -0.06 | 0.23 | 0.12 | 21.39 | -0.18 | 18.28 | -20.1 | 0 | -41.37 | 0 | -37.97 | 0 | -46.06 | 0 | 0.06 | 0.23 | 28.42 | 0.35 | 21.21 | -10.02 | 11.95 | -14.06 | 4.45 | -12.31 | 0
86.016 | -0.23 | 21.45 | -0.41 | 21.45 | -0.35 | 21.5 | 0 | 0 | -32.23 | -0.06 | 0.06 | 0 | 0.06 | 0.06 | 21.33 | 0.18 | 21.74 | -18.87 | -0.06 | -39.55 | -0.06 | -36.33 | -0.06 | -45.35 | 0 | 0 | 0.23 | 28.71 | 0.29 | 29.01 | -8.38 | 19.81 | -12.54 | 12.77 | -10.78 | 5.39
86.048 | -0.18 | 21.5 | -0.41 | 21.45 | -0.23 | 21.5 | 0.12 | 0 | -32.29 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.12 | 21.45 | -17.23 | -0.06 | -38.15 | 0 | -34.45 | 0 | -44.47 | 0 | 0 | 0 | 28.42 | 0.35 | 28.59 | -6.68 | 27.36 | -10.72 | 20.51 | -9.08 | 13.42
86.08 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | 0.29 | -0.06 | -32.23 | -0.06 | 0 | 0.06 | 0 | 0 | 21.39 | 0.29 | 21.15 | -15.18 | -0.06 | -35.39 | 0 | -32.93 | -0.06 | -40.61 | -0.06 | 0 | 0.12 | 28.42 | 0.29 | 28.65 | -4.92 | 28.65 | -9.08 | 28.36 | -7.44 | 21.15
86.112 | -0.06 | 21.45 | -0.29 | 21.45 | -0.35 | 21.45 | 0.41 | -0.06 | -30.41 | 0 | 0 | 0 | 0 | 0 | 21.39 | 0.23 | 21.27 | -13.01 | -0.23 | -33.52 | 0 | -30.24 | -0.06 | -38.91 | 0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.59 | -2.99 | 28.54 | -6.86 | 28.36 | -5.22 | 28.77
86.144 | -0.06 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.47 | -0.06 | -27.89 | 0 | 0 | 0 | 0 | 0.06 | 21.21 | 0.18 | 21.21 | -11.48 | -0.06 | -31.88 | 0 | -28.71 | -0.18 | -37.09 | 0.06 | 0 | -0.06 | 28.54 | 0.29 | 28.54 | -1.93 | 28.48 | -5.22 | 28.59 | -3.87 | 28.48
86.176 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.45 | 0.47 | 0.41 | -25.84 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.39 | -9.49 | -0.06 | -29.59 | 0 | -25.67 | 0 | -35.74 | 0 | 0 | 0 | 28.54 | 0.23 | 28.65 | -1.29 | 28.36 | -3.63 | 28.42 | -2.46 | 28.42
86.208 | -0.18 | 21.45 | -0.29 | 21.45 | -0.41 | 21.45 | 0.41 | 8.03 | -23.91 | 0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -7.85 | -0.06 | -28.13 | 0 | -24.55 | -0.18 | -33.63 | 0 | 0 | 0.06 | 28.48 | 0.12 | 28.59 | -0.94 | 28.48 | -2.52 | 28.48 | -1.7 | 28.42
86.24 | -0.06 | 21.39 | -0.23 | 21.5 | -0.35 | 21.45 | 0.53 | 13.54 | -21.68 | 0 | -31 | 0 | 0 | 0 | 2.46 | 0.23 | 21.45 | -6.09 | -0.18 | -25.37 | -0.06 | -22.85 | -0.06 | -32.34 | 0 | 0 | 0 | 28.54 | 0.18 | 28.48 | -0.47 | 28.48 | -1.76 | 28.65 | -1.11 | 28.42
86.272 | -0.06 | 0.12 | -0.29 | 21.45 | -0.35 | 21.45 | 0.41 | 21.04 | -19.45 | 0 | -63.28 | 0 | 0 | 0.06 | 0.06 | 0.18 | 21.33 | -4.16 | 10.31 | -23.5 | 0 | -19.98 | 0 | -29.24 | 0 | 0 | -0.06 | 28.48 | 0.12 | 29.01 | -0.23 | 28.59 | -0.82 | 28.59 | -0.7 | 28.36
86.304 | -0.06 | 0 | -0.29 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | -17.17 | 0 | -45.29 | -0.06 | 0 | 0 | 0 | 0.18 | 21.33 | -2.99 | 17.81 | -21.5 | 0 | -18.52 | -0.06 | -27.48 | -0.06 | -1.93 | -0.06 | 28.59 | 0.06 | 28.54 | -0.06 | 28.48 | -0.59 | 28.48 | -0.47 | 28.36
86.336 | -0.06 | 0 | -0.23 | 21.45 | -0.18 | 21.45 | 0.35 | 21.39 | -15.53 | -0.06 | -39.96 | -0.06 | 0 | 0 | 0.06 | 0.23 | 21.09 | -1.93 | 25.43 | -19.45 | -0.06 | -16.47 | -0.06 | -25.67 | 0.06 | -91.82 | 0 | 28.54 | 0.12 | 28.42 | -0.06 | 28.3 | -0.53 | 28.48 | -0.06 | 28.42
86.368 | -0.06 | 0 | -0.29 | 21.45 | -0.18 | 21.45 | 0.41 | 21.33 | -13.59 | 0 | -34.28 | 0.06 | 0 | -0.06 | 0 | 0.29 | 21.33 | -1.11 | 28.89 | -17.7 | -0.06 | -14.82 | -0.06 | -24.26 | -0.18 | -64.57 | 0 | 0.29 | 0.12 | 28.36 | 0.06 | 28.48 | -0.18 | 28.42 | -0.18 | 28.48
86.4 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.29 | 21.45 | -11.48 | 0 | -31.88 | 0 | -0.59 | 0 | 0 | 0.35 | 21.15 | -0.76 | 28.42 | -16 | 0 | -13.24 | -0.06 | -21.74 | 0 | -47.93 | 0 | 0.06 | 0 | 28.71 | 0.12 | 28.42 | -0.18 | 28.71 | 0 | 28.42
86.432 | 0.06 | 0 | -0.23 | 21.45 | -0.18 | 21.39 | 0.29 | 21.33 | -8.56 | -0.06 | -32.76 | 0 | -74.3 | 0 | 0 | 0.35 | 21.45 | -0.29 | 28.77 | -14.36 | 6.91 | -11.37 | 0 | -19.1 | -0.18 | -44.36 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.48 | -0.06 | 28.59 | 0.06 | 28.42
86.464 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.23 | 21.39 | -6.33 | 0.12 | -32.23 | -0.06 | -51.68 | 0 | 0 | 0.18 | 21.39 | -0.18 | 28.36 | -12.66 | 14.59 | -9.84 | 3.93 | -17.64 | -0.06 | -38.67 | 0 | 0.06 | 0 | 28.65 | 0.18 | 28.59 | -0.06 | 28.59 | 0.12 | 28.3
86.496 | -0.06 | 0 | -0.18 | 21.45 | -0.35 | 21.33 | 0.29 | 21.39 | -4.16 | 0 | -30.53 | 0 | -30.35 | 0 | 0 | 0.23 | 21.39 | -0.06 | 29.01 | -11.25 | 22.44 | -8.79 | 12.25 | -15.76 | -0.06 | -37.03 | 0 | 0 | 0 | 28.65 | 0.12 | 28.42 | 0.12 | 28.24 | 0.12 | 28.48
86.528 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.06 | 21.39 | -2.4 | 0 | -28.54 | -0.06 | -28.71 | -0.18 | 0 | 0.18 | 21.33 | -0.06 | 28.42 | -9.9 | 29.18 | -7.79 | 19.92 | -14.82 | 0 | -37.79 | 0.06 | 0 | 0 | 28.3 | 0.12 | 28.59 | 0.06 | 28.48 | 0.06 | 28.42
86.56 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.18 | 21.39 | -1.64 | 0 | -26.49 | 0 | -35.39 | -0.06 | 0 | 0.29 | 21.39 | 0 | 28.54 | -8.61 | 28.48 | -6.56 | 27.54 | -12.95 | -0.18 | -37.62 | 0 | 0 | -0.06 | 28.54 | 0.18 | 28.54 | 0.18 | 28.65 | 0.06 | 28.65
86.592 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.12 | 21.39 | -0.59 | 7.21 | -23.61 | 0.12 | -34.86 | -0.18 | 0 | 0.29 | 21.39 | 0.06 | 28.59 | -6.97 | 28.59 | -4.75 | 28.48 | -11.13 | -0.06 | -35.51 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.42 | 0.18 | 28.54 | 0.06 | 28.48
86.624 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.12 | 21.39 | -0.35 | 12.77 | -21.86 | -0.06 | -31.17 | -0.18 | 0 | 0.18 | 21.39 | 0.18 | 28.24 | -5.92 | 28.36 | -4.04 | 28.48 | -9.67 | 0 | -34.4 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | 0.06 | 28.42 | 0.06 | 28.3
86.656 | -0.06 | 0 | -0.06 | 5.1 | -0.18 | 21.5 | 0 | 21.39 | 0 | 18.28 | -19.86 | -0.06 | -28.3 | 0.06 | -6.56 | 0.18 | 21.21 | 0.18 | 28.18 | -5.1 | 28.54 | -3.22 | 28.54 | -8.56 | 0 | -31.88 | 0 | 0 | 0 | 28.42 | 0.12 | 28.48 | 0.12 | 28.59 | 0.06 | 28.48
86.688 | -0.06 | -31.93 | -0.18 | 0.12 | -0.18 | 15.59 | 0.12 | 21.39 | 0.12 | 21.56 | -17.93 | -0.06 | -27.77 | -0.06 | -70.96 | 0.23 | 21.27 | 0.35 | 28.54 | -4.04 | 28.13 | -2.64 | 28.59 | -7.38 | -0.06 | -30.12 | -0.23 | 0 | -0.06 | 28.65 | 0.06 | 28.36 | 0.18 | 28.54 | 0.06 | 28.42
86.72 | -0.06 | -62.23 | -0.18 | 0 | -0.18 | 0.06 | 0.06 | 21.21 | 0.29 | 21.39 | -15.82 | -0.06 | -27.07 | -0.06 | -54.9 | 0.06 | 21.27 | 0.29 | 28.3 | -3.28 | 28.42 | -2.29 | 28.65 | -5.86 | -0.06 | -28.24 | -0.06 | 0 | 0 | 28.65 | 0 | 28.42 | 0.23 | 28.42 | 0.12 | 28.42
86.752 | -0.06 | -41.72 | -0.18 | 0.06 | -0.18 | 0 | -0.06 | 21.33 | 0.18 | 21.45 | -13.3 | 0 | -24.61 | -0.12 | -29.18 | 0.18 | 21.39 | 0.18 | 28.48 | -2.46 | 28.48 | -1.46 | 28.65 | -4.34 | -0.06 | -26.37 | 0 | 0 | 0 | 28.59 | -0.06 | 28.48 | 0.12 | 28.48 | 0.06 | 28.42
86.784 | 0 | -40.08 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -11.66 | -0.06 | -21.86 | -0.06 | -29.06 | 0.12 | 21.27 | 0.12 | 28.3 | -1.88 | 28.48 | -1.05 | 28.42 | -3.63 | -0.06 | -24.08 | -0.06 | -37.44 | 0 | 28.59 | 0 | 28.54 | 0.06 | 28.48 | 0 | 28.42
86.816 | -0.06 | -38.61 | -0.18 | 0 | -0.06 | 0 | 0 | 21.45 | 0.35 | 21.5 | -9.55 | -0.06 | -20.45 | -0.06 | -34.69 | 0.23 | 0.12 | 0.06 | 28.89 | -1.46 | 28.65 | -0.64 | 28.48 | -2.99 | -0.18 | -22.62 | -0.18 | -90.71 | -0.18 | 0.29 | -0.18 | 28.48 | 0.06 | 28.48 | 0.06 | 28.36
86.848 | -0.06 | -34.98 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -7.27 | -0.06 | -19.4 | -0.06 | -35.22 | 0.18 | 0 | 0.18 | 28.36 | -1.23 | 28.42 | -0.64 | 28.83 | -2.46 | 0 | -21.21 | -0.06 | -55.02 | 0 | 0.06 | 0 | 0.7 | 0.12 | 28.48 | 0 | 28.42
86.88 | -0.06 | -33.75 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -5.33 | -0.06 | -17.58 | -0.18 | -32.34 | 0.12 | 0 | 0.29 | 28.71 | -0.76 | 28.71 | -0.47 | 28.42 | -1.93 | -0.06 | -18.99 | -0.18 | -47.93 | 0 | 0 | 0 | 0.12 | -0.06 | 0.29 | 0.12 | 28.36
86.912 | -0.18 | -30.59 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -2.64 | -0.06 | -15.29 | -0.06 | -28.59 | 0 | 0 | 0.23 | 28.59 | -0.64 | 28.36 | -0.29 | 27.95 | -1.05 | 1.93 | -16.29 | 0 | -42.6 | -0.06 | 0 | -0.06 | 0 | 0 | 0.06 | 0.06 | 0.23
86.944 | -0.06 | -29.06 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 21.21 | -2.11 | 1.99 | -13.71 | 0 | -27.13 | 0 | 0 | 0.12 | 28.36 | -0.59 | 28.83 | -0.35 | 28.42 | -0.82 | 11.6 | -14.77 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0.06 | 0 | 0 | 0.12 | 0.06
86.976 | -0.06 | -26.84 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.33 | -0.7 | 6.27 | -11.89 | -0.18 | -26.66 | 0 | 0 | 0.18 | 28.48 | -0.47 | 28.59 | -0.29 | 28.36 | -0.64 | 19.1 | -12.6 | -0.06 | -37.21 | -0.06 | 0 | -0.06 | 0 | 0 | 0.06 | 0 | 0
87.008 | -0.06 | -25.08 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.53 | 12.13 | -10.25 | -0.06 | -24.96 | 0.06 | 0 | 0.12 | 28.42 | -0.41 | 28.48 | -0.29 | 28.59 | -0.53 | 26.95 | -10.49 | -0.06 | -37.74 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0
87.04 | -0.06 | -23.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.45 | -0.18 | 17.58 | -8.61 | -0.06 | -22.62 | -0.06 | 0 | 0.23 | 28.54 | -0.35 | 28.59 | -0.29 | 28.65 | -0.23 | 28.77 | -8.56 | -0.18 | -36.92 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
87.072 | -0.06 | -20.57 | -0.06 | -3.22 | -0.18 | 0 | 0 | 0.23 | 0.18 | 21.39 | 0.06 | 21.45 | -6.39 | 8.79 | -20.74 | 0.12 | 0 | 0.18 | 28.89 | -0.53 | 28.18 | -0.29 | 28.83 | -0.23 | 28.48 | -5.39 | -0.18 | -35.63 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
87.104 | -0.06 | -18.28 | -0.06 | -57.07 | -0.06 | -0.64 | 0 | 0.06 | 0.18 | 21.45 | 0.06 | 21.39 | -4.63 | 14.06 | -19.45 | 0 | 0 | 0.06 | 28.59 | -0.29 | 28.71 | -0.18 | 28.42 | -0.35 | 28.18 | -3.75 | -0.18 | -33.22 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
87.136 | -0.06 | -16.7 | -0.06 | -50.8 | -0.06 | -51.39 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | 21.39 | -3.16 | 19.98 | -17.81 | 0.12 | 0 | 0.29 | 23.26 | -0.35 | 28.48 | -0.23 | 28.42 | -0.35 | 28.65 | -2.11 | 0 | -31.7 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
87.168 | -0.06 | -15 | -0.06 | -40.37 | -0.18 | -52.44 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | -2.29 | 21.45 | -15.7 | 0.12 | 0 | 0.18 | 0.18 | -0.29 | 28.36 | -0.18 | 28.48 | -0.35 | 28.77 | -1.64 | -0.18 | -29.47 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0
87.2 | 0 | -13.36 | -0.18 | -39.2 | -0.06 | -40.37 | 0 | 0 | 0.23 | 21.39 | 0.41 | 21.39 | -1.23 | 21.39 | -14.06 | 0 | 0 | 0.12 | 0 | -0.29 | 28.71 | -0.18 | 28.48 | -0.35 | 28.48 | -0.82 | 0 | -27.13 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
87.232 | -0.06 | -12.01 | -0.06 | -36.15 | -0.18 | -39.44 | 0 | 0 | 0.12 | 21.33 | 0.41 | 21.39 | -0.76 | 21.39 | -12.07 | 0 | -47.81 | 0.12 | 0 | -0.18 | 28.65 | -0.29 | 28.54 | -0.18 | 28.59 | -0.41 | 0 | -25.43 | -0.06 | -36.45 | -0.18 | 0 | -0.06 | 0 | 0 | 0
87.264 | 0 | -10.2 | -0.06 | -34.98 | -0.06 | -37.03 | 0 | 0 | 0.18 | 21.39 | 0.35 | 21.39 | -0.41 | 21.39 | -10.31 | 0 | -62.11 | 0.23 | 0 | -0.23 | 28.65 | -0.18 | 28.48 | -0.23 | 28.42 | -0.18 | -0.06 | -23.97 | -0.06 | -84.91 | -0.06 | -8.44 | -0.06 | 0 | 0.06 | 0
87.296 | 5.68 | -8.91 | -0.06 | -32.81 | -0.18 | -34.86 | 0 | 0 | -0.06 | 21.27 | 0.41 | 21.39 | -0.18 | 21.45 | -9.02 | 0 | -41.49 | 0.18 | 0 | -0.18 | 28.59 | -0.18 | 28.42 | -0.18 | 28.42 | 0 | 0 | -22.09 | -0.18 | -55.31 | -0.06 | -90.12 | 0 | -24.84 | 0 | 0
87.328 | 11.66 | -7.62 | -0.06 | -30.35 | -0.06 | -32.76 | 0.06 | 0 | 0 | 21.27 | 0.35 | 21.39 | -0.18 | 21.39 | -6.86 | 0 | -27.77 | 0.06 | 0 | -0.29 | 0.29 | -0.18 | 28.54 | -0.23 | 28.3 | 0.12 | 5.1 | -20.22 | 0 | -48.52 | -0.06 | -62.11 | -0.18 | -89.18 | -0.06 | -35.68
87.36 | 16.99 | -5.86 | -0.06 | -29.24 | -0.06 | -30.94 | -0.06 | 0 | 0.12 | 21.27 | 0.41 | 21.39 | 0.06 | 21.27 | -5.39 | 0 | -30.29 | 0.12 | 0 | -0.18 | 0.06 | -0.18 | 0.64 | -0.06 | 28.36 | 0.12 | 13.18 | -17.75 | -0.18 | -43.54 | -0.06 | -35.33 | -0.06 | -58.36 | 0 | -85.73
87.392 | 21.62 | -3.87 | -0.06 | -25.67 | -0.06 | -28.24 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.39 | 0 | 21.33 | -3.34 | -0.06 | -35.98 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0.06 | -0.18 | 27.89 | 0.29 | 23.38 | -16.11 | 0 | -37.03 | -0.18 | -32.81 | 0 | -30.7 | -0.06 | -48.46
87.424 | 21.45 | -2.75 | -0.06 | -24.08 | -0.06 | -26.78 | 0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.39 | 0.12 | 21.27 | -2.4 | 0 | -33.34 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.35 | 28.59 | 0.29 | 28.95 | -13.89 | 0 | -37.09 | -0.12 | -42.25 | -0.06 | -33.11 | 0 | -30
87.456 | 21.33 | -1.7 | -0.06 | -21.97 | -0.06 | -24.43 | -0.18 | 0 | 0.12 | 21.39 | 0.23 | 21.27 | 0.23 | 21.09 | -1.41 | -0.06 | -30.29 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 28.54 | 0.29 | 28.71 | -12.07 | -0.18 | -37.5 | -0.18 | -42.01 | -0.06 | -41.49 | -0.06 | -34.04
87.488 | 21.45 | -0.94 | 0 | -20.16 | -0.06 | -22.56 | -0.18 | 0 | 0 | 0.12 | 0.23 | 21.39 | 0.18 | 21.39 | -1.11 | -0.06 | -27.83 | 0.18 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.42 | -10.14 | 0 | -37.21 | -0.18 | -37.44 | -0.06 | -41.37 | -0.06 | -42.36
87.52 | 21.45 | -0.41 | 0 | -17.93 | -0.06 | -19.75 | -0.06 | -67.09 | -0.06 | 0 | 0.18 | 21.45 | 0.23 | 21.39 | -0.23 | 0 | -27.42 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.3 | 0.41 | 28.48 | -7.32 | 0 | -36.04 | -0.18 | -33.75 | -0.06 | -37.56 | 0 | -40.9
87.552 | 21.45 | -0.18 | -0.06 | -15.7 | -0.06 | -17.7 | 0 | -52.74 | 0 | 0 | 0.06 | 21.39 | 0.29 | 21.27 | -0.18 | -0.06 | -25.61 | 0.12 | -0.12 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.71 | 0.41 | 28.71 | -5.39 | -0.06 | -32.99 | -0.06 | -32.46 | -0.06 | -33.28 | -0.18 | -35.86
87.584 | 21.39 | -0.06 | -0.06 | -14.41 | -0.06 | -16.23 | -0.06 | -43.3 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | 0 | -0.06 | -23.61 | 0.18 | -63.11 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.42 | -3.57 | -0.06 | -31 | -0.18 | -32.7 | 0 | -32.4 | -0.06 | -33.11
87.616 | 21.45 | -0.18 | -0.06 | -12.31 | -0.06 | -14.41 | -0.06 | -37.38 | -0.06 | 0 | 0.06 | 21.33 | 0.29 | 21.39 | -0.06 | -0.06 | -21.21 | 0.18 | -63.87 | 0 | 0 | 0 | 0 | -0.18 | 28.59 | 0.29 | 28.42 | -1.88 | 6.97 | -29.77 | -0.18 | -31.29 | 0 | -32.46 | 0 | -32.52
87.648 | 21.5 | -0.41 | -0.06 | -11.48 | -0.18 | -12.77 | -0.06 | -32.93 | 0 | 0 | -0.06 | 21.21 | 0.29 | 21.33 | 0 | -0.06 | -20.22 | 0.18 | -45.59 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.23 | 28.3 | -0.94 | 14.59 | -27.89 | -0.06 | -28.89 | 0 | -31.7 | -0.06 | -32.34
87.68 | 21.45 | -0.41 | 0.7 | -9.67 | -0.06 | -11.48 | -0.06 | -32.52 | -0.06 | 0.06 | 0.12 | 21.39 | 0.18 | 21.33 | 0.06 | -0.18 | -18.87 | 0.12 | -45.06 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.36 | -0.35 | 22.33 | -26.13 | -0.06 | -27.25 | -0.06 | -29.06 | -0.06 | -31
87.712 | 21.45 | -0.29 | 10.72 | -7.97 | 0 | -9.79 | -0.06 | -31.88 | 0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | 0.35 | 1.05 | -16.47 | 0.23 | -42.25 | 0 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0 | 28.36 | -0.23 | 28.71 | -23.79 | -0.06 | -25.72 | -0.06 | -26.72 | 0.06 | -28.01
87.744 | 21.5 | -0.18 | 15.94 | -6.68 | -0.06 | -8.32 | -0.06 | -31.29 | 0 | 0 | 0.06 | 21.39 | 0.12 | 21.21 | 0.41 | 9.38 | -14.77 | 0.06 | -40.9 | -0.06 | -38.79 | -0.06 | 0 | -0.06 | 28.71 | 0.06 | 28.65 | 0.23 | 28.71 | -22.21 | -0.18 | -24.61 | -0.06 | -25.67 | -0.06 | -26.78
87.776 | 21.45 | -0.23 | 21.74 | -5.22 | -0.06 | -6.91 | -0.06 | -29.59 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.21 | 0.35 | 14.47 | -13.54 | 0.06 | -38.97 | 0 | -73.48 | 0 | -9.32 | -0.06 | 28.42 | 0 | 28.54 | 0.18 | 28.77 | -20.45 | -0.18 | -22.62 | 0 | -24.14 | -0.06 | -25.37
87.808 | 21.45 | -0.18 | 21.5 | -3.87 | -0.06 | -5.33 | -0.18 | -27.36 | 0 | 0 | 0 | 21.45 | 0.06 | 21.33 | 0.29 | 20.16 | -11.6 | 0.06 | -37.68 | -0.06 | -49.46 | -0.06 | -76.94 | 0 | 9.02 | 0.06 | 28.48 | 0.41 | 28.59 | -18.16 | 0 | -21.33 | -0.18 | -22.56 | 0 | -23.85
87.84 | 21.45 | -0.23 | 21.45 | -2.46 | 5.86 | -4.04 | -0.06 | -25.08 | -0.06 | 0 | 0.12 | 21.39 | 0 | 21.21 | 0.35 | 21.45 | -10.14 | 0.06 | -35.57 | -0.18 | -45.88 | -0.06 | -56.31 | -0.06 | 0.12 | -0.06 | 28.59 | 0.35 | 28.42 | -16.41 | -0.06 | -19.63 | -0.06 | -20.63 | 0 | -22.38
87.872 | 21.45 | -0.18 | 21.45 | -1.17 | 11.66 | -2.52 | -0.06 | -22.97 | -0.06 | 0 | 0 | 0.06 | 0.12 | 21.39 | 0.35 | 21.39 | -7.62 | 0.06 | -33.34 | -0.06 | -43.6 | -0.06 | -46.12 | -0.06 | 0 | -0.06 | 28.65 | 0.41 | 28.42 | -13.77 | -0.18 | -17.7 | 0 | -18.93 | -0.06 | -19.81
87.904 | 21.45 | -0.18 | 21.45 | -0.53 | 17.17 | -1.41 | 5.86 | -20.74 | -0.06 | -59.06 | 0 | 0 | 0.06 | 21.33 | 0.35 | 21.39 | -5.8 | 0.06 | -31.29 | -0.18 | -40.9 | -0.06 | -44.77 | 0 | 0 | 0.12 | 28.36 | 0.47 | 28.59 | -12.07 | -0.18 | -16.29 | -0.06 | -17.52 | 0 | -18.93
87.936 | 21.39 | -0.23 | 21.45 | -0.18 | 21.74 | -0.94 | 11.72 | -19.28 | -0.06 | -64.22 | -0.06 | 0 | 0 | 11.48 | 0.29 | 21.33 | -4.45 | 0.06 | -29.83 | -0.06 | -39.96 | -0.06 | -42.78 | -0.18 | 0 | -0.06 | 28.42 | 0.41 | 28.36 | -9.96 | 7.73 | -14.59 | 0 | -15.94 | 0 | -17.4
87.968 | 21.45 | -0.23 | 21.39 | -0.06 | 21.5 | -0.35 | 17.29 | -16.93 | -0.06 | -44.88 | -0.06 | 0 | 0.06 | 0.06 | 0.29 | 21.21 | -2.99 | 0 | -27.36 | -0.06 | -38.15 | -0.06 | -41.6 | 0 | 0 | 0.06 | 28.42 | 0.41 | 28.54 | -7.79 | 15.06 | -12.77 | -0.06 | -14.36 | 3.34 | -15.94
88 | 21.39 | -0.29 | 21.45 | -0.23 | 21.39 | -0.18 | 21.8 | -15.35 | 0 | -39.9 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.27 | -2.4 | -0.06 | -25.84 | -0.18 | -36.21 | -0.06 | -38.5 | -0.06 | 0 | 0.06 | 28.42 | 0.47 | 28.13 | -6.15 | 22.97 | -11.13 | 1.93 | -12.54 | 11.54 | -14.06
88.032 | 21.45 | -0.18 | 21.45 | -0.29 | 21.45 | -0.06 | 21.45 | -12.77 | -0.06 | -33.34 | -0.06 | 0 | 0 | 0.06 | 0.35 | 21.45 | -1.11 | 0 | -23.32 | -0.06 | -34.28 | 0 | -36.45 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.42 | -3.46 | 28.83 | -8.73 | 13.83 | -10.84 | 21.86 | -12.13
88.064 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | -10.66 | -0.06 | -32.58 | -0.06 | 0 | -0.06 | 0 | 0.35 | 21.39 | -0.76 | -0.06 | -21.33 | -0.06 | -31.76 | 0 | -34.98 | -0.06 | 0 | 0.06 | 28.83 | 0.18 | 28.42 | -1.82 | 28.48 | -7.56 | 21.5 | -8.79 | 29.01 | -10.08
88.096 | 21.45 | -0.23 | 21.45 | -0.41 | 21.45 | -0.29 | 21.39 | -8.2 | 0.06 | -32.34 | -0.06 | 0 | 0.06 | 0 | 0.23 | 21.33 | -0.53 | -0.06 | -19.69 | -0.06 | -30.53 | 0 | -32.76 | -0.06 | 0 | -0.06 | 28.65 | 0.18 | 28.42 | -1.46 | 28.48 | -5.51 | 28.89 | -7.38 | 28.42 | -8.91
88.128 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.47 | 21.39 | -6.21 | -0.06 | -32.23 | -0.06 | 0 | 0 | 0 | 0.29 | 21.39 | -0.18 | 0 | -17.23 | 0 | -28.48 | -0.06 | -31.41 | 0 | 0 | -0.06 | 28.42 | 0.12 | 28.48 | -0.64 | 28.54 | -4.34 | 28.65 | -5.98 | 28.48 | -7.21
88.16 | 21.45 | -0.06 | 21.45 | -0.29 | 21.45 | -0.41 | 21.45 | -3.93 | 0 | -30.7 | -0.06 | 0 | 0 | 0 | 0.18 | 21.33 | -0.06 | -0.06 | -15.94 | -0.18 | -26.43 | -0.06 | -28.71 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | -0.29 | 28.59 | -2.75 | 28.54 | -4.28 | 28.65 | -5.57
88.192 | 0.12 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -1.76 | 0.06 | -27.77 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.21 | -0.06 | 0 | -13.89 | 0 | -23.97 | -0.06 | -26.43 | -0.06 | 0 | -0.18 | 28.54 | 0.12 | 28.42 | -0.18 | 28.42 | -1.7 | 28.48 | -2.64 | 28.3 | -3.63
88.224 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -1.46 | -0.06 | -25.61 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | -0.06 | -12.54 | 0 | -21.91 | -0.06 | -24.67 | -0.06 | -8.03 | 0 | 0.29 | 0.12 | 28.48 | 0.23 | 28.59 | -0.94 | 28.42 | -1.93 | 28.65 | -2.23
88.256 | 0 | -0.18 | 21.45 | -0.23 | 21.5 | -0.35 | 21.45 | -0.53 | 0.06 | -23.61 | 0.06 | 0 | 0 | 0 | 0.12 | 21.33 | 0.12 | 4.98 | -11.31 | -0.06 | -20.39 | 0 | -22.21 | 0 | -92.29 | 0 | 0.06 | 0.06 | 28.59 | 0.18 | 28.42 | -0.76 | 28.36 | -1 | 28.24 | -1.76
88.288 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.18 | 21.39 | -0.35 | -0.06 | -21.39 | -0.06 | -41.54 | -0.06 | 0 | 0.06 | 21.33 | 0 | 10.78 | -9.61 | -0.06 | -18.22 | -0.06 | -20.74 | 0 | -61.64 | -0.06 | 0 | 0.06 | 28.65 | 0.29 | 28.42 | -0.35 | 28.36 | -0.7 | 28.54 | -1.17
88.32 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -0.06 | 0 | -19.63 | -0.18 | -66.57 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 16.47 | -8.03 | -0.06 | -16.17 | 0.06 | -18.69 | 0 | -48.52 | -0.06 | 0 | 0.06 | 28.42 | 0.29 | 28.65 | -0.47 | 28.65 | -0.53 | 28.42 | -0.35
88.352 | 0 | -0.06 | 21.5 | -0.23 | 21.45 | -0.23 | 21.39 | 0.18 | -0.18 | -17.4 | -0.18 | -43.19 | 0 | -0.82 | 0 | 21.33 | 0.06 | 21.56 | -6.8 | -0.06 | -14.77 | -0.06 | -16.23 | 0 | -44.06 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.24 | -0.06 | 28.42 | -0.29 | 28.42 | -0.18
88.384 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.12 | -0.06 | -15.23 | -0.06 | -38.79 | -0.06 | -69.08 | 0 | 21.33 | 0.12 | 21.39 | -5.8 | -0.06 | -12.89 | 0 | -15 | 0 | -37.91 | -0.06 | 0 | 0 | 28.59 | 0.41 | 28.24 | -0.06 | 28.42 | -0.23 | 28.59 | -0.18
88.416 | 0 | -0.06 | 21.5 | -0.23 | 21.39 | -0.35 | 21.33 | 0.23 | -0.06 | -13.54 | -0.06 | -33.75 | -0.06 | -57.19 | -0.06 | 21.33 | 0.06 | 21.39 | -4.8 | 0 | -11.6 | -0.06 | -13.83 | 0.06 | -36.51 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.77 | 0.12 | 28.48 | 0 | 28.48 | 0
88.448 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | -0.06 | -11.37 | -0.18 | -32.4 | -0.06 | -34.16 | -0.06 | 21.39 | 0.18 | 21.15 | -4.16 | -0.06 | -10.31 | 1.52 | -12.19 | 0 | -37.56 | -0.06 | 0 | 0.06 | 28.36 | 0.35 | 28.42 | 0.06 | 28.65 | 0.06 | 28.48 | 0
88.48 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.33 | 0.29 | -0.06 | -9.32 | -0.06 | -32.81 | -0.06 | -28.13 | -0.06 | 21.39 | 0.18 | 21.45 | -3.4 | 0.06 | -8.91 | 11.48 | -10.72 | 0 | -37.79 | -0.06 | 0 | 0.06 | 28.48 | 0.41 | 28.42 | 0.18 | 28.65 | 0.06 | 28.48 | 0.18
88.512 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.29 | 21.39 | 0.47 | 0 | -6.62 | -0.06 | -32.23 | -0.06 | -34.98 | -0.06 | 21.33 | 0.06 | 21.39 | -2.4 | 8.32 | -7.44 | 21.62 | -8.91 | -0.06 | -36.15 | -0.06 | 0 | -0.06 | 0.18 | 0.18 | 28.54 | 0.06 | 28.3 | 0.12 | 28.36 | 0.12
88.544 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.12 | 21.39 | 0.35 | 6.91 | -4.22 | -0.06 | -30.53 | 0 | -34.63 | 0 | 21.21 | 0.12 | 21.27 | -1.88 | 15.65 | -6.39 | 29.18 | -7.44 | 0 | -34.1 | -0.18 | 0 | 0.06 | 0 | 0.18 | 28.54 | 0.12 | 28.01 | 0.12 | 28.48 | 0.18
88.576 | 0.06 | -0.06 | 0.12 | -0.18 | 21.39 | -0.18 | 21.39 | 0.41 | 12.36 | -2.29 | -0.06 | -28.01 | -0.18 | -31.93 | 0.06 | 21.39 | 0.18 | 21.33 | -1.52 | 23.26 | -5.04 | 28.71 | -6.68 | 0.06 | -32.23 | -0.18 | 0 | 0 | 0 | 0.23 | 28.65 | 0.12 | 28.59 | 0.06 | 28.42 | 0.06
88.608 | -30.47 | -0.06 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 17.7 | -1.76 | -0.06 | -26.13 | -0.06 | -28.83 | -0.06 | 3.4 | 0.23 | 21.27 | -1.29 | 29.12 | -4.16 | 28.77 | -5.1 | 6.86 | -30 | -0.06 | 0 | -0.06 | 0 | 0 | 28.3 | 0.06 | 28.48 | 0.06 | 28.36 | 0.18
88.64 | -60.71 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 21.62 | -0.64 | -0.06 | -24.32 | 0.12 | -27.6 | -0.06 | 0.06 | 0.18 | 21.39 | -1.05 | 28.48 | -3.57 | 28.24 | -4.45 | 14.59 | -28.59 | -0.18 | -28.07 | -0.06 | 0 | 0.06 | 28.48 | 0 | 28.54 | 0.12 | 28.54 | 0.06
88.672 | -40.37 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.39 | 0.35 | 21.45 | -0.23 | -0.06 | -21.68 | -0.06 | -26.54 | -0.06 | 0 | 0.06 | 21.33 | -0.7 | 28.71 | -2.52 | 28.54 | -3.52 | 24.84 | -25.84 | 0 | -78.81 | -0.06 | 0 | 0.06 | 28.48 | 0.06 | 28.54 | 0.12 | 28.48 | 0.23
88.704 | -39.44 | -0.06 | 0 | -0.06 | 21.45 | -0.06 | 21.39 | 0.23 | 21.39 | -0.23 | -0.06 | -19.92 | -0.06 | -25.37 | -0.06 | 0.06 | 0.18 | 21.39 | -0.41 | 28.59 | -2.05 | 28.89 | -2.7 | 28.54 | -24.38 | 0 | -51.33 | 0 | 0 | 0.06 | 28.36 | 0 | 28.48 | 0.06 | 28.48 | 0.18
88.736 | -37.74 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 0 | -17.75 | -0.18 | -23.44 | -0.06 | 0 | 0.12 | 21.33 | -0.23 | 28.42 | -1.64 | 28.65 | -1.99 | 28.65 | -22.74 | 0 | -47.93 | -0.06 | 0 | 0.06 | 28.48 | 0 | 28.42 | 0.18 | 28.48 | 0.23
88.768 | -35.51 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | -0.06 | -15.76 | -0.06 | -20.92 | -0.06 | 0 | 0.12 | 21.39 | -0.47 | 28.83 | -1.17 | 28.24 | -1.46 | 28.71 | -20.92 | 0 | -42.19 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.48 | 0.06 | 28.48 | 0.18
88.8 | -32.99 | 0 | 0 | -0.18 | 0 | -0.18 | 0.18 | 0.06 | 21.39 | 0.23 | -0.06 | -13.71 | -0.06 | -19.45 | -0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.89 | -0.94 | 28.54 | -1.05 | 28.65 | -18.75 | 0 | -37.15 | 0 | 0 | 0 | 28.83 | -0.06 | 28.42 | 0.06 | 28.54 | 0.12
88.832 | -30.94 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.45 | 0.41 | -0.06 | -11.13 | -0.06 | -17.46 | 0 | 0 | 0.06 | 21.39 | -0.29 | 28.71 | -0.7 | 28.48 | -0.64 | 28.36 | -16.29 | -0.06 | -37.5 | -0.12 | 0 | 0 | 0.18 | 0 | 28.48 | 0.12 | 28.42 | 0.12
88.864 | -29.12 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | -0.06 | -8.91 | 0 | -16.17 | 0 | 0 | 0.06 | 21.39 | -0.29 | 28.54 | -0.64 | 28.65 | -0.59 | 28.59 | -14.41 | 0 | -37.38 | 0 | 0 | -0.06 | 0 | 0 | 28.48 | 0.18 | 0.23 | 0.12
88.896 | -27.13 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.45 | 0.47 | -0.06 | -6.68 | 0.06 | -14.18 | -0.18 | 0 | 0.12 | 21.39 | -0.41 | 28.48 | -0.41 | 28.13 | -0.23 | 28.65 | -12.77 | -0.06 | -36.86 | 0 | 0.06 | -0.06 | 0 | 0 | 0.23 | 0.06 | 0.12 | 0.29
88.928 | -24.9 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.41 | 4.45 | -4.63 | 0.06 | -12.25 | 0 | 0 | 0.18 | 21.33 | -0.41 | 28.65 | -0.35 | 28.48 | -0.29 | 28.59 | -10.31 | -0.06 | -34.75 | 0 | -50.98 | -0.06 | 0 | -0.06 | 0.06 | 0.06 | 0 | 0.23
88.96 | -23.32 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 10.49 | -2.64 | 0.06 | -10.78 | -0.06 | 0 | 0.06 | 21.39 | -0.23 | 28.42 | -0.29 | 28.42 | -0.18 | 28.48 | -8.38 | -0.18 | -33.34 | -0.06 | -82.09 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0.06 | 0.12
88.992 | -20.86 | 0 | -22.5 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 17.7 | -1.7 | -0.06 | -8.73 | -0.06 | 0 | 0.18 | 21.39 | -0.23 | 28.36 | -0.35 | 28.54 | -0.23 | 28.48 | -5.92 | 2.4 | -31.29 | -0.06 | -48.87 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.18
89.024 | -18.69 | 0.06 | -62.81 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.62 | -0.64 | -0.06 | -7.15 | -0.06 | -5.45 | 0.23 | 21.33 | -0.23 | 28.65 | -0.23 | 28.65 | -0.29 | 28.42 | -3.69 | 10.31 | -28.89 | -0.06 | -45.7 | 0 | 0 | -0.06 | 0 | 0.18 | 0 | 0.23
89.056 | -17.29 | 0 | -46.47 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -0.23 | 0 | -5.33 | -0.06 | -69.55 | 0 | 21.33 | -0.29 | 28.42 | -0.23 | 28.42 | -0.29 | 28.42 | -2.23 | 17.81 | -27.36 | -0.06 | -41.08 | -0.06 | 0 | -0.06 | 0 | 0.18 | 0 | 0.06
89.088 | -15.59 | 0 | -40.02 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -0.23 | 0.18 | -3.52 | 0 | -55.02 | 0.06 | 21.39 | -0.29 | 28.48 | -0.18 | 28.3 | -0.23 | 28.59 | -1.46 | 25.55 | -25.67 | 0.06 | -36.45 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | -0.06
89.12 | -13.54 | 0 | -38.85 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.23 | 21.33 | 0.18 | 7.62 | -2.29 | -0.06 | -34.98 | 0.18 | 21.39 | -0.18 | 28.77 | -0.18 | 28.59 | -0.23 | 28.59 | -0.76 | 28.65 | -24.08 | -0.06 | -37.68 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
89.152 | -11.66 | -0.06 | -36.15 | -0.06 | -31.7 | -0.06 | 0 | -0.18 | 21.39 | 0.29 | 21.33 | 0.29 | 14.94 | -1.23 | 0 | -29.71 | 0.06 | 0.12 | -0.23 | 28.48 | -0.23 | 28.48 | -0.18 | 28.3 | -0.41 | 28.77 | -21.74 | 0 | -38.32 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06
89.184 | -10.31 | -0.06 | -34.16 | -0.06 | -59.42 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.21 | 0.29 | 20.74 | -0.59 | -0.18 | -34.92 | 0.12 | 0 | -0.18 | 28.59 | -0.23 | 28.65 | -0.18 | 28.48 | 0 | 28.48 | -19.63 | 0 | -36.92 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12
89.216 | -8.79 | -0.06 | -32.29 | -0.06 | -44.18 | -0.06 | -40.61 | -0.06 | 21.39 | 0.12 | 21.33 | 0.41 | 21.39 | -0.41 | -0.06 | -35.33 | 0.12 | 0 | -0.23 | 28.48 | -0.23 | 28.18 | -0.18 | 28.3 | 0.12 | 28.3 | -17.93 | -0.18 | -34.98 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
89.248 | -7.44 | 0 | -30.24 | -0.06 | -39.85 | -0.06 | -69.32 | -0.06 | 21.39 | 0 | 21.39 | 0.41 | 21.45 | -0.35 | -0.06 | -32.4 | 0.06 | 0 | -0.18 | 28.54 | -0.29 | 28.65 | -0.18 | 28.36 | 0.18 | 28.42 | -15.82 | 0 | -32.4 | 0 | -49.69 | -0.06 | 0 | -0.06 | 0 | 0.06
89.28 | -6.09 | -0.06 | -27.83 | -0.06 | -38.15 | 0 | -48.22 | -0.06 | 21.39 | 0.06 | 21.39 | 0.35 | 21.39 | 0 | -0.06 | -29.3 | 0.06 | 0 | -0.18 | 28.65 | -0.06 | 28.36 | -0.18 | 28.48 | 0.35 | 28.48 | -14.06 | -0.06 | -31.06 | -0.06 | -81.1 | -0.06 | 0 | 0 | -7.09 | 0
89.312 | -4.1 | -0.06 | -25.31 | -0.06 | -34.98 | -0.06 | -39.85 | 0 | 21.45 | -0.06 | 21.39 | 0.35 | 21.45 | 0 | 0 | -27.31 | 0.06 | 0 | -0.18 | 28.59 | -0.18 | 28.48 | -0.18 | 28.13 | 0.35 | 28.54 | -11.48 | 0 | -29.06 | -0.06 | -42.54 | -0.06 | -37.62 | 0 | -82.56 | 0.06
89.344 | -2.75 | -0.06 | -23.15 | -0.06 | -33.93 | -0.06 | -34.4 | -0.06 | 21.39 | -0.06 | 21.27 | 0.53 | 21.39 | -0.06 | 2.46 | -26.43 | 0.06 | 0 | -0.18 | 28.71 | -0.18 | 0.23 | -0.12 | 28.83 | 0.35 | 28.42 | -9.38 | 0 | -27.48 | 0 | -27.89 | 0 | -82.8 | -0.06 | -54.61 | 0.12
89.376 | -1.76 | 4.69 | -21.45 | -0.06 | -31.88 | -0.06 | -31.41 | -0.06 | 21.33 | -0.06 | 21.39 | 0.47 | 21.39 | 0.23 | 8.44 | -24.84 | 0.12 | 0 | -0.18 | 28.59 | -0.18 | 0.12 | -0.06 | 28.42 | 0.35 | 28.42 | -7.32 | -0.06 | -25.55 | 0 | -36.33 | 0 | -55.31 | -0.18 | -31.41 | 0
89.408 | -1.11 | 10.78 | -19.45 | -0.06 | -29.3 | -0.06 | -32.52 | -0.06 | 21.39 | -0.06 | 21.39 | 0.41 | 21.39 | 0.18 | 13.83 | -22.85 | 0.12 | 0 | -0.18 | 0.29 | -0.18 | 0 | -0.18 | 28.65 | 0.35 | 28.65 | -5.22 | -0.06 | -23.61 | -0.06 | -42.54 | -0.18 | -32.76 | -0.18 | -31.47 | 0
89.44 | -0.59 | 16.06 | -18.16 | 0 | -26.9 | -0.18 | -32.93 | -0.18 | 0.12 | -0.06 | 21.39 | 0.29 | 21.39 | 0.23 | 19.63 | -21.74 | 0.06 | 0 | -0.18 | 0.06 | -0.18 | 0.06 | -0.18 | 28.48 | 0.29 | 28.48 | -3.28 | 0.06 | -22.15 | -0.06 | -40.08 | -0.06 | -31.52 | 0 | -41.13 | 0
89.472 | -0.18 | 21.74 | -15.76 | -0.18 | -25.31 | -0.06 | -30.59 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.39 | 0.18 | 21.27 | -19.22 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 26.6 | 0.29 | 28.59 | -1.88 | 5.1 | -19.4 | -0.06 | -35.22 | -0.06 | -42.25 | 0 | -40.43 | -0.06
89.504 | -0.23 | 21.5 | -13.89 | -0.06 | -23.15 | 0 | -28.59 | 0 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.23 | 21.45 | -17.4 | 0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.23 | 0.35 | 28.07 | -1.05 | 13.24 | -17.7 | -0.06 | -33.17 | -0.06 | -41.08 | 0 | -36.27 | 0
89.536 | -0.18 | 21.5 | -12.42 | -0.06 | -21.15 | -0.06 | -26.31 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.33 | 0.23 | 21.45 | -16.06 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.41 | 28.42 | -0.53 | 20.92 | -15.47 | -0.18 | -32.23 | -0.06 | -37.5 | 0 | -33.69 | -0.06
89.568 | -0.29 | 21.45 | -10.72 | -0.06 | -18.87 | -0.06 | -24.14 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | 0.29 | 21.27 | -14.06 | -0.06 | -31.06 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.35 | 28.3 | -0.23 | 28.42 | -13.48 | 0.06 | -31.82 | -0.06 | -33.69 | -0.06 | -32.76 | -0.18
89.6 | -0.41 | 21.45 | -9.43 | -0.06 | -16.76 | -0.06 | -22.33 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.29 | 21.21 | -12.83 | 0.06 | -69.38 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | 0.23 | 28.42 | 0 | 28.65 | -11.78 | -0.06 | -30.41 | -0.18 | -32.7 | -0.06 | -32.58 | 0
89.632 | -0.35 | 21.5 | -7.62 | -0.06 | -15.35 | -0.06 | -19.81 | 0 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | 0.35 | 21.39 | -10.49 | 0.12 | -39.49 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.29 | 28.36 | 0.12 | 28.42 | -9.08 | 0 | -28.01 | -0.06 | -32.46 | 0 | -30.94 | -0.06
89.664 | -0.41 | 21.45 | -6.45 | -0.06 | -13.36 | 0.06 | -17.58 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.23 | 21.45 | -8.91 | -0.06 | -27.31 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.42 | 0.12 | 28.65 | -6.56 | -0.18 | -26.31 | -0.06 | -30.94 | 0 | -28.59 | -0.18
89.696 | -0.29 | 21.45 | -4.45 | -0.06 | -11.78 | -0.06 | -15.59 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.27 | 0.23 | 21.39 | -7.03 | 0 | -30.82 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 28.54 | 0.29 | 28.54 | -4.75 | 0 | -25.43 | -0.06 | -28.36 | 0 | -27.01 | 0
89.728 | -0.18 | 21.45 | -3.57 | -0.06 | -10.78 | -0.06 | -13.83 | 0 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | 0.12 | 21.39 | -5.16 | 0 | -35.39 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.59 | 0.29 | 28.18 | -2.99 | -0.06 | -23.79 | -0.06 | -27.25 | 0 | -26.08 | -0.06
89.76 | -0.18 | 21.45 | -2.29 | -0.06 | -8.61 | -0.06 | -11.89 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.33 | 0.23 | 21.15 | -3.69 | 0 | -34.75 | -0.06 | 0 | -0.06 | -6.74 | -0.18 | 0 | 0.18 | 28.59 | 0.35 | 28.65 | -1.99 | -0.18 | -22.21 | -0.18 | -26.25 | 0 | -24.26 | 0
89.792 | -0.18 | 21.45 | -1.11 | 8.79 | -7.62 | 0.29 | -8.73 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.45 | 0.06 | 21.33 | -2.4 | -0.06 | -30.53 | -0.06 | 0.06 | 0 | -74.42 | -0.06 | 0 | 0.06 | 28.24 | 0.41 | 28.36 | -0.7 | -0.18 | -19.86 | 0 | -24.43 | 0 | -22.21 | 0
89.824 | -0.29 | 21.45 | -0.59 | 14.24 | -6.09 | 7.91 | -6.86 | -0.06 | 0.06 | 0 | 0.12 | -0.06 | 21.39 | 0.06 | 21.33 | -1.46 | -0.06 | -27.72 | -0.06 | -41.54 | -0.06 | -51.51 | -0.06 | 0 | 0.06 | 28.42 | 0.35 | 28.42 | -0.7 | -0.06 | -18.63 | -0.18 | -22.85 | 0.06 | -20.74 | 0
89.856 | -0.23 | 21.45 | -0.23 | 19.92 | -4.63 | 13.48 | -4.8 | 0 | -45.35 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | -0.94 | 0 | -26.66 | -0.18 | -69.49 | 0 | -45.76 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.42 | -0.18 | -0.06 | -17.29 | -0.18 | -20.68 | 0.06 | -19.45 | 0
89.888 | -0.29 | 21.45 | -0.18 | 21.45 | -3.16 | 18.87 | -2.7 | -0.06 | -64.69 | 0 | 0 | 0.12 | 21.39 | 0.12 | 21.27 | -0.53 | -0.06 | -26.31 | 0 | -49.16 | -0.06 | -45.35 | -0.06 | 0 | 0.06 | 22.97 | 0.35 | 28.54 | 0 | -0.06 | -15.29 | 0 | -19.22 | 0 | -17.93 | 0
89.92 | -0.35 | 21.45 | -0.23 | 21.45 | -1.88 | 21.56 | -1.99 | -0.06 | -46.82 | 0.12 | 0.06 | -0.06 | 21.33 | 0 | 21.45 | -0.41 | 4.45 | -24.08 | -0.06 | -46.35 | 0 | -42.07 | -0.06 | -39.08 | 0 | 0.18 | 0.23 | 28.59 | 0.18 | 2.64 | -13.83 | 0 | -18.28 | -0.06 | -16.29 | 0
89.952 | -0.18 | 21.45 | -0.41 | 21.5 | -1.11 | 21.39 | -0.41 | 0 | -38.85 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.21 | -0.18 | 12.13 | -21.62 | -0.06 | -43.36 | -0.06 | -40.14 | -0.18 | -63.75 | 0 | 0 | 0.23 | 28.3 | 0.23 | 13.18 | -11.54 | 1.7 | -15.82 | -0.06 | -14.18 | -0.06
89.984 | -0.18 | 21.5 | -0.35 | 21.5 | -0.64 | 21.39 | -0.41 | 0 | -33.58 | -0.06 | 0 | -0.06 | 0.59 | -0.06 | 21.33 | 0 | 17.7 | -20.45 | -0.06 | -41.25 | -0.06 | -38.09 | 0 | -46.58 | -0.06 | 0.06 | 0.29 | 28.48 | 0.35 | 20.57 | -9.84 | 11.19 | -14 | 3.4 | -12.72 | 0
90.016 | -0.23 | 21.39 | -0.41 | 21.45 | -0.35 | 21.45 | 0.06 | 0.12 | -32.05 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.56 | -0.06 | 21.56 | -18.75 | -0.06 | -39.79 | 0 | -36.74 | -0.06 | -46.35 | 0 | 0 | 0.18 | 28.3 | 0.35 | 28.77 | -8.56 | 18.69 | -12.31 | 11.78 | -11.02 | 5.16
90.048 | -0.06 | 21.45 | -0.41 | 21.39 | -0.23 | 21.39 | 0.06 | 0 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0 | 21.45 | -17.58 | -0.06 | -37.38 | -0.06 | -35.45 | 0 | -43.89 | 0.06 | 0 | 0 | 28.54 | 0.35 | 28.48 | -6.68 | 26.37 | -10.61 | 19.28 | -9.38 | 13.24
90.08 | -0.06 | 21.45 | -0.29 | 21.5 | -0.18 | 21.39 | 0.29 | 0 | -32.4 | 0 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.45 | -15.7 | -0.06 | -35.92 | 0 | -31.99 | -0.06 | -41.49 | -0.18 | 0 | 0.06 | 28.48 | 0.23 | 28.59 | -5.22 | 28.54 | -9.2 | 26.9 | -7.56 | 20.8
90.112 | -0.18 | 21.45 | -0.35 | 21.45 | -0.47 | 21.39 | 0.23 | -0.06 | -30.59 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.21 | 0.06 | 21.39 | -13.3 | -0.06 | -34.16 | 0 | -30.82 | -0.06 | -39.61 | 0 | 0 | 0.06 | 28.59 | 0.23 | 28.95 | -3.16 | 28.48 | -7.44 | 28.65 | -5.57 | 28.83
90.144 | -0.06 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | 0.47 | 0 | -27.77 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.12 | 21.33 | -11.6 | -0.06 | -32.58 | -0.06 | -28.59 | -0.06 | -38.15 | 0 | 0 | 0 | 28.36 | 0.29 | 28.48 | -2.23 | 28.65 | -5.39 | 28.48 | -4.04 | 28.77
90.176 | -0.06 | 21.5 | -0.23 | 21.45 | -0.23 | 21.33 | 0.41 | -0.06 | -25.9 | -0.06 | 0 | -0.06 | 0 | 0 | 21.45 | 0.18 | 21.39 | -10.2 | -0.06 | -29.65 | 0.06 | -26.49 | -0.18 | -35.57 | 0 | 0 | 0.06 | 28.71 | 0.29 | 28.54 | -1.46 | 28.42 | -3.69 | 28.65 | -2.64 | 28.42
90.208 | -0.06 | 21.39 | -0.23 | 21.45 | -0.35 | 21.39 | 0.41 | 7.09 | -23.85 | -0.06 | 0 | -0.06 | 0 | 0 | 21.15 | 0.06 | 21.39 | -8.32 | 0 | -27.72 | 0 | -24.55 | 0 | -34.69 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.42 | -0.7 | 28.54 | -2.52 | 28.54 | -2.05 | 28.42
90.24 | -0.06 | 21.39 | -0.23 | 21.45 | -0.41 | 21.33 | 0.53 | 12.83 | -22.38 | 0.06 | -29.94 | -0.18 | 0 | -0.06 | 13.65 | 0.12 | 21.27 | -6.15 | -0.06 | -26.31 | -0.06 | -22.44 | 0 | -31.99 | -0.06 | 0 | -0.06 | 28.65 | 0.12 | 28.71 | -0.64 | 28.42 | -1.7 | 28.42 | -0.94 | 28.36
90.272 | -0.06 | 0.12 | -0.18 | 21.45 | -0.35 | 21.39 | 0.35 | 20.27 | -19.63 | -0.18 | -66.21 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.33 | -4.45 | 9.67 | -23.09 | 0 | -19.98 | 0 | -29.77 | -0.06 | 0 | 0 | 28.24 | 0.18 | 28.42 | -0.18 | 28.48 | -0.94 | 28.48 | -0.59 | 28.48
90.304 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.39 | 0.41 | 21.45 | -17.52 | -0.18 | -46.12 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | -2.93 | 17.11 | -21.74 | -0.06 | -18.16 | -0.06 | -27.6 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.71 | -0.06 | 28.83 | -0.53 | 28.36 | -0.53 | 28.36
90.336 | -0.18 | 0 | -0.35 | 21.45 | -0.23 | 21.39 | 0.29 | 21.45 | -15.82 | 0 | -40.43 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.33 | -1.99 | 24.61 | -19.81 | 0 | -16.76 | -0.06 | -25.78 | -0.06 | -85.32 | -0.06 | 28.59 | 0.06 | 28.42 | -0.06 | 28.48 | -0.41 | 28.71 | -0.18 | 28.24
90.368 | -0.18 | 0.06 | -0.29 | 21.45 | -0.23 | 21.33 | 0.29 | 21.39 | -13.71 | -0.18 | -34.86 | -0.06 | 0 | -0.18 | 0 | 0.23 | 21.39 | -1.29 | 28.54 | -18.16 | 0.06 | -15 | -0.06 | -24.08 | -0.06 | -70.2 | 0 | 0.29 | 0.06 | 28.42 | 0.12 | 28.48 | -0.35 | 28.48 | -0.06 | 28.42
90.4 | -0.06 | 0 | -0.29 | 21.45 | -0.18 | 21.39 | 0.29 | 21.39 | -11.95 | -0.06 | -31.99 | 0 | -0.06 | -0.06 | 0 | 0.18 | 21.33 | -1.05 | 28.65 | -16.06 | -0.06 | -13.65 | -0.06 | -21.91 | 0 | -48.52 | 0 | 0.06 | 0.06 | 28.3 | 0.12 | 28.42 | 0 | 28.42 | -0.18 | 28.65
90.432 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | 0.29 | 21.45 | -8.85 | -0.18 | -32.7 | -0.06 | -73.66 | -0.06 | 0 | 0.06 | 21.39 | -0.18 | 28.71 | -14.12 | 6.09 | -11.6 | -0.06 | -19.04 | -0.18 | -44.77 | -0.06 | 0 | 0 | 29.01 | 0.18 | 28.48 | -0.06 | 28.71 | -0.06 | 28.48
90.464 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | 0.29 | 21.39 | -6.62 | -0.06 | -32.23 | -0.06 | -53.97 | 0.06 | 0 | 0.23 | 21.33 | -0.18 | 28.77 | -12.77 | 14.06 | -10.37 | 2.99 | -17.58 | -0.06 | -38.38 | 0 | 0.06 | -0.06 | 28.48 | 0.12 | 28.3 | 0 | 28.42 | 0 | 28.54
90.496 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -4.63 | -0.06 | -30.94 | -0.06 | -31.58 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 28.54 | -11.31 | 21.62 | -8.91 | 11.31 | -15.7 | -0.06 | -35.92 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.3 | 0.06 | 28.48 | 0 | 28.3
90.528 | -0.18 | 0 | -0.18 | 21.5 | -0.23 | 21.33 | 0.12 | 21.39 | -2.46 | -0.18 | -28.48 | -0.06 | -28.36 | -0.18 | 0 | 0.12 | 21.15 | -0.06 | 28.36 | -10.14 | 29.06 | -7.68 | 18.87 | -14.65 | 0 | -37.33 | 0 | 0 | -0.06 | 28.36 | 0.12 | 28.59 | 0.06 | 28.48 | 0.18 | 28.42
90.56 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.06 | 21.39 | -1.7 | -0.06 | -26.78 | -0.06 | -34.92 | -0.18 | 0 | 0.18 | 21.33 | 0.06 | 28.59 | -8.85 | 28.59 | -6.39 | 26.72 | -12.77 | 0 | -38.32 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.18 | 0.18 | 28.48 | 0.18 | 28.54
90.592 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.33 | -0.06 | 21.45 | -0.82 | 6.27 | -23.97 | -0.06 | -34.86 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 28.83 | -7.21 | 28.65 | -4.92 | 28.77 | -11.48 | 0.06 | -36.1 | 0.06 | 0 | 0 | 28.36 | 0.12 | 28.48 | 0.12 | 28.59 | 0.06 | 28.65
90.624 | -0.06 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | 0.12 | 21.39 | -0.23 | 11.84 | -21.86 | -0.18 | -31.93 | -0.18 | 0 | 0.18 | 21.33 | 0.18 | 28.48 | -6.09 | 28.83 | -4.1 | 28.48 | -9.9 | -0.18 | -34.22 | 0.06 | 0 | 0 | 28.59 | 0 | 28.3 | 0.18 | 28.24 | 0.06 | 28.48
90.656 | -0.06 | 0 | -0.06 | 10.55 | -0.06 | 21.56 | 0.06 | 21.45 | -0.06 | 17.52 | -19.86 | 0 | -29.01 | 0 | -1.05 | 0.18 | 21.27 | 0.18 | 28.3 | -5.04 | 28.83 | -3.16 | 28.65 | -8.56 | 0 | -32.29 | -0.06 | 0 | -0.06 | 28.65 | 0.06 | 28.54 | 0.23 | 28.42 | 0 | 28.42
90.688 | -0.06 | -24.67 | -0.06 | 0.12 | -0.18 | 17.52 | 0.18 | 21.39 | -0.06 | 21.68 | -18.05 | -0.06 | -27.95 | -0.18 | -64.87 | 0.18 | 21.21 | 0.18 | 28.36 | -3.98 | 28.36 | -2.87 | 28.65 | -7.38 | -0.06 | -30 | -0.06 | 0 | 0 | 28.54 | 0 | 28.48 | 0.06 | 28.42 | 0 | 28.42
90.72 | -0.06 | -63.69 | -0.18 | 0 | -0.18 | 0.12 | -0.06 | 21.39 | 0.18 | 21.39 | -16.17 | 0.06 | -26.78 | -0.06 | -59.47 | 0.18 | 21.5 | 0.18 | 28.65 | -3.28 | 28.65 | -2.05 | 28.54 | -6.09 | -0.18 | -28.77 | 0 | 0 | -0.06 | 28.42 | 0.06 | 28.48 | 0.18 | 28.48 | 0.06 | 28.42
90.752 | -0.06 | -43.36 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | -13.71 | 0 | -24.9 | -0.23 | -31.99 | 0.12 | 21.33 | 0.23 | 28.54 | -2.64 | 28.42 | -1.52 | 28.42 | -4.45 | -0.06 | -26.31 | -0.06 | 0 | -0.06 | 28.48 | -0.06 | 28.48 | 0.06 | 28.42 | 0.06 | 28.36
90.784 | -0.06 | -40.61 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | -11.66 | -0.06 | -22.62 | -0.18 | -28.54 | 0.12 | 21.33 | 0.12 | 28.83 | -2.17 | 28.77 | -1.11 | 28.3 | -3.69 | -0.06 | -24.55 | -0.06 | -31.82 | -0.06 | 28.24 | 0 | 28.54 | 0.23 | 28.48 | 0.12 | 28.48
90.816 | -0.06 | -39.08 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.47 | 21.39 | -9.55 | 0.06 | -20.92 | -0.18 | -33.99 | 0.06 | 0.12 | 0.12 | 28.36 | -1.64 | 28.48 | -0.76 | 28.48 | -2.87 | -0.06 | -22.62 | 0 | -92.87 | 0 | 0.29 | 0 | 28.42 | 0.18 | 28.48 | 0.06 | 28.48
90.848 | 0 | -35.74 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.53 | 21.39 | -7.56 | -0.06 | -19.75 | -0.18 | -35.33 | 0.06 | 0 | 0.18 | 28.48 | -1.29 | 28.42 | -0.64 | 28.54 | -2.46 | -0.06 | -21.27 | -0.18 | -56.25 | 0 | 0.12 | -0.06 | 14.06 | 0.06 | 28.48 | 0.06 | 28.48
90.88 | -0.06 | -34.28 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -5.51 | 0 | -17.7 | -0.06 | -33.17 | 0.06 | 0 | 0.12 | 28.42 | -0.76 | 28.18 | -0.47 | 28.65 | -1.7 | -0.06 | -19.22 | 0 | -47.76 | -0.06 | 0 | -0.06 | 0.18 | 0.06 | 3.81 | 0.06 | 28.48
90.912 | -0.06 | -31.23 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -2.75 | 0 | -15.53 | -0.18 | -28.71 | 0.06 | 0 | 0.18 | 28.65 | -0.7 | 28.54 | -0.18 | 28.65 | -1.11 | 1.41 | -16.52 | -0.06 | -43.01 | 0 | 0.06 | -0.06 | 0 | 0.18 | 0.12 | 0.06 | 0.23
90.944 | -0.06 | -29.53 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.39 | -1.76 | 0 | -13.54 | -0.06 | -27.36 | 0.06 | 0 | 0.18 | 28.65 | -0.53 | 28.36 | -0.18 | 28.36 | -0.7 | 11.25 | -14.82 | -0.18 | -36.92 | -0.18 | 0 | 0 | 0.06 | -0.06 | 0 | 0 | 0
90.976 | 0 | -27.36 | -0.06 | 0 | -0.12 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | -0.94 | 5.45 | -12.13 | -0.06 | -26.78 | 0.06 | 0 | 0.18 | 28.59 | -0.59 | 28.65 | -0.29 | 28.59 | -0.53 | 18.93 | -12.72 | 0 | -36.74 | 0 | 0 | -0.06 | 0 | 0.12 | 0.06 | -0.06 | 0
91.008 | -0.06 | -25.08 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.35 | 21.21 | -0.47 | 11.37 | -10.31 | -0.06 | -25.2 | 0.12 | 0 | 0.18 | 28.59 | -0.53 | 28.59 | -0.35 | 28.48 | -0.29 | 26.43 | -10.96 | 0 | -37.44 | -0.18 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06 | 0.06
91.04 | 0 | -23.5 | -0.18 | 0 | -0.06 | 0 | 0 | 21.27 | 0.41 | 21.33 | -0.18 | 16.76 | -8.73 | -0.06 | -23.15 | 0 | 0 | 0.12 | 28.65 | -0.41 | 28.48 | -0.23 | 28.42 | -0.29 | 28.42 | -8.79 | -0.06 | -36.97 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
91.072 | -0.06 | -21.15 | -0.06 | -2.11 | -0.06 | 0 | -0.06 | 4.34 | 0.18 | 21.39 | 0.12 | 21.62 | -6.21 | 8.61 | -20.98 | 0.12 | 0 | 0.06 | 28.48 | -0.47 | 28.71 | -0.23 | 28.59 | -0.23 | 28.71 | -5.92 | 0 | -34.98 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
91.104 | -0.06 | -19.1 | -0.06 | -55.2 | -0.18 | -0.18 | -0.18 | 0.06 | 0.12 | 21.33 | 0.12 | 21.45 | -4.75 | 13.89 | -19.45 | 0 | 0 | 0.12 | 28.54 | -0.35 | 28.3 | -0.23 | 28.48 | -0.29 | 28.54 | -4.1 | -0.18 | -33.28 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
91.136 | 0 | -16.99 | -0.06 | -50.74 | -0.18 | -47.52 | -0.06 | 0 | 0.23 | 21.33 | 0.23 | 21.45 | -3.11 | 19.57 | -17.99 | 0 | 0 | 0.23 | 27.77 | -0.35 | 28.59 | -0.35 | 28.83 | -0.23 | 28.71 | -2.29 | -0.06 | -31.76 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | -0.06 | 0
91.168 | -0.06 | -15.65 | -0.18 | -40.9 | -0.06 | -52.97 | 0.06 | 0 | 0.23 | 21.33 | 0.53 | 21.39 | -2.23 | 21.39 | -16.29 | -0.06 | 0 | 0.12 | 0.18 | -0.23 | 28.3 | -0.18 | 28.48 | -0.23 | 28.89 | -1.7 | -0.06 | -30 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
91.2 | 0 | -13.71 | -0.06 | -39.61 | -0.06 | -40.78 | -0.06 | 0 | 0.18 | 21.39 | 0.47 | 21.39 | -1.41 | 21.33 | -14.77 | 0 | 0 | 0.12 | 0 | -0.35 | 28.65 | -0.18 | 28.71 | -0.35 | 28.42 | -0.94 | -0.06 | -27.19 | 0 | 0 | -0.06 | 0 | 0 | 0 | -0.06 | 0
91.232 | 0 | -11.84 | -0.06 | -37.21 | -0.06 | -39.44 | 0 | 0 | 0.12 | 21.39 | 0.47 | 21.39 | -0.7 | 21.39 | -12.25 | 0.06 | -42.36 | 0.12 | 0 | -0.23 | 28.48 | -0.18 | 28.48 | -0.29 | 28.65 | -0.47 | -0.06 | -26.08 | -0.06 | -26.19 | -0.18 | 0 | -0.06 | 0 | 0 | 0
91.264 | -0.06 | -10.02 | -0.06 | -34.75 | -0.06 | -36.74 | -0.18 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -0.41 | 21.39 | -11.07 | 0 | -63.93 | 0.12 | 0 | -0.18 | 28.65 | -0.18 | 28.36 | -0.29 | 28.42 | -0.18 | 0.06 | -24.08 | -0.18 | -90.3 | -0.23 | -1.05 | -0.18 | 0 | 0 | 0
91.296 | 5.45 | -9.02 | -0.06 | -32.93 | -0.18 | -35.33 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.33 | -0.23 | 21.5 | -9.08 | 0 | -42.6 | 0.12 | 0 | -0.29 | 28.59 | -0.18 | 28.65 | -0.23 | 28.65 | 0.06 | -0.06 | -22.21 | -0.06 | -57.01 | -0.18 | -88.13 | -0.18 | -9.14 | 0 | 0
91.328 | 11.43 | -7.44 | 0 | -30.41 | -0.06 | -33.22 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | -0.18 | 21.39 | -7.21 | -0.06 | -27.54 | 0.18 | 0 | -0.23 | 0.23 | -0.18 | 28.65 | -0.06 | 28.42 | 0.23 | 4.22 | -19.98 | -0.06 | -47.99 | -0.18 | -67.21 | 0 | -88.3 | -0.18 | -30.76
91.36 | 16.88 | -6.33 | 0 | -28.83 | -0.06 | -30.76 | -0.06 | 0 | 0.12 | 21.39 | 0.29 | 21.33 | 0.06 | 21.33 | -5.86 | 0.06 | -30.59 | 0.12 | 0 | -0.18 | 0.06 | -0.18 | 13.77 | -0.18 | 28.42 | 0.18 | 12.48 | -17.99 | -0.06 | -44.36 | -0.06 | -38.38 | -0.18 | -63.75 | -0.18 | -88.3
91.392 | 21.68 | -4.28 | -0.18 | -26.02 | -0.06 | -28.36 | -0.06 | 0 | -0.18 | 21.33 | 0.35 | 21.39 | 0.06 | 21.39 | -3.52 | 0 | -35.86 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.12 | -0.18 | 28.36 | 0.35 | 22.62 | -15.94 | 0 | -36.21 | -0.18 | -31.93 | -0.06 | -32.11 | -0.18 | -50.1
91.424 | 21.56 | -2.99 | -0.06 | -24.08 | -0.06 | -26.43 | -0.06 | 0 | -0.18 | 21.39 | 0.35 | 21.33 | 0.06 | 21.15 | -2.4 | -0.06 | -33.63 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.41 | 29.01 | -14.12 | -0.06 | -37.21 | -0.06 | -40.84 | -0.06 | -32.17 | 0 | -30.35
91.456 | 21.45 | -1.88 | -0.06 | -22.33 | -0.06 | -24.84 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | 0.12 | 21.33 | -1.7 | -0.06 | -31 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.06 | -0.18 | 28.36 | 0.41 | 28.65 | -12.19 | 0 | -37.91 | -0.06 | -42.19 | -0.18 | -41.08 | -0.18 | -33.69
91.488 | 21.45 | -1.29 | -0.06 | -20.1 | -0.06 | -22.09 | -0.18 | -0.06 | -0.06 | 0.12 | 0.18 | 21.33 | 0.23 | 21.39 | -0.94 | -0.06 | -28.07 | 0.12 | 0 | -0.18 | 0 | 0 | 0 | -0.06 | 28.36 | 0.29 | 28.48 | -10.49 | 0 | -37.91 | -0.18 | -38.15 | -0.18 | -41.84 | 0 | -41.25
91.52 | 21.45 | -0.59 | -0.06 | -18.34 | 0 | -19.98 | 0 | -56.37 | -0.06 | 0 | 0.18 | 21.21 | 0.35 | 21.45 | -0.76 | -0.06 | -27.31 | 0.18 | 0 | 0 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.29 | 28.54 | -7.27 | -0.06 | -36.04 | 0 | -34.45 | -0.06 | -38.2 | -0.18 | -40.72
91.552 | 21.39 | -0.18 | -0.06 | -16 | -0.06 | -17.7 | -0.06 | -55.55 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.09 | -0.18 | 0.06 | -25.55 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.23 | 28.59 | -5.22 | 0 | -33.34 | -0.18 | -32.99 | -0.06 | -33.11 | 0 | -36.21
91.584 | 21.45 | -0.18 | -0.06 | -14.59 | -0.06 | -15.76 | -0.06 | -43.3 | -0.06 | 0 | 0.12 | 21.33 | 0.18 | 21.39 | -0.06 | -0.06 | -23.38 | 0.12 | -58.65 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.35 | 28.65 | -3.22 | -0.06 | -31.17 | -0.18 | -32.93 | 0 | -32.58 | -0.06 | -33.81
91.616 | 21.39 | -0.35 | -0.06 | -12.77 | -0.18 | -14.71 | -0.06 | -38.44 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.39 | -0.18 | -0.06 | -21.74 | 0.18 | -65.45 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.29 | 28.65 | -2.05 | 6.09 | -30 | -0.06 | -31.47 | -0.18 | -32.4 | -0.18 | -32.46
91.648 | 21.45 | -0.29 | -0.06 | -11.25 | -0.06 | -13.01 | -0.06 | -33.69 | 0 | 0 | 0 | 21.33 | 0.12 | 21.33 | 0 | -0.06 | -20.39 | 0.18 | -45.7 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.23 | 28.65 | -1.29 | 14 | -28.01 | -0.06 | -29.06 | -0.06 | -32.17 | 0 | -32.7
91.68 | 21.45 | -0.18 | 0.59 | -10.02 | 0 | -11.66 | -0.06 | -31.99 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.33 | 0.12 | -0.06 | -19.69 | 0.12 | -45.12 | -0.06 | 0 | 0 | 0 | 0 | 28.83 | 0.23 | 28.59 | -0.7 | 21.8 | -25.9 | -0.06 | -27.25 | -0.06 | -29.71 | -0.18 | -31.17
91.712 | 21.45 | -0.29 | 10.25 | -8.14 | -0.06 | -9.84 | 0 | -32.52 | 0 | 0 | 0.06 | 21.33 | 0.29 | 21.39 | 0.23 | 0.35 | -16.93 | 0.12 | -43.24 | -0.06 | 0 | 0 | 0 | 0 | 28.83 | 0.18 | 28.59 | -0.18 | 28.65 | -23.79 | -0.06 | -26.13 | 0 | -27.42 | -0.18 | -27.83
91.744 | 21.39 | -0.29 | 15.53 | -6.56 | -0.06 | -8.14 | 0 | -31.99 | -0.06 | 0 | 0 | 21.33 | 0.18 | 21.33 | 0.23 | 7.85 | -15 | 0.06 | -40.72 | -0.06 | -35.45 | -0.18 | 0 | -0.06 | 28.48 | 0.12 | 28.71 | -0.18 | 28.65 | -22.62 | -0.18 | -24.61 | 0 | -26.08 | -0.06 | -26.84
91.776 | 21.45 | -0.29 | 21.45 | -5.1 | -0.06 | -6.86 | -0.06 | -30.94 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.15 | 0.18 | 13.42 | -13.24 | 0.12 | -39.08 | -0.18 | -75.53 | -0.06 | -2.52 | -0.06 | 28.65 | 0.06 | 28.42 | 0.23 | 28.48 | -20.51 | 0 | -22.91 | 0 | -24.61 | -0.18 | -26.25
91.808 | 21.45 | -0.18 | 21.5 | -3.98 | -0.06 | -5.39 | -0.18 | -28.07 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | 0.12 | 18.99 | -11.72 | 0.06 | -36.97 | -0.06 | -50.33 | -0.06 | -71.9 | -0.06 | 13.54 | 0.06 | 28.54 | 0.06 | 28.48 | -18.87 | -0.18 | -21.33 | -0.06 | -22.62 | -0.06 | -24.32
91.84 | 21.45 | -0.18 | 21.45 | -2.64 | 5.33 | -3.93 | -0.06 | -25.55 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.23 | 21.39 | -10.14 | 0.06 | -35.16 | -0.06 | -45.88 | 0 | -58.71 | 0 | 0.12 | 0.06 | 28.48 | 0.29 | 28.54 | -16.52 | -0.12 | -20.39 | -0.06 | -21.33 | 0.06 | -22.44
91.872 | 21.45 | -0.29 | 21.45 | -1.35 | 11.25 | -2.23 | -0.06 | -23.2 | 0.06 | 0 | -0.06 | 0.12 | 0.06 | 21.33 | 0.29 | 21.5 | -7.62 | 0.06 | -33.69 | -0.06 | -43.89 | -0.06 | -46.06 | -0.06 | 0 | 0 | 28.36 | 0.18 | 28.42 | -13.83 | -0.06 | -17.75 | 0 | -19.51 | 0 | -20.39
91.904 | 21.45 | -0.23 | 21.45 | -0.76 | 16.7 | -1.29 | 4.92 | -20.68 | -0.18 | -45.24 | -0.06 | 0 | 0.18 | 21.39 | 0.29 | 21.33 | -5.98 | 0.12 | -31.29 | -0.18 | -41.02 | 0 | -45.35 | -0.18 | 0.06 | 0.06 | 28.48 | 0.35 | 28.59 | -12.01 | -0.12 | -16.29 | 0.06 | -17.99 | -0.06 | -18.99
91.936 | 21.45 | -0.18 | 21.45 | -0.35 | 21.86 | -0.76 | 10.96 | -19.28 | -0.06 | -67.62 | -0.06 | 0 | -0.06 | 13.42 | 0.23 | 21.39 | -4.45 | 0.12 | -29.24 | -0.06 | -40.14 | -0.06 | -42.19 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.18 | -10.25 | 6.33 | -14.53 | -0.06 | -16.11 | -0.06 | -17.64
91.968 | 21.45 | -0.18 | 21.5 | -0.35 | 21.45 | -0.23 | 16.35 | -17.4 | 0 | -47.23 | -0.18 | 0 | 0 | 0.06 | 0.18 | 21.27 | -2.99 | 0.12 | -27.77 | -0.06 | -38.32 | -0.06 | -41.02 | -0.06 | 0 | -0.06 | 28.42 | 0.29 | 28.42 | -8.32 | 14.18 | -13.3 | -0.06 | -14.59 | 1.11 | -15.82
92 | 21.45 | -0.23 | 21.62 | -0.06 | 21.45 | 0 | 21.74 | -15.53 | -0.06 | -41.6 | 0 | 0 | 0.06 | 0 | 0.23 | 21.33 | -2.23 | 0.06 | -25.9 | -0.06 | -35.98 | -0.06 | -38.91 | -0.06 | 0 | 0.06 | 28.59 | 0.29 | 28.42 | -6.09 | 21.86 | -11.25 | 1.58 | -12.72 | 11.13 | -14.24
92.032 | 21.5 | -0.29 | 21.45 | -0.18 | 21.39 | -0.18 | 21.45 | -12.6 | -0.06 | -33.99 | 0 | 0 | 0 | 0.06 | 0.18 | 21.39 | -1 | 0.12 | -22.85 | 0 | -35.16 | 0 | -36.97 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.71 | -3.4 | 28.65 | -9.43 | 13.65 | -10.78 | 21.15 | -12.25
92.064 | 21.39 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -10.72 | -0.06 | -32.7 | 0 | 0 | 0 | 0 | 0.23 | 21.27 | -0.59 | 0.12 | -21.27 | 0 | -31.99 | 0 | -35.16 | 0 | 0 | 0 | 28.65 | 0.23 | 28.48 | -2.29 | 28.71 | -7.79 | 21.39 | -9.2 | 28.65 | -10.66
92.096 | 21.45 | -0.18 | 21.33 | -0.18 | 21.45 | -0.35 | 21.39 | -8.56 | -0.18 | -32.64 | 0 | 0 | -0.06 | 0 | 0.12 | 21.39 | -0.59 | -0.06 | -19.86 | -0.06 | -30.47 | 0 | -33.52 | 0 | 0 | -0.06 | 28.65 | 0.23 | 28.42 | -1.11 | 28.48 | -5.86 | 28.95 | -7.56 | 28.59 | -9.32
92.128 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.39 | -6.15 | 0.06 | -32.05 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.23 | 0.06 | -17.23 | 0 | -28.48 | 0 | -31.06 | 0 | 0 | -0.06 | 28.65 | 0.12 | 28.59 | -0.82 | 28.65 | -4.22 | 28.59 | -5.92 | 28.13 | -7.44
92.16 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | -0.29 | 21.39 | -4.1 | 0 | -30.53 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | -0.06 | 0.06 | -15.65 | 0 | -26.37 | -0.06 | -29.06 | 0 | 0 | -0.06 | 28.54 | 0.12 | 28.42 | -0.23 | 28.65 | -2.87 | 28.59 | -4.51 | 28.59 | -5.57
92.192 | 0.12 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 21.39 | -1.93 | -0.18 | -28.3 | -0.06 | 0 | 0.06 | 0 | 0.12 | 21.39 | -0.18 | 0 | -14.24 | 0 | -23.85 | -0.06 | -27.36 | -0.06 | 0 | -0.18 | 28.65 | 0.06 | 28.48 | -0.06 | 28.3 | -1.7 | 28.48 | -2.7 | 28.42 | -3.69
92.224 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -1.46 | 0 | -25.67 | -0.06 | 0 | 0.06 | 0 | 0.06 | 21.39 | 0.06 | -0.06 | -11.95 | -0.06 | -22.44 | -0.06 | -24.84 | 0 | -4.69 | -0.06 | 0.29 | 0.12 | 28.54 | 0.06 | 28.65 | -1.05 | 28.48 | -1.82 | 28.83 | -2.52
92.256 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | -0.53 | -0.06 | -23.97 | 0.06 | 0 | -0.06 | 0 | 0.12 | 21.15 | 0.12 | 4.16 | -10.9 | 0 | -20.16 | -0.06 | -22.85 | 0 | -90.41 | -0.06 | 0.06 | -0.06 | 28.48 | 0.35 | 28.65 | -0.7 | 28.36 | -1.29 | 28.36 | -1.76
92.288 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.06 | 21.39 | -0.35 | -0.18 | -21.74 | -0.06 | -34.63 | 0 | 0 | 0.23 | 21.39 | 0.12 | 10.2 | -9.9 | -0.06 | -18.16 | 0 | -20.92 | 0 | -63.75 | -0.06 | 0 | 0.06 | 28.59 | 0.23 | 28.42 | -0.47 | 28.48 | -0.59 | 28.48 | -1.11
92.32 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.5 | 0.12 | 0 | -20.22 | -0.06 | -69.73 | -0.06 | 0 | 0.12 | 21.27 | 0.18 | 15.59 | -8.85 | 0 | -16.88 | 0 | -19.22 | 0.06 | -49.04 | -0.06 | 0.06 | 0.12 | 28.54 | 0.18 | 28.48 | -0.29 | 28.48 | -0.7 | 28.3 | -0.76
92.352 | 0 | -0.06 | 21.39 | -0.29 | 21.45 | -0.18 | 21.33 | 0.29 | -0.06 | -17.58 | -0.06 | -44.77 | -0.06 | -0.23 | 0.06 | 21.39 | 0.18 | 21.62 | -6.91 | 0 | -14.59 | 0 | -16.82 | -0.06 | -45.12 | -0.18 | 0 | -0.06 | 28.36 | 0.29 | 28.42 | -0.18 | 28.42 | -0.18 | 28.59 | -0.23
92.384 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | 0.41 | 0.06 | -15.41 | 0 | -40.02 | -0.18 | -66.15 | 0.06 | 21.21 | 0.12 | 21.27 | -5.8 | -0.06 | -12.95 | 0 | -15.29 | -0.06 | -38.61 | 0 | 0 | -0.18 | 28.65 | 0.41 | 28.36 | -0.06 | 28.59 | -0.18 | 28.48 | -0.18
92.416 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.35 | 0.06 | -13.48 | -0.06 | -34.22 | -0.06 | -59.94 | 0.18 | 21.21 | 0.06 | 21.33 | -4.63 | -0.06 | -11.6 | 0 | -13.65 | -0.06 | -36.1 | -0.06 | 0 | -0.06 | 28.42 | 0.35 | 28.59 | 0.06 | 28.65 | -0.06 | 28.42 | -0.06
92.448 | 0 | -0.06 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | 0.41 | -0.06 | -11.72 | -0.06 | -32.23 | 0 | -34.98 | 0 | 21.15 | 0.18 | 21.33 | -3.69 | 0 | -10.31 | 2.99 | -12.48 | 0 | -37.21 | 0.06 | 0 | -0.06 | 28.3 | 0.35 | 28.59 | 0.18 | 28.42 | 0 | 28.65 | -0.18
92.48 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | -0.18 | 21.45 | 0.47 | -0.18 | -9.55 | 0 | -32.76 | -0.18 | -26.78 | -0.06 | 21.39 | 0.23 | 21.45 | -3.16 | -0.06 | -8.85 | 11.13 | -10.96 | -0.06 | -37.5 | -0.18 | 0 | -0.06 | 28.65 | 0.35 | 28.59 | 0.12 | 28.65 | 0.06 | 28.48 | 0.12
92.512 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | -0.06 | -6.62 | 0 | -31.64 | -0.06 | -34.4 | -0.06 | 21.45 | 0.35 | 21.27 | -2.46 | 6.91 | -7.44 | 21.39 | -9.32 | -0.06 | -36.27 | -0.06 | 0 | -0.06 | 0.18 | 0.18 | 28.48 | 0.06 | 28.48 | 0.06 | 27.95 | 0
92.544 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 6.15 | -4.63 | -0.06 | -31.11 | 0.06 | -35.27 | 0 | 21.39 | 0.18 | 21.27 | -2.05 | 14.71 | -6.21 | 29.42 | -7.91 | 0 | -34.69 | 0 | 0 | 0 | 0 | 0.23 | 28.42 | 0.18 | 28.3 | 0.18 | 28.48 | 0.06
92.576 | -0.06 | -0.06 | 0.23 | -0.18 | 21.45 | -0.18 | 21.45 | 0.47 | 11.89 | -2.46 | -0.06 | -28.36 | -0.06 | -32.34 | 0.12 | 21.39 | 0.12 | 21.33 | -1.64 | 22.15 | -4.92 | 28.77 | -6.8 | 0 | -32.52 | 0 | 0 | 0 | 0 | 0.18 | 28.65 | 0.18 | 28.36 | 0.12 | 28.71 | 0.12
92.608 | -28.01 | -0.06 | 0.06 | -0.18 | 21.45 | -0.18 | 21.33 | 0.53 | 17.58 | -1.88 | -0.06 | -26.13 | 0.12 | -29.06 | 0.06 | 18.11 | 0.12 | 21.33 | -1.23 | 29.18 | -4.16 | 28.65 | -5.45 | 6.09 | -30.29 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.59 | 0.06 | 28.42 | 0.18 | 28.65 | 0.12
92.64 | -61.58 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.29 | 21.74 | -1 | -0.06 | -24.32 | -0.06 | -27.77 | 0.06 | 0.12 | 0.29 | 21.33 | -1 | 28.59 | -3.4 | 28.48 | -4.63 | 14.06 | -28.83 | 0 | -18.22 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.36 | 0.06 | 28.48 | 0.18
92.672 | -40.55 | -0.18 | 0.06 | -0.18 | 21.45 | -0.18 | 21.27 | 0.41 | 21.45 | -0.35 | -0.06 | -21.33 | 0 | -26.72 | 0.12 | 0 | 0.18 | 21.39 | -0.7 | 28.48 | -2.52 | 28.3 | -3.34 | 24.14 | -25.43 | 0 | -82.27 | -0.18 | 0 | 0.12 | 28.48 | 0 | 28.36 | 0.12 | 28.59 | 0.06
92.704 | -39.38 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.35 | 21.39 | -0.18 | -0.06 | -19.45 | -0.06 | -25.08 | -0.06 | 0 | 0.18 | 21.21 | -0.53 | 28.65 | -1.93 | 28.65 | -2.64 | 28.48 | -24.14 | -0.18 | -52.97 | -0.06 | 0 | -0.06 | 28.54 | 0.12 | 28.48 | 0 | 28.42 | 0.23
92.736 | -37.68 | -0.06 | 0 | -0.06 | 0.12 | -0.18 | 21.39 | 0.23 | 21.33 | 0.18 | -0.06 | -17.81 | -0.06 | -23.44 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.59 | -1.46 | 28.59 | -1.93 | 28.59 | -22.62 | -0.06 | -47.99 | -0.06 | 0 | 0.06 | 28.3 | 0.06 | 28.71 | 0.12 | 28.42 | 0.12
92.768 | -36.1 | 0 | 0 | -0.06 | 0 | -0.12 | 21.39 | 0.29 | 21.45 | 0.23 | 0 | -15.94 | -0.06 | -21.45 | 0.06 | 0 | 0.18 | 21.15 | -0.41 | 28.42 | -1.11 | 28.48 | -1.52 | 29.01 | -20.92 | -0.06 | -43.13 | -0.06 | 0 | -0.06 | 28.42 | 0 | 28.54 | 0.06 | 28.42 | 0.23
92.8 | -34.22 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.12 | 0.06 | 21.5 | 0.29 | 0 | -13.36 | -0.06 | -19.92 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.48 | -0.94 | 28.48 | -1 | 28.48 | -18.99 | -0.18 | -37.62 | -0.06 | 0 | 0.12 | 28.59 | 0.18 | 28.48 | 0.18 | 28.42 | 0.18
92.832 | -31.06 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.41 | -0.06 | -11.13 | -0.06 | -17.58 | 0 | 0 | 0.12 | 21.39 | -0.23 | 28.65 | -0.59 | 28.48 | -0.64 | 28.65 | -16.29 | -0.18 | -37.5 | -0.06 | 0 | -0.06 | 0.23 | 0.06 | 28.48 | 0.06 | 28.42 | 0.12
92.864 | -29.24 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 21.45 | 0.35 | -0.06 | -9.32 | -0.18 | -16.17 | -0.06 | 0 | 0.12 | 21.39 | -0.29 | 28.42 | -0.35 | 28.42 | -0.35 | 28.65 | -14.53 | -0.06 | -38.15 | -0.06 | 0 | 0 | 0 | 0.12 | 28.48 | 0.18 | 0.29 | 0.12
92.896 | -27.36 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | 0.53 | -0.06 | -6.86 | -0.06 | -14.65 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.48 | -0.35 | 28.3 | -0.41 | 28.48 | -12.77 | -0.06 | -37.15 | -0.06 | 0 | 0 | 0 | -0.06 | 0.29 | 0.12 | 0.12 | 0.18
92.928 | -25.14 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 4.22 | -4.8 | -0.06 | -12.83 | 0 | 0 | 0.12 | 21.33 | -0.35 | 28.83 | -0.29 | 28.42 | -0.35 | 28.3 | -10.55 | 0.06 | -35.16 | -0.06 | -44.12 | -0.18 | 0.06 | -0.06 | 0.12 | 0.12 | 0 | 0.18
92.96 | -23.61 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.47 | 10.31 | -2.87 | 0 | -11.31 | 0 | 0 | 0.06 | 21.39 | -0.41 | 28.48 | -0.35 | 28.13 | -0.29 | 28.59 | -8.44 | -0.06 | -33.81 | 0 | -86.37 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0.06 | 0.12
92.992 | -20.16 | 0.06 | -15.06 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.47 | 17.52 | -1.52 | 0 | -8.91 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.83 | -0.23 | 28.89 | -0.35 | 28.24 | -5.92 | 0 | -31 | 0 | -49.75 | -0.18 | 0 | -0.18 | 0.06 | 0.06 | 0 | 0.23
93.024 | -18.69 | -0.06 | -63.11 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.74 | -0.82 | -0.06 | -7.09 | 0 | -0.12 | 0 | 21.33 | -0.23 | 28.59 | -0.23 | 28.71 | -0.23 | 28.59 | -3.75 | 9.43 | -29.06 | -0.06 | -46.82 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06
93.056 | -17.23 | 0.06 | -47.05 | -0.18 | 0 | -0.18 | 0 | 0 | 21.45 | 0.35 | 21.39 | -0.47 | 0 | -5.39 | 0.06 | -56.19 | 0.06 | 21.39 | -0.29 | 28.65 | -0.29 | 28.83 | -0.23 | 28.36 | -2.23 | 16.7 | -27.6 | -0.06 | -41.78 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
93.088 | -15.18 | -0.18 | -40.37 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | -0.06 | -0.06 | -3.75 | 0 | -60 | 0.06 | 21.21 | -0.29 | 28.3 | -0.29 | 28.48 | -0.18 | 28.59 | -1.7 | 24.38 | -25.55 | -0.06 | -36.21 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.18
93.12 | -14 | -0.06 | -39.2 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.15 | 0.06 | 7.09 | -2.52 | -0.06 | -39.26 | 0.06 | 21.39 | -0.29 | 28.48 | -0.23 | 28.83 | -0.23 | 28.42 | -0.64 | 28.83 | -24.26 | 0 | -36.86 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.18
93.152 | -11.78 | 0.06 | -35.74 | 0 | -27.66 | -0.18 | 0 | 0.06 | 21.45 | 0.29 | 21.39 | 0.23 | 14.47 | -1.35 | -0.06 | -28.24 | 0.12 | 0.12 | -0.18 | 28.65 | -0.29 | 28.59 | -0.23 | 28.59 | -0.59 | 28.54 | -21.91 | -0.06 | -38.5 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
93.184 | -10.25 | -0.06 | -34.45 | 0 | -61.82 | -0.18 | 0 | 0.06 | 21.45 | 0.23 | 21.39 | 0.29 | 20.16 | -0.94 | 0 | -34.4 | 0.06 | 0 | -0.23 | 28.59 | -0.23 | 28.24 | -0.18 | 28.42 | 0.29 | 28.48 | -19.86 | 0 | -37.38 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0
93.216 | -8.85 | -0.06 | -32.23 | -0.06 | -45.29 | 0 | -34.75 | 0 | 21.5 | 0.18 | 21.39 | 0.59 | 21.5 | -0.76 | -0.06 | -35.51 | 0.06 | 0 | -0.18 | 28.48 | -0.29 | 28.42 | -0.18 | 28.48 | 0.18 | 28.71 | -17.75 | -0.06 | -35.33 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.06
93.248 | -7.27 | 0 | -30.06 | -0.06 | -40.2 | -0.06 | -73.77 | -0.06 | 21.45 | 0.12 | 21.39 | 0.53 | 21.33 | -0.41 | 0.06 | -33.17 | 0.06 | 0 | -0.18 | 28.65 | -0.18 | 28.77 | -0.18 | 28.48 | 0.18 | 28.48 | -16.11 | -0.06 | -32.76 | -0.06 | -39.79 | -0.06 | 0 | 0 | 0 | 0.18
93.28 | -6.15 | -0.06 | -28.48 | 0 | -37.97 | -0.06 | -50.16 | -0.06 | 21.56 | 0.12 | 21.33 | 0.41 | 21.33 | -0.18 | 0 | -29.71 | 0.23 | 0 | -0.18 | 28.59 | -0.18 | 28.59 | -0.18 | 28.48 | 0.29 | 28.59 | -14.06 | 0 | -30.94 | 0 | -85.32 | -0.06 | 0 | 0 | -2.34 | 0
93.312 | -4.28 | -0.06 | -25.72 | -0.06 | -35.39 | -0.06 | -40.14 | 0 | 21.39 | 0 | 21.39 | 0.47 | 21.21 | -0.18 | -0.06 | -27.42 | 0 | 0 | -0.23 | 28.48 | -0.18 | 28.54 | -0.23 | 28.65 | 0.29 | 28.48 | -12.07 | -0.06 | -28.77 | -0.06 | -46.88 | -0.06 | -32.81 | -0.06 | -86.19 | 0.06
93.344 | -2.64 | -0.06 | -23.5 | -0.06 | -34.1 | -0.06 | -34.75 | -0.18 | 21.56 | 0.18 | 21.27 | 0.41 | 21.39 | 0 | 0.06 | -26.84 | 0.12 | 0 | -0.18 | 28.42 | -0.18 | 0.23 | -0.06 | 28.42 | 0.35 | 28.59 | -9.61 | -0.06 | -26.72 | -0.06 | -29.65 | -0.06 | -84.14 | 0 | -57.01 | 0.12
93.376 | -1.7 | 3.75 | -22.15 | 0 | -31.93 | -0.06 | -31.99 | -0.06 | 21.39 | -0.06 | 21.39 | 0.41 | 21.33 | 0.18 | 7.68 | -25.25 | 0.12 | 0 | -0.06 | 28.42 | -0.18 | 0.12 | -0.18 | 28.48 | 0.29 | 28.42 | -7.79 | -0.06 | -25.37 | -0.06 | -34.98 | -0.06 | -56.55 | 0 | -33.11 | 0
93.408 | -1 | 9.79 | -19.86 | -0.06 | -29.83 | -0.06 | -32.52 | -0.06 | 21.45 | 0 | 21.33 | 0.29 | 21.39 | 0.29 | 13.07 | -23.44 | 0.06 | 0 | -0.18 | 0.23 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.48 | -5.45 | -0.06 | -23.91 | -0.06 | -43.07 | -0.18 | -32.64 | 0 | -31.06 | 0.06
93.44 | -0.47 | 15.35 | -17.99 | 0 | -27.01 | -0.06 | -31.93 | -0.18 | 0.18 | 0.06 | 21.33 | 0.35 | 21.33 | 0.12 | 18.75 | -21.68 | 0.12 | 0 | -0.18 | 0.06 | -0.18 | 0.06 | 0 | 28.71 | 0.35 | 28.3 | -3.57 | 0 | -21.91 | -0.06 | -40.55 | -0.18 | -31.41 | 0 | -41.19 | 0.12
93.472 | -0.18 | 21.8 | -15.53 | -0.06 | -25.37 | -0.18 | -30.29 | -0.06 | 0 | 0 | 21.21 | 0.23 | 21.33 | 0.12 | 21.45 | -20.16 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.48 | -2.05 | 3.87 | -19.45 | -0.06 | -35.63 | -0.06 | -41.84 | 0 | -41.13 | -0.06
93.504 | -0.18 | 21.5 | -14.18 | -0.06 | -23.09 | 0 | -28.89 | 0 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | 0.12 | 21.27 | -18.05 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.23 | 0.29 | 28.42 | -0.82 | 12.25 | -17.29 | 0 | -33.17 | -0.06 | -41.08 | -0.06 | -36.68 | -0.06
93.536 | -0.35 | 21.5 | -12.36 | -0.06 | -21.56 | 0 | -26.78 | -0.06 | 0 | -0.06 | 21.21 | 0.18 | 21.27 | 0.18 | 21.39 | -16.47 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.23 | 28.54 | -0.53 | 19.86 | -15.82 | -0.06 | -32.64 | -0.06 | -36.8 | -0.18 | -33.63 | -0.06
93.568 | -0.29 | 21.5 | -10.96 | -0.06 | -19.04 | 0 | -24.55 | -0.06 | 0 | -0.18 | 21.27 | 0.23 | 21.45 | 0.23 | 21.33 | -14.24 | -0.06 | -17.64 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.29 | 28.65 | -0.18 | 27.54 | -13.95 | 0 | -32.4 | -0.18 | -34.04 | 0 | -33.34 | -0.06
93.6 | -0.35 | 21.5 | -9.49 | 0.06 | -16.7 | -0.18 | -22.44 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.12 | 21.39 | -13.07 | -0.06 | -73.54 | -0.18 | 0.06 | -0.18 | 0 | -0.06 | 0.06 | 0.29 | 28.54 | 0 | 28.42 | -12.01 | 0 | -30.82 | 0 | -32.81 | 0 | -32.76 | -0.06
93.632 | -0.41 | 21.45 | -8.09 | 0 | -14.94 | 0 | -19.98 | -0.06 | 0 | -0.18 | 21.33 | 0.23 | 21.33 | 0.12 | 21.45 | -11.25 | 0 | -43.54 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.48 | 0.12 | 28.48 | -9.2 | -0.06 | -27.6 | 0 | -32.58 | 0 | -31 | 0
93.664 | -0.41 | 21.45 | -6.15 | 0 | -13.54 | -0.06 | -18.28 | -0.18 | 0 | -0.06 | 21.33 | 0.12 | 21.45 | 0.18 | 21.39 | -9.08 | -0.06 | -28.13 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.48 | 0.29 | 28.59 | -7.03 | 0 | -26.49 | 0 | -31.29 | 0.06 | -28.83 | -0.06
93.696 | -0.29 | 21.39 | -4.69 | -0.06 | -12.25 | -0.06 | -16 | -0.18 | 0 | -0.06 | 21.33 | 0.06 | 21.39 | 0.29 | 21.39 | -7.56 | 0 | -29.59 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 28.48 | 0.41 | 28.48 | -5.04 | 0 | -25.43 | 0 | -28.77 | -0.18 | -27.01 | 0
93.728 | -0.18 | 21.45 | -3.4 | -0.06 | -10.84 | 0 | -14.06 | -0.06 | 0 | -0.06 | 21.45 | 0.06 | 21.39 | 0.12 | 21.33 | -5.51 | -0.06 | -35.1 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.65 | 0.35 | 28.77 | -2.99 | 0 | -23.91 | -0.06 | -27.01 | 0 | -26.25 | 0
93.76 | -0.23 | 21.39 | -2.29 | -0.06 | -8.91 | 0 | -12.07 | -0.06 | 0 | 0.06 | 21.39 | 0 | 21.15 | 0.12 | 21.21 | -3.98 | -0.06 | -35.16 | 0 | 0 | 0 | -3.93 | -0.18 | 0 | 0.06 | 28.71 | 0.35 | 28.48 | -1.99 | -0.06 | -22.68 | -0.06 | -25.9 | -0.18 | -25.02 | 0
93.792 | -0.18 | 21.45 | -1.35 | 7.79 | -7.44 | -0.06 | -9.08 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.12 | 21.39 | -2.46 | 0 | -30.53 | -0.18 | 0.06 | 0 | -74.53 | -0.18 | 0 | 0.06 | 28.24 | 0.35 | 28.36 | -0.94 | -0.06 | -20.45 | -0.18 | -24.43 | -0.06 | -22.38 | -0.06
93.824 | -0.23 | 21.45 | -0.76 | 13.48 | -5.92 | 7.15 | -7.03 | -0.18 | 0.06 | -0.18 | 0.18 | 0 | 21.39 | 0.18 | 21.33 | -1.82 | 0.06 | -28.3 | -0.06 | -34.63 | -0.18 | -51.92 | 0 | 0 | 0 | 28.77 | 0.41 | 28.59 | -0.59 | -0.06 | -18.46 | -0.06 | -22.62 | -0.06 | -20.92 | 0
93.856 | -0.29 | 21.45 | -0.29 | 19.04 | -4.45 | 12.83 | -4.98 | -0.18 | -43.19 | -0.18 | 0 | 0.06 | 21.39 | 0.06 | 21.39 | -0.94 | -0.06 | -27.07 | -0.18 | -71.55 | -0.06 | -46.23 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.48 | -0.18 | -0.06 | -17.23 | -0.06 | -21.27 | -0.06 | -19.86 | -0.06
93.888 | -0.23 | 21.45 | -0.23 | 21.5 | -3.16 | 18.34 | -2.87 | -0.06 | -65.69 | -0.06 | 0 | 0.06 | 21.33 | 0.12 | 21.39 | -0.64 | 0 | -26.6 | 0 | -51.27 | -0.06 | -45.35 | 0 | 0 | 0 | 28.48 | 0.29 | 28.42 | -0.18 | -0.06 | -15.65 | -0.06 | -19.75 | -0.06 | -18.05 | 0
93.92 | -0.18 | 21.39 | -0.23 | 21.45 | -1.46 | 21.68 | -2.11 | -0.06 | -47.52 | -0.06 | 0 | 0.06 | 21.39 | 0 | 21.45 | -0.41 | 4.22 | -24.43 | -0.06 | -45.82 | -0.06 | -41.66 | -0.06 | -34.51 | 0 | 0.18 | 0.41 | 28.59 | 0.06 | 0.41 | -13.95 | -0.18 | -18.16 | 0 | -16.29 | 0
93.952 | -0.29 | 21.45 | -0.29 | 21.5 | -0.94 | 21.45 | -0.35 | -0.06 | -39.08 | -0.06 | 0 | 0.06 | 21.33 | 0.06 | 21.39 | 0 | 12.13 | -21.97 | -0.06 | -43.36 | -0.06 | -40.14 | -0.06 | -66.21 | -0.06 | 0 | 0.18 | 28.24 | 0.29 | 12.66 | -11.95 | 0.82 | -15.7 | -0.18 | -13.83 | -0.06
93.984 | -0.18 | 21.45 | -0.41 | 21.45 | -0.53 | 21.5 | -0.41 | -0.06 | -33.17 | -0.06 | 0 | -0.06 | 6.8 | 0.06 | 21.33 | 0.12 | 17.52 | -20.57 | -0.18 | -41.6 | 0 | -37.97 | -0.06 | -47.05 | 0 | 0 | 0.12 | 28.48 | 0.29 | 20.51 | -9.96 | 10.49 | -13.71 | 0.94 | -12.48 | 0
94.016 | -0.18 | 21.39 | -0.47 | 21.45 | -0.18 | 21.45 | 0.06 | -0.18 | -31.47 | -0.06 | 0 | 0 | 0.06 | 0.06 | 21.39 | 0.12 | 21.62 | -18.99 | -0.06 | -39.96 | 0 | -36.39 | -0.18 | -46.53 | 0 | 0 | 0.18 | 28.42 | 0.29 | 28.54 | -8.67 | 18.05 | -12.31 | 10.78 | -10.37 | 3.87
94.048 | -0.18 | 21.39 | -0.41 | 21.45 | -0.06 | 21.5 | 0.06 | -0.18 | -32.23 | -0.18 | 0 | 0 | 0 | -0.06 | 21.39 | 0.12 | 21.45 | -17.52 | -0.06 | -38.15 | 0 | -34.81 | 0 | -44.3 | 0 | 0 | 0.29 | 28.36 | 0.35 | 28.48 | -7.03 | 25.72 | -10.84 | 18.11 | -9.26 | 11.95
94.08 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | 0.35 | -0.06 | -32.17 | 0.06 | 0 | -0.18 | 0 | 0 | 21.21 | 0 | 21.39 | -15.65 | -0.06 | -36.1 | -0.06 | -33.17 | -0.06 | -42.31 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.54 | -5.45 | 28.65 | -9.61 | 25.9 | -7.68 | 19.75
94.112 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.23 | -0.06 | -30.53 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 21.39 | -13.71 | -0.06 | -34.04 | -0.06 | -30.29 | -0.18 | -38.97 | 0 | 0 | -0.06 | 28.59 | 0.29 | 28.36 | -3.46 | 28.59 | -7.21 | 28.54 | -5.74 | 29.3
94.144 | -0.18 | 21.45 | -0.18 | 21.45 | -0.41 | 21.45 | 0.41 | -0.18 | -28.01 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.39 | -11.95 | 0 | -31.93 | 0 | -28.77 | -0.06 | -37.5 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.36 | -2.23 | 28.71 | -5.45 | 28.48 | -4.04 | 28.48
94.176 | -0.18 | 21.45 | -0.23 | 21.5 | -0.47 | 21.45 | 0.47 | 1.99 | -25.9 | 0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.27 | -10.25 | -0.06 | -29.94 | 0 | -26.72 | 0.06 | -36.45 | 0 | 0 | 0.06 | 28.59 | 0.18 | 28.65 | -1.64 | 28.48 | -3.93 | 28.42 | -2.99 | 28.54
94.208 | -0.06 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | 0.47 | 6.62 | -23.67 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.27 | -8.44 | -0.06 | -28.65 | 0 | -24.84 | -0.06 | -33.81 | -0.06 | 0 | 0 | 28.59 | 0.23 | 28.54 | -0.82 | 28.54 | -2.7 | 28.54 | -1.82 | 28.65
94.24 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.47 | 12.19 | -21.86 | -0.18 | -21.56 | 0 | 0 | 0 | 19.81 | 0.18 | 21.45 | -6.68 | -0.06 | -26.13 | -0.06 | -22.56 | -0.06 | -32.23 | 0 | 0 | 0 | 28.48 | 0.12 | 28.42 | -0.59 | 28.48 | -1.7 | 28.42 | -1.29 | 28.59
94.272 | -0.06 | 0.12 | -0.23 | 21.5 | -0.23 | 21.39 | 0.41 | 19.51 | -19.28 | -0.06 | -68.32 | -0.06 | 0 | 0.06 | 0.12 | 0.23 | 21.45 | -4.75 | 9.43 | -23.61 | 0 | -20.22 | 0.06 | -29.94 | -0.06 | 0 | -0.18 | 28.48 | 0.12 | 28.59 | -0.29 | 28.48 | -0.94 | 28.59 | -0.35 | 28.42
94.304 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | -17.58 | 0 | -45.94 | 0 | 0 | 0.12 | 0 | 0.23 | 21.33 | -3.46 | 16.7 | -21.86 | 0 | -18.52 | -0.06 | -28.59 | 0 | 0 | 0 | 28.3 | 0 | 28.48 | -0.18 | 29.01 | -0.59 | 28.65 | -0.47 | 28.36
94.336 | 0 | 0 | -0.29 | 21.45 | -0.29 | 21.39 | 0.41 | 21.39 | -15.94 | -0.18 | -41.84 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.15 | -2.23 | 24.55 | -19.98 | -0.06 | -16.17 | -0.06 | -26.72 | 0 | -74.77 | 0 | 28.65 | 0.12 | 28.48 | -0.18 | 28.48 | -0.35 | 28.65 | -0.23 | 28.59
94.368 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.35 | 21.39 | -13.71 | -0.06 | -35.16 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | -1.7 | 28.77 | -18.05 | -0.06 | -15.53 | -0.06 | -24.26 | 0 | -72.42 | 0 | 0.7 | 0 | 28.42 | 0.06 | 28.59 | -0.18 | 28.36 | -0.06 | 28.48
94.4 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.47 | 21.39 | -11.66 | -0.06 | -31.58 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.33 | -0.94 | 28.83 | -16.47 | 0 | -13.54 | 0 | -21.62 | -0.06 | -49.87 | -0.06 | 0.06 | 0 | 28.59 | 0.06 | 28.54 | -0.18 | 28.65 | 0 | 28.83
94.432 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.35 | 21.33 | -9.08 | -0.06 | -32.34 | -0.06 | -71.66 | 0 | 0 | 0.12 | 21.15 | -0.29 | 28.59 | -14.77 | 5.68 | -11.72 | 0 | -19.86 | -0.06 | -45.35 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.48 | 0 | 28.65 | -0.06 | 28.24
94.464 | -0.06 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | 0.29 | 21.33 | -6.5 | -0.18 | -32.76 | -0.06 | -56.25 | -0.06 | 0 | 0.12 | 21.33 | -0.23 | 28.65 | -12.77 | 13.83 | -10.37 | 2.46 | -17.64 | -0.18 | -39.44 | 0 | 0 | 0 | 28.54 | 0.06 | 28.13 | 0.12 | 28.83 | 0.06 | 28.54
94.496 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.23 | 21.39 | -4.45 | -0.06 | -31.06 | -0.06 | -32.76 | -0.06 | 0 | 0.18 | 21.39 | -0.18 | 28.59 | -11.6 | 21.21 | -9.2 | 10.31 | -15.94 | 0 | -36.33 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | 0.12 | 28.48 | 0.18 | 28.42
94.528 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.12 | 21.39 | -2.46 | 0 | -28.95 | -0.06 | -27.6 | -0.06 | 0 | 0.23 | 21.45 | 0 | 28.42 | -10.25 | 29.18 | -7.56 | 17.81 | -14.53 | -0.06 | -37.33 | -0.06 | 0 | -0.06 | 28.54 | 0.06 | 28.42 | 0.06 | 28.54 | 0.23 | 28.65
94.56 | 0 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | 0.12 | 21.39 | -1.88 | -0.18 | -26.72 | -0.06 | -34.1 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 28.54 | -8.91 | 28.77 | -6.86 | 25.49 | -13.01 | 0 | -37.74 | -0.06 | 0 | -0.23 | 28.42 | 0.12 | 28.59 | 0.06 | 28.48 | 0.18 | 28.24
94.592 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.18 | 21.39 | -0.94 | 5.33 | -24.43 | 0.12 | -34.86 | -0.18 | 0 | 0.23 | 21.39 | 0 | 28.71 | -7.32 | 28.54 | -4.98 | 28.65 | -11.13 | -0.18 | -36.15 | 0 | 0 | -0.23 | 28.59 | 0.12 | 28.42 | 0.12 | 28.07 | 0.18 | 28.48
94.624 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.5 | 0.18 | 21.33 | -0.35 | 11.31 | -22.09 | -0.18 | -31.99 | 0 | 0 | 0.18 | 21.39 | 0.18 | 28.77 | -6.15 | 28.59 | -4.22 | 28.65 | -10.08 | -0.06 | -34.51 | 0 | 0 | 0 | 28.48 | 0.12 | 28.59 | 0.23 | 28.42 | 0.23 | 28.3
94.656 | -0.06 | 0 | -0.06 | 17.7 | -0.18 | 21.39 | 0.12 | 21.39 | -0.06 | 16.7 | -20.16 | -0.06 | -29.24 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 28.54 | -4.98 | 28.54 | -3.52 | 28.54 | -8.61 | -0.06 | -32.4 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | 0.23 | 28.42 | 0.12 | 28.48
94.688 | -0.18 | -15.12 | -0.18 | 0.12 | -0.18 | 21.45 | 0 | 21.39 | 0.18 | 21.86 | -18.16 | -0.06 | -27.42 | 0 | -56.37 | 0.12 | 21.21 | 0.18 | 28.71 | -4.16 | 28.89 | -2.75 | 28.36 | -7.27 | -0.06 | -30.41 | -0.06 | 0 | 0.06 | 28.42 | 0.06 | 28.65 | 0.18 | 28.48 | 0 | 28.36
94.72 | -0.06 | -63.28 | -0.18 | 0 | -0.18 | 0.12 | -0.06 | 21.33 | 0.29 | 21.39 | -16.29 | -0.18 | -26.72 | 0 | -62.35 | 0.12 | 21.39 | 0.23 | 28.48 | -3.46 | 28.59 | -2.05 | 28.18 | -5.92 | 0 | -28.83 | -0.06 | 0 | 0 | 28.36 | 0 | 28.48 | 0.23 | 28.59 | 0.18 | 28.42
94.752 | -0.06 | -44.01 | -0.18 | 0.06 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -13.77 | -0.06 | -24.84 | -0.06 | -33.75 | 0.06 | 21.33 | 0.23 | 28.59 | -2.64 | 28.36 | -1.46 | 28.07 | -4.63 | 0 | -26.66 | 0 | 0 | -0.06 | 28.42 | 0 | 28.42 | 0.12 | 28.3 | 0.06 | 28.42
94.784 | -0.06 | -40.61 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.33 | 0.41 | 21.39 | -11.72 | -0.06 | -22.62 | -0.06 | -28.3 | 0.12 | 21.33 | 0.18 | 28.3 | -2.05 | 28.36 | -1.05 | 28.36 | -3.75 | 0 | -24.79 | 0 | -15.53 | 0 | 28.48 | -0.06 | 28.42 | 0.06 | 28.65 | 0.18 | 28.48
94.816 | 0 | -38.61 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -9.61 | -0.06 | -21.09 | 0 | -33.28 | 0.18 | 0.12 | 0.23 | 28.42 | -1.46 | 28.3 | -0.7 | 28.36 | -2.99 | -0.06 | -22.79 | -0.18 | -92.76 | 0 | 0.29 | 0 | 28.48 | 0.06 | 28.48 | 0.18 | 28.48
94.848 | -0.18 | -36.04 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.47 | 21.33 | -7.62 | -0.06 | -19.75 | -0.06 | -35.57 | 0.12 | 0 | 0.23 | 28.3 | -1.29 | 28.54 | -0.59 | 28.48 | -2.52 | -0.06 | -21.68 | 0 | -59.42 | 0 | 0.12 | 0 | 20.8 | 0.23 | 28.48 | 0.12 | 28.42
94.88 | -0.06 | -34.69 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.47 | 21.39 | -5.63 | -0.06 | -18.52 | 0 | -33.63 | 0.06 | 0 | 0.23 | 28.48 | -1 | 28.48 | -0.29 | 28.71 | -1.76 | 0 | -19.51 | -0.06 | -47.93 | -0.18 | 0 | -0.06 | 0.18 | 0.06 | 18.4 | 0.23 | 28.42
94.912 | -0.06 | -31.7 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -2.99 | -0.06 | -15.59 | 0 | -29.06 | 0.18 | 0 | 0.18 | 28.83 | -0.7 | 28.48 | -0.35 | 28.71 | -1.11 | 2.4 | -17.17 | -0.18 | -43.19 | -0.06 | 0.06 | -0.06 | 0 | 0 | 0.12 | 0.06 | 0.29
94.944 | 0 | -29.36 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.46 | 0 | -14.18 | 0 | -27.48 | 0.12 | 0 | 0.18 | 28.89 | -0.59 | 28.36 | -0.35 | 28.42 | -0.82 | 10.37 | -15.29 | -0.18 | -37.56 | -0.06 | 0 | 0.06 | 0.06 | 0 | 0 | 0.18 | 0.06
94.976 | -0.06 | -27.31 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -1.05 | 4.63 | -12.36 | 0 | -26.78 | 0 | 0 | 0.18 | 28.42 | -0.47 | 28.07 | -0.35 | 28.3 | -0.64 | 17.81 | -12.77 | 0 | -36.97 | -0.06 | 0 | -0.06 | 0 | 0 | 0.06 | 0.12 | 0
95.008 | 0 | -25.55 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -0.64 | 10.66 | -10.61 | -0.06 | -25.37 | 0 | 0 | 0.18 | 28.42 | -0.47 | 28.18 | -0.29 | 28.54 | -0.41 | 25.2 | -11.13 | -0.18 | -37.5 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0.06
95.04 | 0 | -23.91 | -0.06 | 0 | 0 | 0 | -0.06 | 21.39 | 0.35 | 21.33 | -0.18 | 16 | -9.08 | 0 | -23.2 | 0 | 0 | 0.12 | 28.65 | -0.35 | 28.3 | -0.29 | 28.65 | -0.29 | 28.54 | -9.14 | 0 | -37.09 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0
95.072 | -0.06 | -20.57 | -0.18 | -0.23 | -0.06 | 0 | -0.06 | 11.07 | 0.29 | 21.39 | 0.18 | 21.5 | -6.62 | 7.56 | -21.09 | 0.06 | 0 | 0.23 | 28.48 | -0.41 | 28.65 | -0.23 | 28.42 | -0.18 | 28.65 | -6.45 | 0 | -35.27 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0 | 0 | 0
95.104 | 0 | -18.63 | 0 | -46.23 | -0.06 | 0 | 0 | 0.12 | 0.29 | 21.39 | 0.23 | 21.39 | -4.98 | 13.24 | -19.69 | 0.12 | 0 | 0.18 | 28.48 | -0.41 | 28.48 | -0.29 | 28.36 | -0.23 | 28.24 | -4.28 | -0.18 | -33.69 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0.06
95.136 | -0.06 | -17.05 | -0.06 | -53.09 | -0.06 | -38.97 | 0 | 0 | 0.23 | 21.39 | 0.35 | 21.33 | -3.28 | 18.81 | -17.99 | 0 | 0 | 0.18 | 28.59 | -0.35 | 28.48 | -0.18 | 28.42 | -0.41 | 28.65 | -2.29 | -0.18 | -31.99 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
95.168 | -0.06 | -15.35 | 0 | -41.43 | -0.06 | -55.84 | -0.06 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -2.46 | 21.56 | -16.23 | 0.06 | 0 | 0.18 | 0.18 | -0.29 | 28.42 | -0.18 | 28.24 | -0.35 | 28.54 | -1.76 | -0.06 | -29.77 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 0
95.2 | 0 | -14.12 | -0.06 | -39.38 | -0.06 | -42.6 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | 21.39 | -1.41 | 21.39 | -14.41 | 0 | 0 | 0.18 | 0 | -0.29 | 28.36 | -0.18 | 28.48 | -0.41 | 28.65 | -0.94 | -0.06 | -27.89 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
95.232 | -0.06 | -12.25 | -0.06 | -36.97 | -0.18 | -39.55 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | -0.94 | 21.39 | -12.6 | 0.06 | -39.85 | 0.18 | 0 | -0.35 | 28.48 | -0.18 | 29.01 | -0.23 | 28.59 | -0.18 | -0.06 | -25.72 | 0 | -20.51 | -0.06 | 0 | 0 | 0 | -0.06 | 0
95.264 | 0.06 | -10.37 | -0.06 | -35.1 | -0.06 | -37.38 | -0.06 | 0 | 0.06 | 21.45 | 0.53 | 21.45 | -0.35 | 21.5 | -10.72 | -0.18 | -65.04 | 0.18 | 0 | -0.18 | 28.54 | -0.18 | 28.36 | -0.29 | 28.48 | -0.18 | -0.06 | -24.38 | 0 | -91.7 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0
95.296 | 4.63 | -9.43 | -0.06 | -32.93 | -0.06 | -35.86 | 0 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -0.29 | 21.33 | -9.2 | -0.06 | -44.59 | 0.18 | 0 | -0.18 | 28.48 | -0.18 | 28.48 | -0.23 | 28.48 | 0 | -0.06 | -22.56 | 0 | -58.65 | -0.06 | -83.5 | -0.06 | -1.05 | 0.06 | 0
95.328 | 10.72 | -7.79 | -0.06 | -30.94 | -0.06 | -32.58 | 0.06 | 0 | -0.06 | 21.33 | 0.35 | 21.33 | -0.06 | 21.33 | -7.44 | -0.06 | -28.24 | 0.18 | 0 | -0.23 | 0.23 | -0.18 | 28.48 | -0.23 | 28.42 | 0.06 | 3.87 | -20.74 | 0 | -48.52 | -0.06 | -70.61 | -0.18 | -80.39 | 0 | -16.76
95.36 | 16.17 | -6.39 | 0 | -29.06 | -0.18 | -30.59 | 0 | 0 | 0 | 21.27 | 0.29 | 21.39 | -0.06 | 21.39 | -5.8 | 0 | -30.29 | 0.12 | 0 | -0.18 | 0.12 | -0.18 | 23.15 | -0.18 | 28.54 | 0.23 | 12.31 | -17.99 | 0 | -45 | -0.06 | -39.79 | -0.06 | -69.14 | 0 | -89.01
95.392 | 21.68 | -4.51 | -0.18 | -26.84 | -0.06 | -28.59 | -0.18 | 0 | 0.12 | 21.27 | 0.29 | 21.45 | 0 | 21.39 | -3.75 | 0 | -35.57 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0.12 | -0.18 | 28.3 | 0.29 | 22.5 | -16.29 | 0 | -37.44 | -0.18 | -30.59 | 0 | -34.22 | -0.06 | -52.15
95.424 | 21.5 | -3.16 | -0.06 | -24.84 | -0.06 | -26.9 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.06 | 21.27 | -2.46 | 0 | -34.04 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.35 | 29.01 | -14.06 | -0.06 | -37.56 | -0.06 | -40.2 | -0.06 | -30.76 | -0.06 | -30.94
95.456 | 21.45 | -2.46 | -0.12 | -22.21 | -0.06 | -24.84 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.06 | 21.39 | -1.82 | -0.06 | -30.82 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.24 | 0.35 | 28.65 | -12.77 | -0.06 | -37.91 | -0.06 | -42.31 | -0.18 | -40.14 | 0 | -32.46
95.488 | 21.45 | -1.17 | -0.06 | -20.27 | -0.06 | -22.44 | -0.06 | 0.06 | 0.06 | 0.12 | 0.23 | 21.33 | 0.12 | 21.27 | -1 | 0 | -28.3 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.29 | 28.42 | -10.37 | 0 | -37.33 | 0 | -38.32 | 0 | -42.13 | 0 | -41.54
95.52 | 21.45 | -0.59 | -0.06 | -18.52 | -0.18 | -20.63 | -0.06 | -50.51 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 21.39 | -0.7 | -0.18 | -27.01 | 0.18 | 0 | -0.23 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.35 | 28.42 | -7.68 | 0 | -36.04 | -0.18 | -34.22 | 0 | -38.91 | 0 | -41.66
95.552 | 21.39 | -0.29 | -0.06 | -16.29 | -0.06 | -17.81 | -0.06 | -57.6 | -0.06 | 0 | 0.18 | 21.33 | 0.29 | 21.39 | -0.41 | -0.06 | -26.13 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 28.42 | -5.86 | -0.06 | -33.52 | -0.18 | -33.28 | 0 | -33.81 | -0.06 | -36.39
95.584 | 21.45 | -0.18 | -0.06 | -14.82 | -0.06 | -16.7 | 0 | -43.6 | -0.06 | 0 | 0.18 | 21.33 | 0.29 | 21.33 | -0.06 | -0.06 | -23.56 | 0.18 | -54.32 | -0.18 | 0 | 0 | 0 | -0.06 | 28.48 | 0.35 | 28.48 | -3.69 | -0.06 | -31.64 | -0.18 | -32.46 | 0 | -32.99 | 0 | -33.34
95.616 | 21.45 | -0.18 | 0 | -13.01 | 0 | -14.94 | 0.06 | -37.97 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -0.06 | 0 | -21.91 | 0.18 | -66.8 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.18 | 28.89 | -2.23 | 5.68 | -29.77 | -0.06 | -31.82 | -0.06 | -32.52 | 0 | -32.93
95.648 | 21.45 | -0.18 | -0.06 | -11.48 | -0.06 | -13.13 | 0 | -33.34 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | 0.06 | 0.06 | -20.16 | 0.18 | -46.58 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.18 | 28.42 | -1.46 | 13.71 | -28.01 | -0.06 | -29.3 | -0.06 | -31.82 | -0.18 | -32.64
95.68 | 21.45 | -0.29 | 0.29 | -10.08 | -0.06 | -11.48 | -0.06 | -32.81 | 0 | 0 | 0.06 | 21.45 | 0.23 | 21.27 | 0 | -0.06 | -19.45 | 0.18 | -44.88 | -0.06 | 0 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.54 | -0.76 | 21.27 | -26.37 | 0 | -27.54 | 0 | -29.3 | 0 | -31.76
95.712 | 21.39 | -0.35 | 9.96 | -8.56 | -0.06 | -9.9 | -0.18 | -32.81 | -0.06 | 0 | 0.12 | 21.45 | 0.23 | 21.27 | 0.06 | 2.11 | -17.4 | 0.12 | -43.89 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.65 | 0.18 | 28.48 | -0.18 | 28.83 | -24.38 | 0 | -26.08 | -0.06 | -27.6 | 0 | -28.48
95.744 | 21.45 | -0.29 | 15.35 | -7.03 | -0.06 | -8.38 | -0.06 | -31.58 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.27 | 0.29 | 7.68 | -15 | 0.18 | -40.9 | -0.06 | -29.71 | -0.18 | 0 | -0.06 | 28.48 | 0.12 | 28.24 | -0.18 | 28.48 | -21.91 | -0.18 | -24.61 | 0 | -26.25 | 0 | -26.6
95.776 | 21.45 | -0.29 | 21.09 | -5.45 | -0.06 | -7.21 | -0.06 | -29.77 | -0.18 | 0 | -0.06 | 21.27 | 0.18 | 21.15 | 0.35 | 13.24 | -13.24 | 0.12 | -39.85 | 0 | -76.17 | 0 | -0.06 | 0 | 28.42 | 0.12 | 28.48 | 0.06 | 28.65 | -20.27 | -0.18 | -22.91 | -0.06 | -24.79 | 0 | -25.9
95.808 | 21.45 | -0.35 | 21.62 | -3.93 | -0.06 | -5.45 | 0 | -27.6 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.33 | 0.23 | 18.87 | -11.72 | 0.12 | -37.85 | 0 | -52.15 | 0 | -62.35 | -0.06 | 23.09 | 0 | 28.48 | 0.18 | 28.65 | -18.34 | -0.06 | -21.27 | 0 | -22.79 | 0 | -24.32
95.84 | 21.45 | -0.29 | 21.5 | -2.7 | 4.92 | -4.22 | -0.06 | -25.37 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 21.45 | -10.08 | 0.18 | -35.98 | -0.18 | -45.7 | 0 | -62.52 | -0.18 | 0.18 | 0 | 28.48 | 0.29 | 28.54 | -16.7 | -0.06 | -19.98 | -0.06 | -21.39 | -0.06 | -22.56
95.872 | 21.45 | -0.29 | 21.5 | -1.52 | 11.13 | -2.4 | -0.06 | -23.03 | -0.06 | 0 | -0.06 | 0.12 | 0 | 21.39 | 0.35 | 21.33 | -8.2 | 0 | -33.22 | 0 | -43.83 | 0 | -46.06 | 0 | 0 | 0.06 | 28.59 | 0.18 | 28.48 | -14.24 | 0 | -17.75 | 0 | -19.4 | 0 | -20.68
95.904 | 21.45 | -0.29 | 21.45 | -0.7 | 16.41 | -1.52 | 4.16 | -22.21 | -0.06 | -38.79 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.33 | -6.15 | -0.06 | -31.06 | 0 | -42.07 | 0 | -45.18 | -0.18 | 0 | 0 | 28.65 | 0.29 | 28.48 | -12.07 | -0.18 | -16.41 | 0 | -17.75 | 0 | -19.16
95.936 | 21.45 | -0.29 | 21.5 | -0.41 | 21.8 | -0.94 | 10.31 | -19.45 | -0.18 | -70.55 | -0.18 | 0 | -0.06 | 19.63 | 0.29 | 21.27 | -4.63 | -0.06 | -29.83 | 0 | -40.43 | 0 | -42.54 | -0.06 | 0 | 0.12 | 28.65 | 0.41 | 28.48 | -10.37 | 5.04 | -14.59 | -0.06 | -16.23 | 0 | -17.64
95.968 | 21.45 | -0.23 | 21.45 | -0.23 | 21.5 | -0.47 | 15.59 | -17.4 | -0.06 | -47.99 | -0.06 | 0.06 | 0.06 | 0.12 | 0.23 | 21.33 | -3.16 | -0.06 | -27.6 | -0.06 | -37.85 | -0.06 | -41.54 | 0 | 0 | -0.06 | 28.48 | 0.35 | 28.42 | -8.32 | 13.24 | -12.72 | 0 | -14.53 | 0.64 | -15.94
96 | 21.45 | -0.18 | 21.45 | -0.18 | 21.5 | -0.18 | 21.62 | -15.23 | 0 | -42.13 | -0.06 | 0 | 0 | 0 | 0.18 | 21.33 | -2.29 | -0.06 | -25.55 | -0.06 | -36.21 | 0 | -39.44 | -0.18 | 0 | 0 | 28.42 | 0.23 | 28.36 | -5.92 | 20.74 | -11.31 | 2.46 | -13.36 | 10.66 | -14.3
96.032 | 21.39 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.39 | -12.6 | 0 | -33.87 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.33 | -1.29 | -0.06 | -24.08 | -0.06 | -34.1 | -0.06 | -36.62 | -0.06 | 0 | -0.06 | 28.59 | 0.23 | 28.48 | -3.22 | 28.83 | -9.2 | 12.77 | -10.78 | 20.86 | -12.36
96.064 | 21.45 | -0.18 | 21.45 | -0.23 | 21.5 | -0.35 | 21.39 | -10.61 | -0.18 | -31.47 | -0.06 | 0 | 0 | 0 | 0.29 | 21.33 | -0.41 | -0.06 | -20.92 | 0.06 | -32.76 | 0 | -34.75 | -0.06 | 0 | 0 | 28.54 | 0.12 | 28.83 | -2.23 | 28.77 | -7.62 | 20.33 | -9.79 | 28.89 | -10.55
96.096 | 21.45 | -0.18 | 21.45 | -0.53 | 21.45 | -0.23 | 21.39 | -8.85 | -0.06 | -32.4 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.33 | -0.41 | -0.06 | -19.34 | -0.06 | -30.53 | 0 | -33.63 | 0 | 0 | -0.18 | 28.54 | 0.12 | 28.71 | -1 | 28.48 | -5.74 | 27.95 | -7.79 | 28.65 | -9.2
96.128 | 21.45 | -0.18 | 21.45 | -0.47 | 21.45 | -0.23 | 21.33 | -6.8 | -0.06 | -32.81 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.15 | -0.18 | 0.06 | -18.05 | -0.06 | -28.59 | 0 | -31.47 | 0 | 0 | -0.18 | 28.42 | 0.23 | 28.48 | -0.7 | 28.59 | -4.45 | 28.48 | -6.27 | 28.42 | -7.38
96.16 | 21.39 | -0.18 | 21.45 | -0.41 | 21.39 | -0.23 | 21.39 | -3.98 | -0.18 | -30.94 | -0.06 | 0 | 0 | 0 | 0.12 | 21.39 | -0.06 | 0.06 | -15.65 | -0.06 | -27.07 | 0 | -29.94 | -0.06 | 0 | -0.06 | 28.65 | 0.12 | 28.48 | -0.59 | 28.48 | -2.93 | 28.48 | -4.69 | 28.65 | -5.63
96.192 | 0.23 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | -1.99 | -0.06 | -28.13 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 0.06 | -14.24 | 0 | -24.08 | 0 | -27.07 | -0.06 | 0 | 0 | 28.42 | 0 | 28.13 | 0.06 | 28.48 | -2.11 | 28.65 | -2.75 | 28.59 | -3.87
96.224 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.29 | 21.39 | -1.52 | -0.06 | -26.25 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.39 | 0.23 | 0.06 | -12.95 | 0.06 | -21.91 | -0.06 | -25.9 | 0 | -0.23 | -0.18 | 0.29 | 0.06 | 28.54 | -0.06 | 28.42 | -0.76 | 28.18 | -2.17 | 28.65 | -2.58
96.256 | 0 | -0.18 | 21.5 | -0.18 | 21.39 | -0.29 | 21.39 | -0.59 | -0.06 | -24.08 | -0.06 | 0 | 0.06 | 0 | 0.23 | 21.39 | 0.18 | 3.34 | -11.6 | 0 | -20.16 | -0.06 | -22.97 | -0.06 | -82.03 | -0.06 | 0.06 | 0.06 | 28.48 | 0.23 | 28.48 | -0.64 | 28.65 | -1.17 | 28.42 | -1.88
96.288 | 0.06 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | -0.29 | -0.18 | -22.15 | -0.06 | -25.67 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 9.55 | -10.02 | -0.06 | -18.52 | -0.06 | -21.33 | -0.06 | -67.5 | -0.06 | 0 | 0.06 | 28.59 | 0.12 | 28.3 | -0.59 | 28.48 | -0.76 | 28.59 | -1.11
96.32 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | 0.06 | -0.06 | -20.74 | -0.06 | -73.36 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 14.71 | -8.91 | -0.06 | -16.58 | -0.06 | -19.22 | -0.18 | -48.87 | -0.18 | 0.06 | 0 | 28.71 | 0.23 | 28.48 | -0.23 | 28.65 | -0.53 | 28.48 | -0.7
96.352 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | 0.29 | -0.18 | -17.52 | 0 | -45 | 0.12 | 0.06 | 0.06 | 21.33 | 0.18 | 21.62 | -6.97 | 0.06 | -14.94 | 0.06 | -17.23 | 0 | -45.29 | -0.18 | 0 | -0.06 | 28.36 | 0.29 | 28.48 | 0 | 28.42 | -0.18 | 28.48 | -0.64
96.384 | 0.06 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.21 | 0.29 | -0.06 | -15.59 | 0 | -40.72 | 0.06 | -58.19 | 0.12 | 21.21 | 0.23 | 21.45 | -6.09 | 0 | -12.95 | 0 | -15.41 | 0 | -39.08 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.48 | -0.06 | 28.42 | -0.06 | 28.42 | -0.18
96.416 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.53 | -0.18 | -13.71 | -0.06 | -34.16 | -0.06 | -61.06 | 0.06 | 21.39 | 0.23 | 21.27 | -4.75 | 0.06 | -11.78 | 0 | -14.12 | 0 | -37.15 | -0.06 | 0.06 | -0.18 | 28.48 | 0.29 | 28.48 | 0.12 | 28.65 | 0 | 28.54 | -0.18
96.448 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.41 | 0 | -11.78 | -0.06 | -31.7 | -0.06 | -38.03 | 0.12 | 21.39 | 0.23 | 21.39 | -4.16 | 0 | -10.55 | 2.46 | -12.36 | -0.06 | -37.56 | -0.18 | 0 | 0 | 28.48 | 0.29 | 28.48 | 0.12 | 28.48 | 0.12 | 28.89 | 0
96.48 | 0 | -0.06 | 21.5 | -0.29 | 21.5 | -0.18 | 21.39 | 0.35 | -0.06 | -10.02 | -0.06 | -33.22 | -0.06 | -27.6 | 0.06 | 21.33 | 0.18 | 21.27 | -3.16 | -0.06 | -9.2 | 10.31 | -10.96 | 0.06 | -38.56 | 0 | 0 | 0 | 28.42 | 0.29 | 28.48 | 0 | 28.36 | 0.06 | 28.59 | 0.12
96.512 | 0 | -0.06 | 21.45 | -0.23 | 21.39 | -0.29 | 21.33 | 0.35 | -0.06 | -6.97 | -0.06 | -32.29 | -0.06 | -34.16 | 0 | 21.39 | 0.18 | 21.39 | -2.58 | 5.63 | -7.44 | 20.33 | -9.02 | -0.06 | -37.38 | 0 | 0 | 0 | 0.18 | 0.18 | 28.42 | 0.06 | 28.48 | 0.23 | 28.65 | 0
96.544 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.47 | 5.33 | -5.1 | -0.06 | -31 | -0.06 | -35.39 | 0.06 | 21.39 | 0.23 | 21.09 | -1.99 | 13.71 | -5.98 | 28.01 | -8.09 | 0 | -34.4 | -0.06 | 0 | 0 | 0 | 0.23 | 28.59 | 0.06 | 28.48 | 0.12 | 28.48 | 0.18
96.576 | 0 | 0 | 0.23 | -0.06 | 21.39 | -0.18 | 21.39 | 0.35 | 11.37 | -2.87 | -0.06 | -28.54 | -0.18 | -32.93 | 0.18 | 21.33 | 0.23 | 21.39 | -1.52 | 21.21 | -5.39 | 28.83 | -6.8 | -0.06 | -32.23 | 0 | 0 | -0.18 | 0 | 0.06 | 28.59 | 0.06 | 28.59 | 0.18 | 28.36 | 0.12
96.608 | -20.39 | -0.06 | 0.06 | -0.06 | 21.45 | -0.06 | 21.33 | 0.35 | 16.7 | -1.99 | 0 | -26.72 | -0.06 | -30.06 | 0.06 | 19.4 | 0.29 | 21.39 | -1.41 | 28.95 | -4.28 | 28.71 | -5.63 | 5.68 | -30.7 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.48 | 0.18 | 28.48 | 0.23 | 28.95 | 0.18
96.64 | -63.69 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.74 | -0.94 | 0 | -24.32 | -0.18 | -28.07 | 0 | 0.12 | 0.23 | 21.27 | -1.05 | 28.54 | -3.46 | 28.71 | -4.45 | 13.89 | -28.65 | -0.06 | -9.43 | 0 | 0 | 0.06 | 28.65 | 0.12 | 28.48 | 0.23 | 28.48 | 0.23
96.672 | -42.25 | -0.06 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.41 | 21.39 | -0.18 | -0.06 | -21.74 | -0.06 | -26.6 | 0.06 | 0 | 0.12 | 21.33 | -0.64 | 28.36 | -2.7 | 28.83 | -3.69 | 24.02 | -26.13 | 0 | -84.38 | 0 | 0 | 0 | 28.48 | 0.12 | 28.48 | 0.23 | 28.42 | 0.18
96.704 | -40.08 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 21.39 | -0.18 | -0.06 | -19.98 | 0 | -25.55 | 0 | 0 | 0.12 | 21.33 | -0.53 | 28.83 | -2.23 | 28.59 | -2.81 | 28.83 | -24.84 | -0.18 | -53.21 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.48 | 0.23 | 28.48 | 0.18
96.736 | -38.32 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 21.39 | 0.23 | 21.33 | 0.12 | -0.06 | -17.75 | -0.18 | -23.2 | 0 | 0 | 0.18 | 21.39 | -0.64 | 28.13 | -1.7 | 28.24 | -2.23 | 28.48 | -23.38 | 0 | -48.58 | -0.18 | 0 | 0.06 | 28.48 | 0.12 | 28.42 | 0.23 | 28.48 | 0.18
96.768 | -35.51 | -0.18 | 0 | -0.18 | 0.06 | -0.06 | 21.27 | 0.23 | 21.33 | 0.29 | 0 | -16.17 | 0 | -21.39 | 0 | 0 | 0.23 | 21.39 | -0.53 | 28.24 | -1.29 | 28.54 | -1.76 | 28.65 | -20.92 | 0 | -43.01 | -0.06 | 0 | -0.06 | 28.13 | 0 | 28.71 | 0.18 | 28.42 | 0.23
96.8 | -34.4 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0.23 | 0.18 | 21.39 | 0.29 | -0.06 | -14.41 | 0 | -19.86 | -0.06 | 0 | 0.29 | 21.39 | -0.47 | 28.59 | -1 | 28.59 | -1.23 | 28.54 | -19.16 | 0 | -37.44 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.48 | 0.23 | 28.48 | 0.23
96.832 | -31.11 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.47 | -0.06 | -11.48 | -0.18 | -17.87 | -0.06 | 0 | 0.12 | 21.39 | -0.41 | 28.65 | -0.64 | 28.3 | -0.64 | 28.71 | -16.64 | -0.06 | -36.86 | -0.18 | 0 | 0 | 0.18 | -0.06 | 28.48 | 0.12 | 28.42 | 0.18
96.864 | -29.71 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.41 | -0.06 | -9.67 | 0 | -16.47 | -0.06 | 0 | 0.12 | 21.33 | -0.41 | 28.54 | -0.64 | 28.65 | -0.53 | 28.48 | -14.77 | -0.06 | -38.32 | 0 | 0 | -0.18 | 0 | 0 | 28.48 | 0.06 | 0.29 | 0.18
96.896 | -27.42 | 0 | 0 | -0.06 | 0 | -0.18 | 0.06 | 0.12 | 21.39 | 0.53 | -0.06 | -7.09 | -0.06 | -14.53 | -0.06 | 0 | 0.23 | 21.39 | -0.47 | 28.65 | -0.41 | 28.59 | -0.35 | 28.65 | -12.89 | -0.06 | -37.68 | -0.23 | 0 | 0.06 | 0 | -0.06 | 0.29 | 0 | 0.12 | 0.18
96.928 | -25.78 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.47 | 3.34 | -5.16 | -0.18 | -12.95 | -0.18 | 0 | 0.12 | 21.39 | -0.29 | 28.48 | -0.23 | 28.48 | -0.23 | 28.54 | -10.66 | -0.18 | -35.92 | -0.23 | -40.61 | -0.18 | 0.06 | 0 | 0.12 | 0.23 | 0 | 0.06
96.96 | -23.67 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 9.73 | -2.81 | -0.06 | -11.31 | -0.06 | 0 | 0.12 | 21.39 | -0.29 | 28.54 | -0.29 | 28.48 | -0.18 | 28.48 | -8.67 | 0 | -34.45 | 0.06 | -86.78 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.23
96.992 | -20.86 | -0.06 | -11.78 | 0 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.41 | 16.7 | -1.46 | -0.06 | -9.02 | 0 | 0 | 0.18 | 21.39 | -0.29 | 28.89 | -0.35 | 28.77 | -0.35 | 28.54 | -5.92 | -0.06 | -31.29 | -0.06 | -51.68 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 0 | 0.23
97.024 | -18.46 | -0.06 | -62.11 | -0.18 | 0 | 0 | 0 | -0.06 | 21.27 | 0.47 | 21.74 | -0.94 | 0 | -7.21 | 0 | 0 | 0.12 | 21.39 | -0.29 | 28.65 | -0.18 | 28.36 | -0.29 | 28.36 | -3.87 | 8.85 | -29.47 | -0.06 | -46.7 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.18
97.056 | -16.88 | 0 | -48.69 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.45 | -0.59 | -0.06 | -5.39 | 0 | -51.86 | 0.18 | 21.33 | -0.23 | 28.83 | -0.35 | 28.54 | -0.35 | 28.42 | -2.52 | 16.06 | -27.83 | -0.06 | -41.66 | 0 | 0 | 0 | 0 | 0 | 0 | 0.18
97.088 | -15.35 | -0.06 | -40.43 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.47 | 21.39 | -0.06 | -0.06 | -3.93 | -0.06 | -60.65 | 0.18 | 21.33 | -0.23 | 28.42 | -0.23 | 28.65 | -0.23 | 28.54 | -1.64 | 23.67 | -26.37 | -0.18 | -36.45 | -0.18 | 0 | -0.12 | 0 | 0 | 0 | 0
97.12 | -13.71 | -0.18 | -39.32 | -0.06 | 0 | 0 | 0 | 0 | 21.39 | 0.35 | 21.39 | -0.06 | 6.86 | -2.64 | -0.06 | -40.26 | 0.23 | 21.39 | -0.29 | 28.48 | -0.18 | 28.42 | -0.18 | 28.48 | -1 | 28.77 | -24.43 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.12
97.152 | -11.31 | -0.06 | -35.98 | -0.06 | -24.67 | 0 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | 0.23 | 14.3 | -1.7 | -0.06 | -28.13 | 0.06 | 0.18 | -0.29 | 28.77 | -0.23 | 28.42 | -0.18 | 28.42 | -0.18 | 28.65 | -22.68 | -0.18 | -38.32 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
97.184 | -10.37 | -0.06 | -34.22 | 0 | -62.17 | -0.06 | 0 | -0.18 | 21.45 | 0.23 | 21.39 | 0.23 | 20.04 | -0.82 | 0 | -33.34 | 0.06 | 0.06 | -0.18 | 28.36 | -0.18 | 29.12 | -0.18 | 28.59 | 0 | 28.48 | -20.16 | -0.18 | -36.68 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.06
97.216 | -9.08 | 0 | -32.17 | -0.18 | -46 | -0.06 | -23.38 | -0.06 | 21.39 | 0.23 | 21.39 | 0.35 | 21.56 | -0.59 | -0.06 | -35.57 | 0.06 | 0 | -0.18 | 28.59 | -0.23 | 28.48 | -0.18 | 28.36 | 0 | 28.54 | -18.22 | -0.18 | -34.98 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | 0.18
97.248 | -8.03 | -0.06 | -30.53 | -0.06 | -40.78 | -0.06 | -77.29 | 0 | 21.33 | 0.06 | 21.39 | 0.47 | 21.39 | -0.18 | -0.06 | -33.52 | 0.06 | 0.06 | -0.18 | 28.54 | -0.23 | 28.3 | -0.18 | 28.59 | 0.06 | 28.42 | -16.29 | -0.06 | -33.17 | -0.18 | -25.84 | -0.06 | 0 | 0 | 0 | 0.12
97.28 | -6.45 | 0 | -28.77 | -0.06 | -37.97 | 0 | -51.74 | -0.06 | 21.39 | 0.12 | 21.39 | 0.35 | 21.39 | -0.23 | 0 | -30.41 | 0.12 | 0 | -0.18 | 28.48 | -0.18 | 28.65 | -0.18 | 28.42 | 0.23 | 28.48 | -14.53 | -0.06 | -31.47 | -0.18 | -89.83 | -0.18 | 0 | 0 | -0.82 | 0.12
97.312 | -4.04 | -0.06 | -25.61 | -0.06 | -35.74 | -0.06 | -40.72 | -0.06 | 21.45 | 0.06 | 21.21 | 0.41 | 21.39 | -0.06 | 0 | -27.54 | 0.12 | 0 | -0.18 | 28.59 | -0.06 | 28.83 | -0.06 | 28.71 | 0.29 | 28.48 | -11.95 | -0.06 | -29.24 | -0.06 | -50.74 | -0.06 | -16.99 | -0.06 | -88.71 | 0.06
97.344 | -3.28 | 0 | -23.15 | 0 | -33.46 | 0 | -35.98 | -0.06 | 21.56 | 0.18 | 21.39 | 0.41 | 21.21 | 0.06 | 0 | -26.72 | 0.12 | 0 | -0.18 | 28.48 | -0.18 | 0.29 | -0.18 | 28.36 | 0.35 | 28.42 | -9.79 | -0.06 | -27.19 | -0.18 | -30.06 | -0.06 | -89.48 | -0.06 | -59.18 | 0
97.376 | -2.05 | 3.22 | -21.74 | -0.06 | -31.64 | -0.06 | -31.93 | -0.06 | 21.5 | 0.06 | 21.33 | 0.41 | 21.39 | 0.12 | 6.97 | -25.55 | 0.06 | 0 | -0.18 | 28.48 | -0.18 | 0.12 | -0.23 | 28.65 | 0.53 | 29.01 | -7.79 | -0.06 | -25.67 | -0.18 | -33.22 | -0.06 | -60.82 | -0.06 | -33.63 | -0.06
97.408 | -1.23 | 9.49 | -19.98 | -0.06 | -30 | 0 | -32.7 | 0 | 21.39 | -0.06 | 21.39 | 0.29 | 21.39 | 0.18 | 12.72 | -23.67 | 0.06 | 0 | -0.18 | 0.29 | -0.06 | 0 | -0.06 | 28.42 | 0.47 | 28.59 | -5.16 | -0.06 | -23.44 | -0.18 | -42.25 | -0.06 | -34.45 | 0.06 | -30.76 | 0
97.44 | -0.59 | 14.82 | -17.46 | -0.06 | -26.84 | -0.06 | -32.34 | -0.18 | 0.12 | 0 | 21.33 | 0.41 | 21.33 | 0.18 | 18.22 | -21.68 | 0.12 | 0 | -0.23 | 0.12 | -0.18 | 0 | -0.18 | 28.54 | 0.41 | 28.48 | -3.57 | 0 | -22.33 | -0.18 | -41.9 | -0.18 | -30.06 | -0.18 | -40.9 | 0
97.472 | -0.23 | 21.74 | -15.65 | -0.06 | -25.37 | -0.06 | -31.29 | -0.06 | 0 | -0.06 | 21.21 | 0.29 | 21.15 | 0.18 | 21.56 | -19.98 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.41 | 28.48 | -1.82 | 2.99 | -19.22 | -0.06 | -35.74 | -0.18 | -41.84 | -0.06 | -41.43 | 0
97.504 | -0.06 | 21.5 | -13.83 | 0 | -23.56 | -0.06 | -29.53 | 0 | 0 | -0.06 | 21.15 | 0.29 | 21.39 | 0.23 | 21.39 | -18.28 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | -0.06 | 0.29 | 0.35 | 28.54 | -1.17 | 11.37 | -16.99 | -0.06 | -33.22 | -0.18 | -41.78 | -0.06 | -37.21 | 0
97.536 | -0.06 | 21.45 | -12.48 | -0.06 | -21.74 | -0.06 | -27.01 | -0.06 | 0 | -0.06 | 21.27 | 0.23 | 21.39 | 0.18 | 21.27 | -16.7 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.29 | 28.18 | -0.7 | 18.81 | -15.82 | -0.06 | -32.7 | -0.06 | -37.74 | 0 | -33.75 | 0
97.568 | -0.23 | 21.45 | -11.31 | -0.06 | -19.16 | 0.06 | -24.9 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.29 | 21.33 | -14.94 | 0 | -7.85 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.41 | 28.48 | -0.23 | 26.43 | -13.89 | -0.06 | -32.93 | -0.18 | -33.87 | 0 | -33.22 | 0
97.6 | -0.29 | 21.5 | -9.32 | 0 | -16.7 | -0.06 | -22.62 | 0 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.23 | 21.39 | -13.18 | 0.12 | -73.13 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | 0.29 | 28.42 | -0.06 | 28.54 | -11.78 | -0.06 | -31.47 | 0 | -33.05 | -0.18 | -32.81 | -0.06
97.632 | -0.47 | 21.45 | -7.44 | -0.06 | -15.35 | -0.06 | -20.16 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.33 | 0.23 | 21.5 | -11.02 | -0.06 | -45.35 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.23 | 28.36 | 0.12 | 28.42 | -9.43 | 0 | -28.59 | -0.06 | -32.76 | -0.18 | -30.94 | 0
97.664 | -0.29 | 21.45 | -6.56 | -0.06 | -13.71 | 0 | -18.34 | 0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | 0.23 | 21.33 | -9.2 | -0.06 | -29.47 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 28.48 | 0.18 | 28.42 | -7.38 | -0.06 | -26.66 | -0.06 | -31.64 | -0.18 | -28.65 | -0.06
97.696 | -0.35 | 21.5 | -4.86 | -0.18 | -12.6 | 0 | -16.29 | 0.06 | 0 | -0.06 | 21.33 | 0 | 21.39 | 0.12 | 21.39 | -7.79 | -0.06 | -29.53 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 28.42 | 0.41 | 28.65 | -4.98 | -0.06 | -25.37 | -0.18 | -29.42 | 0 | -27.13 | -0.06
97.728 | -0.23 | 21.45 | -3.57 | 0 | -10.66 | -0.06 | -14.36 | -0.06 | 0 | -0.06 | 21.15 | -0.06 | 21.62 | 0.18 | 21.33 | -5.39 | 0 | -34.69 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 28.71 | 0.35 | 28.42 | -3.16 | -0.06 | -24.08 | -0.06 | -27.42 | -0.18 | -26.31 | -0.06
97.76 | -0.35 | 21.45 | -2.29 | -0.06 | -8.91 | -0.06 | -12.13 | -0.06 | 0 | -0.06 | 21.33 | 0 | 21.27 | 0.23 | 21.39 | -4.04 | -0.06 | -35.22 | -0.06 | 0 | -0.06 | -0.18 | -0.18 | 0 | 0 | 28.54 | 0.53 | 28.83 | -2.17 | -0.06 | -22.68 | -0.06 | -26.13 | 0 | -24.84 | 0
97.792 | -0.18 | 21.45 | -1.35 | 7.03 | -7.68 | 1.99 | -9.43 | -0.06 | 0 | 0 | 21.33 | -0.06 | 21.33 | 0.18 | 21.33 | -2.29 | 0.06 | -31.17 | -0.18 | 0 | 0 | -78.87 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.42 | -1.11 | -0.06 | -20.57 | -0.06 | -24.79 | -0.06 | -22.56 | 0
97.824 | -0.29 | 21.45 | -0.7 | 12.72 | -5.92 | 6.27 | -7.38 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.27 | 0.06 | 21.39 | -1.52 | 0 | -28.48 | -0.18 | -25.61 | -0.06 | -56.02 | -0.06 | 0 | 0.06 | 28.59 | 0.35 | 28.48 | -0.23 | 0 | -18.69 | -0.06 | -22.62 | -0.06 | -21.15 | -0.06
97.856 | -0.35 | 21.45 | -0.29 | 18.34 | -4.45 | 12.07 | -4.92 | -0.06 | -36.74 | -0.06 | 0.06 | -0.18 | 21.39 | 0.18 | 21.21 | -1.05 | 0.12 | -27.6 | -0.18 | -74.01 | 0 | -45.53 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.42 | -0.18 | -0.06 | -17.29 | -0.06 | -21.09 | -0.06 | -19.75 | -0.06
97.888 | -0.23 | 21.39 | -0.18 | 21.68 | -3.16 | 17.58 | -3.52 | 0.06 | -69.26 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.45 | -0.64 | -0.06 | -26.6 | -0.06 | -52.79 | -0.06 | -45.53 | 0 | 0 | 0 | 28.65 | 0.29 | 28.48 | 0.06 | -0.18 | -15.53 | -0.06 | -19.51 | 0.06 | -18.46 | 0
97.92 | -0.18 | 21.39 | -0.18 | 21.5 | -1.7 | 21.74 | -1.88 | 0 | -48.69 | -0.06 | 0.06 | -0.06 | 21.33 | 0.12 | 21.09 | -0.59 | 3.34 | -24.61 | -0.06 | -45.47 | 0 | -43.07 | -0.06 | -31.29 | 0 | 0.23 | 0.35 | 28.48 | 0.18 | -0.06 | -14.24 | -0.06 | -18.4 | -0.06 | -17.05 | 0
97.952 | -0.18 | 21.39 | -0.35 | 21.45 | -1.23 | 21.5 | -0.7 | 0.06 | -39.44 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0 | 11.31 | -21.68 | -0.18 | -43.89 | 0 | -40.67 | 0 | -66.62 | -0.06 | 0 | 0.23 | 28.42 | 0.35 | 11.66 | -11.84 | 0.59 | -16.17 | -0.18 | -14.47 | -0.18
97.984 | -0.23 | 21.45 | -0.35 | 21.45 | -0.47 | 21.45 | -0.59 | 0 | -33.81 | 0.06 | 0 | 0 | 10.37 | 0.06 | 21.33 | -0.06 | 16.7 | -20.74 | -0.06 | -41.25 | -0.06 | -38.91 | 0 | -47.64 | 0 | 0 | 0.35 | 28.42 | 0.41 | 19.28 | -10.37 | 10.2 | -14.3 | 2.34 | -12.72 | 0
98.016 | -0.18 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | 0 | -0.06 | -31.88 | 0.06 | 0 | 0.12 | 0.06 | 0.06 | 21.39 | -0.06 | 21.74 | -19.1 | -0.06 | -40.55 | 0 | -36.45 | 0 | -45.35 | -0.06 | 0 | 0.06 | 28.54 | 0.41 | 27.07 | -8.67 | 17.7 | -12.89 | 10.08 | -11.02 | 1.46
98.048 | -0.18 | 21.5 | -0.41 | 21.45 | -0.18 | 21.5 | 0.06 | -0.06 | -32.64 | -0.18 | 0 | 0 | 0 | 0.06 | 21.45 | -0.06 | 21.45 | -17.99 | -0.18 | -38.67 | -0.18 | -35.57 | 0 | -45.06 | -0.06 | 0 | 0.12 | 28.42 | 0.47 | 28.54 | -7.09 | 25.37 | -10.78 | 17.4 | -9.79 | 11.19
98.08 | -0.23 | 21.45 | -0.41 | 21.45 | -0.23 | 21.45 | 0.18 | -0.06 | -32.46 | 0 | 0 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.39 | -15.82 | -0.06 | -36.15 | -0.06 | -33.17 | -0.06 | -41.37 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.71 | -5.39 | 28.54 | -9.55 | 25.31 | -8.2 | 18.69
98.112 | -0.18 | 21.45 | -0.35 | 21.45 | -0.29 | 21.45 | 0.47 | 0 | -30.29 | -0.06 | 0 | 0.12 | 0 | 0.12 | 21.39 | 0.18 | 21.39 | -13.71 | -0.06 | -34.75 | 0 | -30.82 | -0.06 | -40.08 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.48 | -3.4 | 28.54 | -7.09 | 28.65 | -5.92 | 28.95
98.144 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.45 | 0.35 | -0.06 | -28.01 | 0 | 0 | 0 | 0 | 0 | 21.39 | 0.12 | 21.39 | -12.31 | 0 | -32.05 | 0 | -28.65 | -0.18 | -37.68 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.65 | -2.4 | 28.48 | -5.86 | 28.65 | -4.39 | 28.65
98.176 | -0.18 | 21.45 | -0.29 | 21.39 | -0.35 | 21.39 | 0.53 | 1.99 | -26.49 | 0.06 | 0 | -0.06 | 0 | 0.06 | 21.21 | 0.12 | 21.56 | -10.61 | 0 | -30.76 | -0.06 | -27.01 | -0.06 | -36.1 | -0.06 | 0 | 0.23 | 28.59 | 0.35 | 28.36 | -1.41 | 28.89 | -4.04 | 28.42 | -2.93 | 28.65
98.208 | -0.18 | 21.45 | -0.29 | 21.5 | -0.29 | 21.45 | 0.53 | 6.27 | -23.91 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.45 | -8.56 | -0.18 | -29.06 | -0.06 | -25.37 | 0 | -34.28 | -0.06 | 0 | 0 | 28.36 | 0.18 | 28.42 | -1 | 28.48 | -2.75 | 28.48 | -2.29 | 28.54
98.24 | -0.06 | 21.39 | -0.29 | 21.39 | -0.35 | 21.45 | 0.53 | 12.01 | -22.56 | 0 | -11.25 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.45 | -6.8 | 0 | -26.6 | -0.06 | -22.68 | -0.18 | -32.81 | -0.18 | 0 | 0 | 28.36 | 0.23 | 28.59 | -0.7 | 28.42 | -1.82 | 28.54 | -1.41 | 28.65
98.272 | -0.18 | 0.12 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 19.28 | -19.86 | -0.06 | -72.13 | 0 | 0 | -0.06 | 0.06 | 0.18 | 21.33 | -4.75 | 8.44 | -24.26 | -0.06 | -20.39 | -0.06 | -29.94 | 0 | 0 | 0 | 28.54 | 0.12 | 28.48 | -0.29 | 28.48 | -1.05 | 28.42 | -0.59 | 28.65
98.304 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.35 | 21.39 | -9.38 | 0.06 | -32.64 | -0.18 | -69.96 | -0.06 | 0 | 0.18 | 21.33 | -0.47 | 28.65 | -14.47 | 5.04 | -12.07 | -0.06 | -19.51 | -0.18 | -46.47 | -0.18 | 0 | -0.18 | 28.36 | 0.12 | 28.42 | 0 | 28.48 | -0.06 | 28.59
98.336 | -0.06 | 0.06 | -0.23 | 21.45 | -0.23 | 21.39 | 0.29 | 21.27 | -6.86 | -0.06 | -33.05 | -0.06 | -58.01 | 0 | 0 | 0.18 | 21.39 | -0.29 | 28.54 | -13.3 | 13.18 | -10.25 | 0.23 | -17.58 | -0.06 | -40.55 | -0.06 | 0.06 | 0.18 | 28.48 | 0.12 | 28.48 | 0.06 | 28.48 | 0.06 | 28.89
98.368 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.23 | 21.39 | -5.1 | 0 | -31 | -0.06 | -33.93 | -0.06 | 0 | 0.18 | 21.39 | -0.18 | 28.71 | -11.02 | 20.51 | -8.85 | 9.55 | -15.94 | -0.18 | -36.97 | 0 | 0 | 0 | 28.48 | 0.12 | 28.24 | 0.06 | 28.42 | 0.06 | 28.65
98.4 | -0.06 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | 0.18 | 21.39 | -2.52 | -0.06 | -28.48 | -0.06 | -29.06 | -0.06 | 0 | 0.12 | 21.39 | 0 | 28.24 | -9.79 | 29.01 | -7.62 | 19.63 | -14.41 | -0.06 | -37.68 | 0 | 0 | 0.06 | 28.48 | 0.18 | 28.71 | 0.18 | 28.54 | 0.06 | 28.42
98.432 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -1.7 | 0.12 | -26.31 | 0 | -34.98 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.42 | -8.61 | 28.59 | -6.39 | 27.42 | -12.77 | -0.06 | -37.97 | 0 | 0 | 0 | 28.59 | 0.12 | 28.36 | 0.18 | 28.48 | 0.23 | 28.48
98.464 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.18 | 21.21 | -1.11 | 4.69 | -24.08 | -0.06 | -35.57 | -0.06 | 0 | 0.18 | 21.39 | 0 | 28.36 | -7.27 | 28.59 | -5.33 | 28.59 | -11.48 | 0 | -36.86 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.54 | 0.06 | 28.3 | 0.12 | 28.54
98.496 | 0 | 0 | -0.18 | 21.39 | -0.23 | 21.5 | 0.06 | 21.39 | -0.41 | 10.9 | -22.09 | -0.06 | -32.23 | 0 | 0 | 0.12 | 21.39 | 0.12 | 28.54 | -5.92 | 28.24 | -4.34 | 28.54 | -10.14 | -0.06 | -34.63 | -0.06 | 0 | 0.12 | 28.48 | 0.06 | 28.42 | 0.23 | 28.65 | 0.18 | 28.48
98.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0 | 21.45 | -0.06 | 16.29 | -20.22 | -0.06 | -29.06 | -0.06 | 0 | 0.18 | 21.21 | 0.12 | 28.42 | -5.33 | 28.42 | -3.28 | 28.65 | -8.38 | -0.06 | -32.52 | -0.06 | 0 | -0.06 | 28.59 | 0 | 28.42 | 0.23 | 28.24 | 0.18 | 28.01
98.56 | 0 | -28.48 | -0.18 | 0.06 | -0.18 | 13.83 | 0.06 | 21.27 | 0.06 | 21.56 | -17.58 | -0.06 | -27.48 | -0.06 | -68.21 | 0.12 | 21.21 | 0.06 | 28.59 | -4.04 | 28.48 | -2.64 | 28.59 | -6.97 | -0.18 | -30.12 | -0.18 | 0 | 0 | 28.48 | 0 | 28.36 | 0.23 | 28.59 | 0.23 | 28.42
98.592 | -0.18 | -61.88 | -0.18 | 0 | -0.06 | 0.12 | 0 | 21.33 | 0.35 | 21.39 | -15.76 | 0 | -26.54 | -0.18 | -56.02 | 0.18 | 21.33 | 0.18 | 28.65 | -3.4 | 28.54 | -1.99 | 28.54 | -6.09 | -0.18 | -28.07 | 0 | 0 | 0 | 28.48 | 0.12 | 28.48 | 0.06 | 28.3 | 0.12 | 28.3
98.624 | 0 | -44.77 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.35 | 21.33 | -14.18 | 0.06 | -25.02 | -0.06 | -34.81 | 0.06 | 21.39 | 0.23 | 28.42 | -2.64 | 28.71 | -1.46 | 28.42 | -4.86 | -0.06 | -26.31 | 0.06 | 0 | -0.06 | 28.65 | 0 | 28.48 | 0.12 | 28.48 | 0.29 | 28.48
98.656 | -0.06 | -40.31 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.35 | 21.39 | -11.66 | -0.06 | -22.91 | -0.06 | -27.77 | 0.06 | 21.39 | 0.23 | 28.13 | -1.93 | 28.54 | -0.94 | 28.42 | -3.93 | -0.06 | -24.9 | -0.06 | -4.45 | 0 | 28.18 | -0.06 | 28.48 | 0.18 | 28.42 | 0.23 | 28.48
98.688 | -0.18 | -38.97 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.33 | -10.02 | -0.06 | -20.92 | -0.06 | -32.99 | 0.12 | 0.12 | 0.12 | 28.65 | -1.52 | 28.65 | -0.59 | 28.48 | -3.16 | -0.18 | -23.32 | 0 | -89.59 | -0.06 | 3.81 | 0 | 28.48 | 0.18 | 28.48 | 0.23 | 28.42
98.72 | -0.06 | -35.68 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 21.39 | -7.09 | -0.06 | -18.87 | -0.18 | -35.27 | 0.12 | 0 | 0.18 | 28.42 | -1 | 28.54 | -0.53 | 28.59 | -2.4 | -0.18 | -20.92 | 0 | -55.84 | -0.18 | 0.06 | 0 | 0.29 | 0.12 | 28.48 | 0.12 | 28.42
98.752 | -0.06 | -33.4 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.45 | 0.53 | 21.33 | -4.63 | -0.18 | -17.64 | -0.18 | -32.23 | 0.23 | 0 | 0.18 | 28.83 | -1 | 28.59 | -0.29 | 28.42 | -1.7 | -0.18 | -18.87 | 0 | -48.52 | 0 | 0 | 0 | 0.12 | 0.12 | 0.35 | 0.12 | 28.42
98.784 | 0 | -32.46 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.45 | -3.22 | 0 | -15.94 | -0.06 | -29.36 | 0 | 0 | 0.23 | 28.89 | -0.76 | 28.59 | -0.29 | 28.48 | -1.23 | 0.12 | -16.99 | 0 | -44.06 | -0.06 | 0.06 | 0 | 0 | 0.12 | 0.12 | 0.12 | 0.29
98.816 | -0.06 | -29.36 | -0.06 | 0 | -0.06 | 0 | 0 | 21.21 | 0.47 | 21.33 | -1.76 | -0.06 | -14.24 | -0.06 | -27.72 | 0.06 | 0 | 0.23 | 28.48 | -0.7 | 28.36 | -0.29 | 28.89 | -0.7 | 9.49 | -15.18 | 0 | -37.91 | -0.18 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06 | 0.12
98.848 | 0 | -28.13 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.47 | 21.33 | -1.41 | 3.69 | -12.42 | -0.06 | -26.84 | 0 | 0 | 0.23 | 28.48 | -0.53 | 28.65 | -0.35 | 28.3 | -0.53 | 16.64 | -13.36 | -0.06 | -37.33 | -0.06 | 0 | 0 | 0 | 0.12 | 0 | 0.12 | 0
98.88 | -0.06 | -25.14 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.39 | -0.35 | 11.54 | -10.55 | -0.06 | -25.2 | 0.12 | 0 | 0.18 | 28.71 | -0.35 | 28.65 | -0.29 | 28.83 | -0.35 | 27.01 | -10.66 | -0.06 | -38.2 | -0.18 | 0 | -0.06 | 0 | 0.12 | 0 | 0.12 | 0.06
98.912 | -0.06 | -23.56 | -0.18 | 0 | -0.18 | 0 | 0 | 21.21 | 0.29 | 21.39 | -0.18 | 17.17 | -8.56 | 0.06 | -22.91 | 0.06 | 0 | 0.12 | 28.48 | -0.47 | 28.48 | -0.29 | 28.42 | -0.35 | 28.65 | -8.56 | -0.06 | -37.44 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 0
98.944 | 0.06 | -21.09 | -0.06 | 0 | -0.18 | 0 | -0.06 | 13.42 | 0.18 | 21.39 | 0.06 | 21.74 | -6.68 | 6.62 | -21.09 | 0 | 0 | 0.18 | 28.48 | -0.35 | 28.42 | -0.23 | 28.48 | -0.23 | 28.65 | -6.21 | -0.06 | -35.45 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 0
98.976 | -0.06 | -19.22 | -0.06 | -37.85 | -0.06 | -0.06 | 0 | 0.12 | 0.18 | 21.39 | 0.12 | 21.39 | -5.16 | 12.36 | -19.45 | 0.12 | 0 | 0.18 | 28.3 | -0.29 | 28.48 | -0.29 | 28.59 | -0.35 | 28.42 | -4.45 | -0.18 | -34.16 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
99.008 | 0 | -17.46 | -0.06 | -55.67 | -0.06 | -33.69 | -0.18 | 0 | 0.23 | 21.39 | 0.29 | 21.27 | -3.28 | 17.87 | -18.4 | 0.06 | 0 | 0.18 | 28.24 | -0.23 | 28.65 | -0.23 | 28.77 | -0.23 | 28.54 | -2.64 | -0.06 | -32.29 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
99.04 | -0.06 | -15.18 | -0.18 | -40.2 | -0.06 | -53.32 | -0.06 | 0.06 | 0.18 | 21.39 | 0.47 | 21.39 | -2.29 | 21.45 | -16 | 0.12 | 0 | 0.18 | 0.12 | -0.29 | 28.65 | -0.18 | 28.54 | -0.23 | 28.42 | -1.52 | -0.06 | -29.42 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
99.072 | -0.06 | -13.77 | -0.06 | -39.79 | -0.06 | -40.61 | -0.06 | 0 | 0.18 | 21.33 | 0.47 | 21.39 | -1.17 | 21.39 | -14.3 | 0 | 0 | 0.18 | 0 | -0.35 | 28.48 | -0.18 | 28.36 | -0.29 | 28.65 | -0.7 | 0.06 | -27.6 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
99.104 | 0.29 | -11.84 | -0.06 | -37.21 | -0.18 | -39.44 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -1.05 | 21.39 | -12.48 | 0.18 | -33.05 | 0.18 | 0 | -0.41 | 28.48 | -0.18 | 28.77 | -0.29 | 28.42 | -0.29 | -0.18 | -26.13 | -0.06 | -5.92 | -0.18 | 0 | -0.18 | 0 | 0 | 0
99.136 | 9.84 | -10.66 | -0.18 | -34.63 | -0.06 | -38.79 | -0.06 | 0 | 0.12 | 21.45 | 0.47 | 21.21 | -0.35 | 21.39 | -11.02 | 0.06 | -68.15 | 0.18 | 0 | -0.23 | 28.3 | -0.18 | 28.65 | -0.23 | 28.3 | -0.06 | 0 | -24.2 | -0.18 | -91.59 | -0.18 | 0 | 0 | 0 | -0.18 | 0
99.168 | 15.29 | -9.32 | -0.06 | -33.17 | -0.06 | -35.1 | -0.06 | 0 | 0.18 | 21.27 | 0.47 | 21.39 | -0.29 | 21.33 | -9.32 | -0.06 | -47.46 | 0.12 | 0 | -0.18 | 28.54 | -0.18 | 28.48 | -0.18 | 28.42 | 0 | -0.06 | -23.09 | 0 | -62.64 | -0.06 | -80.39 | -0.06 | -0.18 | -0.18 | 0
99.2 | 20.92 | -7.79 | -0.18 | -30.29 | -0.06 | -33.4 | -0.06 | 0 | 0.18 | 21.15 | 0.41 | 21.39 | 0.06 | 21.33 | -6.86 | -0.06 | -27.77 | 0.18 | 0 | -0.18 | 0.23 | -0.06 | 28.48 | -0.18 | 28.54 | 0.18 | 5.68 | -20.1 | -0.06 | -48.46 | 0 | -58.48 | -0.18 | -89.12 | -0.06 | -32.17
99.232 | 21.5 | -6.15 | -0.06 | -29.01 | -0.06 | -30.29 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0 | 21.33 | -5.39 | 0 | -31.17 | 0.12 | 0 | -0.23 | 0 | -0.18 | 1.35 | -0.18 | 28.65 | 0.23 | 13.71 | -18.22 | -0.06 | -43.65 | -0.06 | -35.45 | 0 | -60.53 | 0 | -83.62
99.264 | 21.5 | -4.63 | -0.18 | -27.01 | 0 | -29.06 | -0.06 | 0 | 0.06 | 21.33 | 0.35 | 21.39 | 0.06 | 21.39 | -3.98 | 0 | -35.51 | 0.18 | 0 | -0.23 | 0 | 0 | 0.12 | -0.18 | 28.65 | 0.41 | 21.27 | -16.41 | -0.06 | -37.21 | -0.18 | -30.18 | -0.06 | -35.33 | -0.18 | -56.84
99.296 | 21.45 | -2.99 | -0.06 | -24.96 | -0.06 | -26.9 | -0.06 | 0 | 0 | 21.27 | 0.35 | 21.39 | 0.06 | 21.39 | -2.52 | -0.06 | -34.22 | 0.23 | 0 | -0.18 | 0.06 | -0.06 | 0 | -0.18 | 28.83 | 0.35 | 28.83 | -14.36 | 0 | -36.51 | -0.06 | -40.2 | -0.06 | -30.18 | -0.18 | -32.52
99.328 | 21.5 | -1.99 | -0.06 | -22.38 | 0 | -25.25 | -0.06 | 0 | -0.06 | 21.15 | 0.23 | 21.39 | 0.18 | 21.33 | -1.93 | -0.06 | -31.7 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.41 | 28.54 | -12.42 | -0.06 | -38.32 | -0.18 | -42.66 | 0 | -39.32 | -0.18 | -31.11
99.36 | 21.5 | -0.94 | -0.06 | -19.86 | -0.06 | -22.21 | -0.06 | -0.53 | -0.06 | 0.12 | 0.18 | 21.39 | 0.18 | 21.15 | -1.17 | 0 | -28.36 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.59 | -9.84 | -0.06 | -37.44 | -0.18 | -36.68 | 0 | -41.13 | -0.06 | -42.01
99.392 | 21.5 | -0.59 | -0.06 | -18.16 | -0.06 | -20.22 | 0 | -66.74 | -0.06 | 0 | 0.23 | 21.39 | 0.23 | 21.33 | -0.23 | -0.06 | -27.48 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.42 | 0.29 | 28.65 | -7.79 | -0.06 | -35.57 | -0.18 | -33.52 | -0.06 | -37.33 | 0 | -41.66
99.424 | 21.45 | -0.18 | -0.06 | -16.41 | -0.06 | -18.46 | -0.06 | -59.18 | 0 | 0 | 0.18 | 21.39 | 0.18 | 21.33 | -0.23 | 0 | -26.25 | 0.23 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.35 | 28.71 | -5.74 | -0.06 | -33.52 | 0 | -33.22 | -0.06 | -33.93 | -0.06 | -37.44
99.456 | 21.5 | -0.06 | -0.18 | -15.29 | 0 | -16.88 | -0.06 | -43.89 | -0.06 | 0 | 0.12 | 21.21 | 0.29 | 21.39 | -0.23 | -0.18 | -23.85 | 0.23 | -43.83 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.36 | 0.29 | 28.48 | -3.75 | 0 | -31.93 | -0.18 | -32.76 | -0.18 | -32.99 | -0.06 | -33.69
99.488 | 21.45 | -0.06 | -0.06 | -13.54 | -0.06 | -14.77 | -0.06 | -38.67 | -0.06 | 0 | 0.18 | 21.39 | 0.18 | 21.45 | 0.06 | -0.06 | -22.62 | 0.23 | -70.32 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.35 | 28.3 | -2.29 | 4.45 | -30.24 | -0.06 | -31.88 | -0.06 | -32.7 | 0 | -33.28
99.52 | 21.45 | -0.29 | -0.06 | -11.54 | 0 | -13.01 | -0.06 | -32.46 | 0 | 0 | 0 | 21.39 | 0.29 | 21.15 | 0.06 | -0.18 | -20.27 | 0.12 | -45.35 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.71 | 0.29 | 28.65 | -1.29 | 15.06 | -28.07 | -0.18 | -29.12 | -0.06 | -31.11 | -0.06 | -32.17
99.552 | 21.45 | -0.35 | 3.05 | -10.02 | -0.06 | -11.6 | 0 | -32.29 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.27 | 0.12 | -0.06 | -18.75 | 0.18 | -44.77 | -0.18 | 0 | -0.06 | 0 | 0 | 28.42 | 0.23 | 28.65 | -0.53 | 23.2 | -25.84 | -0.18 | -27.6 | 0 | -29.01 | -0.06 | -30.53
99.584 | 21.45 | -0.29 | 9.2 | -8.44 | -0.18 | -10.31 | -0.06 | -33.17 | -0.06 | 0 | 0 | 21.45 | 0.23 | 21.39 | 0.18 | -0.12 | -17.23 | 0.12 | -43.3 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.12 | 28.48 | -0.47 | 29.01 | -23.97 | -0.06 | -26.02 | 0 | -27.19 | -0.06 | -28.77
99.616 | 21.45 | -0.29 | 15 | -7.38 | -0.06 | -8.73 | -0.06 | -31.99 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.23 | 6.62 | -15.23 | 0.12 | -41.66 | -0.06 | -24.67 | -0.18 | 0 | -0.06 | 28.48 | 0.06 | 28.48 | 0 | 28.65 | -22.38 | -0.06 | -24.9 | -0.06 | -26.08 | 0 | -27.6
99.648 | 21.39 | -0.29 | 20.22 | -5.45 | -0.06 | -7.32 | -0.18 | -30.35 | 0.06 | 0 | 0 | 21.33 | 0.18 | 21.33 | 0.29 | 12.42 | -13.83 | 0.12 | -39.55 | -0.18 | -78.81 | 0.06 | -0.06 | -0.06 | 28.89 | 0.06 | 28.65 | 0 | 28.48 | -20.45 | -0.18 | -23.32 | -0.06 | -25.14 | -0.06 | -26.13
99.68 | 21.5 | -0.18 | 21.5 | -3.57 | -0.06 | -5.39 | -0.06 | -27.07 | 0 | 0 | 0 | 21.15 | 0.12 | 21.39 | 0.23 | 19.98 | -11.37 | 0.18 | -37.33 | 0 | -47.81 | 0 | -78.69 | 0 | 4.57 | 0.06 | 28.36 | 0.18 | 28.48 | -17.93 | -0.23 | -20.68 | -0.06 | -22.68 | 0 | -24.08
99.712 | 21.45 | -0.23 | 21.45 | -2.7 | 4.16 | -3.87 | -0.06 | -24.84 | 0 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.23 | 21.39 | -9.84 | 0.18 | -35.74 | -0.06 | -45.88 | -0.18 | -57.66 | -0.06 | 0.12 | 0.06 | 28.48 | 0.35 | 28.48 | -16.29 | -0.06 | -19.34 | -0.06 | -20.92 | -0.06 | -22.68
99.744 | 21.39 | -0.23 | 21.45 | -1.7 | 10.49 | -2.75 | 0 | -23.61 | -0.18 | 0 | 0 | 0.23 | 0 | 21.33 | 0.18 | 21.39 | -8.56 | 0.18 | -33.52 | -0.06 | -44.3 | -0.06 | -45.18 | -0.06 | 0 | 0.06 | 29.01 | 0.35 | 28.48 | -14.18 | -0.06 | -18.16 | 0 | -19.69 | 0 | -21.15
99.776 | 21.39 | -0.18 | 21.45 | -0.82 | 15.65 | -1.52 | 3.57 | -21.68 | 0 | -35.63 | 0 | 0 | 0 | 21.15 | 0.35 | 21.39 | -6.5 | 0.12 | -31.29 | 0 | -41.9 | 0 | -45.06 | -0.06 | 0 | -0.06 | 28.65 | 0.35 | 28.48 | -12.36 | -0.06 | -16.47 | 0 | -18.22 | -0.06 | -19.63
99.808 | 21.5 | -0.18 | 21.45 | -0.29 | 21.45 | -0.76 | 9.96 | -19.51 | -0.06 | -73.36 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | -4.75 | 0.06 | -30.18 | -0.06 | -40.67 | 0.06 | -43.07 | -0.18 | 0 | 0 | 28.65 | 0.35 | 28.3 | -10.55 | 4.22 | -14.71 | -0.06 | -16.41 | -0.06 | -18.28
99.84 | 21.39 | -0.23 | 21.45 | -0.18 | 21.5 | -0.35 | 17.52 | -16.58 | -0.18 | -46 | -0.06 | 0 | -0.18 | 0.12 | 0.23 | 21.45 | -3.16 | 0.12 | -27.72 | -0.06 | -37.5 | -0.06 | -40.55 | 0 | 0 | 0.06 | 28.71 | 0.35 | 28.48 | -7.68 | 14.94 | -12.6 | 0 | -14.24 | 3.87 | -15.82
99.872 | 21.45 | -0.18 | 21.5 | -0.06 | 21.45 | -0.18 | 21.5 | -15 | 0 | -40.37 | -0.06 | 0 | 0 | 0 | 0.23 | 21.27 | -2.4 | 0 | -25.37 | -0.06 | -36.74 | -0.06 | -39.26 | -0.18 | 0 | 0 | 28.89 | 0.29 | 28.59 | -5.98 | 22.79 | -11.13 | 3.34 | -12.72 | 12.01 | -14.47
99.904 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -12.95 | 0 | -34.45 | 0.06 | 0 | 0 | 0 | 0.29 | 21.33 | -1.17 | 0 | -23.97 | -0.06 | -34.86 | -0.06 | -36.86 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.3 | -3.63 | 29.01 | -9.38 | 11.72 | -10.84 | 19.86 | -12.72
99.936 | 21.45 | -0.29 | 21.39 | -0.29 | 21.5 | -0.18 | 21.39 | -10.9 | 0.06 | -31.93 | 0.06 | 0 | -0.06 | 0 | 0.18 | 21.15 | -0.82 | 0 | -21.33 | -0.06 | -32.93 | 0 | -34.81 | 0 | 0 | -0.06 | 28.48 | 0.23 | 28.48 | -2.17 | 28.59 | -8.14 | 19.22 | -9.32 | 27.31 | -11.37
99.968 | 21.45 | -0.23 | 21.45 | -0.41 | 21.5 | -0.18 | 21.39 | -8.56 | 0.12 | -32.46 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.33 | -0.41 | 0.06 | -19.63 | -0.06 | -31 | 0 | -33.81 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | -1.23 | 28.48 | -6.33 | 26.95 | -8.14 | 28.65 | -9.61
100 | 0.29 | -0.23 | 21.45 | -0.41 | 21.45 | -0.29 | 21.33 | -6.09 | -0.06 | -32.4 | -0.06 | 0 | 0 | 0 | 0.18 | 21.33 | -0.23 | 0 | -16.88 | -0.18 | -28.24 | 0.06 | -31.29 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.65 | -0.59 | 28.48 | -4.22 | 28.48 | -5.92 | 28.3 | -7.56
100.032 | 0.06 | -0.18 | 21.45 | -0.29 | 21.45 | -0.41 | 21.39 | -3.98 | -0.18 | -30.7 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.12 | -0.06 | -15.47 | -0.18 | -26.25 | -0.06 | -29.01 | 0 | 0 | 0 | 28.48 | 0.23 | 28.54 | -0.35 | 28.59 | -2.81 | 28.48 | -4.28 | 28.42 | -5.63
100.064 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.33 | -2.11 | 0.06 | -28.18 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0.12 | 0 | -14.3 | 0 | -24.61 | 0.06 | -27.36 | 0 | 0 | -0.06 | 28.36 | 0.12 | 28.42 | -0.23 | 28.59 | -2.05 | 28.54 | -2.87 | 28.71 | -3.93
100.096 | 0.06 | -0.18 | 21.39 | -0.29 | 21.5 | -0.18 | 21.39 | -1.7 | 0 | -26.13 | 0.06 | 0 | -0.06 | 0 | 0.23 | 21.33 | 0.06 | 0.12 | -12.72 | 0 | -22.56 | 0 | -25.55 | -0.06 | 0.06 | 0 | 0.23 | 0 | 28.65 | 0 | 28.42 | -1.17 | 28.48 | -2.23 | 28.42 | -2.99
100.128 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.33 | -0.53 | 0 | -24.26 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 2.93 | -11.6 | 0 | -20.45 | 0.06 | -23.61 | -0.06 | -67.74 | -0.06 | 0.12 | 0.06 | 28.65 | 0.12 | 28.65 | -0.7 | 28.48 | -1.11 | 28.48 | -2.23
100.16 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.33 | -0.35 | 0.06 | -21.86 | 0.06 | -42.07 | 0.06 | 0 | 0.06 | 21.39 | 0.29 | 10.9 | -9.79 | -0.18 | -18.52 | 0.06 | -20.86 | 0 | -64.57 | -0.06 | 0 | 0.06 | 28.42 | 0.23 | 28.59 | -0.35 | 28.24 | -0.53 | 28.24 | -1.46
100.192 | 0 | -0.18 | 21.39 | -0.23 | 21.5 | -0.29 | 21.21 | 0.12 | 0.06 | -19.69 | -0.06 | -69.44 | 0 | 0 | 0.12 | 21.39 | 0.35 | 16.29 | -8.09 | 0 | -15.94 | 0 | -18.93 | -0.06 | -48.69 | 0 | 0 | 0 | 28.54 | 0.29 | 28.59 | -0.35 | 28.65 | -0.53 | 28.54 | -0.41
100.224 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | -0.06 | -17.75 | 0 | -46.88 | -0.06 | 0 | 0.06 | 21.21 | 0.23 | 21.5 | -7.15 | 0 | -14.82 | 0.06 | -16.88 | 0.06 | -45.53 | 0 | 0 | 0 | 28.18 | 0.23 | 28.36 | 0 | 28.48 | -0.18 | 28.54 | -0.47
100.256 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.18 | 21.33 | 0.23 | -0.18 | -16.11 | -0.06 | -41.78 | -0.06 | -46.35 | -0.06 | 21.39 | 0.35 | 21.27 | -6.15 | 0 | -13.36 | -0.06 | -15.35 | -0.06 | -39.2 | 0 | 0 | 0 | 28.77 | 0.23 | 28.42 | 0 | 28.65 | -0.18 | 28.65 | -0.23
100.288 | 0 | -0.18 | 21.5 | -0.29 | 21.5 | -0.35 | 21.33 | 0.47 | 0.06 | -14.3 | -0.06 | -34.98 | 0 | -65.04 | 0.06 | 21.27 | 0.18 | 21.39 | -4.92 | 0 | -11.84 | 0.06 | -14 | 0 | -36.27 | -0.06 | 0 | -0.18 | 28.36 | 0.35 | 28.48 | 0 | 28.48 | 0 | 28.42 | -0.06
100.32 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | 0.41 | 0 | -11.37 | -0.06 | -32.29 | 0.06 | -34.98 | 0.12 | 21.39 | 0.29 | 21.39 | -3.87 | 0 | -10.14 | 3.34 | -12.07 | 0 | -37.5 | -0.06 | 0 | 0.12 | 28.65 | 0.23 | 28.59 | 0.18 | 28.48 | 0.06 | 28.71 | 0
100.352 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | -0.29 | 21.33 | 0.47 | 0.12 | -8.91 | -0.18 | -32.64 | -0.06 | -27.6 | 0.06 | 21.15 | 0.29 | 21.56 | -2.99 | -0.06 | -8.67 | 11.72 | -10.84 | 0 | -38.09 | -0.18 | 0 | 0 | 28.48 | 0.29 | 28.71 | 0.06 | 28.48 | 0.18 | 28.83 | 0.06
100.384 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 0.06 | -7.09 | -0.18 | -32.7 | -0.18 | -32.99 | 0.06 | 21.33 | 0.29 | 21.45 | -2.4 | 4.92 | -7.38 | 19.22 | -9.2 | 0 | -36.74 | 0 | 0 | -0.06 | 0.18 | 0.18 | 28.59 | 0.23 | 28.54 | 0.23 | 28.59 | 0.12
100.416 | -9.9 | -0.06 | 21.39 | -0.18 | 21.45 | -0.23 | 21.21 | 0.47 | 4.75 | -4.86 | -0.06 | -30.76 | -0.06 | -35.57 | 0.06 | 21.33 | 0.23 | 21.21 | -2.11 | 13.13 | -6.33 | 26.95 | -7.91 | 0 | -34.51 | -0.18 | 0 | -0.18 | 0.06 | 0.12 | 28.48 | 0.23 | 28.65 | 0.18 | 28.48 | 0.12
100.448 | -64.05 | -0.06 | 3.52 | -0.18 | 21.39 | -0.18 | 21.33 | 0.41 | 10.78 | -3.11 | 0.06 | -28.95 | -0.18 | -33.46 | 0.06 | 21.39 | 0.18 | 21.39 | -1.52 | 20.63 | -5.22 | 28.48 | -6.97 | 0 | -32.34 | 0 | 0 | 0 | 0 | 0.06 | 28.42 | 0.23 | 28.48 | 0.12 | 28.48 | 0.06
100.48 | -48.4 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.41 | 18.16 | -1.82 | -0.06 | -26.02 | -0.18 | -29.59 | 0 | 10.66 | 0.35 | 21.33 | -1.41 | 28.83 | -3.93 | 28.65 | -5.33 | 7.79 | -30.29 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.59 | 0.18 | 28.59 | 0.18 | 28.65 | 0.18
100.512 | -40.37 | -0.18 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.47 | 21.62 | -0.7 | -0.06 | -24.43 | -0.18 | -27.19 | 0 | 0.06 | 0.29 | 21.39 | -1 | 28.65 | -3.16 | 28.18 | -4.34 | 15.06 | -28.13 | -0.06 | -38.85 | 0 | 0 | 0 | 28.48 | 0.12 | 28.59 | 0.29 | 28.59 | 0.12
100.544 | -39.26 | -0.18 | 0 | -0.06 | 21.45 | -0.18 | 21.27 | 0.41 | 21.39 | -0.53 | 0 | -22.44 | -0.18 | -27.31 | 0.06 | 0 | 0.29 | 21.39 | -0.76 | 28.59 | -2.75 | 28.77 | -3.4 | 22.91 | -26.54 | -0.18 | -87.54 | -0.18 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | 0.23 | 28.48 | 0.06
100.576 | -36.21 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.33 | 0.29 | 21.39 | -0.06 | 0 | -20.74 | -0.18 | -25.61 | -0.06 | 0 | 0.35 | 21.45 | -0.59 | 28.48 | -1.99 | 28.3 | -2.99 | 28.89 | -24.79 | -0.06 | -56.02 | 0 | 0 | 0 | 28.59 | 0.18 | 28.42 | 0.18 | 28.48 | 0.18
100.608 | -34.28 | -0.06 | 0 | -0.18 | 0.23 | -0.18 | 21.27 | 0.29 | 21.39 | -0.06 | -0.06 | -18.28 | -0.18 | -23.91 | -0.18 | 0 | 0.18 | 21.39 | -0.59 | 28.65 | -1.46 | 28.71 | -2.4 | 28.65 | -23.38 | -0.06 | -48.75 | -0.06 | 0 | 0.12 | 28.83 | 0.06 | 28.89 | 0.12 | 28.42 | 0.18
100.64 | -33.17 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | 0.23 | 0 | -16 | -0.18 | -21.45 | -0.06 | 0 | 0.18 | 21.39 | -0.41 | 28.42 | -1.23 | 28.13 | -1.46 | 28.65 | -20.86 | -0.06 | -42.6 | 0 | 0 | 0 | 28.42 | 0.12 | 28.59 | 0.12 | 28.48 | 0.12
100.672 | -30.53 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.18 | 0.23 | 21.33 | 0.18 | -0.06 | -13.77 | -0.06 | -19.98 | -0.06 | 0 | 0.06 | 21.33 | -0.35 | 28.42 | -0.94 | 28.59 | -1.05 | 28.59 | -18.87 | -0.06 | -36.21 | -0.18 | 0 | -0.06 | 28.48 | 0 | 28.48 | 0.12 | 28.48 | 0.23
100.704 | -27.95 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.41 | 0 | -11.78 | -0.06 | -18.34 | -0.06 | 0 | 0.29 | 21.39 | -0.35 | 28.59 | -0.76 | 28.36 | -0.82 | 28.42 | -17.23 | -0.06 | -36.92 | -0.06 | 0 | 0.06 | 0.23 | 0.06 | 28.18 | 0.12 | 28.48 | 0.23
100.736 | -26.08 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | 0.41 | 0.06 | -9.67 | -0.06 | -16.7 | 0 | 0 | 0.12 | 21.33 | -0.35 | 29.24 | -0.53 | 28.65 | -0.64 | 28.65 | -15.35 | -0.06 | -38.15 | 0 | 0 | -0.18 | 0.06 | 0 | 28.48 | 0.18 | 0.7 | 0.18
100.768 | -23.09 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0 | 21.33 | 0.47 | -0.06 | -7.56 | -0.06 | -14.82 | -0.06 | 0 | 0.18 | 21.33 | -0.41 | 28.65 | -0.41 | 28.65 | -0.35 | 28.77 | -13.48 | -0.06 | -37.09 | 0 | 0 | 0.06 | 0 | 0.06 | 4.63 | 0 | 0.12 | 0.23
100.8 | -21.74 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.29 | 4.57 | -4.63 | -0.18 | -12.77 | 0 | 0 | 0.18 | 21.33 | -0.29 | 28.65 | -0.35 | 28.48 | -0.18 | 28.59 | -10.61 | 0 | -35.16 | 0 | -52.09 | 0 | 0 | 0 | 0.12 | 0.06 | 0 | 0.12
100.832 | -19.63 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 10.72 | -3.11 | -0.06 | -10.9 | 0 | 0 | 0.23 | 21.33 | -0.29 | 28.42 | -0.35 | 28.65 | -0.18 | 28.42 | -8.38 | -0.06 | -33.22 | 0 | -77.41 | 0 | 0 | 0 | 0 | 0.12 | 0.06 | 0.18
100.864 | -17.52 | -0.06 | -3.98 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 16.06 | -1.52 | -0.06 | -9.2 | -0.06 | 0 | 0.23 | 21.33 | -0.29 | 28.48 | -0.35 | 28.48 | -0.35 | 28.65 | -6.45 | 0 | -31.29 | 0 | -52.44 | 0 | 0 | -0.06 | 0.06 | 0 | 0 | 0.12
100.896 | -15.82 | -0.06 | -57.83 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.47 | 21.8 | -1.29 | -0.06 | -7.44 | -0.06 | 0.06 | 0.18 | 21.33 | -0.29 | 28.42 | -0.29 | 28.59 | -0.29 | 28.48 | -4.34 | 8.44 | -29.94 | -0.06 | -48.11 | -0.18 | 0 | 0 | 0 | 0.12 | 0 | 0.23
100.928 | -13.89 | -0.06 | -50.28 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.35 | 21.45 | -0.41 | -0.06 | -5.98 | -0.06 | -42.48 | 0.12 | 21.33 | -0.29 | 28.83 | -0.29 | 28.65 | -0.29 | 28.42 | -2.58 | 15.65 | -28.13 | 0 | -42.54 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.18
100.96 | -12.42 | 0 | -40.31 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.23 | 21.39 | -0.06 | 2.29 | -3.75 | -0.06 | -58.3 | 0.18 | 21.27 | -0.29 | 28.95 | -0.23 | 28.36 | -0.23 | 28.48 | -1.7 | 26.25 | -25.9 | -0.06 | -36.27 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
100.992 | -11.07 | -0.06 | -39.26 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.33 | 0.29 | 21.39 | 0.18 | 8.03 | -2.52 | -0.06 | -34.16 | 0.23 | 21.33 | -0.23 | 28.36 | -0.35 | 28.54 | -0.29 | 29.01 | -0.64 | 28.71 | -23.97 | -0.06 | -37.68 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.23
101.024 | -9.79 | 0 | -37.03 | -0.18 | -14.06 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | 0.23 | 13.54 | -1.46 | 0 | -27.6 | 0.18 | 0.12 | -0.23 | 28.65 | -0.18 | 28.54 | -0.18 | 28.65 | -0.59 | 28.71 | -21.97 | -0.06 | -38.2 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.12
101.056 | -7.97 | 0 | -34.92 | -0.06 | -64.28 | 0 | 0 | -0.06 | 21.21 | 0.23 | 21.45 | 0.35 | 19.16 | -1.29 | -0.18 | -33.17 | 0.18 | 0 | -0.18 | 28.77 | -0.18 | 28.89 | -0.18 | 28.59 | 0 | 28.48 | -19.63 | -0.18 | -37.44 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.18
101.088 | -6.62 | -0.06 | -32.23 | -0.06 | -47.99 | -0.06 | -15 | -0.06 | 21.39 | 0.18 | 21.39 | 0.53 | 21.5 | -0.59 | 0.12 | -35.39 | 0.12 | 0 | -0.23 | 28.48 | -0.23 | 28.59 | -0.23 | 28.71 | 0.06 | 28.59 | -18.4 | -0.06 | -35.86 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0
101.12 | -4.92 | -0.06 | -29.94 | -0.06 | -40.14 | -0.06 | -69.08 | -0.06 | 21.39 | 0 | 21.39 | 0.47 | 21.33 | -0.29 | -0.06 | -33.22 | 0.12 | 0.06 | -0.23 | 28.42 | -0.18 | 28.42 | -0.18 | 28.42 | 0.18 | 28.36 | -16.17 | 0 | -32.64 | -0.18 | -45.12 | -0.06 | 0 | 0 | 0 | 0.12
101.152 | -3.57 | 0 | -28.01 | -0.06 | -38.38 | -0.06 | -49.16 | 0.06 | 21.33 | 0.06 | 21.39 | 0.47 | 21.33 | -0.18 | -0.06 | -29.65 | 0.12 | 0 | -0.18 | 28.48 | -0.06 | 28.89 | -0.18 | 28.65 | 0.18 | 28.59 | -14 | -0.06 | -30.76 | -0.06 | -80.22 | -0.06 | 0 | 0 | -7.91 | 0.23
101.184 | -2.46 | -0.06 | -26.08 | -0.06 | -36.62 | -0.06 | -41.25 | -0.06 | 21.45 | 0.12 | 21.39 | 0.53 | 21.39 | 0 | -0.06 | -27.77 | 0.12 | 0 | -0.18 | 28.48 | -0.06 | 28.42 | -0.18 | 28.48 | 0.18 | 28.42 | -12.07 | -0.06 | -29.24 | 0 | -52.74 | -0.06 | -3.52 | -0.06 | -89.07 | 0
101.216 | -1.11 | -0.06 | -24.43 | -0.06 | -34.51 | -0.06 | -36.68 | 0.06 | 21.39 | 0.06 | 21.27 | 0.41 | 21.33 | 0.12 | 0 | -27.01 | 0.12 | 0.06 | -0.18 | 28.59 | -0.18 | 4.69 | -0.18 | 28.59 | 0.35 | 28.42 | -10.02 | 0.06 | -27.42 | 0 | -29.59 | -0.06 | -88.6 | -0.18 | -62.58 | 0
101.248 | -0.53 | 1.29 | -22.15 | -0.06 | -32.58 | -0.06 | -31.88 | -0.06 | 21.5 | 0 | 21.27 | 0.41 | 21.33 | 0.12 | 6.86 | -26.13 | 0.06 | 0 | -0.18 | 28.42 | -0.18 | 0.12 | -0.06 | 28.59 | 0.29 | 28.36 | -7.68 | -0.06 | -26.13 | 0 | -32.4 | 0.06 | -65.22 | -0.06 | -35.86 | 0
101.28 | -0.23 | 11.13 | -19.28 | 0 | -29.77 | -0.06 | -32.23 | 0 | 21.39 | 0 | 21.39 | 0.35 | 21.39 | 0.23 | 14.36 | -23.03 | 0 | 0 | -0.18 | 0.23 | -0.06 | 0 | -0.18 | 28.83 | 0.35 | 28.42 | -5.04 | 0 | -23.5 | 0 | -43.07 | -0.06 | -32.99 | -0.18 | -33.99 | 0
101.312 | -0.18 | 16.47 | -17.58 | 0 | -27.95 | -0.06 | -32.52 | -0.06 | 0.12 | 0 | 21.39 | 0.35 | 21.39 | 0.23 | 19.92 | -20.92 | 0 | 0 | -0.06 | 0.06 | -0.23 | 0.06 | -0.06 | 28.65 | 0.35 | 28.54 | -2.99 | -0.18 | -21.68 | 0 | -40.31 | -0.06 | -31.76 | -0.18 | -41.78 | 0
101.344 | -0.18 | 21.91 | -16.06 | -0.06 | -26.13 | 0 | -31.11 | 0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.18 | 21.39 | -19.63 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.89 | 0.29 | 28.48 | -2.23 | 0.76 | -19.98 | 0 | -35.74 | -0.06 | -40.72 | -0.18 | -41.49 | 0
101.376 | -0.35 | 21.5 | -14.77 | 0 | -23.67 | -0.06 | -29.65 | 0 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | 0.29 | 21.27 | -18.4 | 0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.29 | 0.29 | 28.48 | -0.94 | 10.55 | -17.93 | -0.18 | -33.81 | -0.18 | -42.42 | -0.06 | -37.74 | 0.06
101.408 | -0.35 | 21.45 | -12.77 | -0.06 | -21.68 | 0 | -27.01 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.29 | 21.39 | -16.29 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.29 | 29.01 | -0.64 | 18.16 | -16.29 | 0.06 | -33.28 | -0.18 | -38.15 | 0 | -34.63 | 0
101.44 | -0.41 | 21.5 | -10.55 | -0.06 | -19.04 | -0.06 | -24.38 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.21 | 0.35 | 21.27 | -14.53 | -0.06 | -33.46 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.35 | 28.65 | 0 | 28.42 | -13.83 | 0 | -32.23 | -0.06 | -33.52 | -0.06 | -32.99 | 0
101.472 | -0.41 | 21.5 | -9.61 | 0 | -16.93 | -0.06 | -22.33 | -0.06 | 0 | 0 | 21.45 | 0.29 | 21.5 | 0.23 | 21.39 | -12.54 | -0.06 | -69.85 | -0.18 | 0.06 | -0.18 | 0 | -0.18 | 0 | 0.23 | 28.83 | 0 | 28.59 | -11.66 | -0.06 | -30.65 | -0.06 | -32.93 | -0.06 | -32.52 | 0
101.504 | -0.29 | 21.45 | -7.97 | -0.06 | -15.94 | 0 | -20.22 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.39 | 0.29 | 21.39 | -10.96 | -0.06 | -48.05 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.29 | 28.59 | 0.06 | 28.71 | -9.61 | 0 | -28.48 | 0 | -32.29 | -0.12 | -31.47 | -0.06
101.536 | -0.23 | 21.45 | -6.68 | -0.06 | -13.95 | 0 | -18.69 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.23 | 21.39 | -9.61 | 0 | -30.35 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.23 | 28.42 | 0.35 | 28.65 | -7.56 | -0.06 | -27.36 | -0.18 | -31.47 | -0.06 | -29.24 | 0
101.568 | -0.18 | 21.45 | -5.1 | 0 | -12.42 | -0.06 | -16.52 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.29 | 21.39 | -7.62 | 0 | -28.83 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 28.42 | 0.18 | 28.59 | -5.51 | -0.06 | -25.9 | -0.06 | -29.24 | -0.06 | -27.36 | -0.18
101.6 | -0.18 | 21.45 | -3.46 | 0 | -11.07 | -0.06 | -13.71 | 0 | 0 | -0.06 | 21.45 | 0.12 | 21.39 | 0.23 | 21.39 | -5.68 | 0.06 | -35.92 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.42 | 0.35 | 28.48 | -3.11 | -0.18 | -24.26 | -0.18 | -26.84 | 0 | -26.13 | 0
101.632 | -0.29 | 21.45 | -2.29 | 1.99 | -9.32 | -0.06 | -11.95 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.45 | 0.18 | 21.15 | -3.93 | 0 | -34.81 | -0.18 | 0 | -0.18 | -3.22 | -0.06 | 0 | 0.06 | 28.48 | 0.47 | 28.48 | -1.99 | -0.06 | -22.15 | 0 | -25.67 | 0 | -25.02 | -0.06
101.664 | -0.23 | 21.39 | -1.41 | 6.27 | -7.97 | -0.06 | -9.43 | -0.18 | 0 | 0 | 21.39 | 0.06 | 21.33 | 0.18 | 21.33 | -2.64 | -0.06 | -31.41 | -0.06 | 0 | 0 | -75.76 | 0 | 0 | 0.12 | 28.65 | 0.41 | 28.65 | -1 | -0.06 | -20.74 | -0.06 | -24.38 | 0 | -23.09 | -0.18
101.696 | -0.18 | 21.5 | -0.82 | 11.89 | -6.21 | 5.45 | -7.79 | -0.18 | 0 | -0.06 | 0.23 | 0.12 | 21.39 | 0.18 | 21.21 | -1.76 | 0 | -28.89 | -0.06 | -15.41 | -0.18 | -59.24 | -0.06 | 0 | 0 | 28.24 | 0.35 | 28.65 | -0.53 | 0 | -18.99 | -0.18 | -22.91 | 0 | -21.09 | 0
101.728 | -0.18 | 21.45 | -0.47 | 17.7 | -4.69 | 11.25 | -5.39 | -0.06 | -31.41 | 0 | 0.06 | -0.06 | 21.33 | 0.06 | 21.15 | -1.11 | 0.06 | -27.36 | -0.06 | -78.52 | -0.06 | -46 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.42 | -0.18 | 0 | -17.58 | -0.06 | -21.09 | 0 | -19.98 | -0.06
101.76 | -0.18 | 21.39 | -0.06 | 21.5 | -3.28 | 18.63 | -2.87 | 0.06 | -64.98 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.45 | -0.7 | -0.06 | -26.43 | -0.06 | -48.69 | 0 | -45.18 | 0 | 0 | 0 | 17.46 | 0.29 | 28.42 | 0 | 0 | -15.29 | -0.06 | -19.04 | 0 | -17.99 | 0
101.792 | -0.23 | 21.45 | -0.18 | 21.45 | -1.99 | 21.68 | -1.82 | 0 | -45.24 | 0 | 0 | -0.06 | 21.27 | 0.06 | 21.45 | -0.23 | 4.8 | -24.84 | -0.18 | -45.76 | 0 | -42.25 | 0 | -38.97 | 0.06 | 0.12 | 0.18 | 28.48 | 0 | 1.11 | -13.36 | 0 | -17.64 | 0 | -16.29 | -0.06
101.824 | -0.18 | 21.5 | -0.29 | 21.45 | -1.46 | 21.45 | -1.17 | -0.06 | -40.96 | -0.06 | 0 | -0.06 | 21.33 | 0 | 21.39 | -0.23 | 10.78 | -22.44 | -0.06 | -43.6 | -0.18 | -41.13 | -0.18 | -70.73 | -0.06 | 0 | 0.12 | 28.48 | 0.23 | 10.84 | -12.36 | 0 | -16.29 | -0.06 | -14.71 | -0.06
101.856 | -0.18 | 21.45 | -0.47 | 21.39 | -0.76 | 21.45 | -0.47 | -0.06 | -33.93 | -0.06 | 0 | 0 | 17.52 | -0.06 | 21.27 | 0 | 16.29 | -20.92 | -0.18 | -41.54 | 0 | -39.38 | -0.06 | -49.81 | 0 | 0 | 0.18 | 28.83 | 0.29 | 18.22 | -10.61 | 9.38 | -14.41 | 0.18 | -12.77 | -0.06
101.888 | -0.06 | 21.45 | -0.47 | 21.45 | -0.41 | 21.45 | -0.23 | -0.06 | -30.94 | 0 | 0 | -0.06 | 0.12 | -0.06 | 21.33 | 0 | 21.74 | -19.45 | -0.06 | -40.9 | -0.18 | -35.86 | -0.18 | -46.47 | 0 | 0 | 0.12 | 28.42 | 0.35 | 25.9 | -8.79 | 16.7 | -13.13 | 10.14 | -11.25 | 2.64
101.92 | -0.18 | 21.45 | -0.23 | 21.45 | -0.06 | 21.45 | 0.23 | 0 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0 | 21.39 | -17.23 | -0.06 | -37.38 | -0.06 | -35.16 | -0.06 | -44.12 | -0.06 | 0 | 0.06 | 28.3 | 0.47 | 28.71 | -6.62 | 27.01 | -10.96 | 19.86 | -9.08 | 13.07
101.952 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.23 | -0.06 | -32.34 | 0 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.06 | 21.45 | -15.47 | -0.06 | -35.74 | 0.06 | -33.17 | -0.06 | -41.31 | -0.06 | 0 | 0 | 28.48 | 0.29 | 28.48 | -4.75 | 28.54 | -9.38 | 27.42 | -7.79 | 20.68
101.984 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.5 | 0.47 | -0.06 | -30.82 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.12 | 21.45 | -13.89 | -0.06 | -33.99 | -0.06 | -30.82 | 0 | -40.14 | -0.06 | 0 | 0.12 | 28.83 | 0.35 | 28.42 | -3.52 | 28.3 | -7.79 | 28.54 | -5.92 | 28.95
102.016 | -0.06 | 21.45 | -0.23 | 21.45 | -0.41 | 21.56 | 0.53 | -0.06 | -28.77 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | -12.25 | 0.06 | -33.05 | -0.06 | -29.53 | -0.18 | -38.2 | -0.18 | 0 | 0 | 28.65 | 0.35 | 28.42 | -2.11 | 28.65 | -6.15 | 28.65 | -4.45 | 28.65
102.048 | -0.18 | 21.45 | -0.18 | 21.45 | -0.53 | 21.45 | 0.41 | -0.06 | -26.31 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -10.49 | 0 | -30.94 | -0.06 | -26.49 | 0 | -36.62 | -0.18 | 0 | 0.12 | 28.65 | 0.18 | 28.54 | -2.05 | 28.48 | -4.69 | 28.48 | -2.99 | 28.48
102.08 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.5 | 0.41 | 7.62 | -23.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.27 | -7.97 | 0 | -28.07 | -0.06 | -24.67 | -0.18 | -33.81 | 0 | 0 | -0.06 | 28.24 | 0.23 | 28.3 | -0.76 | 28.54 | -2.64 | 28.48 | -1.82 | 28.48
102.112 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.47 | 13.18 | -21.56 | -0.06 | -27.72 | -0.06 | 0 | 0 | 10.43 | 0.23 | 21.33 | -6.33 | 0 | -26.25 | 0 | -23.03 | 0 | -32.23 | 0 | 0 | 0 | 28.42 | 0.18 | 28.54 | -0.29 | 28.54 | -1.93 | 28.42 | -1.29 | 28.48
102.144 | -0.18 | 0.12 | -0.29 | 21.39 | -0.23 | 21.39 | 0.35 | 18.69 | -19.75 | -0.06 | -75.3 | 0 | 0 | -0.06 | 0.06 | 0.18 | 21.27 | -5.1 | 6.97 | -24.32 | -0.06 | -20.98 | 0 | -30.24 | 0 | 0 | 0 | 28.48 | 0.12 | 28.36 | -0.35 | 28.48 | -1.17 | 28.65 | -0.41 | 28.48
102.176 | -0.06 | 0.06 | -0.23 | 21.45 | -0.18 | 21.45 | 0.47 | 21.5 | -17.75 | 0.06 | -50.86 | 0 | 0 | -0.06 | 0 | 0.18 | 21.5 | -3.28 | 14.65 | -22.38 | 0 | -18.99 | -0.06 | -28.48 | -0.06 | 0.12 | -0.06 | 28.48 | 0.12 | 28.48 | -0.06 | 28.48 | -0.7 | 28.48 | -0.29 | 28.59
102.208 | -0.06 | 0 | -0.29 | 21.39 | -0.29 | 21.39 | 0.47 | 21.45 | -16.11 | -0.06 | -43.01 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | -2.23 | 22.33 | -20.22 | -0.06 | -16.47 | 0 | -26.66 | 0 | -59.3 | 0.06 | 28.65 | 0.06 | 28.3 | 0 | 28.48 | -0.47 | 28.54 | -0.18 | 28.48
102.24 | -0.18 | 0 | -0.29 | 21.5 | -0.29 | 21.45 | 0.29 | 21.45 | -13.3 | -0.06 | -35.1 | 0 | 0 | 0 | 0 | 0.23 | 21.27 | -1.41 | 28.65 | -17.64 | 0 | -15.18 | -0.06 | -23.91 | -0.06 | -67.03 | -0.18 | 0.29 | 0.06 | 28.48 | 0.12 | 28.48 | -0.23 | 28.36 | 0 | 28.65
102.272 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | 21.39 | -11.66 | -0.06 | -32.23 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.21 | -0.76 | 28.65 | -16.29 | 0 | -14.24 | -0.06 | -21.91 | 0.06 | -48.81 | 0 | 0.12 | 0 | 28.42 | 0.06 | 28.42 | 0 | 28.42 | 0.06 | 28.65
102.304 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.35 | 21.39 | -9.55 | 0.06 | -32.93 | -0.06 | -62.99 | 0 | 0 | 0.18 | 21.39 | -0.59 | 28.65 | -14.59 | 4.39 | -12.25 | -0.06 | -20.1 | 0 | -47.05 | 0 | 0 | 0.06 | 28.48 | 0.12 | 28.3 | 0 | 28.54 | 0 | 28.59
102.336 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.23 | 21.45 | -7.09 | 0 | -32.64 | -0.06 | -61 | 0.06 | 0 | 0.23 | 21.39 | -0.23 | 28.18 | -13.54 | 12.77 | -10.96 | 0 | -18.34 | -0.06 | -40.43 | -0.18 | 0 | 0.06 | 28.48 | 0.12 | 28.59 | 0.06 | 28.54 | 0.06 | 28.65
102.368 | -0.06 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | 0.18 | 21.39 | -5.1 | 0 | -31.47 | 0.06 | -35.98 | 0 | 0 | 0.18 | 21.27 | -0.06 | 28.48 | -11.84 | 20.33 | -9.08 | 9.43 | -16.29 | -0.18 | -36.33 | 0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.48 | 0.12 | 28.48 | 0.06 | 28.59
102.4 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.06 | 21.33 | -2.75 | -0.06 | -28.89 | 0 | -28.48 | -0.06 | 0 | 0.12 | 21.21 | -0.06 | 28.54 | -9.9 | 29.01 | -7.79 | 19.34 | -14.18 | -0.06 | -37.56 | -0.18 | 0 | 0 | 28.42 | 0.18 | 28.42 | 0.23 | 28.42 | 0.29 | 28.48
102.432 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.18 | 21.39 | -1.64 | 0 | -26.6 | 0 | -34.45 | -0.18 | 0 | 0.18 | 21.39 | 0 | 28.48 | -8.79 | 28.65 | -6.68 | 27.13 | -13.36 | -0.06 | -38.32 | -0.06 | 0 | 0.06 | 28.48 | 0.23 | 28.42 | 0.29 | 28.65 | 0.12 | 28.65
102.464 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.18 | 21.39 | -0.94 | 4.57 | -24.79 | 0 | -35.68 | 0.06 | 0 | 0.23 | 21.39 | 0 | 28.48 | -7.56 | 28.59 | -5.39 | 28.71 | -11.66 | -0.18 | -36.51 | 0 | 0 | -0.18 | 28.3 | 0.18 | 28.42 | 0.23 | 28.42 | 0.06 | 28.48
102.496 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.56 | -0.06 | 21.45 | -0.59 | 10.72 | -22.62 | 0.06 | -32.58 | -0.06 | 0 | 0.29 | 21.39 | 0.12 | 28.48 | -6.45 | 28.48 | -4.63 | 28.3 | -10.14 | -0.06 | -34.69 | -0.06 | 0 | 0 | 28.59 | 0.06 | 28.54 | 0.23 | 28.18 | 0.18 | 28.65
102.528 | -0.06 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | -0.06 | 21.39 | -0.18 | 16.06 | -20.98 | 0.06 | -28.71 | 0 | 0.06 | 0.18 | 21.39 | 0.18 | 28.71 | -5.33 | 28.3 | -3.28 | 28.42 | -9.02 | -0.06 | -32.58 | -0.06 | 0 | 0 | 28.65 | 0.06 | 28.54 | 0.18 | 28.42 | 0.06 | 28.42
102.56 | 0 | -25.61 | -0.18 | 0.12 | -0.18 | 20.04 | 0 | 21.39 | 0.29 | 21.8 | -18.87 | 0 | -27.25 | -0.06 | -65.39 | 0.12 | 21.39 | 0.23 | 28.65 | -4.04 | 28.95 | -2.52 | 28.48 | -7.09 | -0.06 | -30.47 | 0 | 0 | -0.18 | 28.42 | 0.12 | 28.36 | 0.23 | 28.54 | 0.23 | 28.42
102.592 | -0.18 | -62.17 | -0.06 | 0 | -0.18 | 0.12 | 0 | 21.39 | 0.18 | 21.39 | -16.29 | -0.06 | -27.19 | -0.06 | -57.19 | 0.12 | 21.39 | 0.23 | 28.42 | -3.22 | 28.59 | -2.05 | 28.42 | -5.86 | 0 | -29.06 | 0 | 0 | 0.06 | 28.48 | -0.06 | 28.59 | 0.23 | 28.48 | 0.18 | 28.48
102.624 | -0.06 | -45.65 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.45 | -14.3 | -0.06 | -24.9 | -0.06 | -35.86 | 0.18 | 21.33 | 0.29 | 28.77 | -2.4 | 28.65 | -1.64 | 28.48 | -4.75 | 0 | -27.13 | -0.18 | 0 | -0.18 | 28.54 | 0 | 28.59 | 0.12 | 28.3 | 0.23 | 28.71
102.656 | -0.06 | -40.02 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.21 | 0.47 | 21.39 | -12.31 | -0.06 | -23.32 | -0.06 | -27.83 | 0.12 | 21.39 | 0.29 | 28.65 | -2.05 | 28.48 | -1.11 | 28.13 | -4.04 | -0.06 | -25.25 | -0.06 | -0.88 | 0 | 28.59 | -0.06 | 28.48 | 0.12 | 28.48 | 0.06 | 28.48
102.688 | -0.06 | -39.55 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.33 | 0.41 | 21.39 | -10.2 | 0.06 | -21.45 | -0.06 | -33.22 | 0.06 | 0.18 | 0.23 | 28.48 | -1.7 | 28.48 | -0.82 | 28.71 | -3.16 | -0.18 | -23.44 | -0.18 | -85.9 | -0.06 | 18.05 | -0.06 | 28.48 | 0.12 | 28.48 | 0.18 | 28.59
102.72 | -0.06 | -35.22 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.53 | 21.45 | -7.85 | -0.06 | -19.04 | 0 | -35.74 | 0.18 | 0 | 0.18 | 28.42 | -1.17 | 28.42 | -0.53 | 28.3 | -2.23 | 0.06 | -20.92 | -0.06 | -57.42 | -0.06 | 0.12 | 0 | 3.69 | 0.12 | 28.48 | 0.12 | 28.59
102.752 | -0.06 | -33.87 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | -5.45 | 0 | -18.16 | -0.06 | -32.29 | 0.12 | 0 | 0.29 | 28.54 | -0.94 | 28.42 | -0.41 | 28.54 | -1.82 | 0 | -19.16 | -0.18 | -48.75 | -0.06 | 0 | 0 | 0.12 | 0.12 | 0.64 | 0.23 | 28.48
102.784 | -0.06 | -31.82 | -0.06 | 0 | -0.06 | 0.06 | -0.06 | 21.33 | 0.41 | 21.39 | -3.57 | -0.06 | -15.94 | -0.06 | -29.06 | -0.06 | 0.06 | 0.29 | 28.48 | -0.76 | 28.36 | -0.35 | 28.36 | -1.05 | -0.06 | -16.93 | -0.06 | -44.53 | -0.06 | 0 | 0 | 0 | 0.06 | 0.12 | 0.23 | 0.29
102.816 | 0 | -30.35 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.93 | -0.06 | -14.3 | 0.06 | -27.6 | 0.06 | 0 | 0.29 | 28.71 | -0.64 | 28.59 | -0.35 | 28.42 | -0.76 | 8.79 | -15 | -0.06 | -37.97 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.12 | 0.12
102.848 | -0.06 | -28.07 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.47 | 21.5 | -1.23 | 3.22 | -12.31 | 0.06 | -27.07 | 0.06 | 0 | 0.23 | 28.3 | -0.47 | 28.24 | -0.18 | 28.3 | -0.53 | 15.94 | -13.3 | -0.18 | -36.33 | -0.06 | 0 | 0 | 0.06 | 0.06 | 0 | 0.12 | 0
102.88 | -0.06 | -25.2 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.33 | -0.94 | 11.37 | -10.37 | -0.06 | -25.37 | 0.06 | 0 | 0.23 | 28.48 | -0.35 | 28.77 | -0.35 | 28.65 | -0.35 | 26.02 | -10.78 | -0.06 | -38.15 | 0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.12 | 0.06
102.912 | -0.06 | -23.56 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0 | 16.7 | -8.91 | 1.99 | -22.62 | -0.06 | 0 | 0.29 | 28.59 | -0.23 | 28.36 | -0.29 | 28.48 | -0.29 | 28.65 | -8.44 | -0.06 | -37.85 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 0
102.944 | -0.06 | -21.5 | -0.06 | 0 | -0.06 | 0 | -0.06 | 19.81 | 0.29 | 21.39 | 0.06 | 21.74 | -6.68 | 6.15 | -20.92 | 0 | 0 | 0.18 | 28.65 | -0.23 | 28.89 | -0.23 | 28.48 | -0.23 | 28.42 | -6.21 | -0.06 | -36.1 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
102.976 | -0.06 | -19.28 | -0.06 | -32.93 | -0.06 | 0 | 0 | 0.12 | 0.12 | 21.33 | 0.12 | 21.39 | -5.22 | 11.78 | -19.86 | 0 | 0 | 0.18 | 28.54 | -0.23 | 28.54 | -0.18 | 28.3 | -0.23 | 28.65 | -4.16 | -0.18 | -33.69 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
103.008 | 0 | -17.23 | -0.06 | -58.95 | -0.06 | -28.07 | 0 | 0 | 0.29 | 21.33 | 0.23 | 21.39 | -3.57 | 17.46 | -18.46 | 0 | 0 | 0.12 | 28.48 | -0.35 | 28.71 | -0.18 | 28.3 | -0.23 | 28.65 | -2.58 | -0.18 | -31.76 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 0
103.04 | -0.06 | -15.23 | -0.06 | -39.96 | -0.06 | -56.19 | -0.06 | 0 | 0.12 | 21.33 | 0.35 | 21.39 | -2.23 | 21.45 | -16.17 | -0.06 | 0 | 0.18 | 0.18 | -0.23 | 28.48 | -0.18 | 28.54 | -0.23 | 28.48 | -1.52 | -0.06 | -30.24 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
103.072 | -0.06 | -13.95 | -0.06 | -39.38 | -0.06 | -41.72 | 0 | 0 | 0.18 | 21.39 | 0.53 | 21.33 | -1.41 | 21.39 | -14.12 | 0 | 0 | 0.23 | 0 | -0.29 | 28.48 | -0.18 | 28.54 | -0.23 | 28.59 | -0.53 | -0.06 | -28.48 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
103.104 | 2.29 | -12.31 | -0.06 | -38.44 | -0.06 | -39.67 | -0.18 | 0 | 0.23 | 21.33 | 0.53 | 21.33 | -0.94 | 21.45 | -12.6 | -0.06 | -30 | 0.23 | 0 | -0.23 | 28.48 | -0.06 | 28.54 | -0.23 | 28.65 | -0.64 | -0.06 | -26.31 | -0.06 | -0.59 | -0.18 | 0 | -0.18 | 0 | 0.06 | 0
103.136 | 9.26 | -10.66 | -0.18 | -36.1 | -0.06 | -37.74 | -0.06 | 0 | 0.06 | 21.33 | 0.53 | 21.33 | -0.7 | 21.33 | -11.13 | -0.06 | -69.38 | 0.29 | 0 | -0.18 | 28.48 | -0.18 | 28.3 | -0.23 | 28.59 | 0.12 | 0 | -24.38 | -0.06 | -86.37 | -0.06 | 0.06 | -0.06 | 0 | -0.18 | 0
103.168 | 14.94 | -9.38 | -0.06 | -33.46 | -0.18 | -36.33 | -0.06 | 0.06 | 0 | 21.39 | 0.35 | 21.33 | -0.18 | 21.39 | -9.61 | -0.06 | -48.63 | 0.29 | 0 | -0.18 | 28.48 | -0.18 | 28.65 | -0.18 | 28.36 | 0 | -0.18 | -23.15 | -0.18 | -65.8 | -0.06 | -71.14 | 0 | -0.06 | 0.06 | 0
103.2 | 20.63 | -7.38 | -0.06 | -30.59 | -0.06 | -33.34 | -0.06 | 0 | 0.18 | 21.21 | 0.41 | 21.39 | -0.06 | 21.39 | -7.09 | 0.06 | -28.01 | 0.18 | 0 | -0.23 | 0.29 | -0.18 | 28.65 | -0.18 | 28.36 | 0.23 | 4.39 | -20.51 | 0 | -48.22 | -0.06 | -62.35 | -0.06 | -89.3 | -0.06 | -22.56
103.232 | 21.62 | -6.15 | -0.18 | -29.77 | -0.18 | -31.23 | -0.06 | 0 | 0 | 21.33 | 0.35 | 21.33 | -0.06 | 21.33 | -4.75 | 0 | -31.17 | 0.23 | 0 | -0.18 | 0.06 | -0.18 | 3.16 | -0.18 | 28.48 | 0.29 | 12.72 | -18.28 | 0 | -43.54 | -0.18 | -37.79 | 0 | -61.06 | 0 | -88.19
103.264 | 21.5 | -4.69 | 0 | -26.49 | -0.06 | -29.53 | -0.06 | 0 | 0.12 | 21.39 | 0.29 | 21.33 | 0 | 21.33 | -3.87 | -0.06 | -35.22 | 0.18 | 0 | -0.23 | 0 | -0.06 | 0.12 | -0.06 | 28.83 | 0.35 | 20.27 | -16.41 | 0 | -37.91 | -0.06 | -29.42 | 0 | -36.21 | 0 | -59.65
103.296 | 21.45 | -3.46 | -0.06 | -24.55 | -0.06 | -27.48 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.33 | 0.12 | 21.27 | -2.46 | -0.06 | -34.63 | 0.12 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.23 | 28.65 | 0.35 | 28.07 | -14.53 | -0.06 | -36.21 | -0.06 | -39.08 | -0.06 | -29.53 | -0.06 | -34.04
103.328 | 21.45 | -2.23 | 0 | -22.68 | -0.06 | -25.31 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.29 | 21.27 | -1.88 | -0.06 | -31.88 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.71 | 0.29 | 28.65 | -13.01 | -0.06 | -38.09 | -0.06 | -43.07 | -0.12 | -38.97 | 0 | -30.53
103.36 | 21.45 | -1.17 | -0.06 | -20.39 | -0.06 | -22.21 | -0.06 | -0.06 | -0.06 | 0.12 | 0.18 | 21.39 | 0.29 | 21.33 | -1.11 | -0.18 | -28.24 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.23 | 28.48 | -10.31 | 0 | -37.91 | -0.18 | -37.15 | -0.06 | -41.54 | 0 | -41.43
103.392 | 21.45 | -0.76 | -0.06 | -17.99 | -0.06 | -20.63 | -0.18 | -57.48 | -0.06 | 0 | 0.29 | 21.45 | 0.18 | 21.33 | -0.35 | -0.06 | -27.6 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.35 | 28.48 | -8.03 | 0 | -35.92 | 0 | -34.63 | -0.06 | -37.5 | 0 | -41.19
103.424 | 21.45 | -0.29 | -0.06 | -16.23 | 0 | -18.52 | -0.06 | -61.35 | 0 | 0 | 0.12 | 21.15 | 0.23 | 21.45 | -0.41 | 0 | -26.54 | 0.12 | 0.06 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.36 | 0.35 | 28.13 | -6.27 | -0.18 | -33.81 | -0.06 | -33.28 | -0.06 | -34.28 | 0 | -37.91
103.456 | 21.45 | -0.18 | -0.06 | -15.18 | 0 | -16.88 | -0.06 | -44.94 | 0 | 0 | 0.18 | 21.33 | 0.35 | 21.39 | -0.06 | -0.06 | -24.08 | 0.18 | -37.97 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.23 | 28.59 | -3.93 | 0 | -31.99 | -0.18 | -33.05 | -0.06 | -32.99 | 0 | -33.4
103.488 | 21.45 | -0.06 | 0 | -13.36 | -0.06 | -15.59 | -0.06 | -40.02 | -0.06 | 0 | 0.12 | 21.21 | 0.12 | 21.45 | -0.06 | -0.18 | -22.09 | 0.12 | -73.6 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.29 | 28.36 | -2.7 | 2.29 | -30.7 | -0.06 | -32.7 | 0 | -32.46 | 0 | -32.29
103.52 | 21.39 | -0.18 | -0.18 | -11.37 | -0.06 | -13.07 | -0.06 | -33.17 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.06 | -0.06 | -20.16 | 0.12 | -45.65 | 0 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.23 | 28.48 | -1.46 | 14.12 | -28.01 | 0 | -29.24 | 0 | -31.76 | 0 | -32.76
103.552 | 21.45 | -0.35 | 0.35 | -10.2 | -0.06 | -11.48 | -0.06 | -32.23 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.27 | 0.12 | -0.06 | -18.75 | 0.06 | -45.18 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.54 | -0.53 | 21.97 | -26.37 | -0.06 | -27.89 | -0.18 | -29.12 | 0 | -31.06
103.584 | 21.39 | -0.23 | 8.26 | -8.67 | -0.18 | -10.2 | -0.06 | -32.76 | 0 | 0 | 0.06 | 21.33 | 0.18 | 21.27 | 0.12 | 0 | -17.64 | -0.06 | -44.06 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.18 | 28.42 | -0.41 | 28.83 | -24.61 | -0.18 | -26.31 | 0 | -27.25 | -0.06 | -28.48
103.616 | 21.45 | -0.35 | 13.65 | -7.44 | 0 | -8.73 | 0 | -32.29 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | 0.18 | 5.8 | -15.18 | 0.06 | -41.08 | -0.18 | -19.4 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.36 | 0 | 28.71 | -22.74 | -0.06 | -24.84 | -0.06 | -26.02 | 0 | -27.01
103.648 | 21.39 | -0.29 | 19.45 | -5.57 | -0.06 | -7.44 | -0.06 | -30.29 | -0.06 | 0 | 0 | 21.56 | 0.12 | 21.15 | 0.29 | 11.54 | -13.59 | -0.06 | -39.79 | -0.06 | -78.69 | 0 | 0.06 | -0.18 | 28.65 | 0.12 | 28.42 | 0 | 28.65 | -21.15 | -0.18 | -23.56 | 0 | -24.96 | 0 | -26.72
103.68 | 21.5 | -0.18 | 21.56 | -4.04 | -0.06 | -5.8 | 0 | -27.42 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.33 | 0.29 | 19.04 | -11.66 | -0.06 | -37.21 | -0.18 | -48.99 | 0 | -77.46 | -0.06 | 14.94 | 0 | 28.71 | 0.06 | 28.48 | -18.22 | -0.18 | -21.33 | -0.18 | -23.03 | 0 | -24.32
103.712 | 21.45 | -0.18 | 21.5 | -2.93 | 3.4 | -4.28 | -0.06 | -25.31 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.35 | 21.56 | -10.02 | 0 | -35.63 | -0.06 | -46.29 | -0.06 | -59.24 | 0 | 0.12 | 0.06 | 28.42 | 0.35 | 28.48 | -16.82 | 0 | -19.75 | -0.06 | -21.15 | -0.06 | -22.74
103.744 | 21.39 | -0.18 | 21.45 | -1.82 | 9.55 | -3.16 | -0.06 | -23.67 | 0 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0.23 | 21.39 | -8.32 | 0 | -34.75 | 0 | -44.3 | -0.06 | -44.94 | 0 | 0 | 0.06 | 28.3 | 0.47 | 28.3 | -14.41 | -0.18 | -17.99 | -0.06 | -19.45 | 0.06 | -21.15
103.776 | 21.45 | -0.18 | 21.45 | -0.94 | 14.94 | -1.88 | 3.34 | -22.09 | 0 | -31.7 | 0.06 | 0 | 0.18 | 21.39 | 0.23 | 21.33 | -6.39 | 0 | -31.17 | 0 | -41.84 | 0 | -45.59 | 0 | 0 | 0 | 28.48 | 0.35 | 28.59 | -12.72 | 0.06 | -16.58 | -0.18 | -18.16 | 0 | -19.63
103.808 | 21.45 | -0.18 | 21.45 | -0.47 | 20.63 | -1.05 | 9.67 | -20.22 | 0 | -73.83 | -0.06 | 0 | 0.06 | 21.33 | 0.18 | 21.33 | -4.45 | -0.06 | -30.41 | -0.06 | -40.55 | 0 | -43.65 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.42 | -10.49 | 3.87 | -15.06 | -0.18 | -16.47 | 0 | -18.28
103.84 | 21.45 | -0.23 | 21.45 | -0.06 | 21.45 | -0.35 | 16.7 | -17.23 | -0.18 | -46.53 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | -3.16 | -0.06 | -27.19 | -0.06 | -38.15 | -0.06 | -41.08 | -0.06 | 0 | -0.06 | 28.42 | 0.35 | 28.77 | -7.68 | 14.59 | -13.01 | -0.06 | -14.59 | 1.7 | -15.76
103.872 | 21.45 | -0.29 | 21.45 | -0.18 | 21.5 | -0.18 | 21.86 | -15.06 | -0.06 | -40.37 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.33 | -2.17 | 0 | -25.67 | 0 | -36.45 | 0 | -39.26 | -0.18 | 0 | 0.06 | 28.42 | 0.35 | 28.65 | -5.68 | 22.33 | -11.48 | 1.58 | -12.77 | 11.19 | -14.41
103.904 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | -13.01 | -0.18 | -34.63 | 0 | 0 | 0 | 0 | 0.23 | 21.33 | -1.29 | -0.06 | -23.67 | -0.06 | -34.45 | -0.06 | -36.74 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.65 | -4.04 | 29.01 | -10.08 | 10.96 | -11.25 | 18.69 | -12.89
103.936 | 21.5 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -11.31 | -0.06 | -31.17 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.27 | -0.7 | 0.06 | -21.86 | -0.06 | -32.46 | -0.06 | -35.1 | -0.18 | 0 | -0.06 | 28.59 | 0.29 | 28.59 | -2.17 | 28.65 | -8.14 | 18.52 | -9.79 | 26.37 | -10.96
103.968 | 21.45 | -0.18 | 21.39 | -0.29 | 21.45 | -0.29 | 21.39 | -9.08 | 0 | -32.76 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.59 | -0.06 | -19.98 | -0.06 | -31.06 | -0.06 | -33.81 | -0.06 | 0 | 0 | 28.42 | 0.18 | 28.48 | -1.46 | 28.48 | -6.33 | 26.43 | -8.14 | 28.65 | -9.61
104 | 0.29 | -0.18 | 21.56 | -0.29 | 21.39 | -0.41 | 21.39 | -6.15 | -0.18 | -32.23 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -0.23 | -0.06 | -17.05 | 0 | -28.24 | -0.06 | -31.58 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | -0.76 | 28.42 | -4.04 | 28.48 | -5.92 | 28.42 | -7.62
104.032 | 0.06 | -0.06 | 21.39 | -0.41 | 21.45 | -0.41 | 21.39 | -4.16 | -0.18 | -30.53 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | -0.06 | 0 | -15.94 | -0.06 | -25.9 | -0.06 | -29.24 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.65 | -0.23 | 28.42 | -2.99 | 28.65 | -4.34 | 28.42 | -5.68
104.064 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | 21.39 | -2.46 | -0.18 | -29.06 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0.12 | 0 | -14.36 | -0.06 | -23.97 | -0.06 | -27.77 | 0 | 0 | 0 | 28.42 | 0.18 | 28.59 | -0.18 | 28.48 | -2.23 | 28.59 | -2.99 | 28.42 | -4.34
104.096 | 0.06 | -0.06 | 21.45 | -0.35 | 21.45 | -0.29 | 21.39 | -1.88 | 0 | -26.49 | 0 | 0 | 0 | 0 | 0.18 | 21.39 | 0.18 | 0 | -12.77 | 0 | -22.38 | 0 | -25.61 | -0.06 | 0 | -0.18 | 0.29 | 0.12 | 28.65 | -0.06 | 28.42 | -1.29 | 28.54 | -2.05 | 28.65 | -2.99
104.128 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | -0.64 | -0.06 | -24.67 | 0 | 0 | 0 | 0 | 0.18 | 21.27 | 0.18 | 2.7 | -11.6 | -0.06 | -20.57 | 0 | -23.5 | -0.06 | -62.29 | 0 | 0.12 | 0.06 | 28.3 | 0 | 28.42 | -0.7 | 28.54 | -1.35 | 28.42 | -2.11
104.16 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | -0.18 | 21.45 | -0.23 | 0 | -22.09 | -0.06 | -38.09 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 10.66 | -10.02 | -0.06 | -17.93 | -0.06 | -20.74 | 0.06 | -65.8 | 0 | 0 | 0.06 | 28.48 | 0.29 | 28.65 | -0.7 | 28.54 | -0.94 | 28.13 | -1.29
104.192 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | 0 | -0.06 | -20.39 | -0.18 | -72.37 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 16.06 | -8.56 | -0.06 | -15.82 | -0.06 | -18.87 | -0.06 | -48.05 | -0.18 | 0.06 | 0 | 28.3 | 0.35 | 28.48 | -0.06 | 28.42 | -0.47 | 28.42 | -0.7
104.224 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | -0.23 | 21.45 | 0.18 | -0.06 | -18.16 | -0.06 | -47.99 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.5 | -7.56 | 0.06 | -14.82 | 0 | -17.05 | -0.06 | -45.82 | -0.06 | 0 | 0.06 | 28.42 | 0.23 | 28.18 | 0 | 28.54 | -0.35 | 28.42 | -0.59
104.256 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | 0.29 | -0.06 | -16.17 | -0.18 | -41.84 | -0.06 | -42.42 | 0.12 | 21.39 | 0.18 | 21.39 | -6.33 | 0 | -13.48 | 0 | -15.47 | -0.06 | -40.72 | -0.18 | 0 | -0.18 | 28.54 | 0.35 | 28.65 | 0.06 | 28.65 | 0 | 28.54 | -0.29
104.288 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | 0.23 | -0.06 | -14.12 | -0.06 | -35.16 | 0 | -66.27 | 0.06 | 21.39 | 0.35 | 21.39 | -5.39 | 0 | -11.78 | -0.06 | -14.18 | -0.06 | -37.15 | -0.06 | 0 | 0.06 | 28.36 | 0.35 | 28.42 | 0.06 | 28.48 | 0 | 28.48 | -0.18
104.32 | 0 | -0.06 | 21.45 | -0.23 | 21.39 | -0.23 | 21.45 | 0.53 | -0.06 | -11.48 | -0.06 | -32.23 | 0 | -36.51 | 0.18 | 21.33 | 0.23 | 21.39 | -4.04 | -0.06 | -10.08 | 2.87 | -12.25 | 0 | -37.62 | 0 | 0 | 0 | 28.48 | 0.23 | 28.42 | 0 | 28.42 | 0 | 28.48 | -0.18
104.352 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | 0.53 | -0.06 | -9.2 | -0.06 | -32.76 | -0.06 | -27.77 | 0 | 21.39 | 0.23 | 21.39 | -3.16 | -0.06 | -8.38 | 11.19 | -10.96 | 0 | -37.68 | -0.18 | 0 | 0 | 28.3 | 0.23 | 28.18 | 0.23 | 28.71 | 0 | 28.89 | 0.06
104.384 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | -0.06 | -7.27 | 0 | -32.76 | -0.06 | -32.58 | 0.12 | 21.39 | 0.29 | 21.45 | -2.52 | 4.39 | -7.62 | 18.52 | -9.55 | 0 | -36.04 | -0.18 | 0 | 0 | 0.23 | 0 | 28.59 | 0.18 | 28.48 | 0.12 | 28.42 | 0.06
104.416 | -4.34 | -0.06 | 21.45 | -0.18 | 21.39 | -0.18 | 21.33 | 0.47 | 4.51 | -5.04 | 0 | -31.29 | 0 | -35.63 | 0.06 | 21.45 | 0.29 | 21.39 | -2.17 | 12.83 | -6.33 | 26.31 | -8.14 | 0 | -35.16 | -0.18 | 0 | 0 | 0.06 | 0.06 | 28.65 | 0.29 | 28.42 | -0.06 | 27.95 | 0.06
104.448 | -62.35 | -0.06 | 14.3 | -0.18 | 21.39 | -0.23 | 21.39 | 0.35 | 10.55 | -3.22 | 0.06 | -29.24 | -0.06 | -34.22 | -0.06 | 21.33 | 0.29 | 21.15 | -1.52 | 20.39 | -5.16 | 28.54 | -6.97 | 0 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.48 | 0.23 | 28.42 | 0.12 | 28.42 | 0.12
104.48 | -49.87 | -0.06 | 0.06 | -0.18 | 21.5 | -0.06 | 21.39 | 0.41 | 17.87 | -1.7 | -0.06 | -26.13 | -0.06 | -29.77 | 0 | 17.7 | 0.29 | 21.39 | -1.11 | 28.89 | -4.22 | 28.54 | -5.92 | 6.33 | -30.12 | 0 | 0 | 0 | 0 | 0.18 | 28.54 | 0.12 | 28.42 | 0.12 | 28.42 | 0.12
104.512 | -40.14 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.47 | 21.68 | -0.94 | 0 | -24.08 | -0.06 | -27.6 | 0 | 0.12 | 0.29 | 21.45 | -0.94 | 28.71 | -3.22 | 28.48 | -4.45 | 14.12 | -28.77 | -0.06 | -23.79 | -0.06 | 0 | 0 | 28.71 | 0.12 | 28.59 | 0.12 | 28.42 | 0.18
104.544 | -39.61 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.29 | 21.39 | -0.47 | -0.06 | -22.74 | -0.06 | -27.01 | -0.06 | 0 | 0.23 | 21.45 | -0.7 | 28.59 | -2.52 | 28.71 | -3.57 | 21.86 | -27.25 | -0.06 | -92.82 | -0.06 | 0 | 0 | 28.13 | 0.12 | 28.54 | 0.18 | 28.48 | 0
104.576 | -35.86 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.39 | -0.06 | -0.06 | -20.51 | -0.06 | -25.55 | 0 | 0 | 0.23 | 21.45 | -0.47 | 28.59 | -1.82 | 28.59 | -2.93 | 29.18 | -24.9 | -0.18 | -57.72 | -0.18 | 0 | 0 | 28.48 | 0.06 | 28.42 | 0.23 | 28.42 | 0.06
104.608 | -34.51 | -0.18 | 0 | -0.06 | 0.59 | -0.06 | 21.27 | 0.35 | 21.39 | 0.06 | 0 | -18.52 | -0.06 | -23.85 | 0 | 0 | 0.23 | 21.45 | -0.53 | 28.89 | -1.46 | 28.36 | -2.23 | 28.48 | -23.03 | -0.06 | -48.81 | -0.06 | 0 | 0 | 28.42 | 0 | 28.42 | 0.18 | 28.48 | 0.18
104.64 | -32.23 | 0 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | 0.29 | -0.06 | -15.94 | 0 | -21.86 | -0.06 | 0 | 0.23 | 21.21 | -0.41 | 28.48 | -1.29 | 28.89 | -1.46 | 28.59 | -20.86 | -0.06 | -43.13 | 0 | 0 | 0 | 28.48 | 0 | 28.48 | 0.12 | 28.42 | 0.06
104.672 | -30.53 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.12 | 0.29 | 21.39 | 0.35 | -0.06 | -14.3 | -0.06 | -19.81 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.13 | -0.94 | 29.3 | -0.94 | 28.65 | -18.63 | -0.06 | -36.39 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.42 | 0.23 | 28.42 | 0.18
104.704 | -28.48 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.06 | 0.12 | 21.45 | 0.47 | -0.06 | -11.95 | -0.18 | -18.46 | -0.06 | 0 | 0.29 | 21.39 | -0.35 | 28.54 | -0.7 | 28.54 | -0.76 | 28.54 | -16.82 | 0 | -36.45 | -0.06 | 0 | 0 | 0.23 | 0.06 | 28.48 | 0.12 | 28.48 | 0.18
104.736 | -25.55 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 0 | -10.14 | -0.06 | -16.7 | -0.06 | 0 | 0.23 | 21.39 | -0.53 | 28.59 | -0.53 | 28.59 | -0.59 | 28.54 | -14.77 | -0.18 | -37.91 | 0 | 0 | 0 | 0.06 | 0.06 | 28.48 | 0.12 | 14 | 0.12
104.768 | -23.67 | 0 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 21.39 | 0.47 | -0.06 | -7.85 | -0.18 | -14.71 | 0 | 0 | 0.23 | 21.21 | -0.41 | 28.65 | -0.41 | 28.65 | -0.41 | 28.65 | -13.36 | -0.06 | -37.91 | -0.06 | 0 | -0.06 | 0 | -0.06 | 14.94 | 0.06 | 0.12 | 0
104.8 | -22.38 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | 3.98 | -4.98 | -0.06 | -12.72 | 0 | 0 | 0.23 | 21.39 | -0.29 | 28.42 | -0.29 | 28.3 | -0.35 | 28.83 | -10.61 | -0.06 | -36.04 | 0 | -40.61 | -0.06 | 0 | 0.06 | 0.12 | 0.12 | 0 | 0.12
104.832 | -19.98 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 10.2 | -3.22 | 0 | -11.13 | -0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.83 | -0.29 | 28.42 | -0.23 | 28.54 | -8.56 | -0.06 | -33.63 | 0 | -82.5 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0.06 | 0.23
104.864 | -18.16 | -0.06 | -0.18 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.47 | 15.47 | -1.7 | -0.06 | -9.32 | -0.06 | 0 | 0.23 | 21.39 | -0.23 | 28.65 | -0.29 | 28.71 | -0.35 | 28.36 | -6.5 | 0 | -31.47 | -0.18 | -55.14 | -0.18 | 0 | 0 | 0.06 | 0 | 0 | 0.12
104.896 | -15.53 | 0 | -47.46 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -1.29 | -0.06 | -7.68 | -0.06 | 0 | 0.18 | 21.21 | -0.35 | 28.48 | -0.29 | 28.3 | -0.35 | 28.48 | -4.45 | 6.91 | -29.94 | -0.18 | -48.4 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.23
104.928 | -13.95 | -0.06 | -52.91 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.47 | 21.39 | -0.35 | -0.06 | -5.68 | 0.06 | -36.15 | 0.23 | 21.39 | -0.29 | 28.59 | -0.18 | 28.42 | -0.29 | 28.42 | -2.64 | 14.65 | -28.01 | -0.06 | -44.06 | -0.18 | 0 | 0 | 0 | 0.06 | 0 | 0.12
104.96 | -12.72 | -0.06 | -40.14 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.45 | -0.06 | -0.06 | -3.63 | 0.06 | -61.64 | 0.18 | 21.21 | -0.23 | 28.3 | -0.23 | 28.36 | -0.23 | 28.54 | -1.64 | 24.9 | -25.78 | -0.06 | -35.98 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.23
104.992 | -11.02 | 0 | -38.85 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -0.06 | 7.09 | -2.4 | 0 | -38.09 | 0.12 | 21.39 | -0.23 | 28.54 | -0.18 | 28.48 | -0.23 | 28.42 | -0.82 | 28.83 | -23.85 | -0.06 | -37.09 | 0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12
105.024 | -9.84 | -0.06 | -37.38 | -0.18 | -5.57 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.29 | 12.89 | -1.99 | 0 | -28.01 | 0.18 | 0.12 | -0.23 | 28.18 | -0.23 | 28.65 | -0.23 | 28.65 | -0.41 | 28.71 | -22.38 | -0.18 | -38.44 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12
105.056 | -8.14 | -0.06 | -35.33 | -0.06 | -63.11 | 0 | 0 | -0.06 | 21.33 | 0.23 | 21.27 | 0.29 | 18.22 | -1 | 0 | -32.64 | 0.18 | 0.06 | -0.23 | 28.48 | -0.18 | 28.59 | -0.29 | 28.42 | -0.06 | 28.48 | -20.57 | 0 | -37.38 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06
105.088 | -6.62 | -0.06 | -33.17 | -0.06 | -50.1 | 0 | -9.73 | -0.06 | 21.39 | 0.23 | 21.45 | 0.35 | 21.56 | -0.7 | -0.06 | -35.57 | 0.12 | 0 | -0.29 | 28.48 | -0.23 | 28.42 | -0.18 | 28.42 | 0 | 28.42 | -18.75 | -0.06 | -35.74 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | 0.06
105.12 | -4.92 | -0.06 | -30 | 0 | -40.2 | -0.06 | -70.67 | 0 | 21.39 | 0.06 | 21.39 | 0.47 | 21.45 | -0.23 | -0.06 | -33.28 | 0.12 | 0.06 | -0.18 | 28.42 | -0.18 | 28.71 | -0.18 | 28.65 | 0.29 | 28.48 | -16.23 | -0.06 | -33.17 | 0 | -41.6 | 0 | 0 | 0 | 0 | 0
105.152 | -3.75 | -0.06 | -28.77 | 0 | -38.67 | -0.06 | -49.46 | 0 | 21.15 | 0.12 | 21.39 | 0.53 | 21.39 | -0.06 | 0 | -30 | 0.12 | 0 | -0.18 | 28.89 | -0.06 | 28.65 | -0.18 | 28.36 | 0.12 | 28.48 | -14.06 | 0 | -30.76 | -0.18 | -80.92 | -0.18 | 0 | 0.06 | -1.05 | 0.12
105.184 | -2.64 | -0.06 | -25.9 | -0.06 | -36.21 | -0.06 | -41.6 | 0 | 21.39 | 0.18 | 21.45 | 0.47 | 21.39 | -0.06 | 0 | -28.07 | 0.06 | 0 | -0.18 | 28.54 | -0.18 | 28.42 | -0.23 | 28.42 | 0.35 | 28.36 | -12.25 | 0 | -29.59 | -0.18 | -53.03 | 0 | -1.17 | 0 | -87.07 | 0.12
105.216 | -1.35 | 0.06 | -24.43 | -0.06 | -34.86 | -0.06 | -36.15 | 0 | 21.21 | 0.06 | 21.39 | 0.35 | 21.39 | 0.06 | 1.99 | -26.78 | 0.12 | 0.06 | -0.18 | 28.59 | -0.06 | 14.88 | -0.18 | 28.48 | 0.23 | 28.42 | -10.14 | -0.06 | -27.42 | 0 | -31.23 | 0 | -86.9 | 0 | -67.5 | 0
105.248 | -0.82 | 2.58 | -22.68 | -0.06 | -32.23 | 0 | -31.99 | 0 | 21.39 | 0.06 | 21.39 | 0.41 | 21.39 | 0 | 5.8 | -26.13 | 0 | 0 | -0.06 | 28.54 | -0.18 | 0.12 | -0.18 | 28.48 | 0.35 | 28.65 | -8.38 | 0 | -26.13 | 0 | -31.82 | -0.18 | -67.33 | 0.06 | -38.85 | 0.06
105.28 | -0.29 | 10.31 | -20.16 | -0.06 | -30.12 | -0.06 | -32.76 | -0.06 | 21.39 | 0.06 | 21.45 | 0.47 | 21.39 | 0.18 | 13.54 | -23.67 | 0.06 | 0 | -0.18 | 0.23 | -0.06 | 0 | -0.18 | 28.54 | 0.41 | 28.42 | -5.33 | -0.06 | -23.79 | 0 | -43.48 | 0 | -33.17 | -0.06 | -31.99 | 0
105.312 | -0.06 | 15.7 | -17.87 | -0.06 | -28.01 | -0.06 | -32.46 | 0.12 | 0.12 | -0.06 | 21.39 | 0.29 | 21.39 | 0.29 | 19.16 | -21.56 | 0 | 0 | -0.18 | 0.06 | -0.18 | 0.06 | -0.18 | 28.54 | 0.35 | 28.42 | -3.22 | 0 | -21.86 | -0.06 | -39.85 | -0.06 | -31 | 0 | -41.19 | 0
105.344 | -0.18 | 21.39 | -15.94 | 0 | -26.08 | -0.06 | -31.06 | -0.06 | 0 | 0 | 21.45 | 0.29 | 21.45 | 0.29 | 21.33 | -20.39 | 0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.35 | 28.42 | -2.11 | 2.4 | -20.27 | -0.06 | -35.92 | 0 | -40.55 | 0 | -42.42 | 0.06
105.376 | -0.18 | 21.5 | -14.71 | -0.06 | -23.67 | -0.06 | -29.59 | 0.06 | 0 | 0.06 | 21.45 | 0.29 | 21.33 | 0.18 | 21.56 | -18.52 | 0.06 | 0 | -0.18 | 0.06 | 0 | 0 | -0.18 | 0.23 | 0.29 | 28.48 | -1.46 | 10.31 | -18.4 | -0.06 | -33.93 | -0.06 | -42.01 | -0.06 | -38.85 | 0
105.408 | -0.35 | 21.5 | -13.01 | -0.06 | -22.38 | -0.06 | -26.9 | -0.06 | 0.06 | 0 | 21.39 | 0.23 | 21.33 | 0.12 | 21.39 | -17.05 | -0.06 | 0 | -0.18 | 0 | 0 | 0.06 | -0.18 | 0.06 | 0.41 | 28.65 | -0.64 | 17.81 | -16.11 | -0.06 | -33.22 | -0.06 | -38.56 | 0 | -34.28 | 0
105.44 | -0.29 | 21.45 | -11.13 | -0.06 | -19.45 | 0 | -23.97 | 0 | 0 | -0.06 | 21.39 | 0.23 | 21.33 | 0.18 | 21.27 | -14.71 | -0.06 | -28.89 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.23 | 28.83 | -0.35 | 28.13 | -14.06 | -0.06 | -32.93 | -0.06 | -34.22 | -0.18 | -33.17 | 0
105.472 | -0.41 | 21.45 | -9.79 | 0 | -17.81 | -0.06 | -22.21 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | 0.35 | 21.39 | -13.13 | -0.06 | -71.08 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.35 | 28.42 | 0 | 28.71 | -11.78 | -0.06 | -30.7 | -0.06 | -33.11 | -0.06 | -32.64 | -0.06
105.504 | -0.23 | 21.45 | -8.32 | 0 | -15.82 | -0.06 | -20.51 | 0 | 0 | -0.06 | 21.5 | 0.18 | 21.15 | 0.12 | 21.33 | -11.31 | 0 | -49.75 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.42 | 0.12 | 28.59 | -9.84 | 0 | -28.3 | -0.06 | -32.93 | -0.06 | -31.82 | 0.06
105.536 | -0.23 | 21.45 | -7.03 | 0 | -14.12 | -0.06 | -18.63 | 0 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.18 | 21.39 | -9.67 | 0.06 | -30.94 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.65 | 0.29 | 28.48 | -7.68 | -0.06 | -26.84 | -0.06 | -31.82 | 0 | -29.65 | 0
105.568 | -0.18 | 21.45 | -5.39 | -0.06 | -12.83 | -0.06 | -16.29 | -0.06 | 0 | 0 | 21.27 | 0.06 | 21.39 | 0.23 | 21.39 | -7.79 | 0 | -28.36 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0.06 | 28.54 | 0.12 | 28.42 | -5.8 | -0.06 | -25.9 | -0.06 | -29.65 | -0.18 | -27.6 | 0.06
105.6 | -0.29 | 21.5 | -3.57 | 0 | -10.66 | -0.06 | -14 | -0.06 | 0 | 0.06 | 21.45 | 0 | 21.39 | 0.18 | 21.27 | -5.39 | 0 | -35.22 | -0.18 | 0 | 0 | 0 | -0.06 | 0 | 0 | 28.59 | 0.35 | 28.48 | -3.46 | -0.06 | -23.61 | -0.18 | -27.07 | -0.18 | -26.08 | 0
105.632 | -0.29 | 21.45 | -2.4 | 0 | -9.32 | 0 | -12.07 | -0.06 | 0 | 0 | 21.39 | -0.06 | 21.33 | 0.18 | 21.39 | -4.04 | 0 | -34.75 | -0.18 | 0 | 0 | -0.88 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.65 | -2.17 | -0.06 | -22.21 | -0.06 | -25.84 | -0.06 | -25.08 | 0
105.664 | -0.23 | 21.45 | -1.46 | 5.45 | -8.14 | 0 | -9.67 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.39 | 0.23 | 21.27 | -2.87 | 0 | -32.23 | -0.06 | 0 | 0 | -71.78 | -0.06 | 0 | 0.06 | 28.42 | 0.47 | 28.65 | -1.05 | -0.06 | -20.22 | -0.18 | -24.84 | -0.06 | -23.38 | -0.06
105.696 | -0.18 | 21.45 | -0.76 | 11.37 | -6.5 | 5.04 | -8.09 | -0.06 | 0 | -0.06 | 0.23 | 0.12 | 21.33 | 0.06 | 21.45 | -2.05 | 0.06 | -28.54 | -0.18 | -11.89 | -0.06 | -60.94 | -0.18 | 0 | 0.06 | 28.59 | 0.41 | 28.36 | -0.76 | 0.06 | -19.04 | 0 | -23.09 | -0.06 | -21.68 | -0.06
105.728 | -0.29 | 21.45 | -0.29 | 16.76 | -5.22 | 11.13 | -5.68 | -0.06 | -27.07 | -0.06 | 0.06 | 0.06 | 21.39 | 0.18 | 21.33 | -1.35 | 0.06 | -27.54 | -0.18 | -78.52 | -0.06 | -45.76 | 0 | 0 | 0.12 | 28.48 | 0.29 | 28.48 | -0.23 | -0.06 | -17.7 | 0 | -21.15 | -0.06 | -19.63 | -0.06
105.76 | -0.18 | 21.45 | -0.18 | 21.5 | -3.46 | 18.16 | -3.11 | -0.06 | -65.86 | -0.06 | 0 | 0 | 21.45 | 0.18 | 21.39 | -0.59 | 0.06 | -26.66 | -0.18 | -48.34 | -0.06 | -45.24 | -0.06 | 0 | -0.18 | 26.49 | 0.29 | 28.65 | 0 | 0.06 | -15.7 | 0 | -19.75 | -0.06 | -18.4 | 0
105.792 | -0.23 | 21.39 | -0.23 | 21.5 | -2.17 | 21.74 | -1.7 | 0 | -46.35 | 0 | 0.06 | -0.06 | 21.27 | 0.18 | 21.33 | -0.23 | 4.69 | -24.84 | -0.18 | -45.76 | -0.06 | -42.6 | -0.06 | -32.05 | 0 | 0.18 | 0.12 | 28.3 | -0.06 | 2.34 | -13.71 | -0.06 | -18.16 | -0.06 | -16.11 | 0
105.824 | -0.18 | 21.45 | -0.35 | 21.5 | -1.52 | 21.45 | -1.41 | 0.12 | -41.6 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | -0.18 | 10.66 | -22.68 | -0.18 | -44.42 | -0.06 | -40.9 | 0 | -74.18 | 0 | 0 | 0.29 | 28.65 | 0.12 | 10.2 | -12.31 | 0 | -16.7 | -0.06 | -14.41 | 0
105.856 | -0.18 | 21.45 | -0.47 | 21.5 | -0.7 | 21.45 | -0.35 | 0 | -34.45 | -0.06 | 0 | -0.06 | 21.27 | 0.12 | 21.39 | -0.23 | 16.06 | -20.92 | -0.18 | -42.01 | -0.06 | -39.44 | 0 | -52.21 | -0.06 | 0 | 0.23 | 28.95 | 0.23 | 17.52 | -10.31 | 8.56 | -14.71 | -0.06 | -12.89 | 0
105.888 | -0.18 | 21.45 | -0.41 | 21.5 | -0.23 | 21.45 | -0.18 | -0.18 | -32.05 | -0.06 | 0 | -0.18 | 0.12 | 0.12 | 21.39 | 0 | 21.39 | -19.1 | -0.06 | -40.9 | 0 | -36.8 | -0.06 | -46 | -0.06 | 0 | 0.06 | 28.42 | 0.35 | 25.37 | -9.14 | 15.7 | -12.95 | 9.02 | -11.37 | 2.46
105.92 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | -0.06 | 0 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | -17.4 | -0.06 | -37.21 | 0 | -35.16 | -0.06 | -44.3 | -0.06 | 0 | 0.12 | 28.48 | 0.23 | 28.59 | -6.5 | 25.78 | -10.84 | 18.75 | -9.61 | 12.72
105.952 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | 0.18 | 0 | -31.7 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.33 | -15.94 | -0.06 | -35.92 | 0 | -33.22 | 0 | -41.37 | -0.06 | 0 | 0.12 | 28.36 | 0.29 | 28.48 | -4.98 | 28.24 | -9.38 | 26.37 | -7.91 | 20.33
105.984 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.5 | 0.41 | 0 | -31 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.15 | -14.06 | -0.06 | -34.1 | -0.06 | -31.52 | 0 | -40.61 | 0 | 0 | 0 | 28.48 | 0.18 | 28.42 | -3.87 | 28.42 | -7.62 | 28.65 | -6.21 | 28.24
106.016 | -0.06 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | 0.29 | 0.12 | -28.36 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | -12.36 | 0 | -32.7 | -0.06 | -29.77 | -0.18 | -38.09 | 0 | 0 | 0 | 28.48 | 0.29 | 28.48 | -2.52 | 28.18 | -5.92 | 28.59 | -4.28 | 28.65
106.048 | -0.18 | 21.45 | -0.29 | 21.39 | -0.35 | 21.45 | 0.35 | 0 | -26.6 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -10.72 | 0 | -30.53 | -0.06 | -26.6 | 0 | -36.8 | -0.06 | 0 | 0 | 28.54 | 0.18 | 28.48 | -1.93 | 28.48 | -4.51 | 28.48 | -2.93 | 28.65
106.08 | -0.23 | 21.39 | -0.35 | 21.45 | -0.29 | 21.45 | 0.41 | 7.03 | -23.67 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.23 | 21.27 | -8.14 | 0 | -28.24 | 0 | -24.96 | -0.06 | -34.45 | 0 | 0 | 0 | 28.3 | 0.29 | 28.54 | -0.76 | 28.65 | -2.75 | 28.48 | -1.82 | 28.42
106.112 | -0.06 | 21.39 | -0.29 | 21.5 | -0.23 | 21.39 | 0.41 | 12.6 | -21.68 | 0 | -20.39 | -0.06 | 0 | 0.06 | 17.64 | 0.12 | 21.21 | -6.68 | -0.06 | -26.13 | 0 | -23.09 | 0 | -32.4 | 0 | 0 | 0 | 28.59 | 0.06 | 28.42 | -0.7 | 28.59 | -1.88 | 28.48 | -1.17 | 28.48
106.144 | -0.18 | 0.23 | -0.35 | 21.45 | -0.29 | 21.5 | 0.47 | 18.4 | -20.27 | -0.18 | -77.76 | -0.18 | 0 | 0.06 | 0.06 | 0.18 | 21.45 | -4.98 | 5.98 | -24.08 | 0 | -21.09 | -0.06 | -30.59 | 0 | 0 | -0.06 | 28.42 | 0.12 | 28.65 | -0.23 | 28.48 | -1.29 | 28.48 | -0.82 | 28.54
106.176 | -0.18 | 0.06 | -0.23 | 21.45 | -0.23 | 21.39 | 0.47 | 21.68 | -17.75 | -0.06 | -52.09 | 0 | 0 | -0.06 | 0 | 0.29 | 21.27 | -3.69 | 13.95 | -22.85 | -0.06 | -19.1 | -0.06 | -28.83 | -0.06 | 0 | -0.06 | 28.48 | 0.23 | 28.59 | -0.18 | 28.48 | -0.59 | 28.36 | -0.41 | 28.54
106.208 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.5 | 0.47 | 21.39 | -16.23 | -0.06 | -42.6 | 0 | 0 | -0.06 | 0 | 0.29 | 21.33 | -2.23 | 21.74 | -20.45 | 0 | -16.93 | -0.06 | -26.31 | -0.06 | -45.76 | 0 | 28.48 | 0.06 | 28.48 | -0.18 | 28.48 | -0.59 | 28.42 | -0.35 | 28.36
106.24 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.5 | 0.41 | 21.45 | -13.54 | -0.06 | -35.27 | -0.06 | 0 | -0.18 | 0 | 0.29 | 21.27 | -1.64 | 28.89 | -17.99 | 0 | -15.23 | 0 | -24.2 | 0 | -72.42 | 0 | 0.29 | 0.06 | 28.48 | 0.12 | 28.48 | -0.23 | 28.42 | -0.18 | 28.71
106.272 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.35 | 21.39 | -11.48 | -0.06 | -32.17 | 0.06 | 0 | 0 | 0 | 0.18 | 21.21 | -0.59 | 28.65 | -16.17 | 0 | -13.77 | -0.06 | -22.27 | -0.06 | -48.75 | 0 | 0.12 | 0.06 | 28.42 | 0.06 | 28.54 | -0.18 | 28.36 | -0.06 | 28.65
106.304 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.35 | 21.39 | -9.43 | -0.06 | -33.17 | 0 | -54.26 | -0.06 | 0 | 0.18 | 21.39 | -0.64 | 28.65 | -14.71 | 3.34 | -12.19 | 0 | -20.22 | -0.06 | -46.99 | -0.18 | 0 | 0 | 28.42 | 0.18 | 28.54 | 0.06 | 28.48 | 0.12 | 28.36
106.336 | -0.06 | 0.06 | -0.18 | 21.39 | -0.29 | 21.45 | 0.23 | 21.39 | -6.86 | 0 | -32.93 | -0.06 | -63.23 | -0.06 | 0 | 0.23 | 21.39 | -0.18 | 28.48 | -13.77 | 11.78 | -10.78 | 1.99 | -18.34 | -0.06 | -41.54 | 0 | 0 | 0 | 28.42 | 0.23 | 28.48 | 0.06 | 28.65 | 0.06 | 28.36
106.368 | -0.06 | 0 | -0.12 | 21.39 | -0.18 | 21.45 | 0.29 | 21.39 | -5.22 | -0.06 | -31.64 | 0 | -38.85 | -0.06 | 0 | 0.18 | 21.39 | -0.23 | 28.42 | -11.95 | 19.28 | -9.08 | 8.38 | -16.52 | -0.06 | -36.68 | -0.06 | 0 | 0 | 28.54 | 0.18 | 28.65 | 0 | 28.48 | 0.18 | 28.59
106.4 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.18 | 21.39 | -2.46 | -0.18 | -29.01 | 0 | -28.01 | -0.06 | 0 | 0.29 | 21.39 | -0.06 | 28.42 | -10.02 | 29.06 | -7.85 | 18.28 | -14.47 | -0.06 | -37.44 | 0 | 0 | -0.06 | 28.59 | 0.06 | 28.42 | 0.29 | 28.42 | 0.06 | 28.65
106.432 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.23 | 21.27 | -1.7 | -0.06 | -26.72 | -0.18 | -33.52 | -0.06 | 0 | 0.23 | 21.39 | -0.06 | 28.48 | -8.67 | 28.65 | -6.68 | 26.31 | -13.36 | -0.18 | -37.79 | -0.06 | 0 | -0.06 | 28.65 | 0.18 | 28.48 | 0.23 | 28.59 | 0.12 | 28.54
106.464 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.18 | 21.39 | -0.7 | 3.75 | -24.55 | -0.06 | -36.15 | -0.06 | 0 | 0.12 | 21.33 | 0.12 | 28.83 | -7.44 | 28.48 | -5.57 | 28.59 | -11.54 | -0.06 | -36.62 | -0.18 | 0 | 0 | 28.36 | 0.06 | 28.07 | 0.23 | 28.54 | 0.23 | 28.48
106.496 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.06 | 21.39 | -0.41 | 9.9 | -22.56 | 0 | -33.17 | 0 | 0 | 0.12 | 21.45 | 0 | 28.83 | -6.5 | 28.65 | -4.39 | 28.54 | -10.37 | 0 | -35.16 | -0.18 | 0 | -0.06 | 28.48 | 0.06 | 28.65 | 0.23 | 28.42 | 0.12 | 28.59
106.528 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.12 | 21.39 | -0.18 | 15 | -21.09 | -0.06 | -29.59 | 0 | 0 | 0.18 | 21.39 | 0.23 | 29.06 | -5.16 | 28.71 | -3.69 | 28.13 | -8.67 | -0.18 | -33.17 | 0 | 0 | -0.06 | 28.65 | 0.06 | 28.3 | 0.23 | 28.36 | 0.12 | 28.54
106.56 | -0.06 | -16.17 | -0.06 | 0.12 | -0.18 | 21.45 | 0.12 | 21.39 | 0.18 | 21.62 | -18.69 | -0.06 | -27.77 | -0.06 | -57.48 | 0.23 | 21.33 | 0.18 | 28.42 | -3.98 | 28.59 | -2.64 | 28.59 | -7.21 | -0.06 | -31.47 | 0 | 0 | -0.06 | 28.3 | 0 | 28.42 | 0.23 | 28.89 | 0.23 | 28.42
106.592 | -0.18 | -64.16 | -0.06 | 0 | -0.18 | 0.12 | 0.06 | 21.39 | 0.23 | 21.39 | -16.17 | -0.18 | -27.42 | 0 | -59.24 | 0.23 | 21.39 | 0.12 | 28.42 | -3.52 | 28.54 | -2.11 | 28.24 | -5.98 | -0.06 | -29.06 | 0 | 0 | -0.06 | 28.71 | 0.06 | 28.48 | 0.12 | 28.54 | 0.18 | 28.42
106.624 | -0.06 | -46.53 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.23 | 21.39 | -14.71 | -0.06 | -25.55 | 0 | -37.91 | 0.23 | 21.39 | 0.18 | 28.54 | -2.81 | 28.59 | -1.52 | 28.59 | -4.8 | 0 | -27.13 | -0.06 | 0 | -0.06 | 28.89 | 0.06 | 28.48 | 0.12 | 28.48 | 0.06 | 28.48
106.656 | 0 | -40.31 | -0.06 | 0.06 | -0.23 | 0 | 0.06 | 21.39 | 0.53 | 21.39 | -12.25 | 0.06 | -23.44 | 0.06 | -26.9 | 0.12 | 21.33 | 0.18 | 28.54 | -2.05 | 28.54 | -0.94 | 28.48 | -4.28 | -0.06 | -25.14 | -0.18 | -0.41 | -0.06 | 28.48 | 0 | 28.48 | 0.18 | 28.42 | 0.06 | 28.48
106.688 | -0.06 | -39.26 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.53 | 21.33 | -10.37 | -0.06 | -21.86 | 0.06 | -31.58 | 0.23 | 0.12 | 0.12 | 28.59 | -1.64 | 28.48 | -0.76 | 28.54 | -3.28 | -0.06 | -24.08 | 0.06 | -82.97 | 0 | 23.73 | 0 | 28.42 | 0.06 | 28.42 | 0.12 | 28.54
106.72 | 0 | -36.33 | -0.06 | 0 | 0 | 0 | -0.06 | 21.21 | 0.53 | 21.33 | -7.85 | 0.12 | -19.45 | 0 | -36.04 | 0.12 | 0 | 0.18 | 28.48 | -1.29 | 28.65 | -0.59 | 28.48 | -2.4 | -0.18 | -21.68 | 0 | -58.42 | -0.06 | 0.12 | -0.06 | 18.46 | 0.12 | 28.48 | 0.12 | 28.48
106.752 | 0 | -34.51 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.35 | 21.39 | -5.22 | -0.06 | -17.75 | -0.18 | -33.28 | 0 | 0 | 0.18 | 28.65 | -1.05 | 28.59 | -0.35 | 28.59 | -1.88 | -0.06 | -19.34 | -0.06 | -48.75 | -0.12 | 0 | 0 | 0.18 | 0.12 | 13.89 | 0.06 | 28.48
106.784 | -0.06 | -31.93 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.27 | -3.57 | 0 | -16.41 | -0.06 | -29.77 | 0.06 | 0 | 0.06 | 28.42 | -0.76 | 28.48 | -0.23 | 28.59 | -1.35 | -0.06 | -17.4 | -0.06 | -44.53 | -0.06 | 0 | 0 | 0 | 0 | 0.12 | 0 | 0.29
106.816 | -0.06 | -30.53 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -1.7 | 0 | -14.77 | -0.06 | -27.42 | 0.12 | 0 | 0.12 | 28.54 | -0.64 | 28.59 | -0.23 | 28.42 | -0.94 | 8.44 | -15.35 | -0.06 | -38.67 | -0.18 | 0 | 0 | 0 | 0.06 | 0 | 0.06 | 0.12
106.848 | 0 | -28.18 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -1.46 | 2.99 | -12.89 | -0.06 | -27.07 | 0 | 0 | 0.18 | 28.89 | -0.59 | 29.06 | -0.23 | 28.65 | -0.76 | 15.59 | -13.48 | -0.06 | -36.15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
106.88 | -0.06 | -25.61 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -0.41 | 10.96 | -10.55 | 0.06 | -25.37 | 0.12 | 0 | 0.18 | 28.71 | -0.41 | 28.42 | -0.23 | 28.65 | -0.41 | 25.9 | -11.07 | -0.18 | -38.2 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.18 | 0.06
106.912 | -0.06 | -23.97 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.35 | 21.27 | -0.29 | 16.47 | -9.02 | 1.99 | -23.09 | 0.12 | 0 | 0.18 | 28.89 | -0.35 | 28.65 | -0.23 | 28.13 | -0.41 | 28.71 | -8.61 | -0.06 | -37.97 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06 | 0
106.944 | 0.06 | -21.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.39 | 0.18 | 21.86 | -7.44 | 5.8 | -21.91 | 0.06 | 0 | 0.18 | 28.71 | -0.29 | 28.59 | -0.18 | 28.59 | -0.29 | 28.77 | -6.74 | 0 | -36.04 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
106.976 | -0.06 | -19.45 | -0.06 | -27.54 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.27 | 0.12 | 21.39 | -5.51 | 11.54 | -19.63 | 0.12 | 0 | 0.06 | 28.54 | -0.35 | 28.71 | -0.35 | 28.48 | -0.29 | 28.42 | -4.69 | -0.06 | -33.81 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
107.008 | -0.06 | -17.93 | 0 | -61.82 | -0.06 | -25.08 | -0.06 | 0 | 0.23 | 21.39 | 0.35 | 21.45 | -3.93 | 17.23 | -18.69 | 0 | 0 | 0.12 | 28.48 | -0.29 | 28.36 | -0.23 | 28.65 | -0.35 | 28.48 | -3.16 | -0.06 | -32.34 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12 | 0
107.04 | -0.18 | -15.76 | -0.06 | -41.25 | -0.06 | -56.78 | -0.06 | 0 | 0.23 | 21.39 | 0.29 | 21.39 | -2.23 | 21.45 | -16.7 | 0.12 | 0 | 0.12 | 0.23 | -0.29 | 28.65 | -0.23 | 28.65 | -0.35 | 28.42 | -1.7 | 0 | -30.18 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 0
107.072 | -0.06 | -13.71 | -0.06 | -40.37 | -0.06 | -42.42 | 0 | 0 | 0.12 | 21.39 | 0.47 | 21.21 | -1.64 | 21.39 | -14.53 | -0.06 | 0 | 0.06 | 0 | -0.29 | 28.42 | -0.18 | 28.54 | -0.35 | 28.83 | -0.94 | 0 | -27.72 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12 | 0
107.104 | 1.99 | -12.6 | -0.06 | -37.56 | -0.18 | -40.61 | -0.18 | 0 | 0.06 | 21.39 | 0.53 | 21.33 | -1 | 21.39 | -13.18 | 0.06 | -25.84 | 0.06 | 0 | -0.23 | 28.77 | -0.18 | 28.65 | -0.18 | 28.24 | -0.53 | 0 | -26.72 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0.06 | 0
107.136 | 9.08 | -10.96 | -0.06 | -35.16 | -0.18 | -38.67 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -0.47 | 21.39 | -11.54 | 0 | -70.08 | 0.18 | 0 | -0.29 | 28.48 | -0.18 | 28.71 | -0.18 | 28.3 | -0.18 | -0.06 | -24.55 | 0.12 | -77.58 | -0.18 | 0.06 | 0 | 0 | 0 | 0
107.168 | 14.65 | -9.61 | -0.06 | -33.87 | -0.06 | -35.74 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 21.27 | -0.18 | 21.33 | -10.02 | 0.06 | -50.22 | 0.18 | 0 | -0.23 | 28.95 | -0.23 | 28.71 | -0.18 | 29.06 | -0.06 | 0 | -22.85 | -0.18 | -68.5 | -0.06 | -55.2 | -0.18 | 0 | 0 | 0
107.2 | 20.39 | -7.62 | -0.06 | -30.7 | -0.06 | -33.81 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.18 | 21.39 | -7.32 | 0 | -28.36 | 0.12 | 0 | -0.23 | 0.23 | -0.23 | 28.54 | -0.23 | 28.65 | 0.18 | 2.7 | -20.33 | -0.06 | -48.46 | -0.06 | -68.32 | -0.18 | -82.27 | 0 | -16.29
107.232 | 21.5 | -6.33 | -0.18 | -29.59 | -0.18 | -31.17 | -0.06 | 0 | 0.12 | 21.27 | 0.47 | 21.39 | 0.12 | 21.39 | -5.74 | 0 | -31.11 | 0.12 | 0 | -0.23 | 0.12 | -0.06 | 17.7 | -0.23 | 28.59 | 0.12 | 12.13 | -18.05 | -0.06 | -44.83 | -0.18 | -40.67 | 0 | -65.8 | 0 | -89.01
107.264 | 21.5 | -4.63 | -0.06 | -27.54 | 0 | -29.42 | 0 | 0 | 0 | 21.39 | 0.41 | 21.39 | 0.18 | 21.39 | -4.28 | 0 | -35.74 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.12 | -0.29 | 28.65 | 0.23 | 19.63 | -16.7 | 0 | -39.67 | -0.06 | -28.95 | 0 | -40.14 | 0.06 | -60.94
107.296 | 21.45 | -3.57 | -0.06 | -25.84 | -0.06 | -26.78 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | 0.23 | 21.27 | -2.64 | -0.06 | -34.69 | 0.18 | 0 | -0.18 | 0.06 | -0.23 | 0 | -0.18 | 28.65 | 0.35 | 27.31 | -14.94 | 0 | -36.62 | -0.06 | -37.44 | 0 | -28.89 | 0.06 | -34.4
107.328 | 21.45 | -2.4 | -0.06 | -23.09 | -0.06 | -25.67 | 0 | 0 | 0 | 21.5 | 0.18 | 21.39 | 0.18 | 21.39 | -1.88 | 0 | -31.88 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.42 | 0.35 | 28.65 | -12.77 | -0.06 | -37.38 | 0 | -42.83 | -0.06 | -36.97 | -0.06 | -29.42
107.36 | 21.45 | -1.23 | -0.06 | -20.45 | -0.06 | -22.56 | -0.18 | 0 | 0 | 0.18 | 0.29 | 21.39 | 0.35 | 21.33 | -1.11 | -0.06 | -27.95 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.65 | 0.35 | 28.54 | -10.08 | -0.06 | -37.62 | -0.06 | -38.5 | 0 | -42.13 | 0 | -42.19
107.392 | 21.45 | -0.7 | -0.06 | -18.75 | -0.06 | -20.27 | -0.06 | -50.22 | 0.12 | 0 | 0.18 | 21.33 | 0.29 | 21.33 | -0.41 | -0.06 | -27.01 | 0.23 | 0 | -0.18 | 0 | 0 | 0 | -0.18 | 28.18 | 0.41 | 28.48 | -8.14 | 0 | -35.68 | -0.06 | -35.45 | -0.18 | -38.56 | 0 | -41.66
107.424 | 21.45 | -0.18 | -0.06 | -16.88 | -0.06 | -18.87 | -0.06 | -63.75 | -0.06 | 0 | 0.18 | 21.45 | 0.23 | 21.39 | -0.35 | -0.06 | -26.43 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.35 | 28.48 | -6.15 | -0.06 | -34.22 | -0.18 | -33.28 | -0.06 | -34.75 | -0.18 | -37.91
107.456 | 21.45 | -0.06 | 0.06 | -15.7 | 0 | -16.47 | -0.06 | -46.76 | 0 | 0 | 0.12 | 21.39 | 0.29 | 21.09 | -0.06 | 0 | -24.43 | 0.12 | -33.17 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.41 | 28.59 | -3.93 | 0 | -32.05 | -0.18 | -32.76 | -0.18 | -33.34 | -0.06 | -33.69
107.488 | 21.45 | -0.18 | 0 | -13.36 | -0.06 | -15.06 | -0.06 | -40.37 | 0 | 0 | 0.06 | 21.33 | 0.12 | 21.45 | -0.06 | -0.06 | -22.56 | 0.18 | -74.24 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.71 | -2.29 | 2.64 | -30.7 | -0.06 | -32.05 | -0.06 | -32.46 | -0.06 | -33.4
107.52 | 21.45 | -0.23 | 0 | -11.54 | -0.06 | -13.36 | -0.18 | -33.34 | -0.06 | 0 | 0.12 | 21.33 | 0.18 | 21.33 | 0.12 | -0.06 | -20.16 | 0.18 | -47.76 | -0.18 | 0 | 0 | 0 | -0.18 | 28.59 | 0.35 | 28.3 | -1.64 | 13.18 | -27.95 | -0.18 | -29.59 | -0.06 | -31.47 | -0.18 | -32.81
107.552 | 21.45 | -0.35 | 0.29 | -10.31 | -0.06 | -11.66 | -0.06 | -32.05 | 0 | 0 | 0.12 | 21.15 | 0.18 | 21.39 | 0.18 | -0.06 | -19.04 | 0.06 | -44.88 | -0.18 | 0 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.18 | -0.7 | 20.86 | -26.25 | -0.06 | -27.83 | -0.18 | -29.06 | -0.18 | -31.23
107.584 | 21.5 | -0.29 | 7.68 | -8.73 | -0.06 | -10.25 | -0.06 | -32.17 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.15 | 0.18 | -0.18 | -17.75 | 0.12 | -44.59 | -0.18 | 0 | -0.06 | 0 | -0.18 | 28.36 | 0.12 | 28.65 | -0.18 | 28.89 | -24.73 | -0.06 | -26.9 | -0.06 | -27.31 | -0.18 | -29.24
107.616 | 21.45 | -0.35 | 13.24 | -7.32 | -0.06 | -8.91 | -0.06 | -32.29 | -0.18 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.12 | 5.33 | -15.53 | 0.06 | -41.31 | -0.06 | -5.63 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.48 | -0.06 | 28.54 | -23.32 | -0.06 | -25.25 | -0.18 | -26.02 | -0.18 | -27.36
107.648 | 21.45 | -0.29 | 18.87 | -5.86 | -0.06 | -7.56 | 0 | -30.94 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | 0.35 | 11.13 | -13.83 | 0.12 | -40.31 | -0.18 | -75.47 | 0 | 0 | 0 | 28.54 | 0.18 | 28.65 | 0.18 | 28.54 | -21.33 | -0.06 | -23.44 | -0.06 | -25.14 | -0.06 | -26.02
107.68 | 21.45 | -0.23 | 21.56 | -3.98 | -0.06 | -5.57 | 0.12 | -28.01 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.21 | 0.29 | 18.52 | -11.66 | 0.06 | -37.68 | 0 | -51.8 | -0.18 | -71.66 | -0.18 | 17.34 | 0.06 | 28.48 | 0.18 | 28.65 | -18.75 | 0 | -21.33 | -0.18 | -23.44 | -0.18 | -24.43
107.712 | 21.45 | -0.23 | 21.5 | -2.75 | 2.7 | -4.16 | 0.06 | -25.37 | 0 | 0 | -0.06 | 21.5 | 0.18 | 21.5 | 0.18 | 21.56 | -9.84 | 0 | -35.74 | -0.06 | -45.18 | -0.06 | -63.23 | -0.18 | 0.18 | 0 | 28.65 | 0.41 | 28.48 | -16.7 | -0.18 | -20.27 | -0.18 | -21.86 | -0.18 | -22.56
107.744 | 21.45 | -0.23 | 21.45 | -1.7 | 8.73 | -3.16 | -0.06 | -23.32 | 0 | 0 | 0 | 0.12 | 0.12 | 21.45 | 0.35 | 21.39 | -8.44 | 0.06 | -34.28 | -0.06 | -44.83 | -0.06 | -44.24 | -0.06 | 0 | 0.12 | 28.65 | 0.41 | 28.59 | -15 | 0 | -18.22 | -0.06 | -20.51 | -0.18 | -20.74
107.776 | 21.45 | -0.18 | 21.45 | -0.94 | 14.06 | -1.93 | 1 | -21.68 | -0.06 | -19.04 | 0 | 0.06 | 0.12 | 21.33 | 0.29 | 21.39 | -6.33 | 0 | -32.29 | 0 | -42.54 | -0.06 | -45.76 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.59 | -12.89 | -0.06 | -17.17 | -0.18 | -18.4 | -0.18 | -19.69
107.808 | 21.39 | -0.18 | 21.45 | -0.47 | 19.98 | -1.23 | 8.79 | -19.75 | -0.06 | -77.23 | -0.06 | 0 | 0.12 | 21.27 | 0.29 | 21.39 | -5.16 | 0 | -30.41 | -0.06 | -41.02 | -0.06 | -44.01 | 0 | 0 | 0.06 | 28.42 | 0.35 | 28.65 | -10.96 | 1.35 | -15.35 | -0.18 | -16.7 | 0 | -18.63
107.84 | 21.39 | -0.23 | 21.45 | -0.06 | 21.5 | -0.47 | 15.94 | -17.23 | 0 | -48.17 | 0.06 | 0 | 0.06 | 0.12 | 0.29 | 21.39 | -2.99 | 0.06 | -27.83 | -0.06 | -38.44 | 0 | -41.78 | 0 | 0 | 0 | 28.65 | 0.41 | 28.42 | -7.91 | 13.71 | -13.07 | 0 | -14.77 | 0.94 | -16
107.872 | 21.45 | -0.23 | 21.45 | -0.06 | 21.5 | -0.06 | 21.62 | -15.47 | -0.06 | -41.43 | -0.18 | 0 | 0.06 | 0 | 0.29 | 21.39 | -2.05 | 0.12 | -25.43 | -0.06 | -36.92 | -0.06 | -39.61 | 0 | 0 | -0.06 | 28.59 | 0.35 | 28.59 | -6.15 | 21.27 | -11.66 | 1.11 | -12.83 | 10.43 | -14.47
107.904 | 21.39 | -0.29 | 21.45 | -0.18 | 21.45 | -0.06 | 21.56 | -13.54 | -0.06 | -36.04 | -0.18 | 0 | 0 | 0 | 0.35 | 21.39 | -1.17 | 0.06 | -23.79 | 0 | -34.75 | -0.06 | -38.15 | -0.23 | 0 | -0.06 | 28.59 | 0.23 | 28.83 | -4.1 | 29.01 | -10.02 | 10.78 | -11.48 | 18.05 | -12.95
107.936 | 21.45 | -0.18 | 21.45 | -0.23 | 21.5 | -0.18 | 21.39 | -11.37 | -0.06 | -31.52 | -0.18 | 0 | 0 | 0 | 0.29 | 21.39 | -0.82 | 0.06 | -21.86 | -0.06 | -33.46 | -0.06 | -36.1 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.65 | -2.46 | 28.77 | -8.2 | 18.16 | -9.79 | 25.67 | -11.25
107.968 | 21.45 | -0.18 | 21.39 | -0.29 | 21.39 | -0.23 | 21.39 | -9.32 | -0.06 | -32.23 | -0.06 | 0 | 0 | 0 | 0.29 | 21.39 | -0.47 | 0 | -20.16 | 0.06 | -31.11 | -0.06 | -33.69 | 0 | 0 | -0.06 | 28.3 | 0.29 | 28.65 | -1.82 | 28.65 | -6.5 | 25.96 | -8.67 | 28.77 | -9.55
108 | 2.99 | -0.18 | 21.45 | -0.41 | 21.5 | -0.29 | 21.45 | -6.56 | -0.06 | -32.93 | 0 | 0 | 0.06 | 0 | 0.18 | 21.39 | -0.29 | 0.06 | -17.4 | 0 | -28.59 | -0.06 | -31.47 | 0 | 0 | 0 | 28.65 | 0.12 | 27.95 | -0.7 | 28.48 | -4.75 | 28.54 | -6.27 | 28.48 | -7.62
108.032 | 0.06 | -0.18 | 21.5 | -0.41 | 21.5 | -0.12 | 21.39 | -4.34 | -0.06 | -30.76 | -0.18 | 0 | -0.06 | 0 | 0.29 | 21.45 | -0.06 | 0 | -15.7 | -0.06 | -26.37 | -0.06 | -29.65 | 0 | 0 | 0 | 28.54 | 0.12 | 28.65 | -0.29 | 28.54 | -3.16 | 28.48 | -4.45 | 28.65 | -5.63
108.064 | 0 | -0.18 | 21.39 | -0.18 | 21.5 | -0.41 | 21.39 | -2.46 | -0.06 | -29.01 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -0.06 | -0.06 | -14.77 | -0.06 | -25.37 | 0 | -27.6 | 0 | 0 | 0 | 28.59 | 0.18 | 28.59 | 0 | 28.48 | -2.23 | 28.48 | -3.4 | 28.59 | -4.16
108.096 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | -1.7 | 0.06 | -26.49 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0 | 0.06 | -13.24 | 0 | -22.62 | 0 | -26.02 | -0.06 | 0.06 | 0 | 1.35 | 0.12 | 28.59 | 0.06 | 28.48 | -1.29 | 28.48 | -2.23 | 28.36 | -2.87
108.128 | 0 | -0.06 | 21.39 | -0.18 | 21.5 | -0.29 | 21.39 | -0.82 | -0.06 | -25.08 | 0 | 0 | -0.06 | 0 | 0.29 | 21.39 | 0.06 | 2.4 | -11.72 | 0 | -20.74 | 0 | -23.32 | 0 | -57.07 | -0.06 | 0.12 | 0.06 | 28.42 | 0.18 | 28.48 | -0.76 | 28.36 | -1.41 | 28.59 | -2.17
108.16 | 0 | -0.06 | 21.39 | -0.23 | 21.5 | -0.23 | 21.39 | -0.29 | -0.18 | -22.33 | -0.18 | -36.27 | 0 | 0 | 0.12 | 21.21 | 0.18 | 10.14 | -9.84 | 0 | -18.34 | -0.06 | -21.15 | 0 | -68.21 | 0 | 0 | 0.06 | 28.54 | 0.18 | 28.3 | -0.53 | 28.65 | -0.82 | 28.42 | -1.17
108.192 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | -0.18 | 0 | -20.22 | -0.06 | -73.07 | 0.06 | 0 | 0.18 | 21.62 | 0.06 | 15.53 | -8.61 | 0 | -16.47 | 0 | -19.4 | -0.18 | -48.22 | 0 | 0.06 | 0 | 28.42 | 0.35 | 28.3 | -0.35 | 28.3 | -0.23 | 28.65 | -0.59
108.224 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | 0.35 | 0 | -18.05 | -0.06 | -48.87 | 0 | 0 | 0.18 | 21.15 | 0.12 | 21.33 | -6.91 | 0 | -15.06 | 0 | -17.58 | 0 | -46.17 | 0 | 0 | 0 | 28.48 | 0.29 | 28.42 | -0.23 | 28.36 | -0.18 | 28.54 | -0.59
108.256 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.29 | 21.27 | 0.18 | 0.06 | -16.88 | -0.06 | -42.6 | 0 | -37.74 | 0.23 | 21.21 | 0.06 | 21.45 | -5.98 | 0 | -13.71 | -0.18 | -15.7 | 0.06 | -40.37 | -0.18 | 0 | -0.06 | 28.54 | 0.23 | 28.3 | 0.06 | 28.65 | -0.23 | 28.36 | -0.23
108.288 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.18 | 21.39 | 0.23 | -0.06 | -14.77 | -0.06 | -35.92 | -0.06 | -67.5 | 0.06 | 21.45 | 0.06 | 21.45 | -5.16 | 0 | -12.01 | -0.06 | -14.36 | -0.06 | -37.03 | -0.06 | 0 | 0.06 | 28.48 | 0.35 | 28.54 | 0 | 28.42 | 0.06 | 28.36 | -0.18
108.32 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | 0.35 | -0.06 | -12.01 | -0.06 | -32.05 | -0.06 | -36.86 | 0 | 21.33 | 0.23 | 21.39 | -3.75 | -0.06 | -10.37 | 2.64 | -12.54 | 0.06 | -37.44 | -0.18 | 0 | -0.06 | 28.18 | 0.29 | 28.65 | 0.12 | 28.42 | 0.12 | 28.54 | 0
108.352 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | 0.53 | -0.06 | -9.9 | 0 | -32.7 | -0.18 | -27.25 | 0.12 | 21.33 | 0.12 | 21.39 | -3.28 | -0.06 | -8.73 | 10.9 | -11.31 | 0 | -37.33 | -0.18 | 0 | -0.06 | 28.42 | 0.23 | 28.59 | 0.23 | 28.36 | 0.18 | 28.36 | 0
108.384 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.18 | 21.39 | 0.41 | 0.06 | -7.62 | -0.06 | -32.4 | 0 | -32.46 | 0.06 | 21.33 | 0.23 | 21.39 | -2.7 | 3.34 | -7.79 | 18.28 | -9.49 | -0.06 | -36.39 | -0.18 | 0 | -0.06 | 0.29 | 0.23 | 28.48 | 0.12 | 28.36 | 0.23 | 28.3 | 0.12
108.416 | -3.05 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 3.75 | -5.63 | -0.06 | -30.94 | 0 | -35.63 | -0.06 | 21.15 | 0.18 | 21.39 | -2.29 | 11.72 | -6.5 | 25.96 | -8.44 | 0 | -34.86 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 28.65 | 0.12 | 28.48 | 0.12 | 28.36 | 0.23
108.448 | -61.64 | 0 | 18.11 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 10.08 | -3.69 | -0.06 | -29.12 | -0.06 | -33.99 | 0.06 | 21.39 | 0.18 | 21.33 | -1.64 | 19.22 | -5.51 | 28.65 | -6.8 | 0.06 | -32.76 | -0.18 | 0 | -0.06 | 0 | 0.18 | 28.42 | 0.18 | 28.48 | 0.06 | 28.65 | 0.23
108.48 | -50.51 | -0.06 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 17.17 | -2.29 | 0 | -26.13 | -0.06 | -30 | -0.06 | 20.57 | 0.23 | 21.33 | -1.35 | 29.06 | -4.04 | 28.71 | -5.57 | 5.51 | -30 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.07 | 0.23 | 28.48 | 0.12 | 28.48 | 0.12
108.512 | -39.9 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.35 | 21.74 | -0.76 | 0 | -24.08 | 0 | -28.13 | 0 | 0.12 | 0.23 | 21.56 | -0.94 | 28.59 | -3.16 | 28.36 | -4.63 | 13.3 | -28.48 | 0 | -12.95 | 0 | 0 | 0.12 | 28.42 | 0.12 | 28.48 | 0.23 | 28.48 | 0.12
108.544 | -38.91 | 0 | 0 | -0.06 | 21.45 | -0.06 | 21.39 | 0.35 | 21.39 | -0.76 | -0.06 | -22.09 | -0.06 | -27.07 | 0.06 | 0 | 0.18 | 21.33 | -0.7 | 28.3 | -2.46 | 28.77 | -3.63 | 21.27 | -27.42 | -0.06 | -92.93 | -0.18 | 0 | 0.06 | 28.48 | 0.12 | 28.18 | 0.18 | 28.48 | 0.18
108.576 | -36.86 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.35 | 21.39 | -0.06 | -0.06 | -20.39 | -0.18 | -25.78 | -0.06 | 0 | 0.12 | 21.27 | -0.64 | 28.18 | -2.23 | 28.83 | -3.11 | 29.06 | -25.2 | 0 | -59.94 | -0.18 | 0 | 0 | 28.42 | 0.12 | 28.42 | 0.29 | 28.54 | 0.12
108.608 | -34.63 | -0.06 | 0 | -0.06 | 11.07 | -0.18 | 21.45 | 0.18 | 21.39 | -0.06 | 0 | -18.69 | 0 | -24.02 | 0 | 0.06 | 0.12 | 21.39 | -0.53 | 28.42 | -1.76 | 28.48 | -2.23 | 28.48 | -23.44 | -0.06 | -48.17 | -0.18 | 0 | 0 | 28.18 | 0 | 28.3 | 0.12 | 28.42 | 0.06
108.64 | -32.29 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 21.39 | 0.23 | 21.39 | 0.35 | -0.06 | -16.06 | -0.18 | -21.39 | 0 | 0 | 0.12 | 21.39 | -0.35 | 28.95 | -1 | 28.42 | -1.64 | 28.77 | -20.98 | -0.18 | -43.54 | -0.18 | 0 | 0 | 28.18 | 0.06 | 28.42 | 0.12 | 28.48 | 0.12
108.672 | -31.23 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | 0.29 | 0 | -14.12 | -0.18 | -19.92 | -0.06 | 0 | 0.12 | 21.39 | -0.53 | 28.24 | -1.05 | 28.71 | -1.11 | 28.48 | -18.93 | 0 | -37.5 | -0.06 | 0 | 0 | 28.42 | 0 | 28.36 | 0.12 | 28.48 | 0.18
108.704 | -29.06 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.18 | 21.39 | 0.41 | -0.06 | -11.78 | -0.06 | -18.4 | -0.06 | 0 | 0.18 | 21.33 | -0.35 | 28.48 | -0.94 | 28.24 | -0.76 | 28.36 | -16.99 | 0 | -36.74 | -0.18 | 0 | 0 | 0.29 | 0 | 28.18 | 0.18 | 28.42 | 0.18
108.736 | -25.78 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | -0.06 | -10.2 | -0.06 | -17.05 | 0.06 | 0 | 0.23 | 21.39 | -0.35 | 28.48 | -0.53 | 28.36 | -0.7 | 28.42 | -15.23 | 0 | -37.97 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 28.13 | 0.06 | 20.74 | 0.18
108.768 | -23.61 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.35 | -0.06 | -7.62 | -0.06 | -15.18 | -0.18 | 0 | 0.12 | 21.33 | -0.23 | 28.77 | -0.23 | 28.48 | -0.53 | 28.36 | -13.36 | 0 | -37.91 | -0.18 | 0 | 0.06 | 0 | 0 | 17.46 | 0.06 | 0.18 | 0.12
108.8 | -21.74 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 21.39 | 0.47 | 3.69 | -4.86 | -0.06 | -12.77 | -0.06 | 0 | 0.12 | 21.39 | -0.47 | 28.59 | -0.35 | 28.48 | -0.41 | 28.42 | -11.25 | 0 | -35.16 | 0 | -24.2 | 0.06 | 0.06 | 0.06 | 0.12 | 0 | 0 | 0.18
108.832 | -19.92 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 9.96 | -3.28 | -0.06 | -11.07 | 0 | 0 | 0.23 | 21.39 | -0.35 | 28.54 | -0.23 | 28.54 | -0.35 | 28.13 | -8.67 | 0 | -33.34 | -0.06 | -90.88 | -0.12 | 0 | -0.06 | 0 | 0 | 0.06 | 0.12
108.864 | -17.75 | 0 | -0.18 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.35 | 15.29 | -1.88 | 0 | -9.2 | 0.06 | 0 | 0.06 | 21.33 | -0.23 | 28.65 | -0.23 | 28.89 | -0.18 | 28.48 | -6.91 | -0.06 | -31.29 | 0 | -58.54 | -0.18 | 0 | -0.06 | 0.06 | 0 | 0 | 0.12
108.896 | -16 | -0.06 | -44.01 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 20.98 | -1.17 | -0.06 | -7.79 | 0 | 0 | 0.06 | 21.39 | -0.29 | 28.65 | -0.29 | 28.48 | -0.29 | 28.59 | -4.45 | 5.68 | -30 | -0.18 | -48.58 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0
108.928 | -13.83 | -0.06 | -53.85 | -0.06 | 0 | 0 | 0 | -0.18 | 21.39 | 0.29 | 21.45 | -0.64 | -0.06 | -6.15 | 0.06 | -32.29 | 0.23 | 21.39 | -0.23 | 28.48 | -0.29 | 28.77 | -0.41 | 28.42 | -2.75 | 13.48 | -28.48 | -0.18 | -44.42 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
108.96 | -12.72 | -0.06 | -40.37 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.18 | -0.06 | -3.93 | 0 | -63.23 | 0.18 | 21.39 | -0.35 | 28.59 | -0.23 | 28.65 | -0.23 | 28.71 | -2.11 | 23.85 | -26.02 | -0.18 | -36.74 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0.12
108.992 | -11.13 | 0 | -39.08 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | -0.18 | 6.45 | -2.64 | 0.06 | -38.15 | 0.06 | 21.39 | -0.18 | 28.42 | -0.18 | 28.48 | -0.23 | 28.65 | -0.53 | 28.59 | -24.08 | -0.18 | -36.45 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
109.024 | -10.02 | -0.06 | -37.62 | -0.06 | -0.94 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | 0.29 | 12.25 | -1.99 | -0.06 | -27.25 | 0.12 | 0.23 | -0.23 | 28.65 | -0.18 | 28.48 | -0.18 | 28.48 | -0.59 | 28.71 | -22.38 | -0.06 | -37.68 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0
109.056 | -8.14 | 0 | -35.45 | -0.18 | -59.12 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.27 | 0.35 | 17.7 | -1 | 0.12 | -31.52 | 0.23 | 0.06 | -0.23 | 28.36 | -0.18 | 28.65 | -0.23 | 28.95 | -0.06 | 28.48 | -20.45 | 0 | -37.68 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0
109.088 | -6.62 | 0 | -33.28 | -0.06 | -52.15 | -0.06 | -1.93 | -0.06 | 21.27 | 0.06 | 21.33 | 0.29 | 21.62 | -0.64 | -0.06 | -36.1 | 0.06 | 0 | -0.18 | 28.59 | -0.18 | 28.65 | -0.23 | 28.36 | -0.06 | 28.59 | -18.87 | -0.06 | -35.86 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0
109.12 | -5.16 | 0 | -30.29 | 0 | -39.85 | -0.06 | -74.36 | -0.06 | 21.39 | 0.06 | 21.27 | 0.35 | 21.39 | -0.41 | -0.06 | -33.75 | 0.06 | 0.06 | -0.18 | 28.48 | -0.18 | 28.54 | -0.18 | 28.3 | 0.12 | 28.54 | -16.29 | -0.06 | -33.69 | -0.06 | -29.53 | -0.18 | 0 | -0.06 | 0 | 0
109.152 | -3.75 | 0 | -28.95 | 0 | -38.97 | 0 | -51.27 | -0.06 | 21.15 | 0.12 | 21.39 | 0.41 | 21.45 | -0.06 | 0 | -30 | 0.18 | 0 | -0.23 | 28.48 | -0.18 | 28.65 | -0.18 | 28.42 | 0.18 | 28.59 | -14.36 | 0 | -32.17 | 0 | -84.55 | -0.06 | 0 | 0 | 0 | 0
109.184 | -2.64 | -0.06 | -26.13 | 0 | -37.09 | -0.06 | -42.54 | 0 | 21.21 | 0.06 | 21.39 | 0.47 | 21.45 | -0.06 | 0.06 | -27.77 | 0.06 | 0 | -0.06 | 28.48 | -0.06 | 28.65 | -0.06 | 28.36 | 0.29 | 28.71 | -12.6 | -0.18 | -29.71 | -0.18 | -57.07 | 0 | -0.47 | -0.18 | -83.21 | 0
109.216 | -1.23 | 0 | -24.43 | -0.06 | -34.63 | -0.06 | -37.21 | -0.06 | 21.45 | 0.06 | 21.45 | 0.41 | 21.33 | 0 | -0.18 | -27.31 | 0 | 0 | -0.18 | 28.77 | -0.06 | 17.7 | -0.18 | 28.48 | 0.35 | 28.42 | -10.55 | 0 | -28.13 | 0 | -32.4 | -0.18 | -85.08 | 0 | -70.26 | 0
109.248 | -0.76 | 0.18 | -22.38 | -0.06 | -32.99 | 0 | -32.46 | -0.18 | 21.15 | 0.12 | 21.39 | 0.35 | 21.39 | 0.12 | 4.92 | -26.13 | 0.18 | 0 | -0.18 | 28.54 | -0.18 | 0.12 | -0.23 | 28.42 | 0.41 | 28.48 | -8.56 | 0.06 | -27.07 | 0 | -30.82 | -0.06 | -69.38 | 0 | -39.49 | -0.06
109.28 | -0.41 | 9.55 | -19.69 | 0 | -30.06 | -0.06 | -32.52 | 0.12 | 21.39 | 0 | 21.33 | 0.41 | 21.39 | 0.12 | 12.83 | -23.44 | 0.12 | 0 | -0.18 | 0.29 | -0.18 | 0 | -0.06 | 28.54 | 0.35 | 28.65 | -5.51 | 0 | -24.08 | 0 | -43.07 | -0.06 | -33.75 | -0.06 | -30.7 | 0
109.312 | -0.23 | 15.12 | -18.05 | 0 | -27.95 | -0.06 | -33.52 | 0 | 0.18 | 0.06 | 21.45 | 0.29 | 21.39 | 0.12 | 18.28 | -21.5 | 0.12 | 0 | -0.18 | 0.06 | -0.18 | 0.06 | -0.23 | 28.42 | 0.35 | 28.42 | -3.75 | 0 | -22.21 | 0 | -40.78 | -0.18 | -30.94 | 0 | -40.02 | 0
109.344 | -0.23 | 20.51 | -16.11 | -0.06 | -25.9 | -0.06 | -31.82 | -0.18 | 0.06 | 0.06 | 21.45 | 0.23 | 21.33 | 0.18 | 21.39 | -20.22 | 0.06 | 0 | 0 | 0 | -0.23 | 0 | -0.06 | 28.24 | 0.29 | 28.42 | -2.11 | 0 | -20.45 | 0 | -36.39 | -0.18 | -40.2 | -0.18 | -42.01 | 0
109.376 | -0.18 | 21.68 | -14.36 | -0.06 | -24.08 | 0 | -29.83 | -0.18 | 0 | -0.06 | 21.33 | 0.23 | 21.39 | 0.35 | 21.39 | -18.69 | 0.06 | 0 | -0.06 | 0.06 | -0.06 | 0 | -0.06 | 0.29 | 0.29 | 28.42 | -1.64 | 9.38 | -18.46 | -0.18 | -33.69 | -0.06 | -42.48 | -0.06 | -39.02 | -0.06
109.408 | -0.41 | 21.45 | -13.3 | 0.06 | -22.21 | -0.06 | -27.19 | -0.06 | 0 | 0 | 21.33 | 0.23 | 21.39 | 0.29 | 21.27 | -17.29 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.29 | 28.36 | -0.7 | 16.7 | -16.88 | -0.06 | -33.05 | 0 | -38.2 | 0 | -34.63 | 0
109.44 | -0.29 | 21.5 | -11.54 | -0.06 | -19.16 | -0.06 | -24.67 | -0.06 | 0 | 0.12 | 21.39 | 0.23 | 21.33 | 0.18 | 21.33 | -14.65 | 0.18 | -24.38 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.48 | -0.35 | 27.25 | -14.24 | 0.12 | -32.64 | 0 | -33.69 | -0.06 | -32.4 | 0
109.472 | -0.35 | 21.5 | -9.79 | 0 | -17.52 | -0.06 | -22.62 | -0.06 | 0 | 0 | 21.45 | 0.12 | 21.15 | 0.35 | 21.09 | -13.07 | 0.18 | -74.18 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.23 | 28.48 | -0.06 | 28.65 | -12.01 | -0.18 | -30.35 | -0.06 | -33.34 | -0.18 | -32.64 | 0
109.504 | -0.29 | 21.45 | -8.14 | 0.06 | -15.53 | 0 | -20.74 | 0 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.23 | 21.33 | -11.13 | 0.12 | -50.33 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.65 | -0.06 | 28.42 | -10.25 | -0.06 | -29.24 | 0 | -32.81 | 0 | -31.99 | 0
109.536 | -0.18 | 21.45 | -6.91 | -0.06 | -14.12 | -0.06 | -18.75 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.27 | 0.23 | 21.33 | -9.9 | 0.06 | -31.76 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.54 | 0.06 | 28.48 | -8.03 | 0 | -27.89 | 0 | -31.17 | -0.06 | -30.29 | 0
109.568 | -0.29 | 21.45 | -5.33 | 0 | -13.01 | -0.06 | -16.76 | -0.06 | 0 | -0.18 | 21.39 | 0.12 | 21.39 | 0.29 | 21.39 | -8.32 | 0 | -28.54 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.48 | 0.29 | 28.48 | -6.5 | 0 | -26.31 | -0.06 | -29.3 | 0 | -27.6 | 0
109.6 | -0.18 | 21.45 | -3.75 | 0 | -11.13 | 0 | -14.24 | 0 | 0 | 0 | 21.39 | 0.06 | 21.09 | 0.18 | 21.33 | -6.09 | 0 | -35.16 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.23 | 28.54 | 0.29 | 28.48 | -3.46 | -0.18 | -23.97 | -0.06 | -27.13 | -0.06 | -26.13 | -0.06
109.632 | -0.23 | 21.45 | -2.81 | 0 | -9.49 | -0.06 | -12.01 | -0.06 | 0 | 0 | 21.45 | 0.12 | 21.39 | 0.18 | 21.33 | -4.22 | 0 | -34.86 | -0.18 | 0 | -0.06 | -0.12 | -0.06 | 0 | 0.12 | 28.42 | 0.29 | 28.42 | -2.23 | 0 | -22.44 | -0.12 | -25.9 | 0 | -24.43 | -0.06
109.664 | -0.23 | 21.39 | -1.52 | 4.63 | -8.2 | 0.06 | -9.84 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.33 | 0.29 | 21.39 | -2.99 | -0.06 | -31.82 | -0.18 | 0 | -0.06 | -65.74 | 0 | 0 | 0 | 28.59 | 0.35 | 28.59 | -1.11 | -0.18 | -21.09 | -0.06 | -24.84 | 0 | -22.85 | 0
109.696 | -0.18 | 21.5 | -0.82 | 10.78 | -6.68 | 4.69 | -7.97 | -0.18 | 0 | -0.06 | 0.23 | 0 | 21.27 | 0.18 | 21.39 | -2.23 | -0.06 | -28.83 | -0.06 | -3.63 | -0.06 | -63.11 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.65 | -0.94 | 0 | -18.93 | -0.06 | -22.91 | 0 | -21.68 | 0
109.728 | -0.18 | 21.45 | -0.59 | 15.94 | -4.92 | 10.9 | -5.92 | -0.18 | -13.42 | 0.06 | 0.06 | 0 | 21.33 | 0.23 | 21.33 | -1.23 | -0.06 | -27.54 | -0.06 | -77.58 | -0.06 | -45.18 | -0.06 | 0 | 0.06 | 28.54 | 0.35 | 28.48 | -0.06 | -0.06 | -17.7 | -0.06 | -21.21 | -0.06 | -20.16 | 0.06
109.76 | -0.29 | 21.45 | -0.29 | 21.74 | -3.46 | 18.05 | -2.81 | -0.06 | -68.56 | 0 | 0 | -0.06 | 21.33 | 0.23 | 21.27 | -0.41 | -0.06 | -26.31 | -0.06 | -49.81 | -0.06 | -45.7 | -0.06 | 0 | 0 | 28.83 | 0.23 | 28.59 | -0.18 | 0 | -15.76 | -0.18 | -19.69 | -0.06 | -18.52 | 0
109.792 | -0.29 | 21.45 | -0.18 | 21.45 | -2.17 | 21.68 | -2.11 | 0 | -47.35 | 0 | 0 | -0.18 | 21.39 | 0.12 | 21.39 | -0.41 | 3.75 | -24.67 | -0.06 | -45.29 | -0.06 | -42.95 | -0.06 | -24.55 | -0.06 | 0.23 | 0.29 | 28.42 | 0.06 | 0.06 | -13.89 | 0 | -17.93 | 0 | -16.64 | 0
109.824 | -0.18 | 21.45 | -0.29 | 21.45 | -1.52 | 21.5 | -0.76 | -0.18 | -41.66 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.39 | -0.18 | 9.96 | -22.15 | -0.18 | -44.77 | -0.06 | -40.84 | -0.06 | -77.58 | -0.06 | 0 | 0.29 | 28.48 | 0.12 | 9.9 | -12.25 | -0.06 | -16.11 | 0 | -14.36 | 0
109.856 | -0.18 | 21.5 | -0.47 | 21.45 | -0.82 | 21.45 | -0.53 | -0.06 | -35.39 | 0.12 | 0 | 0.12 | 21.39 | 0.12 | 21.33 | -0.06 | 15.35 | -20.68 | 0 | -42.66 | 0 | -39.26 | -0.18 | -53.56 | -0.06 | 0 | 0 | 28.48 | 0.35 | 17.17 | -10.55 | 7.38 | -14.36 | 0 | -13.3 | -0.06
109.888 | -0.18 | 21.45 | -0.29 | 21.45 | -0.41 | 21.45 | -0.23 | -0.18 | -31.58 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 21.33 | 0.06 | 20.92 | -19.63 | 0 | -41.08 | 0 | -36.74 | 0 | -45.35 | -0.18 | 0 | 0.18 | 28.36 | 0.29 | 25.08 | -9.14 | 15.06 | -12.83 | 7.73 | -11.6 | 0
109.92 | -0.06 | 21.45 | -0.47 | 21.45 | -0.23 | 21.5 | 0.12 | -0.06 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -17.7 | -0.06 | -37.5 | 0 | -35.16 | -0.18 | -44.24 | 0.06 | 0 | 0.06 | 28.42 | 0.41 | 28.71 | -7.09 | 25.2 | -10.78 | 17.64 | -9.61 | 11.6
109.952 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.39 | 0.23 | -0.06 | -32.05 | 0 | 0 | -0.18 | 0 | 0 | 21.39 | 0.18 | 21.39 | -15.82 | -0.06 | -36.04 | 0 | -33.99 | -0.06 | -41.08 | -0.06 | 0 | 0.18 | 28.59 | 0.23 | 28.59 | -5.1 | 28.59 | -9.2 | 25.43 | -7.79 | 19.22
109.984 | -0.06 | 21.39 | -0.29 | 21.45 | -0.18 | 21.5 | 0.41 | -0.18 | -31.47 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.27 | -14.12 | -0.06 | -34.75 | -0.06 | -32.05 | 0 | -40.67 | -0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.65 | -3.57 | 28.65 | -7.68 | 28.65 | -6.21 | 27.07
110.016 | -0.18 | 21.45 | -0.29 | 21.45 | -0.35 | 21.39 | 0.41 | 0 | -29.24 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.33 | -12.54 | 0 | -32.93 | 0 | -29.24 | -0.06 | -38.91 | 0 | 0 | 0.06 | 28.65 | 0.23 | 28.59 | -2.52 | 28.65 | -6.33 | 28.42 | -4.8 | 28.65
110.048 | -0.18 | 21.5 | -0.18 | 21.45 | -0.47 | 21.39 | 0.35 | -0.06 | -26.9 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | -10.9 | -0.18 | -30.7 | 0 | -27.13 | -0.06 | -36.97 | 0.06 | 0 | -0.18 | 28.42 | 0.23 | 28.48 | -1.88 | 28.48 | -4.51 | 28.54 | -3.22 | 28.65
110.08 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | 0.47 | 6.68 | -24.43 | 0 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.27 | -8.44 | -0.06 | -28.54 | -0.06 | -25.08 | 0 | -34.45 | 0 | 0 | -0.06 | 28.59 | 0.23 | 28.48 | -1.17 | 28.54 | -2.7 | 28.65 | -2.17 | 28.54
110.112 | -0.06 | 21.45 | -0.18 | 21.39 | -0.23 | 21.45 | 0.35 | 12.31 | -22.33 | 0.06 | -14 | -0.18 | 0 | -0.06 | 21.21 | 0.23 | 21.39 | -6.68 | 0 | -26.9 | -0.06 | -23.15 | 0 | -33.17 | -0.18 | 0 | -0.06 | 28.54 | 0.18 | 28.54 | -0.47 | 28.48 | -2.23 | 28.36 | -1.23 | 28.59
110.144 | -0.06 | 0.64 | -0.23 | 21.5 | -0.35 | 21.39 | 0.41 | 17.87 | -20.22 | -0.06 | -76.47 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.45 | -5.22 | 5.68 | -24.26 | 0 | -21.21 | 0 | -30.29 | 0.06 | 0 | 0 | 28.42 | 0.12 | 28.71 | -0.35 | 28.59 | -1.11 | 28.54 | -0.76 | 28.42
110.176 | -0.06 | 0.06 | -0.29 | 21.39 | -0.29 | 21.39 | 0.35 | 21.62 | -18.28 | 0 | -53.97 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | -3.75 | 13.59 | -22.97 | 0.06 | -18.93 | 0 | -28.48 | -0.06 | 0 | -0.06 | 28.65 | 0.18 | 28.65 | -0.18 | 28.48 | -0.76 | 28.48 | -0.64 | 28.01
110.208 | -0.18 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.41 | 21.45 | -16.58 | 0 | -42.78 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -2.52 | 21.15 | -20.63 | -0.06 | -17.05 | 0.06 | -27.13 | 0 | -30.35 | 0 | 28.65 | 0.06 | 28.48 | 0 | 28.65 | -0.53 | 28.36 | -0.35 | 28.48
110.24 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.23 | 21.39 | -13.95 | -0.06 | -35.98 | 0 | 0 | 0 | 0 | 0.23 | 21.39 | -1.46 | 28.65 | -18.57 | -0.06 | -15.41 | -0.06 | -24.38 | -0.06 | -76.41 | 0 | 3.87 | 0.06 | 28.48 | 0.12 | 28.59 | -0.18 | 28.59 | -0.18 | 28.48
110.272 | -0.18 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.41 | 21.39 | -12.01 | 0.06 | -31.99 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | -0.7 | 28.71 | -16.23 | -0.06 | -13.89 | -0.06 | -22.38 | -0.06 | -50.45 | -0.18 | 0.12 | 0.06 | 28.48 | -0.06 | 28.65 | -0.06 | 28.42 | -0.18 | 28.42
110.304 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.29 | 21.39 | -9.67 | -0.06 | -32.76 | -0.06 | -46.35 | -0.06 | 0 | 0.18 | 21.45 | -0.53 | 28.48 | -15 | 2.93 | -12.25 | -0.18 | -20.22 | 0.06 | -47.7 | 0 | 0 | 0 | 28.71 | 0.18 | 28.59 | 0 | 28.65 | 0 | 28.48
110.336 | 0 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.35 | 21.39 | -7.56 | -0.06 | -32.93 | -0.06 | -65.33 | 0 | 0 | 0.06 | 21.21 | -0.41 | 28.36 | -13.48 | 11.13 | -10.96 | -0.06 | -18.46 | -0.06 | -42.42 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.42 | 0.06 | 28.48 | 0.12 | 28.48
110.368 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.23 | 21.39 | -5.68 | -0.06 | -31.06 | -0.06 | -41.54 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 28.89 | -12.07 | 18.57 | -9.08 | 6.97 | -16.23 | 0 | -36.97 | 0 | 0 | 0 | 28.48 | 0.18 | 28.48 | 0.23 | 28.65 | 0 | 28.54
110.4 | -0.06 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.18 | 21.39 | -2.64 | -0.06 | -29.06 | -0.06 | -27.31 | -0.06 | 0 | 0.18 | 21.45 | -0.06 | 28.36 | -10.08 | 28.95 | -8.14 | 17.11 | -14.65 | 0.06 | -37.56 | 0 | 0 | 0 | 28.3 | 0.23 | 28.42 | 0.12 | 28.42 | 0.18 | 28.42
110.432 | -0.06 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | 0.12 | 21.39 | -2.11 | -0.06 | -26.84 | -0.06 | -33.81 | 0 | 0 | 0.18 | 21.39 | -0.06 | 28.48 | -8.91 | 28.71 | -6.68 | 24.96 | -13.48 | 0 | -38.09 | 0 | 0 | 0 | 28.48 | 0.23 | 28.48 | 0.06 | 28.59 | 0.18 | 29.01
110.464 | -0.06 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | 0.23 | 21.39 | -0.94 | 1.93 | -25.02 | -0.06 | -36.15 | 0.06 | 0 | 0.23 | 21.39 | 0.12 | 28.65 | -7.68 | 28.71 | -5.74 | 28.77 | -11.84 | -0.06 | -36.68 | -0.18 | 0 | -0.06 | 28.54 | 0.06 | 28.24 | 0.12 | 28.42 | 0.18 | 28.65
110.496 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.06 | 21.39 | -0.64 | 9.38 | -23.09 | -0.18 | -33.46 | 0 | 0 | 0.18 | 21.39 | 0.18 | 29.01 | -6.56 | 28.42 | -4.63 | 28.36 | -10.49 | 0 | -35.45 | -0.18 | 0 | 0 | 28.48 | 0.06 | 28.48 | 0.18 | 28.42 | 0.18 | 28.3
110.528 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.06 | 21.39 | -0.18 | 14.47 | -20.92 | -0.06 | -30.12 | -0.06 | 0 | 0.06 | 21.45 | 0.12 | 28.65 | -5.57 | 28.65 | -3.69 | 28.48 | -8.85 | -0.06 | -33.4 | 0 | 0 | -0.06 | 28.42 | 0 | 28.48 | 0.18 | 28.24 | 0.12 | 28.42
110.56 | -0.06 | -6.21 | -0.18 | 0.12 | -0.18 | 21.45 | 0.06 | 21.45 | 0.06 | 21.68 | -18.52 | -0.06 | -27.6 | -0.06 | -44.59 | 0.18 | 21.39 | 0.18 | 28.54 | -4.39 | 28.59 | -2.93 | 28.13 | -7.15 | 0 | -30.7 | -0.18 | 0 | -0.06 | 28.42 | 0.06 | 28.48 | 0.23 | 28.54 | 0.06 | 28.42
110.592 | 0 | -61.88 | -0.18 | 0 | -0.06 | 0.12 | 0.06 | 21.33 | 0.23 | 21.39 | -16.76 | -0.06 | -26.84 | 0 | -63.17 | 0.23 | 21.39 | 0.18 | 28.54 | -3.46 | 28.54 | -2.11 | 28.24 | -6.21 | -0.18 | -29.24 | 0 | 0 | -0.06 | 28.48 | 0 | 28.42 | 0.23 | 28.48 | 0.18 | 28.42
110.624 | -0.06 | -50.04 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.45 | 0.23 | 21.39 | -15 | -0.18 | -25.9 | 0 | -41.43 | 0.18 | 21.39 | 0.06 | 28.65 | -2.64 | 28.48 | -1.64 | 28.65 | -5.33 | 0 | -27.72 | -0.06 | 0 | 0 | 28.48 | 0 | 28.42 | 0.18 | 28.48 | 0.12 | 28.42
110.656 | -0.06 | -40.67 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -12.42 | -0.06 | -23.91 | 0.12 | -27.89 | 0.12 | 21.33 | 0.18 | 28.71 | -2.23 | 28.59 | -1.29 | 28.95 | -4.1 | 0 | -25.55 | -0.18 | 0 | 0 | 28.42 | 0.06 | 28.59 | 0.18 | 28.48 | 0.12 | 28.42
110.688 | -0.06 | -39.14 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | -10.66 | -0.06 | -21.91 | 0 | -30.7 | 0.12 | 0.23 | 0.06 | 28.77 | -1.88 | 28.59 | -0.94 | 28.48 | -3.46 | 0.12 | -23.91 | -0.06 | -69.08 | -0.06 | 25.9 | 0 | 28.65 | 0.12 | 28.54 | 0.23 | 28.42
110.72 | -0.06 | -36.45 | -0.18 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.33 | -7.97 | 0.06 | -19.45 | -0.06 | -35.57 | 0.12 | 0 | 0.12 | 28.3 | -1.46 | 28.59 | -0.47 | 28.59 | -2.4 | 0 | -21.62 | -0.06 | -63.23 | 0.06 | 0.12 | 0 | 23.85 | 0 | 28.48 | 0.18 | 28.48
110.752 | -0.06 | -34.45 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -5.45 | -0.06 | -18.4 | 0 | -33.87 | 0.12 | 0 | 0.12 | 28.59 | -0.94 | 28.65 | -0.23 | 28.59 | -1.82 | -0.12 | -19.28 | -0.18 | -48.17 | 0 | 0 | 0.06 | 0.18 | 0 | 23.26 | 0.18 | 28.42
110.784 | -0.06 | -32.52 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.27 | 0.53 | 21.39 | -3.69 | 0.06 | -16.47 | 0 | -30.29 | 0.18 | 0 | 0.18 | 28.59 | -0.7 | 28.42 | -0.47 | 28.77 | -1.17 | -0.18 | -17.7 | -0.18 | -45.94 | 0.06 | 0 | 0 | 0 | 0 | 0.18 | 0.06 | 3.81
110.816 | 0 | -30.47 | 0 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.53 | 21.33 | -2.17 | 0 | -14.59 | -0.06 | -27.77 | 0.06 | 0 | 0.18 | 28.59 | -0.7 | 28.71 | -0.29 | 28.65 | -0.94 | 6.86 | -15.76 | -0.06 | -39.32 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0.12
110.848 | 0 | -28.07 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -1.46 | 1.05 | -12.95 | 0 | -27.36 | 0.06 | 0 | 0.12 | 28.59 | -0.47 | 28.77 | -0.29 | 28.48 | -0.59 | 14.71 | -14.06 | 0 | -36.92 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 0
110.88 | -0.06 | -25.31 | 0 | 0 | -0.12 | 0 | 0 | 21.33 | 0.35 | 21.33 | -0.53 | 10.2 | -11.02 | 0.06 | -25.84 | 0.12 | 0 | 0.18 | 28.65 | -0.41 | 28.54 | -0.18 | 28.65 | -0.35 | 24.84 | -11.31 | 0 | -37.91 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0.06
110.912 | -0.06 | -23.91 | -0.06 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | -0.18 | 15.59 | -9.08 | 0.06 | -23.85 | 0 | 0 | 0.18 | 28.59 | -0.29 | 28.48 | -0.23 | 28.48 | -0.29 | 28.83 | -9.43 | 0 | -37.85 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 0
110.944 | -0.06 | -21.62 | -0.18 | 0 | -0.06 | 0 | -0.23 | 21.39 | 0.23 | 21.39 | 0 | 21.27 | -7.44 | 4.92 | -21.68 | 0 | 0 | 0.23 | 28.48 | -0.29 | 28.65 | -0.29 | 29.06 | -0.23 | 28.48 | -7.32 | -0.18 | -36.04 | -0.06 | 0 | -0.06 | 0.06 | 0 | 0 | 0.06 | 0
110.976 | -0.06 | -19.86 | -0.06 | -19.63 | -0.06 | 0 | -0.06 | 0.18 | 0.29 | 21.39 | 0.12 | 21.39 | -5.74 | 10.9 | -19.86 | 0.12 | 0 | 0.18 | 28.54 | -0.35 | 28.65 | -0.18 | 28.65 | -0.29 | 28.42 | -5.1 | -0.18 | -34.22 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 0
111.008 | 0 | -17.81 | -0.18 | -63.28 | -0.06 | -22.74 | -0.06 | 0 | 0.18 | 21.33 | 0.29 | 21.39 | -4.39 | 16.35 | -18.87 | 0 | 0 | 0.06 | 28.59 | -0.35 | 28.65 | -0.29 | 28.54 | -0.35 | 28.3 | -3.11 | 0 | -32.52 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
111.04 | -0.06 | -15.41 | -0.06 | -42.54 | -0.06 | -57.37 | -0.06 | 0 | 0.18 | 21.27 | 0.29 | 21.39 | -2.64 | 21.39 | -16.76 | 0.06 | 0 | 0.06 | 0.23 | -0.29 | 28.59 | -0.23 | 28.42 | -0.29 | 28.42 | -2.11 | -0.18 | -30.47 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0.12 | 0
111.072 | 0 | -14 | -0.06 | -40.14 | -0.12 | -42.89 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.39 | -1.7 | 21.33 | -14.47 | -0.06 | 0 | 0.06 | 0 | -0.23 | 28.36 | -0.18 | 28.3 | -0.23 | 28.65 | -0.7 | -0.06 | -28.95 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0 | 0
111.104 | -0.06 | -12.54 | 0 | -37.5 | -0.06 | -39.38 | -0.06 | 0 | 0.12 | 21.27 | 0.41 | 21.33 | -0.82 | 21.33 | -13.24 | 0.06 | -12.72 | 0.06 | 0 | -0.18 | 28.54 | -0.18 | 28.77 | -0.29 | 28.42 | -0.64 | 0 | -26.84 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
111.136 | 8.14 | -11.25 | 0 | -35.63 | -0.06 | -38.32 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -0.41 | 21.33 | -11.48 | 0.06 | -73.83 | 0.23 | 0 | -0.29 | 28.59 | -0.18 | 28.65 | -0.29 | 28.65 | -0.35 | 0.06 | -24.79 | -0.06 | -73.36 | -0.18 | 0 | -0.18 | 0 | 0.12 | 0
111.168 | 13.83 | -9.79 | -0.06 | -34.04 | -0.06 | -35.39 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.15 | -0.53 | 21.39 | -9.61 | 0 | -53.09 | 0.12 | 0 | -0.23 | 28.59 | -0.18 | 28.65 | -0.23 | 28.65 | -0.06 | -0.06 | -23.15 | -0.18 | -70.73 | -0.06 | -50.92 | -0.18 | 0 | 0.06 | 0
111.2 | 19.45 | -7.79 | -0.06 | -31.41 | -0.06 | -33.17 | -0.06 | 0 | 0.12 | 21.33 | 0.35 | 21.33 | -0.18 | 21.39 | -7.79 | -0.06 | -29.71 | 0 | 0 | -0.18 | 0.29 | -0.18 | 28.3 | -0.18 | 28.42 | 0.06 | 3.34 | -20.74 | 0 | -48.05 | -0.06 | -70.32 | 0 | -73.77 | 0 | -5.1
111.232 | 21.56 | -6.33 | -0.06 | -30 | -0.18 | -31.41 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.33 | -0.06 | 21.33 | -5.51 | 0 | -30.06 | 0.23 | 0 | -0.23 | 0.12 | -0.18 | 26.37 | -0.18 | 28.42 | 0.12 | 11.6 | -18.63 | -0.18 | -45.35 | -0.06 | -43.3 | 0 | -71.02 | 0 | -89.12
111.264 | 21.45 | -4.92 | 0 | -27.54 | -0.06 | -29.47 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.39 | -0.06 | 21.33 | -4.1 | 0.06 | -34.98 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0.18 | -0.18 | 28.24 | 0.29 | 19.34 | -16.88 | -0.18 | -39.55 | -0.06 | -29.24 | 0 | -43.83 | 0.06 | -63.87
111.296 | 21.5 | -3.63 | -0.06 | -25.78 | -0.06 | -27.6 | 0 | 0 | 0 | 21.33 | 0.35 | 21.39 | 0.29 | 21.39 | -2.81 | -0.06 | -34.63 | 0.12 | 0 | -0.18 | 0.06 | -0.06 | 0 | -0.18 | 28.48 | 0.29 | 27.01 | -15.06 | 0 | -36.45 | 0 | -36.45 | 0 | -28.48 | 0 | -36.74
111.328 | 21.5 | -2.46 | -0.06 | -22.91 | -0.18 | -25.55 | -0.06 | 0 | -0.18 | 21.27 | 0.23 | 21.33 | 0.18 | 21.15 | -2.05 | 0 | -32.34 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.89 | 0.35 | 28.83 | -13.24 | -0.18 | -37.91 | -0.06 | -42.78 | 0 | -35.92 | 0 | -29.77
111.36 | 21.45 | -1.11 | -0.18 | -20.86 | -0.06 | -22.91 | -0.06 | 0 | -0.06 | 0.12 | 0.23 | 21.39 | 0.23 | 21.33 | -1.05 | -0.06 | -28.24 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.3 | 0.23 | 28.59 | -10.66 | -0.06 | -37.5 | -0.06 | -38.2 | 0 | -43.07 | 0 | -40.9
111.392 | 21.45 | -0.76 | -0.06 | -18.69 | -0.06 | -21.09 | -0.06 | -45.24 | 0 | 0 | 0.18 | 21.27 | 0.06 | 21.21 | -0.59 | 0 | -27.13 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.54 | 0.41 | 28.42 | -8.56 | -0.06 | -36.62 | 0 | -35.63 | -0.06 | -38.85 | 0 | -41.6
111.424 | 21.45 | -0.29 | 0 | -17.17 | -0.06 | -18.93 | -0.06 | -65.69 | 0.06 | 0 | 0.18 | 21.39 | 0.18 | 21.39 | -0.53 | 0 | -26.37 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.77 | 0.35 | 28.65 | -6.5 | -0.18 | -34.22 | 0 | -34.22 | -0.06 | -35.27 | 0 | -38.15
111.456 | 21.45 | -0.29 | -0.06 | -15.18 | -0.06 | -16.93 | -0.06 | -46.76 | -0.06 | 0.06 | 0.06 | 21.15 | 0.12 | 21.39 | -0.06 | -0.06 | -24.67 | 0.06 | -30.06 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.54 | 0.41 | 28.83 | -4.39 | -0.06 | -32.17 | -0.18 | -33.69 | -0.06 | -32.7 | -0.18 | -34.16
111.488 | 21.45 | -0.18 | -0.06 | -13.95 | -0.06 | -15.35 | -0.06 | -40.2 | -0.06 | 0 | 0.12 | 21.39 | 0.29 | 21.27 | -0.18 | -0.06 | -22.68 | 0.12 | -75.06 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.23 | 28.65 | -2.7 | 0.35 | -30.53 | -0.06 | -31.93 | -0.06 | -32.64 | -0.18 | -32.46
111.52 | 21.45 | -0.23 | -0.06 | -11.48 | -0.06 | -13.36 | 0 | -33.22 | -0.06 | 0 | -0.06 | 21.27 | 0.29 | 21.33 | 0.06 | -0.06 | -20.57 | 0.06 | -47.17 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.3 | 0.23 | 28.42 | -1.7 | 12.48 | -28.95 | -0.06 | -29.24 | -0.06 | -31.99 | -0.18 | -32.52
111.552 | 21.45 | -0.23 | 0 | -10.31 | -0.06 | -12.01 | -0.06 | -31.82 | 0.06 | 0 | 0.06 | 21.09 | 0.23 | 21.39 | 0.18 | -0.06 | -19.45 | 0.12 | -45.24 | -0.06 | 0 | 0 | 0 | -0.18 | 28.42 | 0.18 | 28.54 | -0.7 | 20.33 | -26.37 | -0.06 | -27.77 | -0.06 | -29.71 | -0.18 | -31.99
111.584 | 21.45 | -0.35 | 7.5 | -9.02 | -0.06 | -10.55 | 0 | -32.81 | -0.06 | 0 | 0.12 | 21.33 | 0.06 | 21.39 | 0.06 | -0.06 | -18.22 | 0.12 | -45.12 | 0 | 0 | -0.06 | 0 | 0 | 28.54 | 0.12 | 28.42 | -0.41 | 27.95 | -25.02 | -0.12 | -26.49 | -0.18 | -27.83 | -0.18 | -29.42
111.616 | 21.45 | -0.23 | 13.13 | -7.85 | -0.06 | -9.14 | -0.06 | -32.52 | -0.18 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.18 | 5.04 | -16.06 | 0.12 | -41.31 | -0.18 | -3.22 | -0.18 | 0 | -0.18 | 28.59 | 0.18 | 28.48 | -0.18 | 28.71 | -23.44 | -0.18 | -25.37 | -0.06 | -26.25 | -0.18 | -27.01
111.648 | 21.45 | -0.29 | 18.69 | -6.39 | -0.06 | -7.85 | 0 | -30.94 | -0.06 | 0 | 0.06 | 21.27 | 0.12 | 21.15 | 0.18 | 10.96 | -14.24 | 0.12 | -40.84 | -0.06 | -72.66 | -0.06 | 0 | 0 | 28.48 | 0.12 | 28.42 | 0.12 | 28.65 | -21.5 | 0 | -23.56 | -0.06 | -25.61 | -0.18 | -25.9
111.68 | 21.45 | -0.35 | 21.5 | -4.34 | -0.06 | -5.74 | -0.06 | -28.89 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.27 | 0.23 | 18.34 | -12.13 | 0.12 | -38.67 | -0.18 | -52.33 | -0.06 | -63.23 | -0.18 | 26.13 | 0.06 | 28.48 | 0.18 | 28.48 | -18.69 | 0 | -21.91 | 0 | -23.32 | -0.18 | -24.38
111.712 | 21.45 | -0.23 | 21.45 | -2.87 | 2.46 | -4.45 | -0.06 | -26.43 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.33 | 0.35 | 21.62 | -10.2 | 0.18 | -35.57 | -0.06 | -45.06 | -0.06 | -65.86 | 0 | 0.18 | 0.12 | 28.24 | 0.35 | 28.42 | -16.47 | 0 | -20.27 | -0.06 | -21.56 | -0.06 | -22.91
111.744 | 21.45 | -0.18 | 21.5 | -1.76 | 8.09 | -2.99 | 0 | -23.67 | -0.06 | 0 | 0 | 0.23 | 0.06 | 21.39 | 0.18 | 21.39 | -8.56 | 0.12 | -34.16 | -0.06 | -45.06 | -0.06 | -45.29 | -0.06 | 0 | 0.23 | 28.48 | 0.35 | 28.95 | -14.59 | -0.18 | -18.52 | 0 | -20.1 | -0.06 | -21.33
111.776 | 21.45 | -0.18 | 21.45 | -1.23 | 13.77 | -1.99 | 2.4 | -22.21 | 0 | -5.57 | -0.06 | 0.06 | 0.06 | 21.21 | 0.23 | 21.39 | -6.86 | 0.12 | -32.76 | -0.06 | -41.54 | -0.06 | -45.53 | -0.06 | 0 | 0 | 28.77 | 0.29 | 28.65 | -13.01 | -0.06 | -17.23 | 0.06 | -18.28 | -0.06 | -19.98
111.808 | 21.45 | -0.23 | 21.5 | -0.35 | 19.28 | -1.11 | 8.2 | -19.98 | 0 | -74.42 | -0.06 | 0 | 0.12 | 21.33 | 0.29 | 21.27 | -5.22 | 0.06 | -29.94 | -0.06 | -40.78 | 0 | -44.24 | -0.18 | 0 | -0.18 | 28.54 | 0.29 | 28.54 | -11.13 | 2.46 | -15.65 | 0.06 | -16.93 | -0.18 | -18.22
111.84 | 21.5 | -0.23 | 21.45 | -0.18 | 21.5 | -0.35 | 15.41 | -16.93 | 0.06 | -49.63 | -0.06 | 0.06 | -0.06 | 0.12 | 0.18 | 21.33 | -3.4 | 0.06 | -27.77 | 0 | -38.32 | -0.06 | -41.66 | -0.18 | 0 | -0.18 | 28.48 | 0.29 | 28.48 | -8.32 | 12.66 | -13.36 | 0 | -15 | 0.53 | -16.11
111.872 | 21.39 | -0.18 | 21.5 | -0.18 | 21.5 | -0.18 | 21.21 | -15.35 | -0.06 | -41.84 | 0 | 0 | 0.06 | 0 | 0.23 | 21.39 | -2.23 | 0.12 | -25.78 | -0.06 | -37.15 | -0.06 | -39.9 | -0.06 | 0 | -0.06 | 28.3 | 0.29 | 28.48 | -6.09 | 20.27 | -11.37 | 0.29 | -13.07 | 10.25 | -14.53
111.904 | 21.45 | -0.23 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | -13.24 | -0.06 | -36.51 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -1.52 | -0.18 | -23.91 | 0 | -35.45 | -0.06 | -37.68 | -0.06 | 0 | -0.18 | 28.65 | 0.35 | 28.48 | -4.34 | 27.6 | -9.84 | 9.96 | -11.66 | 17.7 | -12.89
111.936 | 21.39 | -0.18 | 21.45 | -0.35 | 21.45 | -0.23 | 21.39 | -11.78 | -0.18 | -32.34 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | -0.82 | -0.06 | -22.44 | -0.18 | -33.05 | 0 | -36.1 | -0.06 | 0 | -0.18 | 28.48 | 0.12 | 28.54 | -2.58 | 28.48 | -8.56 | 17.05 | -9.9 | 25.37 | -11.02
111.968 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | -0.35 | 21.33 | -9.2 | 0 | -32.52 | -0.06 | 0 | 0 | 0 | 0.12 | 21.33 | -0.53 | -0.06 | -19.69 | -0.06 | -31.06 | -0.06 | -33.93 | 0 | 0 | -0.06 | 28.36 | 0.18 | 28.36 | -1.76 | 28.3 | -6.8 | 24.84 | -8.56 | 28.65 | -9.61
112 | 13.95 | -0.23 | 21.5 | -0.41 | 21.5 | -0.47 | 21.33 | -6.62 | 0 | -32.23 | -0.18 | 0 | 0 | 0 | 0.23 | 21.33 | -0.23 | -0.18 | -17.75 | -0.06 | -28.65 | 0 | -31.52 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.42 | -0.47 | 28.48 | -4.45 | 28.71 | -6.21 | 28.48 | -7.97
112.032 | 0.12 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | -4.28 | -0.06 | -31.41 | -0.06 | 0 | 0 | 0 | 0.29 | 21.27 | -0.18 | 0.06 | -16.23 | 0 | -27.07 | -0.06 | -30.06 | 0 | 0 | 0 | 28.59 | 0.18 | 28.48 | -0.53 | 28.48 | -3.16 | 28.59 | -4.75 | 28.59 | -6.15
112.064 | 0 | -0.18 | 21.45 | -0.41 | 21.45 | -0.18 | 21.39 | -2.4 | -0.06 | -29.59 | -0.06 | 0 | -0.06 | 0 | 0.23 | 21.39 | 0.06 | 0.12 | -14.24 | 0 | -25.02 | -0.06 | -28.01 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.54 | -0.06 | 28.48 | -2.46 | 28.48 | -3.57 | 28.42 | -4.63
112.096 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | -1.7 | 0 | -27.6 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | 0.06 | 0 | -12.83 | 0 | -22.79 | 0 | -26.31 | -0.06 | 0 | -0.18 | 3.16 | 0.12 | 28.48 | 0 | 28.48 | -1.05 | 28.54 | -2.4 | 28.59 | -2.99
112.128 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | -1 | -0.06 | -25.37 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | 0 | 2.34 | -11.66 | -0.06 | -20.98 | 0 | -23.91 | 0 | -45.82 | 0 | 0.12 | 0.12 | 28.48 | 0.23 | 28.59 | -0.82 | 28.3 | -1.46 | 28.59 | -1.99
112.16 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | -0.23 | 21.39 | -0.29 | -0.06 | -22.38 | 0 | -22.68 | 0 | 0 | 0.18 | 21.21 | 0.12 | 9.79 | -9.9 | -0.06 | -18.22 | -0.06 | -21.62 | 0 | -71.84 | -0.06 | 0 | 0.12 | 28.54 | 0.23 | 28.36 | -0.64 | 28.48 | -0.94 | 28.48 | -1.23
112.192 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | -0.18 | 21.45 | 0.18 | -0.06 | -20.39 | -0.18 | -76.88 | 0 | 0 | 0.18 | 21.21 | 0.18 | 15.18 | -8.56 | 0 | -16.52 | -0.06 | -19.28 | 0 | -48.75 | 0 | 0.06 | 0.06 | 28.42 | 0.29 | 28.48 | -0.18 | 28.3 | -0.47 | 28.48 | -0.7
112.224 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | 0.23 | 0 | -18.69 | -0.18 | -51.92 | -0.06 | 0 | 0.12 | 21.33 | 0.23 | 20.98 | -7.44 | -0.06 | -15 | -0.06 | -17.4 | -0.06 | -46.41 | 0 | 0 | 0.06 | 28.3 | 0.29 | 28.42 | -0.18 | 28.77 | -0.35 | 28.48 | -0.41
112.256 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.29 | -0.06 | -16.82 | 0.06 | -42.25 | -0.18 | -34.22 | 0.06 | 21.39 | 0.23 | 21.56 | -6.21 | -0.06 | -13.71 | -0.06 | -15.94 | -0.06 | -42.07 | -0.18 | 0 | 0 | 28.36 | 0.35 | 28.48 | -0.06 | 28.48 | -0.18 | 28.65 | -0.18
112.288 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.23 | 21.39 | 0.35 | -0.06 | -14.77 | -0.06 | -36.86 | -0.06 | -69.91 | 0.06 | 21.15 | 0.12 | 21.15 | -5.22 | 0 | -12.25 | -0.06 | -14.53 | 0 | -36.1 | 0 | 0 | 0 | 28.65 | 0.29 | 28.42 | 0 | 28.42 | 0 | 28.48 | -0.06
112.32 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.39 | 0.53 | -0.06 | -12.13 | -0.06 | -31.47 | -0.18 | -39.61 | 0.06 | 21.39 | 0.18 | 21.45 | -4.04 | -0.06 | -10.37 | 2.23 | -12.36 | 0.06 | -37.5 | 0 | 0 | 0 | 28.54 | 0.29 | 28.48 | 0.18 | 28.48 | 0.12 | 28.71 | 0.06
112.352 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.53 | -0.06 | -10.25 | 0 | -32.34 | -0.06 | -27.54 | -0.06 | 21.15 | 0.12 | 21.45 | -3.16 | 0.06 | -9.2 | 9.96 | -11.25 | -0.06 | -38.15 | 0 | 0 | -0.06 | 28.36 | 0.29 | 28.36 | 0 | 28.65 | 0.18 | 28.89 | 0
112.384 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.53 | 0 | -7.79 | -0.06 | -32.64 | -0.06 | -31.06 | 0.12 | 21.33 | 0.12 | 21.45 | -2.64 | 2.64 | -7.97 | 17.17 | -9.61 | -0.06 | -37.21 | -0.06 | 0 | -0.06 | 0.35 | 0 | 28.36 | 0.12 | 28.13 | 0.12 | 28.59 | 0.12
112.416 | -0.12 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.53 | 2.99 | -5.92 | -0.06 | -31.29 | -0.06 | -35.51 | -0.06 | 21.21 | 0.18 | 21.39 | -1.88 | 10.84 | -6.33 | 24.96 | -8.32 | -0.06 | -35.51 | -0.06 | 0 | -0.06 | 0.06 | 0.23 | 28.59 | 0.35 | 28.83 | 0.23 | 28.3 | 0.18
112.448 | -54.2 | -0.06 | 19.81 | -0.06 | 21.45 | -0.18 | 21.39 | 0.53 | 9.38 | -3.69 | -0.06 | -29.24 | -0.06 | -34.22 | 0 | 21.21 | 0.23 | 21.33 | -1.64 | 18.28 | -5.68 | 28.65 | -7.27 | -0.06 | -33.17 | 0 | 0 | -0.06 | 0 | 0.18 | 28.54 | 0.23 | 28.42 | 0.18 | 28.83 | 0.12
112.48 | -53.38 | -0.18 | 0.06 | -0.18 | 21.45 | -0.18 | 21.39 | 0.47 | 16.41 | -1.76 | -0.06 | -26.66 | -0.06 | -30.18 | -0.06 | 21.21 | 0.12 | 21.45 | -1.29 | 28.77 | -4.22 | 28.71 | -5.39 | 5.04 | -30.82 | 0 | 0 | -0.06 | 0.06 | 0.06 | 28.48 | 0.06 | 28.65 | 0.12 | 28.65 | 0.06
112.512 | -40.84 | -0.18 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.47 | 21.68 | -1.23 | -0.06 | -24.43 | -0.06 | -28.59 | -0.06 | 0.12 | 0.23 | 21.39 | -0.94 | 28.89 | -3.34 | 28.95 | -4.45 | 13.18 | -28.48 | 0 | -7.79 | 0 | 0 | 0.06 | 28.42 | 0.18 | 28.42 | 0.23 | 28.48 | 0.12
112.544 | -39.85 | -0.18 | 0 | -0.06 | 21.39 | -0.18 | 21.39 | 0.47 | 21.45 | -0.53 | 0 | -22.68 | -0.06 | -27.07 | -0.06 | 0 | 0.23 | 20.98 | -0.76 | 28.71 | -2.93 | 28.48 | -3.69 | 20.98 | -27.01 | 0 | -92.29 | -0.06 | 0 | 0.06 | 28.42 | 0.12 | 28.18 | 0.18 | 28.59 | 0.12
112.576 | -37.38 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.35 | 21.39 | -0.29 | -0.06 | -20.86 | -0.06 | -25.61 | -0.06 | 0 | 0.18 | 21.45 | -0.47 | 28.71 | -2.11 | 28.83 | -3.11 | 28.48 | -25.43 | -0.06 | -61.41 | 0 | 0 | 0.06 | 28.42 | 0.12 | 28.54 | 0.18 | 28.42 | 0.18
112.608 | -35.27 | -0.06 | 0 | -0.06 | 14.82 | -0.18 | 21.39 | 0.35 | 21.39 | 0.12 | -0.06 | -18.87 | 0 | -23.91 | -0.06 | 0.06 | 0.12 | 21.39 | -0.7 | 28.65 | -1.7 | 28.65 | -2.23 | 28.71 | -24.08 | 0.06 | -47.99 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.48 | 0.23 | 28.48 | 0.18
112.64 | -32.99 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.33 | 0.29 | 21.39 | 0.35 | 0 | -16.47 | -0.06 | -21.5 | 0.06 | 0 | 0.18 | 21.39 | -0.53 | 28.65 | -1.17 | 28.42 | -1.46 | 28.42 | -20.74 | 0 | -43.65 | 0 | 0 | 0 | 28.59 | 0 | 28.36 | 0.12 | 28.48 | 0.06
112.672 | -30.82 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0.23 | 0.23 | 21.39 | 0.35 | 0 | -14.24 | -0.06 | -19.69 | -0.06 | 0 | 0.18 | 21.39 | -0.35 | 28.48 | -0.82 | 28.48 | -1.11 | 28.48 | -18.99 | 0 | -37.85 | 0 | 0 | -0.06 | 28.48 | 0.06 | 28.42 | 0.29 | 28.48 | 0.23
112.704 | -29.06 | 0 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.12 | 21.45 | 0.29 | 0 | -12.07 | -0.06 | -18.46 | -0.06 | 0 | 0.18 | 21.33 | -0.35 | 28.71 | -0.76 | 28.42 | -0.94 | 28.48 | -17.17 | 0 | -36.74 | -0.18 | 0 | -0.06 | 0.29 | 0.06 | 28.48 | 0.12 | 28.42 | 0
112.736 | -26.25 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.23 | 21.27 | 0.41 | -0.06 | -10.25 | -0.06 | -16.88 | 0 | 0 | 0.06 | 21.39 | -0.35 | 28.77 | -0.47 | 28.36 | -0.59 | 28.42 | -15.65 | 0 | -38.32 | 0.06 | 0 | 0.06 | 0.12 | 0.12 | 28.36 | 0.12 | 23.03 | 0.18
112.768 | -23.79 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | -0.06 | -7.97 | 0.06 | -15.23 | 0.06 | 0 | 0.18 | 21.39 | -0.29 | 28.89 | -0.35 | 28.3 | -0.29 | 28.59 | -13.48 | 0 | -37.79 | -0.06 | 0 | -0.06 | 0 | 0 | 26.02 | 0.18 | 0.18 | 0.12
112.8 | -21.97 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 1.58 | -5.33 | -0.06 | -12.95 | -0.06 | 0 | 0.12 | 21.39 | -0.35 | 28.48 | -0.23 | 28.59 | -0.23 | 28.59 | -11.48 | -0.06 | -36.27 | -0.06 | -18.22 | -0.06 | 0 | -0.06 | 0.12 | 0.12 | 0 | 0.12
112.832 | -20.39 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 9.26 | -3.22 | -0.18 | -11.37 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.65 | -0.29 | 28.83 | -0.47 | 28.95 | -9.2 | -0.06 | -33.93 | 0.06 | -92.23 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.06
112.864 | -18.28 | -0.06 | -0.06 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.45 | 0.41 | 14.47 | -1.7 | -0.06 | -9.67 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.42 | -0.23 | 28.48 | -0.41 | 28.42 | -6.86 | -0.06 | -32.17 | -0.06 | -59.24 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.12
112.896 | -16.23 | 0 | -40.43 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 20.33 | -1.46 | -0.06 | -8.2 | 0.06 | 0 | 0.23 | 21.39 | -0.18 | 28.48 | -0.18 | 28.89 | -0.18 | 28.24 | -4.8 | 4.92 | -29.77 | -0.18 | -49.22 | -0.18 | 0 | -0.06 | 0 | 0.23 | 0 | 0.06
112.928 | -14.71 | -0.06 | -54.73 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -0.35 | 0 | -6.33 | -0.06 | -30.29 | 0.18 | 21.39 | -0.29 | 28.59 | -0.35 | 28.48 | -0.35 | 28.59 | -3.22 | 13.18 | -28.24 | -0.06 | -44.24 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06
112.96 | -13.07 | 0 | -39.9 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -0.18 | -0.06 | -4.1 | 0.06 | -63.75 | 0.18 | 21.39 | -0.35 | 28.48 | -0.29 | 28.42 | -0.29 | 29.06 | -1.82 | 23.2 | -25.9 | 0.06 | -37.79 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06
112.992 | -11.48 | -0.06 | -38.85 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | 0.12 | 6.27 | -2.93 | 0.06 | -39.9 | 0.23 | 21.39 | -0.18 | 28.42 | -0.18 | 28.65 | -0.35 | 28.24 | -0.76 | 28.83 | -24.08 | -0.18 | -37.15 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0
113.024 | -10.08 | -0.06 | -37.03 | -0.06 | -0.18 | -0.06 | 0 | -0.06 | 21.45 | 0.23 | 21.39 | 0.12 | 12.07 | -1.41 | -0.06 | -27.95 | 0.12 | 0.29 | -0.18 | 28.48 | -0.18 | 28.18 | -0.23 | 28.83 | -0.59 | 28.65 | -22.62 | 0 | -37.91 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.06
113.056 | -8.32 | -0.06 | -35.16 | -0.06 | -54.67 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.21 | 0.18 | 17.52 | -1.41 | -0.06 | -31.29 | 0.18 | 0.12 | -0.23 | 28.42 | -0.23 | 28.36 | -0.23 | 28.59 | -0.06 | 28.65 | -20.27 | 0 | -37.5 | 0 | 0 | -0.18 | 0 | 0.06 | 0 | -0.06
113.088 | -6.97 | -0.06 | -33.69 | -0.06 | -53.26 | -0.06 | -0.47 | -0.06 | 21.56 | 0.23 | 21.33 | 0.41 | 21.62 | -0.59 | 0.06 | -35.57 | 0.12 | 0 | -0.18 | 28.42 | -0.18 | 28.54 | -0.23 | 28.54 | 0 | 28.48 | -18.69 | 0 | -35.92 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0
113.12 | -5.39 | -0.06 | -30.41 | -0.06 | -40.9 | -0.06 | -77.23 | 0 | 21.39 | 0.12 | 21.33 | 0.47 | 21.39 | -0.35 | -0.06 | -33.75 | 0.12 | 0.06 | -0.18 | 28.3 | -0.18 | 28.42 | -0.23 | 28.48 | 0.23 | 28.42 | -16.29 | 0 | -33.63 | 0 | -15.76 | -0.18 | 0 | 0 | 0 | -0.06
113.152 | -4.1 | -0.06 | -28.83 | -0.06 | -38.79 | 0 | -52.44 | -0.18 | 21.39 | 0.12 | 21.39 | 0.47 | 21.5 | -0.18 | 0 | -29.83 | 0.18 | 0 | -0.18 | 28.59 | -0.18 | 28.89 | -0.06 | 28.42 | 0.35 | 28.48 | -14.41 | 0 | -31.29 | 0 | -89.65 | -0.18 | 0 | 0 | -0.06 | 0
113.184 | -2.75 | 0 | -26.54 | -0.06 | -36.74 | -0.06 | -43.13 | 0 | 21.33 | 0.12 | 21.39 | 0.41 | 21.15 | -0.06 | -0.06 | -28.01 | 0.18 | 0 | -0.18 | 28.48 | -0.18 | 28.3 | -0.18 | 28.48 | 0.29 | 28.54 | -12.42 | 0 | -29.71 | -0.18 | -60.88 | -0.06 | 0 | 0 | -80.75 | 0
113.216 | -1.46 | 0 | -25.08 | -0.06 | -35.22 | 0 | -37.68 | -0.06 | 21.39 | 0.23 | 21.33 | 0.35 | 21.33 | -0.06 | 0 | -27.36 | 0.12 | 0.06 | -0.18 | 28.42 | -0.18 | 26.19 | -0.18 | 28.42 | 0.47 | 28.42 | -10.61 | -0.06 | -27.54 | -0.06 | -34.1 | -0.06 | -76 | 0 | -70.73 | -0.06
113.248 | -0.76 | 1.99 | -22.62 | 0 | -32.7 | -0.06 | -32.76 | 0.06 | 21.45 | 0.06 | 21.39 | 0.35 | 21.39 | 0.06 | 4.16 | -26.25 | 0.18 | 0 | -0.18 | 28.59 | -0.18 | 0.12 | -0.18 | 28.77 | 0.47 | 28.65 | -8.79 | 0 | -26.37 | -0.06 | -29.94 | -0.18 | -71.96 | 0 | -40.61 | 0
113.28 | -0.47 | 8.85 | -19.92 | -0.06 | -30.53 | -0.06 | -32.81 | -0.06 | 21.21 | 0.06 | 21.15 | 0.35 | 21.39 | 0.18 | 11.84 | -23.32 | 0.12 | 0 | -0.18 | 0.29 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.59 | -5.39 | 0 | -23.91 | 0 | -42.42 | -0.06 | -35.92 | 0 | -30 | 0.06
113.312 | -0.06 | 14.3 | -17.93 | -0.06 | -28.18 | -0.06 | -32.64 | -0.06 | 0.12 | 0 | 21.39 | 0.35 | 21.33 | 0.18 | 17.52 | -21.45 | 0.18 | 0 | -0.18 | 0.12 | -0.18 | 0 | -0.06 | 28.48 | 0.35 | 28.48 | -3.4 | -0.06 | -22.44 | -0.18 | -41.49 | -0.06 | -29.77 | 0 | -39.55 | -0.06
113.344 | -0.06 | 20.04 | -16 | 0 | -26.25 | -0.06 | -31.29 | 0 | 0.06 | 0 | 21.33 | 0.35 | 21.27 | 0.18 | 21.62 | -19.86 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.59 | 0.41 | 28.3 | -2.05 | -0.06 | -20.45 | 0 | -37.03 | -0.18 | -39.38 | 0 | -42.25 | 0
113.376 | -0.18 | 21.5 | -14.59 | -0.06 | -24.32 | -0.06 | -30.12 | -0.06 | 0 | 0 | 21.33 | 0.29 | 21.33 | 0.18 | 21.45 | -18.75 | 0.06 | 0 | -0.18 | 0.06 | -0.18 | 0 | -0.18 | 0.29 | 0.35 | 28.89 | -1.46 | 8.44 | -18.16 | -0.06 | -33.93 | -0.18 | -42.72 | -0.06 | -39.32 | -0.18
113.408 | -0.35 | 21.45 | -13.07 | -0.06 | -22.74 | 0 | -27.89 | -0.18 | 0 | 0 | 21.45 | 0.23 | 21.39 | 0.18 | 21.39 | -17.29 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0.12 | 0.35 | 28.77 | -0.7 | 15.59 | -16.52 | 0.06 | -32.93 | -0.18 | -38.79 | -0.06 | -34.45 | 0
113.44 | -0.35 | 21.45 | -10.96 | 0 | -19.75 | -0.06 | -24.61 | 0 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | 0.23 | 21.39 | -14.71 | 0.06 | -18.05 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.35 | 28.59 | -0.35 | 26.08 | -14.06 | -0.06 | -32.81 | -0.18 | -33.93 | -0.06 | -32.93 | -0.06
113.472 | -0.41 | 21.45 | -10.02 | -0.06 | -17.29 | -0.06 | -23.2 | 0 | 0 | -0.06 | 21.33 | 0.18 | 21.15 | 0.29 | 21.09 | -13.01 | 0 | -73.77 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.23 | 28.42 | 0 | 28.71 | -12.07 | -0.06 | -31.52 | -0.06 | -34.04 | 0 | -33.28 | 0.06
113.504 | -0.29 | 21.45 | -8.56 | -0.06 | -16.29 | 0 | -20.98 | 0.06 | 0 | -0.06 | 21.27 | 0.12 | 21.15 | 0.18 | 21.39 | -11.54 | 0.12 | -52.44 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.29 | 28.3 | 0.06 | 28.48 | -10.55 | 0 | -29.01 | -0.06 | -33.28 | -0.06 | -31.82 | 0
113.536 | -0.29 | 21.5 | -7.21 | -0.06 | -14.53 | -0.06 | -18.87 | 0 | 0 | -0.06 | 21.33 | 0.06 | 21.04 | 0.23 | 21.33 | -9.84 | 0.06 | -32.34 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.23 | 28.36 | 0 | 28.65 | -8.32 | -0.06 | -27.54 | 0 | -31.82 | 0 | -30.18 | 0
113.568 | -0.29 | 21.45 | -5.39 | -0.06 | -12.83 | -0.06 | -16.47 | -0.06 | 0 | 0 | 21.27 | 0 | 21.39 | 0.23 | 21.33 | -8.32 | 0 | -28.18 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 28.36 | 0.12 | 28.65 | -6.09 | 0 | -26.43 | -0.06 | -30.18 | 0 | -27.42 | 0
113.6 | -0.23 | 21.45 | -3.75 | -0.06 | -11.13 | -0.06 | -14.24 | -0.06 | 0 | 0 | 21.39 | 0 | 21.39 | 0.23 | 21.39 | -5.86 | 0.12 | -34.81 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.65 | 0.23 | 28.18 | -3.75 | -0.18 | -24.08 | -0.06 | -27.01 | 0 | -26.13 | 0
113.632 | -0.18 | 21.45 | -2.7 | -0.06 | -9.43 | 0 | -12.13 | -0.06 | 0 | -0.06 | 21.33 | 0 | 21.39 | 0.18 | 21.33 | -4.28 | 0.06 | -34.98 | -0.18 | 0 | -0.06 | 0.06 | -0.06 | 0 | 0.18 | 28.48 | 0.29 | 28.59 | -1.99 | 0 | -22.38 | -0.06 | -26.31 | -0.06 | -25.02 | 0
113.664 | -0.29 | 21.45 | -1.46 | 3.75 | -8.61 | -0.06 | -10.31 | -0.06 | 0 | -0.06 | 21.27 | -0.06 | 21.27 | 0.18 | 21.33 | -3.22 | -0.06 | -32.58 | -0.18 | 0 | -0.06 | -56.9 | -0.06 | 0 | 0.12 | 28.77 | 0.47 | 28.18 | -1.46 | -0.06 | -21.15 | -0.18 | -25.02 | 0.06 | -22.91 | 0
113.696 | -0.23 | 21.5 | -0.94 | 10.14 | -7.03 | 3.75 | -8.32 | -0.06 | 0 | -0.06 | 2.93 | 0 | 21.27 | 0.12 | 21.33 | -2.11 | 0.06 | -28.48 | -0.18 | -0.12 | 0 | -65.22 | -0.06 | 0 | 0.06 | 28.54 | 0.18 | 28.42 | -0.82 | -0.06 | -19.45 | -0.18 | -23.67 | 0 | -21.86 | 0
113.728 | -0.23 | 21.39 | -0.35 | 15.18 | -5.39 | 10.14 | -5.92 | 0 | -2.81 | 0.06 | 0.06 | -0.06 | 21.15 | 0.23 | 21.39 | -1.17 | -0.06 | -27.25 | -0.06 | -70.43 | -0.06 | -44.3 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.18 | -0.35 | -0.18 | -17.99 | 0 | -21.91 | 0 | -19.92 | -0.06
113.76 | -0.23 | 21.45 | -0.18 | 21.86 | -3.57 | 17.29 | -3.34 | 0.12 | -73.25 | 0.06 | 0 | -0.06 | 21.15 | 0.18 | 21.33 | -0.47 | -0.06 | -26.54 | -0.06 | -53.44 | 0 | -45.18 | -0.06 | 0 | 0.12 | 28.42 | 0.23 | 28.42 | 0.06 | 0 | -16 | -0.06 | -19.75 | 0 | -18.05 | -0.06
113.792 | -0.23 | 21.45 | -0.29 | 21.5 | -2.4 | 21.91 | -1.76 | 0.06 | -49.87 | 0 | 0 | 0 | 21.27 | 0.23 | 21.33 | -0.41 | 3.16 | -24.43 | -0.18 | -44.59 | 0 | -43.3 | -0.06 | -19.16 | 0 | 0.18 | 0.18 | 28.48 | 0 | -0.06 | -14.3 | -0.18 | -17.81 | 0 | -16.7 | 0.06
113.824 | -0.29 | 21.45 | -0.29 | 21.5 | -1.41 | 21.5 | -1.35 | -0.06 | -41.84 | 0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.33 | -0.53 | 9.55 | -22.09 | -0.18 | -44.71 | 0 | -41.66 | -0.06 | -79.05 | 0 | 0 | 0.18 | 28.48 | 0.12 | 9.14 | -12.89 | 0 | -16.52 | -0.06 | -14.59 | -0.18
113.856 | -0.18 | 21.33 | -0.18 | 21.5 | -0.7 | 21.5 | -0.59 | -0.06 | -37.03 | 0 | 0 | -0.06 | 21.39 | 0 | 21.45 | -0.06 | 14.71 | -20.98 | -0.06 | -42.36 | -0.06 | -39.32 | -0.18 | -53.79 | -0.06 | 0 | 0.12 | 28.65 | 0.23 | 16.17 | -11.13 | 7.09 | -15.18 | 0 | -13.36 | -0.06
113.888 | -0.18 | 21.5 | -0.41 | 21.45 | -0.29 | 21.5 | -0.23 | -0.06 | -32.7 | 0.06 | 0 | 0.06 | 0.18 | 0 | 21.33 | 0 | 20.45 | -19.63 | -0.06 | -41.54 | -0.18 | -36.15 | 0 | -45.59 | 0 | 0 | 0.12 | 28.54 | 0.35 | 23.85 | -9.79 | 14.77 | -13.65 | 6.74 | -11.48 | 1.99
113.92 | -0.18 | 21.45 | -0.41 | 21.5 | -0.23 | 21.45 | 0.23 | -0.18 | -32.93 | -0.06 | 0 | 0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | -17.58 | -0.06 | -38.15 | -0.06 | -35.33 | -0.06 | -44.59 | 0 | 0 | 0 | 28.48 | 0.35 | 28.71 | -7.03 | 24.84 | -11.13 | 16.93 | -9.61 | 10.72
113.952 | -0.18 | 21.45 | -0.35 | 21.45 | -0.18 | 21.45 | 0.23 | -0.06 | -32.7 | -0.06 | 0 | 0 | 0 | 0.06 | 21.39 | 0.12 | 21.15 | -15.65 | -0.18 | -35.98 | -0.06 | -34.1 | -0.06 | -41.6 | 0 | 0 | 0.06 | 28.54 | 0.35 | 28.71 | -5.16 | 28.65 | -9.38 | 24.61 | -8.03 | 18.16
113.984 | -0.18 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | 0.47 | 0 | -31.17 | 0 | 0 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | -14.24 | -0.06 | -35.1 | -0.06 | -31.88 | 0 | -40.37 | -0.06 | 0 | -0.06 | 28.42 | 0.23 | 28.54 | -3.93 | 28.48 | -8.14 | 28.83 | -6.09 | 26.13
114.016 | -0.18 | 21.39 | -0.29 | 21.45 | -0.23 | 21.39 | 0.35 | -0.06 | -28.59 | -0.06 | 0 | 0.12 | 0 | -0.06 | 21.15 | 0.12 | 21.33 | -12.72 | 0 | -33.05 | 0 | -29.12 | 0 | -38.38 | 0.06 | 0 | -0.06 | 28.65 | 0.35 | 28.48 | -2.58 | 28.48 | -6.39 | 28.65 | -4.51 | 28.77
114.048 | -0.18 | 21.5 | -0.18 | 21.45 | -0.35 | 21.45 | 0.47 | -0.06 | -27.01 | 0 | 0 | -0.06 | 0 | 0 | 21.21 | 0.06 | 21.39 | -10.66 | -0.06 | -30.94 | -0.06 | -27.77 | 0 | -36.92 | 0 | 0 | 0.06 | 28.42 | 0.35 | 28.36 | -1.99 | 28.48 | -4.8 | 28.65 | -3.22 | 28.48
114.08 | -0.06 | 21.45 | -0.18 | 21.5 | -0.35 | 21.45 | 0.53 | 5.8 | -24.38 | -0.06 | 0 | 0.12 | 0 | 0.06 | 21.39 | 0.18 | 21.45 | -8.32 | -0.06 | -28.18 | 0 | -25.55 | 0 | -34.45 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.48 | -1 | 28.48 | -2.75 | 28.48 | -2.11 | 28.24
114.112 | -0.06 | 21.39 | -0.29 | 21.45 | -0.41 | 21.45 | 0.53 | 11.72 | -22.21 | 0 | -5.39 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.33 | -6.97 | -0.06 | -26.37 | -0.06 | -22.85 | 0 | -32.7 | -0.06 | 0 | 0 | 28.65 | 0.23 | 28.59 | -0.41 | 28.65 | -2.05 | 28.48 | -1.11 | 28.36
114.144 | -0.06 | 7.68 | -0.35 | 21.45 | -0.23 | 21.39 | 0.53 | 17.23 | -20.39 | -0.06 | -73.54 | 0 | 0 | 0 | 0.12 | 0.12 | 21.45 | -5.1 | 4.51 | -24.84 | 0.06 | -21.21 | 0 | -31.06 | 0 | 0 | 0.06 | 28.65 | 0.12 | 28.13 | -0.35 | 28.24 | -1.17 | 28.48 | -1 | 28.3
114.176 | -0.06 | 0.12 | -0.29 | 21.45 | -0.18 | 21.45 | 0.41 | 21.68 | -18.99 | -0.06 | -55.55 | -0.06 | 0 | 0 | 0 | 0.18 | 21.21 | -3.75 | 12.83 | -22.85 | 0 | -19.51 | 0 | -28.95 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.59 | -0.23 | 28.48 | -0.94 | 28.18 | -0.29 | 28.54
114.208 | -0.06 | 0 | -0.29 | 21.45 | -0.23 | 21.45 | 0.53 | 21.45 | -16.76 | 0.06 | -42.78 | 0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -2.52 | 20.22 | -20.74 | 0 | -16.82 | 0 | -26.84 | -0.18 | -28.89 | -0.06 | 28.71 | 0.18 | 28.65 | -0.06 | 28.59 | -0.53 | 28.3 | -0.35 | 28.42
114.24 | -0.18 | 0.06 | -0.23 | 21.45 | -0.29 | 21.45 | 0.53 | 21.39 | -14.18 | -0.06 | -36.27 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.21 | -1.52 | 29.01 | -18.69 | -0.06 | -15.35 | -0.06 | -24.43 | 0 | -79.69 | -0.06 | 18.4 | 0.06 | 28.42 | 0.06 | 28.65 | -0.06 | 28.59 | -0.06 | 28.48
114.272 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.29 | 21.39 | -12.25 | -0.06 | -32.05 | 0 | 0 | 0 | 0 | 0.12 | 21.21 | -0.94 | 28.71 | -16.76 | 0 | -13.83 | -0.06 | -22.74 | -0.06 | -51.97 | -0.06 | 0.18 | 0 | 28.48 | -0.06 | 28.48 | -0.06 | 28.65 | -0.18 | 28.42
114.304 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.23 | 21.39 | -10.14 | 0 | -32.93 | 0.12 | -43.77 | 0 | 0 | 0.23 | 21.21 | -0.35 | 28.65 | -15.06 | 0.64 | -12.6 | -0.06 | -20.16 | 0 | -48.17 | 0 | 0 | 0.06 | 28.59 | 0.23 | 28.65 | -0.06 | 28.83 | 0.06 | 28.65
114.336 | 0 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.23 | 21.33 | -7.91 | 0 | -32.58 | -0.06 | -65.69 | 0 | 0 | 0.23 | 21.33 | -0.18 | 28.42 | -13.71 | 10.9 | -10.78 | 0 | -18.4 | 0 | -42.89 | -0.18 | 0 | 0.12 | 28.54 | 0.12 | 28.48 | -0.06 | 27.89 | 0 | 28.42
114.368 | -0.06 | 0 | -0.18 | 21.45 | -0.35 | 21.39 | 0.29 | 21.33 | -5.74 | -0.06 | -31.64 | -0.06 | -42.54 | -0.06 | 0 | 0.23 | 21.15 | -0.18 | 28.77 | -12.48 | 18.22 | -9.2 | 6.04 | -16.93 | 0.06 | -36.8 | -0.06 | 0 | 0 | 28.36 | 0.06 | 28.48 | 0.12 | 28.48 | 0.12 | 28.71
114.4 | -0.06 | 0 | -0.18 | 21.45 | -0.35 | 21.39 | 0.18 | 21.39 | -3.16 | 0.06 | -29.24 | -0.06 | -27.36 | -0.06 | 0 | 0.23 | 21.33 | 0 | 28.54 | -10.72 | 28.89 | -8.2 | 16.41 | -14.82 | 0.06 | -36.51 | -0.18 | 0 | -0.06 | 28.48 | 0.29 | 28.59 | 0.29 | 28.59 | 0.12 | 28.42
114.432 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.12 | 21.39 | -1.52 | -0.06 | -27.01 | 0 | -32.52 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.83 | -9.32 | 28.65 | -6.68 | 24.32 | -13.36 | 0.12 | -38.26 | 0 | 0 | 0 | 28.42 | 0.18 | 28.42 | 0.35 | 28.48 | 0.23 | 28.48
114.464 | -0.18 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | 0.12 | 21.39 | -1.46 | 2.81 | -25.2 | 0 | -35.98 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.65 | -7.62 | 28.65 | -5.86 | 28.65 | -12.07 | 0 | -37.21 | -0.06 | 0 | -0.06 | 28.18 | 0.18 | 28.42 | 0.18 | 28.42 | 0.23 | 28.59
114.496 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.06 | 21.21 | -0.35 | 8.85 | -22.85 | 0.06 | -34.22 | -0.06 | 0 | 0.18 | 21.33 | 0.12 | 28.48 | -7.03 | 28.48 | -4.63 | 28.65 | -10.55 | -0.06 | -36.04 | 0 | 0 | -0.18 | 28.48 | 0.18 | 28.48 | 0.06 | 28.48 | 0.23 | 28.59
114.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.39 | -0.18 | 14.12 | -21.27 | -0.18 | -30.53 | -0.18 | 0 | 0.18 | 21.39 | 0.18 | 28.3 | -5.57 | 28.59 | -3.4 | 28.71 | -9.26 | 0 | -33.52 | -0.06 | 0 | 0 | 28.59 | 0.18 | 28.42 | 0.23 | 28.65 | 0.29 | 28.42
114.56 | -0.06 | -2.81 | -0.06 | 0.18 | -0.18 | 21.45 | 0.12 | 21.33 | 0 | 21.5 | -18.93 | -0.06 | -27.77 | -0.06 | -37.97 | 0.18 | 21.39 | 0.23 | 28.42 | -4.1 | 28.65 | -2.7 | 28.48 | -7.62 | 0 | -30.82 | -0.06 | 0 | 0 | 28.36 | 0 | 28.36 | 0.12 | 28.65 | 0.23 | 28.3
114.592 | -0.06 | -58.07 | -0.06 | 0 | -0.23 | 0.12 | 0.12 | 21.39 | 0.23 | 21.39 | -16.7 | 0.06 | -26.84 | -0.06 | -65.57 | 0.12 | 21.27 | 0.18 | 28.54 | -3.4 | 28.48 | -2.05 | 28.83 | -6.15 | -0.18 | -29.06 | -0.18 | 0 | -0.18 | 28.59 | 0 | 28.48 | 0.18 | 28.24 | 0.23 | 28.42
114.624 | -0.06 | -50.45 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.27 | 0.35 | 21.39 | -15.47 | -0.06 | -25.9 | -0.06 | -44.65 | 0.12 | 21.39 | 0.12 | 28.59 | -2.7 | 28.42 | -1.7 | 28.54 | -5.04 | 0 | -27.36 | -0.06 | 0 | -0.06 | 28.59 | 0 | 28.42 | 0.12 | 28.65 | 0.18 | 28.42
114.656 | -0.06 | -40.02 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -12.83 | -0.06 | -23.85 | -0.06 | -27.95 | 0.12 | 21.33 | 0.06 | 28.48 | -2.23 | 28.65 | -1.11 | 28.48 | -4.22 | 0.06 | -25.72 | -0.06 | 0 | -0.06 | 28.48 | 0 | 28.89 | 0 | 28.3 | 0.12 | 28.48
114.688 | -0.06 | -39.9 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.39 | 0.47 | 21.39 | -11.13 | 0 | -21.33 | -0.06 | -29.42 | 0.06 | 0.88 | 0.12 | 28.95 | -1.7 | 28.48 | -0.59 | 28.65 | -3.16 | -0.06 | -23.56 | -0.06 | -53.44 | -0.06 | 28.71 | 0.06 | 28.42 | 0.06 | 28.48 | 0.18 | 28.54
114.72 | -0.06 | -36.21 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.33 | -8.03 | -0.06 | -19.81 | -0.06 | -35.74 | 0.23 | 0 | 0.18 | 28.42 | -1.29 | 28.54 | -0.59 | 28.36 | -2.46 | 0 | -21.86 | -0.06 | -68.03 | -0.06 | 0.18 | 0.06 | 25.72 | 0.12 | 28.48 | 0.23 | 28.42
114.752 | -0.06 | -34.51 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.39 | -5.74 | 0 | -18.16 | -0.18 | -34.28 | 0.12 | 0 | 0.18 | 28.54 | -1.05 | 28.83 | -0.35 | 28.59 | -1.88 | -0.06 | -19.45 | -0.06 | -47.58 | 0 | 0 | 0 | 0.18 | 0 | 27.66 | 0.18 | 28.48
114.784 | -0.06 | -32.7 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.56 | 0.53 | 21.39 | -3.87 | 0.06 | -16.41 | -0.06 | -31.11 | 0.12 | 0 | 0.12 | 28.48 | -0.82 | 28.42 | -0.29 | 28.77 | -1.29 | 0 | -17.99 | -0.18 | -45.94 | -0.06 | 0 | 0 | 0.06 | 0.06 | 0.18 | 0 | 18.4
114.816 | -0.06 | -30.7 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.47 | 21.39 | -2.23 | 0.12 | -15.18 | -0.06 | -29.06 | 0.06 | 0 | 0.18 | 28.48 | -0.7 | 28.48 | -0.23 | 28.48 | -0.82 | 5.68 | -15.94 | -0.06 | -40.55 | 0 | 0 | 0.06 | 0 | -0.06 | 0 | 0 | 0.18
114.848 | -0.06 | -29.42 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.53 | 21.21 | -1.46 | 2.11 | -12.83 | -0.18 | -27.6 | 0.12 | 0 | 0.18 | 28.36 | -0.59 | 28.48 | -0.29 | 28.89 | -0.59 | 13.59 | -14.06 | -0.18 | -37.15 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
114.88 | 0 | -25.55 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.41 | 21.39 | -0.76 | 9.61 | -10.9 | -0.06 | -25.78 | 0 | 0 | 0.12 | 28.95 | -0.41 | 28.65 | -0.29 | 28.36 | -0.47 | 23.85 | -11.72 | -0.18 | -37.91 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0.06
114.912 | 0 | -24.26 | 0 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.33 | -0.23 | 14.65 | -8.91 | -0.06 | -23.79 | -0.06 | 0 | 0.18 | 28.54 | -0.23 | 28.24 | -0.29 | 28.36 | -0.29 | 28.95 | -9.43 | -0.06 | -38.2 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0
114.944 | -0.06 | -22.38 | -0.18 | 0 | -0.06 | 0 | -0.18 | 21.39 | 0.29 | 21.27 | 0.06 | 20.51 | -7.62 | 4.22 | -21.86 | 0.06 | 0 | 0.18 | 28.59 | -0.41 | 28.77 | -0.23 | 28.59 | -0.35 | 28.77 | -7.56 | -0.18 | -36.62 | 0 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
114.976 | 0 | -19.98 | -0.06 | -6.62 | -0.06 | 0 | -0.18 | 0.12 | 0.35 | 21.33 | 0.18 | 21.39 | -5.92 | 10.25 | -20.16 | 0.06 | 0 | 0.12 | 28.83 | -0.35 | 28.48 | -0.29 | 28.42 | -0.29 | 28.77 | -5.39 | 0 | -34.63 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
115.008 | -0.06 | -17.99 | -0.06 | -63.28 | -0.18 | -10.84 | -0.06 | 0 | 0.35 | 21.15 | 0.18 | 21.39 | -4.16 | 15.65 | -18.87 | 0 | 0 | 0.23 | 28.3 | -0.23 | 28.48 | -0.29 | 28.42 | -0.29 | 28.59 | -3.16 | -0.06 | -32.76 | -0.06 | 0 | 0 | 0 | -0.18 | 0 | 0 | 0
115.04 | -0.06 | -15.7 | -0.06 | -44.06 | -0.18 | -60.47 | -0.06 | 0 | 0.18 | 21.39 | 0.53 | 21.33 | -2.29 | 21.56 | -17.05 | 0 | 0 | 0.06 | 0.29 | -0.29 | 28.54 | -0.23 | 28.71 | -0.23 | 28.48 | -2.05 | -0.06 | -30.53 | 0 | 0 | 0 | 0 | 0 | 0 | -0.06 | 0
115.072 | -0.06 | -14.18 | 0 | -40.08 | -0.18 | -44.3 | -0.18 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -1.76 | 21.39 | -15.47 | 0.06 | 0 | 0.12 | 0.06 | -0.29 | 28.3 | -0.23 | 28.71 | -0.29 | 28.42 | -0.76 | -0.18 | -28.07 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
115.104 | -0.06 | -12.83 | -0.06 | -38.32 | -0.06 | -40.43 | -0.06 | 0 | 0.12 | 21.39 | 0.35 | 21.39 | -1.11 | 21.15 | -13.48 | 0 | -6.5 | 0.18 | 0 | -0.23 | 28.42 | -0.23 | 28.54 | -0.18 | 28.54 | -0.76 | -0.06 | -27.01 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 0
115.136 | 7.38 | -11.13 | -0.06 | -36.04 | -0.06 | -38.85 | -0.06 | 0 | 0.06 | 21.39 | 0.41 | 21.33 | -0.7 | 21.39 | -11.37 | 0 | -73.25 | 0.18 | 0.06 | -0.29 | 28.71 | -0.18 | 28.42 | -0.18 | 28.83 | 0.06 | -0.18 | -25.43 | -0.06 | -58.95 | -0.18 | 0.06 | 0 | 0 | 0 | 0
115.168 | 12.95 | -9.9 | 0 | -34.69 | -0.06 | -35.92 | 0 | 0 | 0 | 21.15 | 0.53 | 21.39 | -0.47 | 21.39 | -10.08 | 0 | -54.61 | 0.12 | 0 | -0.18 | 28.48 | -0.23 | 28.48 | -0.23 | 28.54 | -0.18 | -0.18 | -23.44 | 0 | -74.36 | -0.18 | -48.17 | -0.06 | 0 | 0 | 0
115.2 | 18.93 | -8.09 | -0.06 | -31.47 | -0.18 | -33.93 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 21.33 | -0.06 | 21.33 | -7.68 | 0.06 | -29.77 | 0.12 | 0 | -0.23 | 0.29 | -0.06 | 28.48 | -0.18 | 28.48 | 0.12 | 2.64 | -20.68 | -0.06 | -48.28 | -0.06 | -70.2 | 0 | -66.27 | -0.18 | -0.18
115.232 | 21.62 | -6.68 | -0.06 | -29.94 | -0.06 | -31.52 | -0.18 | 0 | 0 | 21.39 | 0.41 | 21.39 | -0.06 | 21.39 | -5.63 | 0 | -29.12 | 0.06 | 0 | -0.23 | 0.12 | -0.18 | 28.54 | -0.18 | 28.83 | 0.18 | 10.72 | -18.87 | 0 | -46.23 | -0.18 | -42.48 | 0 | -71.66 | -0.06 | -82.56
115.264 | 21.62 | -5.22 | -0.06 | -28.07 | -0.06 | -30 | -0.06 | 0 | 0 | 21.45 | 0.29 | 21.33 | 0.06 | 21.33 | -4.04 | 0 | -34.57 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0.18 | -0.18 | 28.59 | 0.23 | 18.22 | -17.23 | 0 | -40.14 | -0.18 | -28.18 | 0 | -45.12 | 0 | -70.32
115.296 | 21.45 | -3.93 | -0.06 | -26.43 | -0.06 | -27.83 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.21 | 0.12 | 21.5 | -2.81 | 0 | -34.98 | 0.18 | 0 | -0.23 | 0.06 | -0.18 | 0 | -0.18 | 28.54 | 0.23 | 26.08 | -15.18 | 0 | -36.62 | -0.18 | -36.45 | -0.06 | -29.24 | 0.06 | -39.85
115.328 | 21.39 | -2.46 | 0 | -23.67 | -0.06 | -25.55 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.15 | 0.18 | 21.39 | -2.17 | -0.06 | -32.76 | 0.12 | 0 | -0.18 | 0 | -0.06 | 0 | -0.29 | 28.01 | 0.29 | 28.83 | -13.54 | 0 | -37.38 | -0.06 | -42.78 | -0.06 | -34.92 | -0.06 | -28.36
115.36 | 21.45 | -1.41 | -0.06 | -21.27 | -0.06 | -23.15 | 0 | 0.06 | -0.06 | 0.23 | 0.23 | 21.39 | 0.23 | 21.39 | -1.11 | 0.06 | -28.36 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.23 | 28.65 | -10.78 | 0.06 | -38.09 | -0.18 | -38.56 | -0.06 | -42.95 | 0 | -39.79
115.392 | 21.45 | -0.7 | -0.06 | -18.99 | -0.06 | -21.15 | -0.06 | -39.08 | -0.18 | 0.06 | 0.12 | 21.27 | 0.23 | 21.33 | -0.41 | 0.06 | -27.07 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.59 | -8.44 | 0 | -36.92 | -0.06 | -34.98 | -0.06 | -39.85 | -0.18 | -42.01
115.424 | 21.45 | -0.47 | -0.06 | -17.46 | -0.06 | -19.4 | -0.06 | -67.85 | -0.06 | 0 | 0.18 | 21.45 | 0.18 | 21.27 | -0.47 | -0.06 | -26.6 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.29 | 28.65 | -6.68 | -0.06 | -34.22 | -0.06 | -33.11 | -0.06 | -35.51 | -0.18 | -38.91
115.456 | 21.39 | -0.29 | -0.06 | -16.06 | 0 | -17.64 | -0.06 | -48.99 | 0 | 0.06 | 0.23 | 21.39 | 0.23 | 21.39 | 0 | -0.06 | -24.84 | 0.12 | -18.46 | -0.06 | 0 | 0 | 0 | -0.18 | 28.48 | 0.29 | 28.48 | -4.8 | 0 | -32.05 | -0.18 | -32.93 | -0.06 | -32.52 | -0.06 | -34.75
115.488 | 21.45 | -0.18 | -0.06 | -14.18 | -0.06 | -15.7 | -0.18 | -41.08 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 21.33 | -0.06 | -0.06 | -22.85 | 0.06 | -78.23 | -0.06 | 0 | 0 | 0 | -0.06 | 28.36 | 0.29 | 28.42 | -2.93 | 0.29 | -30.7 | -0.06 | -31.99 | -0.06 | -32.34 | -0.18 | -32.99
115.52 | 21.5 | -0.18 | -0.06 | -12.25 | 0 | -13.54 | -0.06 | -33.52 | -0.06 | 0 | 0.12 | 21.56 | 0.18 | 21.39 | -0.06 | 0.06 | -21.09 | 0.12 | -49.28 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.29 | 28.83 | -1.7 | 12.25 | -28.71 | 0 | -29.06 | -0.06 | -32.05 | -0.06 | -32.52
115.552 | 21.45 | -0.35 | -0.06 | -10.37 | -0.06 | -12.07 | -0.06 | -31.99 | -0.06 | 0 | 0.06 | 21.15 | 0.12 | 21.33 | 0.18 | 0.06 | -19.45 | 0.12 | -45.06 | -0.06 | 0 | 0 | 0 | 0 | 28.71 | 0.23 | 28.18 | -0.94 | 19.98 | -26.78 | 0 | -27.31 | -0.06 | -29.77 | 0 | -31.64
115.584 | 21.45 | -0.29 | 6.68 | -9.2 | -0.06 | -10.84 | -0.06 | -32.99 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 21.15 | 0.35 | 0.06 | -17.87 | 0.12 | -44.12 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.48 | -0.53 | 27.25 | -25.02 | -0.06 | -26.54 | -0.06 | -27.72 | -0.06 | -29.71
115.616 | 21.45 | -0.47 | 12.36 | -7.91 | -0.06 | -9.32 | -0.06 | -32.7 | -0.06 | 0 | -0.06 | 21.33 | 0.29 | 21.39 | 0.23 | 4.16 | -15.76 | 0.12 | -41.13 | 0 | -1.29 | -0.06 | 0 | -0.18 | 28.48 | 0.23 | 28.42 | -0.06 | 28.54 | -23.32 | -0.06 | -25.55 | -0.06 | -26.31 | -0.18 | -27.6
115.648 | 21.45 | -0.35 | 17.87 | -6.56 | 0 | -8.14 | -0.06 | -30.94 | 0 | 0 | 0.06 | 21.33 | 0.23 | 21.39 | 0.18 | 10.25 | -14.18 | 0.12 | -39.85 | -0.06 | -69.61 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.59 | -0.06 | 28.48 | -21.56 | 0 | -23.91 | -0.06 | -25.37 | 0 | -26.72
115.68 | 21.45 | -0.23 | 21.5 | -4.63 | 0 | -6.15 | -0.06 | -28.3 | -0.06 | 0 | 0.06 | 21.33 | 0.23 | 21.21 | 0.18 | 17.64 | -12.31 | 0.18 | -38.44 | -0.06 | -54.49 | -0.06 | -55.72 | 0 | 28.36 | 0.06 | 28.13 | 0.23 | 28.36 | -19.04 | 0 | -21.91 | -0.18 | -23.67 | -0.18 | -25.14
115.712 | 21.45 | -0.23 | 21.5 | -3.34 | 2.29 | -4.28 | 0 | -26.31 | -0.06 | 0 | 0 | 21.33 | 0.12 | 21.21 | 0.12 | 21.68 | -10.55 | 0.06 | -35.51 | -0.06 | -44.77 | -0.06 | -67.68 | -0.06 | 0.18 | 0 | 28.83 | 0.18 | 28.42 | -16.99 | 0 | -20.16 | -0.18 | -21.68 | -0.18 | -23.03
115.744 | 21.45 | -0.23 | 21.45 | -1.93 | 7.79 | -3.16 | -0.06 | -23.97 | -0.06 | 0 | -0.06 | 0.64 | 0.12 | 21.33 | 0.23 | 21.39 | -8.79 | 0.18 | -34.45 | -0.06 | -45.47 | -0.06 | -46.99 | 0 | 0.06 | 0 | 28.48 | 0.23 | 28.48 | -15.35 | 0 | -18.52 | -0.18 | -19.98 | -0.06 | -21.33
115.776 | 21.45 | -0.23 | 21.45 | -1.46 | 13.71 | -2.05 | 2.29 | -22.21 | -0.18 | -3.16 | -0.06 | 0.06 | 0.06 | 21.39 | 0.06 | 21.45 | -7.38 | 0 | -32.4 | -0.06 | -42.78 | -0.18 | -45.18 | -0.06 | 0 | 0.12 | 28.65 | 0.29 | 28.42 | -13.13 | -0.06 | -17.58 | -0.06 | -18.46 | -0.06 | -19.98
115.808 | 21.45 | -0.23 | 21.45 | -0.7 | 19.04 | -1.17 | 7.91 | -19.98 | -0.06 | -72.19 | -0.18 | 0 | 0.18 | 21.39 | 0.18 | 21.33 | -5.51 | 0 | -30.94 | 0 | -40.72 | -0.18 | -44.36 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.65 | -11.13 | 0 | -15.7 | -0.18 | -16.99 | 0 | -18.75
115.84 | 21.45 | -0.23 | 21.5 | -0.18 | 21.45 | -0.47 | 15.18 | -17.7 | -0.06 | -50.51 | 0 | 0 | 0 | 0.12 | 0.12 | 21.39 | -3.28 | 0.12 | -28.3 | 0 | -38.79 | 0 | -41.25 | -0.18 | 0 | 0 | 28.48 | 0.35 | 28.36 | -8.56 | 11.43 | -13.3 | -0.06 | -14.94 | -0.06 | -16.47
115.872 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | -0.23 | 20.8 | -15.7 | -0.06 | -42.6 | -0.06 | 0 | 0.12 | 0 | 0.12 | 21.39 | -2.46 | 0 | -25.78 | 0 | -36.39 | 0 | -39.61 | -0.06 | 0 | 0 | 28.13 | 0.35 | 28.48 | -6.09 | 19.16 | -11.48 | -0.06 | -13.48 | 9.26 | -14.47
115.904 | 21.45 | -0.23 | 21.5 | -0.23 | 21.45 | -0.18 | 21.56 | -13.71 | -0.06 | -36.97 | -0.06 | 0 | 0 | 0 | 0.18 | 21.21 | -1.7 | 0.06 | -24.43 | 0 | -34.63 | 0 | -37.62 | -0.06 | 0 | 0 | 28.42 | 0.41 | 28.48 | -4.16 | 26.9 | -9.61 | 9.02 | -11.72 | 16.7 | -13.24
115.936 | 21.5 | -0.29 | 21.39 | -0.35 | 21.39 | -0.18 | 21.39 | -11.95 | -0.18 | -32.4 | -0.06 | 0 | 0.06 | 0 | 0.06 | 21.21 | -0.94 | 0 | -21.86 | 0 | -33.63 | 0 | -36.86 | -0.06 | 0 | 0 | 28.3 | 0.23 | 28.71 | -2.64 | 28.59 | -8.14 | 16.11 | -9.9 | 24.32 | -11.66
115.968 | 21.45 | -0.29 | 21.45 | -0.59 | 21.5 | -0.35 | 21.39 | -9.55 | -0.06 | -32.23 | 0 | 0 | 0 | 0 | 0.12 | 21.39 | -0.76 | 0.12 | -20.92 | 0 | -31.99 | -0.06 | -33.87 | -0.06 | 0 | 0 | 28.48 | 0.35 | 28.48 | -1.76 | 28.42 | -6.86 | 23.97 | -8.91 | 28.83 | -9.79
116 | 20.04 | -0.18 | 21.45 | -0.35 | 21.45 | -0.35 | 21.45 | -6.86 | 0 | -32.29 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -0.35 | 0.06 | -17.7 | -0.06 | -28.65 | 0 | -31.76 | 0 | 0 | 0 | 28.59 | 0.18 | 28.65 | -0.7 | 28.54 | -4.86 | 28.59 | -6.86 | 28.42 | -8.14
116.032 | 0.12 | -0.18 | 21.39 | -0.35 | 21.45 | -0.41 | 21.39 | -4.75 | 0 | -31 | -0.18 | 0 | 0.06 | 0 | 0.18 | 21.45 | -0.06 | 0.06 | -16.47 | -0.06 | -26.9 | 0 | -29.65 | 0 | 0 | 0 | 28.65 | 0.06 | 28.42 | -0.47 | 28.42 | -3.28 | 28.48 | -5.1 | 28.42 | -6.33
116.064 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.35 | 21.39 | -2.75 | -0.06 | -29.47 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | -0.06 | -14.94 | 0 | -24.67 | 0 | -27.77 | -0.06 | 0 | 0 | 28.42 | 0.12 | 28.48 | -0.18 | 28.42 | -2.17 | 28.3 | -3.28 | 28.59 | -4.63
116.096 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.23 | 21.45 | -1.88 | 0 | -26.9 | -0.18 | 0 | -0.06 | 0 | 0.29 | 21.27 | 0.06 | 0 | -13.07 | 0 | -23.32 | -0.06 | -26.37 | 0 | 0 | -0.06 | 17.99 | 0 | 28.59 | -0.18 | 28.48 | -1.41 | 28.48 | -2.23 | 28.71 | -3.52
116.128 | 0 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | -0.94 | -0.06 | -24.84 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 2.11 | -11.89 | 0 | -21.09 | 0 | -24.61 | 0 | -36.39 | 0 | 0.12 | 0.06 | 28.3 | 0.18 | 28.48 | -0.94 | 28.65 | -1.7 | 28.3 | -2.4
116.16 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.27 | -0.53 | -0.06 | -22.44 | -0.06 | -12.83 | 0.12 | 0 | 0.06 | 21.21 | 0.12 | 9.55 | -9.84 | -0.18 | -18.05 | 0 | -21.39 | 0 | -75 | -0.06 | 0 | 0.06 | 28.54 | 0.29 | 28.59 | -0.41 | 28.48 | -0.94 | 28.48 | -1.29
116.192 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.06 | -0.06 | -20.51 | -0.06 | -76.35 | 0 | 0 | 0 | 21.33 | 0.06 | 14.77 | -8.79 | 0 | -16.76 | 0 | -19.98 | 0 | -49.69 | -0.06 | 0.06 | 0.12 | 28.18 | 0.18 | 28.48 | -0.41 | 28.48 | -0.64 | 28.48 | -0.7
116.224 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.18 | 21.39 | 0.06 | 0.06 | -18.34 | 0 | -53.56 | -0.18 | 0 | 0.12 | 21.45 | 0.29 | 20.45 | -7.56 | 0 | -15.06 | 0 | -17.87 | 0 | -47.35 | 0.06 | 0 | 0 | 28.36 | 0.29 | 28.54 | -0.06 | 28.42 | -0.23 | 28.65 | -0.53
116.256 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.29 | 21.39 | 0.53 | 0 | -16.76 | -0.18 | -43.07 | 0 | -21.91 | 0.12 | 21.21 | 0.35 | 21.45 | -6.62 | 0 | -13.71 | 0 | -16.23 | 0.06 | -41.9 | 0 | 0 | 0.06 | 28.59 | 0.29 | 28.59 | -0.06 | 28.77 | -0.18 | 28.59 | -0.29
116.288 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | 0.59 | -0.06 | -15.23 | -0.06 | -37.09 | -0.06 | -74.01 | 0.12 | 21.45 | 0.35 | 21.39 | -5.39 | 0 | -12.36 | 0 | -14.71 | 0 | -37.21 | 0 | 0 | -0.06 | 28.48 | 0.29 | 28.59 | 0 | 28.42 | -0.18 | 28.71 | -0.06
116.32 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | 0.53 | 0.06 | -12.31 | -0.06 | -31.99 | 0 | -42.95 | 0.12 | 21.27 | 0.29 | 21.33 | -4.28 | -0.06 | -10.25 | 0 | -12.72 | -0.06 | -37.38 | -0.06 | 0 | -0.06 | 28.59 | 0.23 | 28.54 | 0 | 28.48 | -0.06 | 28.48 | -0.06
116.352 | 0 | -0.06 | 21.45 | -0.23 | 21.39 | -0.18 | 21.39 | 0.47 | -0.06 | -10.37 | -0.06 | -32.46 | -0.06 | -28.77 | 0 | 21.27 | 0.29 | 21.27 | -3.34 | 0.06 | -9.43 | 9.02 | -11.31 | 0 | -38.32 | -0.06 | 0 | 0 | 28.36 | 0.23 | 28.59 | 0.06 | 28.24 | 0.06 | 28.42 | 0.12
116.384 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | 0.35 | 0.12 | -8.03 | -0.06 | -32.52 | -0.06 | -30.59 | 0.06 | 21.39 | 0.35 | 21.45 | -2.75 | 2.23 | -7.79 | 16.11 | -10.02 | 0.06 | -36.92 | 0 | 0 | 0 | 0.64 | 0.06 | 28.59 | 0.12 | 28.71 | 0.12 | 28.95 | 0.06
116.416 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.23 | 21.33 | 0.53 | 2.46 | -5.98 | 0.06 | -31.52 | 0 | -34.98 | 0.06 | 21.15 | 0.35 | 21.56 | -2.29 | 9.9 | -6.91 | 23.91 | -8.61 | -0.06 | -35.74 | -0.18 | 0 | 0 | 0.12 | 0.12 | 28.65 | 0.18 | 28.48 | 0.12 | 28.48 | 0.18
116.448 | -47.17 | -0.06 | 21.5 | -0.18 | 21.45 | -0.23 | 21.39 | 0.53 | 8.44 | -3.69 | -0.06 | -30.06 | 0 | -34.63 | 0 | 21.39 | 0.35 | 21.33 | -1.88 | 17.11 | -5.57 | 28.83 | -7.38 | -0.06 | -33.52 | -0.06 | 0 | 0 | 0 | 0.12 | 28.65 | 0.06 | 28.48 | 0 | 28.42 | 0.06
116.48 | -54.79 | -0.06 | 0.06 | -0.06 | 21.45 | -0.18 | 21.39 | 0.47 | 15.65 | -1.93 | 0 | -27.13 | -0.06 | -31.17 | 0.12 | 21.33 | 0.41 | 21.15 | -1.29 | 27.36 | -4.39 | 28.71 | -5.86 | 3.81 | -30.59 | -0.06 | 0 | 0 | 0.06 | 0.18 | 28.59 | 0.12 | 28.59 | 0.12 | 28.07 | 0.18
116.512 | -40.96 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.53 | 21.39 | -1.35 | -0.06 | -24.96 | -0.18 | -28.48 | 0 | 0.12 | 0.29 | 21.39 | -0.94 | 28.48 | -3.63 | 28.48 | -4.75 | 12.25 | -29.01 | 0 | -1.17 | 0.06 | 0 | 0 | 28.18 | 0.12 | 28.42 | 0.23 | 28.83 | 0.35
116.544 | -39.9 | -0.18 | 0.06 | -0.06 | 21.5 | -0.18 | 21.39 | 0.47 | 21.45 | -0.53 | -0.06 | -22.97 | 0.06 | -27.01 | 0.06 | 0 | 0.35 | 21.45 | -0.59 | 28.71 | -2.64 | 28.59 | -4.04 | 19.92 | -27.6 | -0.06 | -86.14 | 0 | 0 | 0.18 | 28.48 | 0.06 | 28.48 | 0.18 | 28.36 | 0.18
116.576 | -37.09 | -0.06 | 0 | -0.06 | 21.5 | -0.18 | 21.39 | 0.47 | 21.39 | -0.29 | -0.06 | -21.09 | 0 | -26.13 | -0.06 | 0 | 0.23 | 21.27 | -0.64 | 28.83 | -2.17 | 28.83 | -3.22 | 27.6 | -25.84 | 0 | -65.74 | 0.06 | 0 | 0 | 28.24 | 0.12 | 28.65 | 0.18 | 28.48 | 0.18
116.608 | -34.63 | -0.06 | 0 | -0.06 | 17.52 | -0.18 | 21.39 | 0.35 | 21.39 | 0.18 | 0 | -18.52 | -0.06 | -24.55 | -0.06 | 0.06 | 0.23 | 21.39 | -0.64 | 28.13 | -1.7 | 28.24 | -2.52 | 28.54 | -23.67 | 0 | -47.93 | 0 | 0 | 0 | 28.48 | 0.06 | 28.36 | 0.18 | 28.48 | 0.23
116.64 | -33.17 | -0.06 | 0 | -0.06 | 0.06 | -0.18 | 21.39 | 0.29 | 21.39 | 0.12 | -0.06 | -16.41 | 0 | -21.91 | 0 | 0 | 0.23 | 21.33 | -0.53 | 28.48 | -1.41 | 29.06 | -1.7 | 28.48 | -21.09 | -0.06 | -45.59 | -0.18 | 0 | 0 | 28.42 | 0.12 | 28.42 | 0.12 | 28.42 | 0.18
116.672 | -31.47 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.29 | 0.18 | 21.33 | 0.12 | -0.06 | -14.82 | 0 | -20.39 | 0.06 | 0 | 0.23 | 21.33 | -0.35 | 28.71 | -1 | 28.83 | -1.11 | 28.65 | -19.63 | 0 | -38.15 | -0.06 | 0 | 0 | 28.54 | 0.06 | 28.77 | 0.12 | 28.48 | 0.29
116.704 | -29.24 | 0 | 0 | -0.06 | 0 | -0.18 | 0.06 | 0.12 | 21.39 | 0.35 | -0.06 | -12.48 | 0.06 | -19.04 | 0.06 | 0 | 0.23 | 21.39 | -0.29 | 28.48 | -0.64 | 28.48 | -0.82 | 28.83 | -17.87 | 0 | -36.21 | 0 | 0 | -0.06 | 5.33 | 0.12 | 28.48 | 0.06 | 28.48 | 0.29
116.736 | -26.49 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.53 | -0.06 | -10.49 | -0.06 | -17.4 | 0 | 0 | 0.23 | 21.33 | -0.53 | 28.54 | -0.47 | 28.71 | -0.76 | 28.24 | -15.53 | 0 | -37.79 | -0.06 | 0 | 0 | 0.12 | -0.18 | 28.42 | 0.18 | 28.42 | 0.18
116.768 | -24.2 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.53 | -0.06 | -8.56 | 0 | -15.65 | -0.06 | 0 | 0.23 | 21.33 | -0.35 | 28.48 | -0.53 | 28.18 | -0.53 | 28.54 | -14.06 | 0 | -37.56 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 0.23 | 0.12
116.8 | -22.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.5 | 0.41 | 2.46 | -5.57 | 0.06 | -13.48 | -0.06 | 0 | 0.23 | 21.33 | -0.29 | 28.48 | -0.35 | 28.54 | -0.35 | 28.65 | -11.48 | 0 | -36.39 | -0.06 | -11.37 | -0.06 | 0.06 | 0 | 0.12 | 0 | 0 | 0.18
116.832 | -20.51 | 0 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 8.61 | -3.57 | -0.06 | -11.66 | -0.06 | 0 | 0.23 | 21.33 | -0.23 | 28.54 | -0.29 | 28.48 | -0.29 | 28.42 | -9.14 | -0.18 | -34.22 | 0 | -93.64 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.18
116.864 | -18.57 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.53 | 13.71 | -1.88 | -0.06 | -9.84 | 0 | 0 | 0.29 | 21.27 | -0.29 | 28.59 | -0.29 | 28.89 | -0.35 | 28.59 | -7.21 | -0.18 | -32.46 | -0.06 | -60.94 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.12
116.896 | -15.94 | -0.06 | -35.1 | -0.18 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.41 | 19.63 | -1.46 | -0.06 | -8.32 | 0 | 0 | 0.18 | 21.33 | -0.35 | 28.3 | -0.35 | 28.42 | -0.29 | 28.42 | -5.22 | 4.51 | -30.47 | -0.06 | -48.87 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.18
116.928 | -14.41 | -0.06 | -56.78 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.29 | 21.39 | -0.53 | -0.06 | -6.33 | -0.06 | -17.75 | 0.12 | 21.33 | -0.29 | 28.42 | -0.18 | 28.65 | -0.23 | 28.59 | -3.16 | 12.83 | -28.59 | -0.06 | -45.76 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06
116.96 | -13.18 | -0.06 | -40.26 | 0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -0.06 | -0.06 | -3.98 | 0 | -66.21 | 0.23 | 21.33 | -0.18 | 28.59 | -0.29 | 28.83 | -0.18 | 28.18 | -1.93 | 23.03 | -26.37 | -0.06 | -37.79 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.18
116.992 | -11.66 | -0.06 | -39.96 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.35 | 21.39 | 0 | 5.39 | -2.87 | 0 | -43.19 | 0.18 | 21.33 | -0.18 | 28.54 | -0.41 | 28.95 | -0.18 | 28.48 | -1 | 29.01 | -24.61 | 0 | -37.21 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.12
117.024 | -10.49 | -0.06 | -37.97 | -0.06 | -0.18 | -0.18 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | 0.29 | 11.31 | -1.88 | -0.06 | -28.71 | 0.23 | 0.76 | -0.29 | 28.59 | -0.23 | 28.65 | -0.23 | 28.65 | -0.59 | 28.65 | -23.03 | 0 | -38.67 | 0 | 0 | 0 | 0 | 0 | 0 | 0.06
117.056 | -8.32 | 0 | -35.16 | -0.06 | -52.5 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | 0.41 | 16.82 | -1.46 | 0.12 | -31 | 0.12 | 0.06 | -0.29 | 28.42 | -0.29 | 28.71 | -0.18 | 28.54 | -0.18 | 28.48 | -21.09 | 0 | -37.91 | -0.18 | 0 | -0.18 | 0 | 0 | 0 | 0.12
117.088 | -6.91 | -0.06 | -33.4 | -0.18 | -54.49 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.29 | 21.62 | -0.64 | 0.06 | -35.39 | 0.18 | 0 | -0.23 | 28.59 | -0.18 | 28.65 | -0.18 | 28.59 | 0.06 | 28.54 | -18.69 | -0.06 | -36.33 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0.06
117.12 | -5.45 | -0.18 | -30.06 | -0.06 | -40.08 | -0.06 | -77.35 | -0.06 | 21.39 | 0.06 | 21.39 | 0.47 | 21.39 | -0.41 | 0.06 | -34.1 | 0.06 | 0 | -0.23 | 28.59 | -0.18 | 28.65 | -0.18 | 28.24 | 0.18 | 28.54 | -16.23 | -0.06 | -33.22 | -0.06 | -3.81 | 0 | 0 | 0 | 0 | 0
117.152 | -3.93 | 0 | -29.06 | -0.06 | -38.56 | -0.06 | -54.96 | 0 | 21.45 | 0.12 | 21.39 | 0.35 | 21.39 | -0.18 | -0.06 | -30.29 | 0.06 | 0 | -0.18 | 28.65 | -0.06 | 28.42 | -0.18 | 28.42 | 0.18 | 28.59 | -14.65 | -0.06 | -31.47 | -0.18 | -89.59 | -0.06 | 0 | 0.12 | -0.06 | 0.06
117.184 | -2.7 | -0.06 | -27.13 | -0.06 | -37.44 | -0.06 | -42.6 | 0.06 | 21.39 | 0.12 | 21.39 | 0.35 | 21.33 | -0.06 | -0.18 | -28.18 | -0.06 | 0 | -0.18 | 28.59 | -0.18 | 28.48 | -0.18 | 28.42 | 0.35 | 28.59 | -12.83 | -0.23 | -30.47 | -0.06 | -65.28 | -0.18 | 0.06 | 0 | -68.56 | 0
117.216 | -1.7 | 0 | -25.55 | -0.06 | -35.33 | -0.06 | -38.32 | -0.06 | 21.21 | 0.06 | 21.39 | 0.41 | 21.33 | 0 | -0.18 | -27.77 | 0.12 | 0 | -0.06 | 28.42 | -0.18 | 28.65 | -0.18 | 28.48 | 0.18 | 28.65 | -10.78 | 0 | -28.36 | -0.06 | -36.62 | -0.18 | -62.05 | -0.06 | -75.06 | 0
117.248 | -0.94 | 1.99 | -22.33 | -0.06 | -33.46 | -0.06 | -32.76 | -0.06 | 21.15 | -0.06 | 21.39 | 0.53 | 21.39 | 0 | 3.57 | -26.54 | 0.06 | 0 | -0.18 | 28.54 | -0.18 | 0.12 | -0.18 | 28.48 | 0.29 | 28.36 | -8.85 | -0.18 | -26.72 | -0.06 | -28.42 | -0.06 | -75.59 | 0 | -44.65 | 0.12
117.28 | -0.47 | 8.67 | -20.16 | -0.06 | -30.94 | -0.06 | -32.34 | -0.06 | 21.27 | 0 | 21.33 | 0.47 | 21.33 | 0.18 | 11.43 | -23.44 | 0 | 0 | -0.06 | 0.94 | -0.18 | 0 | -0.06 | 28.59 | 0.35 | 28.65 | -5.86 | -0.18 | -24.43 | -0.06 | -41.84 | -0.06 | -38.91 | -0.06 | -29.71 | -0.18
117.312 | -0.23 | 14.18 | -18.16 | 0 | -28.36 | -0.06 | -32.34 | -0.06 | 0.23 | 0 | 21.39 | 0.29 | 21.33 | 0.06 | 17.11 | -22.15 | 0.06 | 0 | -0.18 | 0.12 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.48 | -3.93 | -0.06 | -22.62 | -0.06 | -41.9 | -0.06 | -29.59 | -0.06 | -38.44 | -0.06
117.344 | -0.23 | 20.1 | -16.17 | 0 | -26.9 | -0.06 | -31.17 | -0.06 | 0.06 | -0.06 | 21.33 | 0.29 | 21.39 | 0.12 | 21.39 | -20.27 | -0.06 | 0 | -0.18 | 0 | -0.12 | 0 | -0.06 | 28.36 | 0.29 | 28.3 | -2.29 | -0.06 | -20.33 | -0.06 | -37.62 | 0 | -38.2 | -0.06 | -42.83 | -0.06
117.376 | -0.23 | 21.56 | -15.18 | 0 | -24.26 | -0.06 | -30.06 | -0.06 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | 0.18 | 21.39 | -18.69 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.7 | 0.23 | 28.3 | -1.7 | 7.44 | -18.52 | 0 | -34.4 | -0.06 | -42.31 | 0 | -40.02 | 0
117.408 | -0.35 | 21.39 | -13.54 | -0.06 | -23.2 | -0.06 | -27.77 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.33 | 0.23 | 21.39 | -17.7 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.41 | 28.42 | -0.76 | 14.94 | -16.76 | 0 | -33.05 | -0.18 | -39.44 | 0 | -35.92 | 0
117.44 | -0.41 | 21.5 | -11.25 | 0 | -19.45 | -0.06 | -24.9 | -0.06 | 0 | -0.06 | 21.21 | 0 | 21.39 | 0.23 | 21.45 | -15.23 | -0.06 | -11.6 | -0.18 | 0 | -0.18 | 0 | -0.18 | 0 | 0.23 | 28.48 | -0.35 | 25.31 | -14.24 | -0.18 | -32.93 | 0 | -34.98 | -0.06 | -33.69 | 0
117.472 | -0.29 | 21.45 | -9.9 | -0.06 | -17.87 | -0.06 | -22.85 | -0.18 | 0 | -0.06 | 21.21 | 0.18 | 21.39 | 0.18 | 21.33 | -13.36 | 0 | -72.66 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.35 | 28.59 | -0.06 | 28.71 | -12.19 | 0 | -31.23 | -0.12 | -33.34 | -0.06 | -32.93 | -0.06
117.504 | -0.41 | 21.5 | -8.56 | 0 | -16 | -0.18 | -21.09 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.21 | 0.23 | 21.39 | -11.89 | -0.06 | -55.08 | -0.23 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18 | 28.48 | -0.06 | 28.54 | -10.25 | -0.06 | -29.94 | -0.18 | -32.76 | 0 | -31.99 | -0.06
117.536 | -0.18 | 21.45 | -7.09 | -0.06 | -14.65 | -0.06 | -18.87 | 0 | 0 | 0 | 21.33 | 0.06 | 21.33 | 0.18 | 21.45 | -9.84 | 0 | -32.76 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.23 | 28.59 | 0.23 | 28.48 | -8.14 | -0.06 | -27.6 | -0.18 | -32.23 | -0.18 | -30.18 | -0.06
117.568 | -0.23 | 21.45 | -5.74 | 0 | -13.36 | -0.06 | -16.93 | -0.18 | 0 | 0 | 21.39 | 0.06 | 21.39 | 0.18 | 21.33 | -8.14 | -0.06 | -27.77 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.06 | 28.42 | 0.29 | 28.59 | -6.09 | -0.06 | -26.37 | -0.18 | -30.35 | 0 | -28.07 | 0
117.6 | -0.23 | 21.45 | -3.87 | -0.06 | -10.84 | -0.06 | -14.24 | -0.06 | 0 | 0 | 21.39 | 0 | 21.39 | 0.18 | 21.39 | -5.92 | -0.06 | -33.99 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.77 | 0.47 | 28.71 | -3.52 | -0.06 | -24.61 | -0.06 | -27.42 | 0 | -26.25 | 0
117.632 | -0.23 | 21.45 | -2.64 | 0 | -9.61 | -0.06 | -12.6 | 0.06 | 0 | -0.06 | 21.39 | 0 | 21.45 | 0.12 | 21.39 | -4.75 | 0 | -34.63 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 28.48 | 0.35 | 28.65 | -2.05 | -0.06 | -22.97 | 0 | -26.13 | -0.06 | -25.08 | 0
117.664 | -0.29 | 21.45 | -1.64 | 3.16 | -8.44 | -0.06 | -10.49 | 0 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.18 | 21.39 | -3.16 | 0 | -32.81 | -0.06 | 0 | -0.06 | -45.65 | -0.06 | 0 | -0.18 | 28.48 | 0.35 | 28.71 | -1.23 | -0.18 | -21.5 | -0.18 | -25.31 | -0.06 | -23.67 | 0.06
117.696 | -0.18 | 21.45 | -0.94 | 9.38 | -6.68 | 3.22 | -8.44 | 0 | 0 | -0.06 | 13.83 | -0.06 | 21.21 | 0.18 | 21.33 | -2.17 | -0.06 | -29.65 | -0.06 | 0.06 | 0 | -67.74 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.54 | -0.76 | -0.06 | -19.63 | -0.18 | -23.09 | 0 | -22.09 | 0
117.728 | -0.29 | 21.45 | -0.59 | 14.77 | -5.57 | 9.55 | -6.21 | 0 | -0.23 | 0 | 0.12 | -0.06 | 21.33 | 0.12 | 21.39 | -1.46 | 0 | -27.42 | 0 | -63.69 | 0 | -44.36 | -0.06 | 0 | -0.06 | 28.65 | 0.23 | 28.42 | -0.29 | -0.06 | -18.4 | -0.06 | -21.62 | 0 | -20.39 | 0
117.76 | -0.29 | 21.45 | -0.29 | 21.91 | -3.4 | 16.52 | -3.16 | -0.06 | -77.46 | -0.06 | 0 | -0.06 | 21.27 | 0.18 | 21.39 | -0.59 | -0.06 | -26.54 | -0.06 | -55.2 | -0.06 | -45.7 | 0 | 0 | -0.06 | 28.54 | 0.18 | 28.65 | -0.06 | -0.18 | -16.23 | -0.18 | -20.1 | -0.18 | -18.28 | 0
117.792 | -0.18 | 21.5 | -0.23 | 21.5 | -2.11 | 21.91 | -1.99 | -0.06 | -52.85 | -0.06 | 0 | -0.06 | 21.21 | 0.06 | 21.33 | -0.53 | 2.46 | -24.84 | 0 | -44.83 | -0.06 | -43.77 | 0 | -6.45 | 0 | 0.18 | 0.29 | 28.59 | 0.12 | 1.99 | -14.47 | -0.06 | -18.34 | 0 | -16.82 | -0.06
117.824 | -0.18 | 21.45 | -0.23 | 21.45 | -1.35 | 21.5 | -1.29 | -0.06 | -42.54 | 0 | 0 | -0.06 | 21.39 | 0.18 | 21.33 | -0.06 | 9.38 | -22.97 | -0.06 | -45.7 | -0.06 | -41.31 | -0.06 | -79.34 | -0.06 | 0 | 0.23 | 28.54 | 0.23 | 7.62 | -12.72 | -0.06 | -16.88 | 0 | -14.82 | 0
117.856 | -0.18 | 21.45 | -0.35 | 21.5 | -0.59 | 21.45 | -0.64 | 0 | -36.92 | 0 | 0 | -0.06 | 21.39 | 0.12 | 21.15 | -0.23 | 14.47 | -21.09 | -0.06 | -43.01 | 0 | -39.96 | -0.06 | -56.78 | -0.06 | 0 | 0.18 | 28.65 | 0.29 | 15.06 | -11.07 | 5.63 | -15.18 | -0.06 | -13.48 | 0
117.888 | -0.18 | 21.39 | -0.41 | 21.45 | -0.23 | 21.45 | -0.18 | -0.06 | -32.81 | -0.18 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0.12 | 20.22 | -19.63 | 0 | -40.43 | 0 | -36.86 | 0 | -45.76 | 0 | 0.06 | 0.06 | 28.65 | 0.35 | 22.97 | -9.49 | 13.65 | -13.59 | 6.33 | -11.95 | -0.18
117.92 | -0.18 | 21.39 | -0.47 | 21.45 | -0.23 | 21.45 | 0 | 0 | -32.76 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -17.75 | -0.06 | -37.97 | 0 | -35.74 | -0.06 | -45.35 | 0 | 0 | 0.12 | 28.59 | 0.47 | 28.71 | -7.56 | 23.85 | -11.48 | 16.82 | -9.84 | 9.9
117.952 | -0.18 | 21.45 | -0.29 | 21.45 | -0.18 | 21.45 | 0.18 | -0.06 | -32.34 | -0.06 | 0 | 0.12 | 0 | 0.06 | 21.33 | 0.18 | 21.39 | -16.17 | -0.06 | -37.09 | 0 | -33.99 | 0 | -42.54 | 0 | 0 | 0.06 | 28.54 | 0.35 | 28.59 | -5.74 | 28.71 | -9.61 | 24.38 | -8.67 | 16.99
117.984 | -0.06 | 21.45 | -0.29 | 21.5 | -0.23 | 21.5 | 0.29 | 0.06 | -32.17 | 0 | 0 | 0 | 0.06 | 0.06 | 21.33 | 0.18 | 21.45 | -14.41 | -0.06 | -35.39 | -0.06 | -32.23 | -0.18 | -40.37 | 0.06 | 0 | 0.12 | 28.54 | 0.35 | 28.42 | -4.22 | 28.54 | -8.09 | 28.71 | -6.62 | 24.73
118.016 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | 0.29 | -0.06 | -29.83 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -12.36 | 0 | -34.04 | 0 | -30.29 | -0.06 | -39.2 | 0 | 0 | 0 | 28.77 | 0.47 | 28.65 | -2.75 | 28.42 | -6.62 | 28.48 | -4.92 | 28.77
118.048 | -0.18 | 21.39 | -0.18 | 21.45 | -0.41 | 21.45 | 0.41 | 0 | -27.54 | -0.06 | 0 | 0 | 0 | 0.06 | 21.21 | 0.18 | 21.39 | -11.25 | -0.06 | -32.05 | 0 | -28.13 | 0 | -37.15 | -0.06 | 0 | -0.06 | 28.83 | 0.29 | 28.59 | -2.29 | 28.48 | -4.8 | 28.54 | -3.46 | 28.65
118.08 | -0.18 | 21.5 | -0.29 | 21.45 | -0.29 | 21.45 | 0.35 | 4.92 | -24.84 | -0.06 | 0 | 0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | -8.73 | -0.06 | -29.01 | 0 | -25.55 | -0.06 | -34.98 | -0.18 | 0 | -0.06 | 28.83 | 0.18 | 28.42 | -1.41 | 28.59 | -2.99 | 28.42 | -2.11 | 28.59
118.112 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.45 | 0.47 | 11.07 | -22.74 | -0.06 | -0.76 | 0 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -7.15 | -0.06 | -27.19 | 0 | -23.61 | -0.06 | -33.63 | -0.06 | 0 | 0 | 28.65 | 0.23 | 28.48 | -0.47 | 28.59 | -2.23 | 28.48 | -1.52 | 28.42
118.144 | -0.18 | 10.66 | -0.29 | 21.45 | -0.35 | 21.45 | 0.53 | 16.35 | -20.92 | -0.06 | -67.09 | -0.06 | 0 | 0.12 | 0.12 | 0.23 | 21.33 | -5.33 | 3.4 | -24.67 | 0 | -21.86 | 0.06 | -31.17 | -0.06 | 0 | 0 | 28.65 | 0 | 28.59 | -0.47 | 28.48 | -1.46 | 28.48 | -1 | 28.36
118.176 | -0.06 | 0.12 | -0.23 | 21.45 | -0.23 | 21.45 | 0.35 | 21.45 | -18.69 | -0.06 | -59.18 | 0 | 0 | 0.06 | 0 | 0.18 | 21.27 | -3.93 | 11.72 | -23.2 | 0 | -19.63 | -0.06 | -29.06 | -0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.48 | -0.06 | 28.59 | -0.76 | 28.65 | -0.53 | 28.3
118.208 | -0.06 | 0 | -0.35 | 21.39 | -0.23 | 21.45 | 0.35 | 21.39 | -17.05 | 0 | -44.12 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | -2.64 | 19.34 | -21.56 | 0 | -17.4 | -0.06 | -26.49 | -0.18 | -20.22 | 0 | 28.42 | 0.06 | 28.3 | -0.06 | 28.71 | -0.59 | 28.59 | -0.41 | 28.42
118.24 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.35 | 21.39 | -14.71 | -0.06 | -37.21 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.33 | -1.7 | 29.18 | -18.34 | 0 | -15.29 | 0 | -24.96 | 0 | -81.57 | 0.12 | 23.91 | 0 | 28.48 | 0.06 | 28.65 | -0.41 | 28.42 | -0.18 | 28.48
118.272 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.5 | 0.35 | 21.39 | -12.25 | -0.06 | -32.46 | -0.06 | 0 | 0 | 0 | 0.18 | 21.39 | -0.94 | 28.83 | -16.82 | 0.06 | -14.18 | 0 | -22.97 | -0.06 | -52.5 | 0 | 0.18 | 0 | 28.65 | 0.06 | 28.59 | -0.18 | 28.83 | -0.06 | 28.65
118.304 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.39 | 0.35 | 21.39 | -10.14 | -0.06 | -32.93 | -0.06 | -36.92 | 0.12 | 0 | 0.29 | 21.21 | -0.7 | 28.42 | -14.82 | 0.18 | -12.83 | 0 | -20.33 | -0.06 | -48.93 | -0.06 | 0 | 0 | 28.83 | 0.18 | 28.42 | -0.06 | 28.54 | -0.06 | 28.42
118.336 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.29 | 21.33 | -7.79 | -0.06 | -33.05 | -0.06 | -69.03 | 0 | 0 | 0.18 | 21.39 | -0.29 | 28.83 | -13.71 | 9.96 | -11.48 | 0.06 | -18.75 | 0 | -42.54 | 0 | 0 | 0.06 | 28.77 | 0.12 | 28.54 | 0 | 28.42 | -0.06 | 28.48
118.368 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | 0.18 | 21.39 | -5.98 | -0.06 | -32.23 | -0.06 | -45.47 | 0 | 0 | 0.29 | 21.39 | -0.29 | 28.54 | -12.6 | 17.23 | -9.61 | 5.68 | -16.82 | 0 | -37.09 | 0 | 0.06 | 0 | 28.3 | 0.23 | 28.42 | 0.12 | 28.18 | 0.23 | 28.24
118.4 | -0.06 | 0 | -0.18 | 21.5 | -0.23 | 21.39 | 0.18 | 21.39 | -3.28 | 0 | -30.06 | -0.06 | -27.31 | 0 | 0 | 0.12 | 21.27 | -0.06 | 28.48 | -10.61 | 27.48 | -8.32 | 16.11 | -14.77 | 0 | -37.21 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.71 | 0.12 | 28.42 | 0.18 | 27.95
118.432 | -0.06 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | 0.12 | 21.33 | -1.93 | -0.06 | -27.83 | 0 | -31.82 | 0 | 0 | 0.18 | 21.33 | 0.06 | 28.48 | -9.14 | 28.89 | -7.32 | 24.02 | -13.36 | 0.06 | -38.2 | -0.18 | 0 | -0.06 | 28.3 | 0.06 | 28.59 | 0.12 | 28.48 | 0.18 | 28.42
118.464 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.5 | 0.12 | 21.15 | -1.29 | 2.46 | -25.37 | 0 | -35.86 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 29.18 | -7.79 | 28.83 | -5.63 | 28.65 | -11.84 | 0.06 | -37.62 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.65 | 0 | 28.48 | 0.23 | 28.48
118.496 | 0 | 0 | -0.18 | 21.45 | -0.23 | 21.39 | 0.06 | 21.33 | -0.7 | 8.61 | -22.85 | 0 | -34.69 | 0 | 0 | 0.23 | 21.39 | 0.12 | 28.36 | -6.5 | 28.36 | -4.8 | 28.65 | -10.72 | 0 | -36.1 | 0.06 | 0 | 0.06 | 28.42 | 0.06 | 28.48 | 0.12 | 28.71 | 0.12 | 28.65
118.528 | -0.06 | 0 | -0.18 | 21.45 | -0.06 | 21.39 | 0.12 | 21.39 | -0.18 | 13.89 | -21.33 | -0.06 | -31 | -0.06 | 0 | 0.12 | 21.39 | 0.18 | 28.59 | -5.86 | 28.42 | -3.52 | 28.42 | -9.08 | -0.06 | -34.16 | -0.18 | 0 | -0.18 | 28.13 | 0.12 | 28.48 | 0.18 | 28.3 | 0.18 | 28.59
118.56 | -0.06 | -1.17 | -0.18 | 0.12 | -0.06 | 21.39 | 0.12 | 21.33 | 0.06 | 21.39 | -18.69 | -0.06 | -27.6 | -0.06 | -36.1 | 0.12 | 21.39 | 0.18 | 28.48 | -4.39 | 28.77 | -2.75 | 28.65 | -7.38 | -0.18 | -31.29 | -0.06 | 0 | -0.06 | 28.65 | 0.06 | 28.42 | 0.18 | 28.42 | 0.18 | 28.42
118.592 | -0.18 | -54.14 | -0.18 | 0 | -0.18 | 0.12 | 0.06 | 21.39 | 0.18 | 21.45 | -16.82 | 0 | -27.42 | -0.06 | -66.74 | 0.12 | 21.39 | 0.12 | 28.48 | -3.63 | 28.42 | -2.4 | 28.65 | -6.45 | 0 | -29.24 | -0.18 | 0 | 0 | 28.18 | 0.06 | 28.48 | 0.12 | 28.54 | 0.06 | 28.42
118.624 | -0.18 | -51.27 | -0.06 | 0 | -0.18 | 0.06 | -0.18 | 21.39 | 0.35 | 21.39 | -15.29 | 0 | -25.9 | 0 | -45.65 | 0.12 | 21.39 | 0.23 | 28.48 | -2.64 | 28.13 | -1.88 | 28.36 | -5.57 | -0.06 | -27.6 | 0 | 0 | 0 | 28.3 | 0.06 | 28.65 | 0.06 | 28.36 | 0.06 | 28.42
118.656 | -0.06 | -40.67 | -0.06 | 0.06 | -0.18 | 0 | 0.06 | 21.15 | 0.41 | 21.39 | -13.01 | -0.06 | -23.91 | -0.06 | -28.24 | 0.12 | 21.39 | 0.18 | 28.48 | -2.11 | 28.59 | -1.41 | 28.3 | -4.22 | -0.06 | -26.25 | -0.18 | 0.06 | 0 | 28.48 | 0.06 | 28.18 | 0.12 | 28.42 | 0.12 | 28.42
118.688 | 0 | -39.85 | -0.06 | 0 | -0.18 | 0.06 | -0.06 | 21.39 | 0.53 | 21.39 | -11.02 | 0.06 | -21.27 | -0.06 | -29.53 | 0.18 | 3.81 | 0.18 | 28.54 | -1.7 | 28.65 | -0.7 | 28.77 | -3.57 | -0.06 | -24.08 | -0.06 | -49.46 | -0.06 | 28.42 | 0 | 28.48 | 0.12 | 28.59 | 0.12 | 28.48
118.72 | -0.06 | -36.62 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.53 | 21.39 | -8.2 | -0.06 | -19.45 | -0.06 | -35.74 | 0.12 | 0.06 | 0.12 | 28.83 | -1.23 | 28.71 | -0.53 | 28.3 | -2.46 | -0.06 | -21.86 | -0.06 | -70.26 | 0 | 0.18 | 0 | 28.48 | 0.06 | 28.48 | 0.12 | 28.59
118.752 | 0 | -34.86 | -0.06 | 0 | -0.18 | 0 | 0 | 21.45 | 0.53 | 21.27 | -5.74 | -0.06 | -18.69 | -0.18 | -34.16 | 0.18 | 0 | 0 | 28.3 | -1.05 | 28.3 | -0.53 | 28.48 | -1.88 | -0.06 | -19.63 | -0.18 | -48.63 | -0.06 | 0 | -0.06 | 0.23 | 0.18 | 28.42 | 0.06 | 28.48
118.784 | -0.06 | -32.76 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.41 | 21.33 | -3.69 | 0 | -16.7 | -0.06 | -30.94 | 0.18 | 0 | 0.12 | 28.89 | -0.94 | 29.01 | -0.41 | 28.65 | -1.35 | -0.06 | -17.93 | -0.18 | -47.35 | -0.06 | 0 | -0.18 | 0 | 0 | 0.18 | 0.12 | 26.6
118.816 | -0.18 | -30.7 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.39 | -2.17 | 0 | -15.18 | -0.06 | -28.3 | 0.12 | 0 | 0.12 | 28.59 | -0.76 | 28.54 | -0.41 | 28.3 | -0.94 | 4.92 | -15.82 | 0 | -40.78 | 0 | 0 | 0 | 0 | 0 | 0 | 0.06 | 0.18
118.848 | -0.06 | -28.83 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.35 | 21.39 | -1.46 | -0.06 | -13.3 | -0.06 | -27.07 | 0.12 | 0 | 0.18 | 28.71 | -0.53 | 28.59 | -0.29 | 28.48 | -0.64 | 13.07 | -14.24 | -0.06 | -36.51 | -0.06 | 0 | 0.06 | 0.06 | -0.06 | 0 | 0.06 | 0
118.88 | -0.06 | -25.9 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.5 | 0.35 | 21.39 | -0.59 | 9.26 | -11.07 | -0.06 | -26.13 | 0.12 | 0 | 0.12 | 28.59 | -0.35 | 28.59 | -0.41 | 28.36 | -0.47 | 23.2 | -11.6 | 0 | -37.85 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0
118.912 | -0.06 | -24.08 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.39 | -0.18 | 14.36 | -9.43 | -0.06 | -23.38 | 0.06 | 0 | 0.06 | 28.42 | -0.35 | 28.59 | -0.18 | 28.48 | -0.41 | 28.83 | -9.43 | 0 | -38.32 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
118.944 | -0.06 | -21.91 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0.06 | 20.22 | -7.56 | 3.34 | -21.68 | 0 | 0 | 0.12 | 28.48 | -0.35 | 28.65 | -0.29 | 28.59 | -0.41 | 28.48 | -7.38 | -0.06 | -36.45 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
118.976 | -0.18 | -20.63 | -0.06 | -3.93 | 0 | 0 | -0.06 | 0.12 | 0.23 | 21.33 | 0.06 | 21.45 | -6.5 | 9.73 | -20.63 | 0 | 0 | 0 | 28.48 | -0.41 | 28.42 | -0.18 | 28.3 | -0.35 | 28.65 | -5.1 | -0.18 | -34.69 | -0.06 | 0 | -0.18 | 0 | 0.12 | 0 | 0 | 0
119.008 | -0.06 | -17.87 | -0.06 | -62.29 | -0.06 | -5.68 | -0.06 | 0.06 | 0.06 | 21.33 | 0.35 | 21.39 | -4.28 | 14.88 | -19.45 | 0.12 | 0 | 0.18 | 28.65 | -0.35 | 28.48 | -0.29 | 28.65 | -0.35 | 28.54 | -3.63 | -0.18 | -33.17 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0 | 0
119.04 | -0.06 | -15.7 | -0.06 | -44.71 | -0.06 | -62.11 | 0.06 | 0 | 0.18 | 21.45 | 0.47 | 21.39 | -2.52 | 21.8 | -16.82 | 0 | 0 | 0.18 | 0.23 | -0.35 | 28.24 | -0.29 | 28.3 | -0.23 | 28.42 | -1.99 | -0.06 | -30.7 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0 | 0
119.072 | -0.06 | -14.53 | -0.06 | -40.78 | -0.06 | -45.24 | -0.18 | 0 | 0.06 | 21.39 | 0.41 | 21.33 | -2.05 | 21.45 | -15.35 | 0 | 0 | 0.12 | 0.12 | -0.29 | 28.59 | -0.23 | 28.48 | -0.29 | 28.65 | -1.23 | -0.18 | -28.71 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0
119.104 | -0.06 | -12.6 | -0.06 | -39.08 | -0.06 | -40.31 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | -1.29 | 21.39 | -13.89 | 0.06 | -3.87 | 0.06 | 0 | -0.29 | 28.59 | -0.18 | 28.48 | -0.18 | 28.36 | -0.47 | -0.12 | -27.19 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0
119.136 | 6.91 | -11.48 | 0 | -36.21 | -0.18 | -39.55 | -0.06 | 0 | 0 | 21.09 | 0.35 | 21.39 | -0.59 | 21.39 | -11.78 | 0.06 | -71.78 | 0.18 | 0.06 | -0.29 | 28.65 | -0.23 | 28.48 | -0.23 | 28.48 | -0.18 | -0.06 | -25.43 | -0.06 | -45.88 | -0.18 | 0 | -0.18 | 0 | 0 | 0
119.168 | 12.77 | -9.61 | -0.06 | -34.28 | -0.06 | -36.04 | -0.06 | 0 | -0.06 | 21.39 | 0.41 | 21.33 | -0.35 | 21.39 | -10.55 | 0.12 | -56.13 | 0.06 | 0 | -0.23 | 28.54 | -0.29 | 28.3 | -0.18 | 28.54 | -0.06 | -0.18 | -23.56 | 0 | -79.81 | 0 | -38.2 | 0 | 0 | 0 | 0
119.2 | 18.22 | -8.32 | -0.06 | -31.47 | 0 | -34.04 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.33 | 0 | 21.39 | -8.14 | 0 | -30 | 0.06 | 0 | -0.18 | 0.7 | -0.23 | 28.36 | -0.18 | 28.48 | 0 | 2.29 | -21.09 | -0.06 | -49.1 | -0.18 | -75.71 | -0.06 | -61.23 | 0 | 0
119.232 | 21.86 | -6.86 | -0.06 | -30.41 | -0.06 | -31.64 | -0.06 | 0 | 0.06 | 21.33 | 0.35 | 21.39 | -0.06 | 21.33 | -6.62 | 0.06 | -29.24 | 0.18 | 0 | -0.23 | 0.12 | -0.18 | 28.42 | -0.18 | 28.71 | 0.35 | 10.08 | -18.93 | -0.06 | -46 | -0.06 | -47.17 | 0 | -73.42 | 0 | -75.41
119.264 | 21.5 | -5.04 | -0.06 | -27.72 | -0.18 | -29.77 | 0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | 0 | 21.39 | -4.63 | 0.06 | -34.22 | 0.18 | 0 | -0.23 | 0 | -0.18 | 0.23 | -0.18 | 28.36 | 0.35 | 17.17 | -17.7 | 0 | -40.26 | -0.18 | -28.89 | 0.06 | -47.11 | 0 | -72.31
119.296 | 21.5 | -4.1 | -0.06 | -26.43 | 0.06 | -28.48 | 0 | 0 | -0.06 | 21.39 | 0.35 | 21.33 | 0.23 | 21.39 | -3.11 | 0.06 | -35.1 | 0.12 | 0 | -0.18 | 0 | -0.12 | 0 | -0.18 | 28.42 | 0.29 | 24.84 | -15.41 | 0 | -35.92 | -0.06 | -34.4 | -0.06 | -29.71 | -0.06 | -42.19
119.328 | 21.5 | -2.75 | 0.06 | -23.67 | -0.06 | -25.84 | -0.06 | 0 | 0 | 21.27 | 0.41 | 21.39 | 0.18 | 21.39 | -2.23 | 0 | -32.99 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.23 | 28.48 | 0.35 | 28.65 | -13.83 | -0.06 | -37.38 | -0.06 | -42.89 | -0.06 | -34.92 | 0 | -29.65
119.36 | 21.45 | -1.46 | -0.06 | -21.27 | 0 | -23.32 | 0 | 0 | -0.18 | 0.23 | 0.12 | 21.33 | 0.23 | 21.39 | -1.29 | -0.06 | -29.06 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.42 | 0.23 | 28.48 | -10.84 | 0.06 | -37.85 | 0 | -39.38 | -0.06 | -42.6 | 0 | -38.91
119.392 | 21.45 | -0.76 | -0.06 | -19.04 | -0.06 | -21.45 | 0 | -30.41 | 0 | 0.06 | 0.06 | 21.39 | 0.12 | 21.09 | -0.41 | -0.06 | -27.01 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.35 | 28.54 | -9.26 | -0.06 | -36.68 | -0.06 | -35.39 | -0.06 | -39.08 | -0.06 | -42.48
119.424 | 21.5 | -0.29 | -0.06 | -17.4 | -0.06 | -18.93 | -0.06 | -73.3 | 0 | 0 | 0.12 | 21.33 | 0.29 | 21.45 | -0.29 | -0.06 | -26.43 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.48 | -6.39 | -0.06 | -34.63 | -0.06 | -33.05 | -0.06 | -35.51 | -0.06 | -39.08
119.456 | 21.45 | -0.18 | 0 | -16.23 | -0.06 | -17.58 | -0.18 | -50.39 | -0.18 | 0.06 | 0.06 | 21.33 | 0.35 | 21.39 | -0.23 | -0.06 | -25.08 | 0.12 | -6.09 | -0.18 | 0 | -0.06 | 0 | 0 | 28.65 | 0.29 | 28.59 | -4.45 | 0 | -32.76 | -0.18 | -32.64 | -0.18 | -32.76 | 0.06 | -34.75
119.488 | 21.45 | -0.18 | -0.06 | -13.89 | -0.06 | -16 | -0.06 | -42.13 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.45 | -0.18 | -0.06 | -22.79 | 0.12 | -74.65 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.29 | 28.48 | -2.93 | 0 | -30.7 | -0.06 | -32.34 | 0 | -32.93 | -0.06 | -33.11
119.52 | 21.45 | -0.23 | -0.06 | -12.07 | -0.06 | -13.54 | -0.18 | -35.16 | -0.06 | 0 | 0 | 21.39 | 0.35 | 21.33 | 0.06 | -0.06 | -20.57 | 0.18 | -51.86 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.42 | 0.29 | 28.59 | -1.64 | 11.19 | -28.65 | -0.06 | -29.77 | 0 | -32.23 | 0 | -32.64
119.552 | 21.45 | -0.29 | -0.06 | -10.61 | -0.06 | -12.48 | -0.06 | -31.99 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | 0.06 | -0.06 | -19.45 | 0.12 | -44.71 | -0.06 | 0 | -0.06 | 0 | 0 | 28.36 | 0.23 | 28.48 | -0.94 | 18.87 | -26.9 | 0 | -27.95 | -0.06 | -29.77 | 0 | -31.58
119.584 | 21.45 | -0.35 | 5.8 | -9.38 | -0.06 | -10.78 | -0.06 | -32.05 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.62 | 0.23 | -0.18 | -18.16 | 0.12 | -44.77 | -0.18 | 0 | -0.18 | 0 | 0 | 28.42 | 0.23 | 28.36 | -0.64 | 26.25 | -25.25 | 0 | -26.9 | -0.18 | -28.18 | 0.06 | -29.3
119.616 | 21.45 | -0.35 | 11.72 | -8.03 | -0.06 | -9.32 | -0.06 | -32.64 | -0.18 | 0 | 0.06 | 20.98 | 0.23 | 21.45 | 0.18 | 3.34 | -16.29 | 0.18 | -42.36 | -0.06 | -0.12 | -0.06 | 0 | -0.18 | 28.95 | 0.06 | 28.59 | -0.18 | 28.48 | -23.56 | 0 | -25.55 | 0 | -26.72 | 0 | -27.6
119.648 | 21.45 | -0.35 | 17.17 | -6.39 | -0.18 | -7.91 | -0.18 | -31.06 | 0 | 0 | 0 | 21.39 | 0.18 | 21.33 | 0.29 | 9.49 | -14.36 | 0.18 | -41.08 | -0.06 | -60.65 | -0.06 | 0 | -0.18 | 28.48 | 0.12 | 28.59 | 0 | 28.59 | -21.97 | -0.06 | -24.38 | -0.18 | -25.67 | 0 | -26.66
119.68 | 21.45 | -0.29 | 21.5 | -4.34 | 0 | -5.92 | -0.06 | -29.06 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | 0.12 | 16.76 | -12.25 | 0.18 | -38.91 | -0.06 | -55.72 | -0.06 | -50.98 | -0.06 | 28.3 | 0.18 | 28.65 | 0.18 | 28.48 | -19.28 | 0 | -22.21 | -0.18 | -23.32 | 0 | -25.08
119.712 | 21.39 | -0.23 | 21.45 | -3.34 | 0 | -4.45 | -0.18 | -26.49 | 0 | 0 | 0 | 21.45 | 0.18 | 21.39 | 0.29 | 21.5 | -10.31 | 0.12 | -36.1 | -0.06 | -45.24 | -0.18 | -68.85 | -0.06 | 0.29 | 0.06 | 28.42 | 0.29 | 28.42 | -17.4 | -0.18 | -20.68 | -0.18 | -21.62 | -0.06 | -22.85
119.744 | 21.45 | -0.23 | 21.45 | -1.93 | 7.15 | -3.16 | 0 | -24.14 | -0.06 | 0 | 0.12 | 7.27 | 0.18 | 21.39 | 0.12 | 21.39 | -9.08 | 0.12 | -34.22 | 0 | -45.53 | 0 | -47.46 | -0.18 | 0.06 | 0.12 | 28.36 | 0.29 | 28.71 | -15.18 | -0.18 | -18.87 | -0.06 | -19.98 | -0.06 | -21.68
119.776 | 21.45 | -0.29 | 21.5 | -1.29 | 12.83 | -2.23 | 1.99 | -22.62 | 0 | -1.29 | -0.18 | 0.12 | 0.12 | 21.33 | 0.12 | 21.45 | -6.97 | 0.18 | -33.34 | 0 | -42.89 | 0 | -46.06 | -0.18 | 0 | 0.12 | 28.48 | 0.29 | 28.65 | -13.59 | -0.18 | -17.29 | -0.06 | -18.87 | -0.06 | -20.27
119.808 | 21.39 | -0.29 | 21.45 | -0.53 | 18.34 | -1.23 | 7.09 | -20.51 | -0.06 | -69.26 | 0 | 0 | 0.06 | 21.39 | 0.23 | 21.33 | -5.57 | 0.12 | -30.7 | -0.06 | -41.43 | 0 | -44.42 | 0 | 0.06 | 0.06 | 28.48 | 0.29 | 28.48 | -11.54 | 1.99 | -15.7 | -0.06 | -17.4 | -0.06 | -19.1
119.84 | 21.45 | -0.29 | 21.5 | -0.06 | 21.62 | -0.64 | 14.53 | -18.4 | 0 | -52.5 | -0.18 | 0 | 0.12 | 0.12 | 0.18 | 21.39 | -3.75 | 0.12 | -28.83 | 0 | -39.08 | 0 | -41.08 | 0.06 | 0 | 0 | 28.65 | 0.29 | 28.42 | -9.02 | 10.78 | -13.95 | -0.18 | -15.18 | -0.06 | -16.58
119.872 | 21.39 | -0.23 | 21.45 | -0.06 | 21.45 | -0.23 | 20.33 | -16.47 | -0.06 | -42.42 | 0 | 0 | 0.06 | 0 | 0.18 | 21.33 | -2.23 | 0.12 | -26.02 | 0 | -37.21 | 0 | -39.96 | -0.18 | 0 | 0 | 28.42 | 0.23 | 28.59 | -7.32 | 18.22 | -12.01 | 0 | -13.65 | 8.85 | -14.82
119.904 | 21.45 | -0.18 | 21.45 | -0.18 | 21.5 | -0.06 | 21.5 | -14 | 0.06 | -36.45 | -0.18 | 0 | 0.06 | 0 | 0.12 | 21.39 | -1.7 | 0 | -24.67 | 0 | -35.27 | 0 | -37.79 | 0 | 0 | 0 | 28.54 | 0.35 | 28.48 | -4.75 | 25.9 | -10.37 | 8.38 | -12.31 | 16.06 | -13.36
119.936 | 21.45 | -0.18 | 21.39 | -0.29 | 21.45 | -0.06 | 21.39 | -12.07 | -0.06 | -32.64 | -0.18 | 0 | 0.12 | 0.06 | 0.29 | 21.27 | -0.82 | 0.12 | -22.62 | -0.06 | -33.28 | 0 | -36.39 | 0 | 0 | 0 | 28.48 | 0.29 | 29.01 | -2.99 | 28.65 | -8.56 | 15.35 | -10.14 | 23.56 | -11.54
119.968 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.23 | 21.39 | -9.55 | -0.06 | -32.29 | 0 | 0 | 0.06 | 0 | 0.18 | 21.39 | -0.7 | 0.06 | -20.51 | -0.06 | -31.47 | 0 | -34.98 | 0 | 0 | 0 | 28.48 | 0.23 | 28.59 | -1.93 | 28.54 | -7.27 | 23.2 | -8.56 | 28.71 | -10.25
120 | 21.45 | -0.18 | 21.39 | -0.35 | 21.45 | -0.53 | 21.39 | -7.15 | -0.06 | -32.4 | -0.18 | 0 | -0.06 | 0 | 0.23 | 21.27 | -0.41 | 0.18 | -18.16 | -0.18 | -28.89 | 0 | -31.47 | 0 | 0 | -0.06 | 28.36 | 0.12 | 28.48 | -1.05 | 28.59 | -5.04 | 28.59 | -6.8 | 28.48 | -8.14
120.032 | 0.12 | -0.06 | 21.45 | -0.35 | 21.45 | -0.29 | 21.39 | -4.8 | -0.06 | -31.47 | 0 | 0 | 0 | 0 | 0.12 | 21.45 | -0.18 | -0.06 | -16.17 | -0.06 | -26.9 | -0.06 | -30.47 | -0.06 | 0 | -0.12 | 28.65 | 0.12 | 29.01 | -0.41 | 28.65 | -3.4 | 28.54 | -4.86 | 28.65 | -6.68
120.064 | 0 | -0.06 | 21.45 | -0.41 | 21.5 | -0.35 | 21.39 | -2.99 | 0 | -29.94 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.04 | 0 | 0.06 | -15.23 | -0.06 | -25.31 | -0.06 | -27.77 | -0.18 | 0 | -0.06 | 28.36 | 0.18 | 28.48 | -0.29 | 28.48 | -2.23 | 28.65 | -3.34 | 28.24 | -4.75
120.096 | 0 | -0.06 | 21.45 | -0.35 | 21.45 | -0.41 | 21.39 | -1.7 | -0.06 | -27.6 | -0.06 | 0 | 0 | 0 | 0.06 | 21.45 | 0.12 | -0.06 | -13.36 | -0.06 | -23.44 | -0.06 | -26.13 | -0.06 | 0 | -0.06 | 26.19 | 0.06 | 28.59 | 0.06 | 28.65 | -1.46 | 28.59 | -2.52 | 28.59 | -3.28
120.128 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | -0.35 | 21.39 | -1 | 0 | -25.14 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.15 | 0.06 | 2.05 | -11.84 | 0 | -21.15 | 0.06 | -24.38 | 0 | -26.08 | 0 | 0.18 | 0.06 | 28.48 | 0.06 | 28.65 | -0.7 | 28.24 | -1.76 | 28.3 | -2.4
120.16 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | -0.23 | 21.27 | -0.7 | -0.06 | -22.62 | -0.06 | -5.68 | 0 | 0 | 0.18 | 20.98 | 0.12 | 9.38 | -10.49 | -0.06 | -18.87 | 0 | -21.33 | 0 | -78.17 | -0.06 | 0 | 0.06 | 28.07 | 0.35 | 28.59 | -0.59 | 28.54 | -0.94 | 28.48 | -1.46
120.192 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.18 | 21.39 | -0.06 | -0.06 | -20.74 | -0.18 | -73.66 | -0.06 | 0 | 0.18 | 21.33 | 0.23 | 14.53 | -9.2 | 0 | -16.93 | -0.06 | -19.63 | -0.06 | -51.8 | -0.06 | 0 | 0.12 | 28.48 | 0.35 | 28.77 | -0.23 | 28.71 | -0.53 | 28.59 | -0.76
120.224 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.23 | 21.5 | 0 | 0 | -19.04 | -0.06 | -55.14 | -0.06 | 0 | 0.18 | 21.21 | 0.06 | 20.63 | -7.85 | -0.06 | -15.41 | -0.06 | -17.93 | -0.06 | -48.28 | 0 | 0 | 0 | 28.65 | 0.35 | 28.42 | -0.23 | 28.65 | -0.41 | 28.36 | -0.76
120.256 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.29 | 21.39 | 0.18 | -0.06 | -17.05 | -0.06 | -43.48 | -0.06 | -11.02 | 0.12 | 21.39 | 0.06 | 21.45 | -6.62 | 0.06 | -14.24 | 0 | -16.29 | -0.06 | -43.13 | -0.06 | 0 | 0 | 28.3 | 0.29 | 28.54 | 0 | 28.42 | -0.06 | 28.48 | 0
120.288 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.41 | 21.39 | 0.29 | 0 | -15 | -0.06 | -38.15 | -0.06 | -72.78 | 0.06 | 21.15 | 0.18 | 21.27 | -5.39 | 0 | -12.6 | 0 | -14.65 | -0.06 | -37.03 | -0.06 | 0 | -0.06 | 28.36 | 0.41 | 28.71 | 0 | 28.48 | -0.18 | 28.36 | 0
120.32 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.23 | 21.39 | 0.41 | -0.06 | -12.36 | -0.06 | -32.4 | -0.06 | -45.12 | 0 | 21.33 | 0.35 | 21.39 | -4.28 | 0 | -10.61 | 1.99 | -12.6 | -0.06 | -37.38 | 0 | 0 | 0 | 28.54 | 0.35 | 28.42 | 0.29 | 28.65 | 0.06 | 28.65 | 0
120.352 | 0 | 0.06 | 21.45 | -0.06 | 21.45 | -0.23 | 21.39 | 0.53 | -0.06 | -10.37 | -0.18 | -32.46 | -0.06 | -29.71 | 0.06 | 21.39 | 0.18 | 21.33 | -3.28 | 0 | -9.43 | 8.32 | -11.48 | -0.06 | -37.56 | 0 | 0 | -0.06 | 28.13 | 0.29 | 28.24 | 0.23 | 28.65 | 0.06 | 28.65 | 0
120.384 | 0 | 0 | 21.45 | -0.23 | 21.39 | -0.18 | 21.45 | 0.41 | -0.06 | -8.14 | -0.06 | -32.52 | -0.06 | -30 | 0 | 21.33 | 0.12 | 21.21 | -2.93 | 0 | -8.09 | 15.35 | -10.02 | -0.06 | -37.15 | 0 | 0 | -0.06 | 13.83 | 0.12 | 28.59 | 0.12 | 28.42 | 0.06 | 28.42 | 0.12
120.416 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | -0.23 | 21.39 | 0.41 | 0.41 | -6.15 | 0 | -31.93 | -0.06 | -35.1 | -0.06 | 21.21 | 0.18 | 21.39 | -2.11 | 9.2 | -6.97 | 23.2 | -8.56 | -0.06 | -35.74 | 0 | 0 | 0.12 | 0.12 | 0.12 | 28.42 | 0.12 | 28.48 | 0.12 | 28.13 | 0.06
120.448 | -41.31 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.29 | 7.91 | -3.93 | -0.06 | -30.06 | 0 | -35.1 | -0.06 | 21.21 | 0.12 | 21.39 | -1.82 | 16.06 | -5.86 | 28.71 | -7.44 | 0 | -33.69 | -0.06 | 0 | -0.18 | 0 | 0.12 | 28.54 | 0.18 | 28.77 | 0 | 28.71 | 0.23
120.48 | -57.01 | -0.06 | 0.12 | -0.06 | 21.39 | -0.18 | 21.27 | 0.41 | 15.06 | -1.88 | -0.06 | -27.07 | -0.06 | -31.23 | 0 | 21.33 | 0.23 | 21.21 | -1.29 | 26.49 | -4.39 | 28.83 | -5.74 | 2.99 | -31.17 | -0.06 | 0 | 0 | 0.06 | 0.18 | 28.54 | 0.29 | 28.42 | 0.23 | 28.65 | 0.23
120.512 | -41.19 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.47 | 20.92 | -1.46 | -0.06 | -24.96 | -0.06 | -28.07 | 0.06 | 0.12 | 0.23 | 21.33 | -1.05 | 28.77 | -3.52 | 28.65 | -4.63 | 11.25 | -29.01 | -0.06 | -0.06 | -0.06 | 0 | 0.06 | 28.48 | 0.06 | 28.18 | 0.18 | 28.48 | 0.12
120.544 | -40.08 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.45 | -0.47 | -0.06 | -23.03 | 0 | -27.13 | -0.06 | 0.06 | 0.18 | 21.39 | -0.82 | 28.65 | -2.75 | 28.42 | -3.75 | 18.81 | -28.07 | 0 | -75.76 | -0.18 | 0 | 0.06 | 28.54 | 0.12 | 28.48 | 0.18 | 28.48 | 0.06
120.576 | -37.09 | 0 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.41 | 21.39 | -0.41 | -0.06 | -21.74 | -0.06 | -26.37 | 0 | 0 | 0.23 | 21.39 | -0.76 | 28.83 | -2.17 | 28.71 | -3.11 | 26.25 | -26.25 | 0 | -69.67 | -0.06 | 0 | 0.18 | 28.36 | 0.12 | 28.24 | 0.29 | 28.48 | 0.06
120.608 | -35.27 | 0 | 0 | -0.18 | 21.39 | -0.06 | 21.39 | 0.18 | 21.45 | 0.06 | -0.06 | -18.93 | -0.06 | -25.14 | 0 | 0 | 0.18 | 21.39 | -0.59 | 28.65 | -1.82 | 28.65 | -2.4 | 28.83 | -24.08 | 0 | -49.75 | 0.06 | 0 | -0.06 | 28.48 | 0 | 28.42 | 0.18 | 28.48 | 0.06
120.64 | -32.81 | -0.06 | 0 | -0.06 | 0.06 | -0.18 | 21.39 | 0.29 | 21.39 | 0.06 | 0 | -16.52 | -0.06 | -21.97 | 0 | 0 | 0.23 | 21.39 | -0.53 | 28.83 | -1.29 | 28.65 | -1.52 | 28.83 | -21.45 | 0 | -44.88 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.48 | 0.23 | 28.48 | 0.12
120.672 | -31.47 | -0.06 | 0 | 0 | 0 | -0.18 | 2.99 | 0.12 | 21.39 | 0.41 | -0.06 | -14.47 | -0.06 | -20.45 | -0.06 | 0 | 0.18 | 21.39 | -0.41 | 28.89 | -0.94 | 28.65 | -1.23 | 28.42 | -19.81 | 0 | -38.38 | -0.06 | 0 | 0 | 28.59 | -0.06 | 28.54 | 0.29 | 28.48 | 0.06
120.704 | -29.77 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0.06 | 0.23 | 21.39 | 0.41 | -0.06 | -12.72 | -0.06 | -18.87 | 0 | 0 | 0.12 | 21.15 | -0.35 | 28.48 | -0.7 | 28.3 | -0.76 | 28.65 | -17.99 | -0.06 | -35.68 | -0.06 | 0 | 0.12 | 17.05 | -0.06 | 28.48 | 0.23 | 28.48 | 0.18
120.736 | -26.9 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.35 | -0.06 | -10.84 | -0.06 | -17.58 | 0 | 0 | 0.23 | 21.39 | -0.41 | 28.89 | -0.41 | 28.3 | -0.59 | 28.89 | -15.7 | 0 | -37.15 | -0.06 | 0 | 0 | 0.18 | -0.06 | 28.54 | 0.18 | 28.48 | 0.12
120.768 | -24.96 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.53 | 0.06 | -8.38 | -0.06 | -15.82 | 0 | 0 | 0.23 | 21.39 | -0.47 | 28.48 | -0.64 | 28.48 | -0.47 | 28.77 | -14 | 0.06 | -38.5 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.29 | 0.18 | 0.18
120.8 | -22.91 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.53 | 2.29 | -5.22 | 0 | -13.3 | 0 | 0 | 0.18 | 21.39 | -0.41 | 28.48 | -0.29 | 28.54 | -0.35 | 28.65 | -11.13 | -0.06 | -36.1 | -0.06 | -1.93 | -0.06 | 0.06 | 0 | 0.12 | 0.18 | 0 | 0
120.832 | -20.92 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.45 | 0.47 | 7.97 | -3.57 | 0 | -11.89 | 0.06 | 0 | 0.06 | 21.27 | -0.35 | 28.65 | -0.29 | 28.65 | -0.23 | 28.59 | -9.32 | 0 | -34.1 | -0.06 | -91.06 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.06
120.864 | -19.04 | -0.06 | -0.06 | -0.18 | 0 | -0.06 | 0 | 0.18 | 21.39 | 0.53 | 13.42 | -2.11 | 0 | -9.84 | 0 | 0 | 0.18 | 21.33 | -0.23 | 28.59 | -0.23 | 28.42 | -0.18 | 28.95 | -7.62 | -0.06 | -32.76 | -0.06 | -64.34 | -0.18 | 0 | 0 | 0 | 0.18 | 0 | 0.12
120.896 | -16.29 | -0.06 | -31 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 18.99 | -1.41 | 0.06 | -8.56 | 0.06 | 0 | 0.18 | 21.39 | -0.23 | 28.83 | -0.23 | 28.71 | -0.18 | 28.36 | -5.22 | 3.34 | -30.82 | 0 | -48.4 | 0 | 0 | -0.18 | 0 | 0.12 | 0 | 0.12
120.928 | -14.41 | -0.06 | -59.94 | 0 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.5 | -0.7 | -0.06 | -6.68 | 0.06 | -6.62 | 0.06 | 21.15 | -0.35 | 28.3 | -0.29 | 28.83 | -0.35 | 28.48 | -3.22 | 11.84 | -28.77 | -0.18 | -45.94 | -0.06 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06
120.96 | -13.36 | -0.06 | -40.26 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.27 | 0.35 | 21.39 | -0.18 | -0.06 | -4.28 | -0.06 | -68.56 | 0.18 | 21.33 | -0.35 | 28.48 | -0.23 | 28.42 | -0.23 | 28.89 | -1.64 | 22.03 | -26.43 | 0 | -37.79 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.12
120.992 | -11.84 | 0 | -39.61 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.21 | -0.06 | 4.57 | -2.99 | 0.06 | -45.94 | 0.18 | 21.33 | -0.23 | 28.54 | -0.23 | 28.59 | -0.29 | 28.48 | -1 | 29.42 | -24.38 | -0.06 | -36.45 | -0.06 | 0 | 0 | 0 | 0.12 | 0 | 0.12
121.024 | -10.2 | -0.06 | -38.09 | 0 | -0.06 | -0.06 | 0 | -0.18 | 21.45 | 0.35 | 21.39 | 0.35 | 10.55 | -2.17 | -0.06 | -29.36 | 0.18 | 2.99 | -0.18 | 28.54 | -0.23 | 28.42 | -0.23 | 28.42 | -0.47 | 28.65 | -23.09 | 0 | -37.56 | -0.18 | 0 | -0.18 | 0 | 0.06 | 0 | 0.06
121.056 | -8.61 | -0.06 | -35.33 | -0.06 | -43.19 | -0.18 | 0 | 0 | 21.39 | 0.29 | 21.39 | 0.29 | 15.94 | -1.23 | -0.06 | -29.71 | 0.12 | 0.06 | -0.29 | 28.54 | -0.18 | 28.42 | -0.23 | 28.59 | -0.18 | 28.65 | -21.09 | 0 | -38.15 | 0 | 0 | 0 | 0 | -0.06 | 0 | 0
121.088 | -7.27 | 0 | -33.81 | 0 | -56.6 | -0.18 | 0 | -0.18 | 21.39 | 0.23 | 21.39 | 0.23 | 21.39 | -0.82 | -0.06 | -34.28 | 0.12 | 0 | -0.23 | 28.65 | -0.18 | 28.48 | -0.23 | 28.59 | -0.06 | 28.59 | -19.22 | -0.06 | -36.62 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | 0.18
121.12 | -5.8 | 0 | -30.94 | -0.06 | -40.67 | -0.06 | -75.82 | -0.18 | 21.27 | 0.12 | 21.39 | 0.47 | 21.39 | -0.53 | 0 | -34.4 | 0.12 | 0 | -0.18 | 28.42 | -0.18 | 28.59 | -0.18 | 28.59 | 0.23 | 28.42 | -17.17 | 0 | -33.52 | -0.06 | -0.18 | -0.06 | 0 | -0.06 | 0 | 0.06
121.152 | -4.04 | -0.06 | -29.24 | 0 | -39.32 | -0.06 | -57.01 | -0.06 | 21.33 | 0.06 | 21.33 | 0.53 | 21.39 | -0.29 | -0.18 | -31.29 | 0.06 | 0 | -0.06 | 28.13 | -0.18 | 28.65 | -0.18 | 28.36 | 0.18 | 28.59 | -14.82 | -0.18 | -31.82 | -0.18 | -85.02 | -0.06 | 0 | -0.06 | 0 | 0.06
121.184 | -2.7 | -0.06 | -26.84 | -0.06 | -38.09 | -0.06 | -43.24 | -0.18 | 21.45 | 0.12 | 21.45 | 0.53 | 21.39 | -0.06 | -0.06 | -28.36 | 0.18 | 0 | -0.23 | 28.83 | -0.29 | 28.42 | -0.18 | 28.65 | 0.35 | 28.48 | -12.89 | -0.18 | -30.82 | -0.18 | -69.85 | -0.18 | 0 | 0 | -61.88 | 0.06
121.216 | -1.7 | 0.06 | -25.14 | -0.06 | -34.81 | -0.06 | -38.32 | -0.06 | 21.39 | 0.06 | 21.39 | 0.53 | 21.33 | 0 | 0 | -27.01 | 0.06 | 0 | -0.18 | 28.59 | -0.18 | 28.42 | -0.18 | 28.65 | 0.29 | 28.59 | -11.02 | -0.06 | -28.65 | -0.18 | -39.02 | -0.06 | -54.67 | -0.18 | -76.94 | 0.06
121.248 | -0.94 | 1.99 | -22.62 | -0.06 | -34.22 | -0.06 | -33.34 | 0 | 21.39 | 0 | 21.45 | 0.41 | 21.39 | 0.23 | 3.34 | -26.66 | 0.12 | 0 | -0.18 | 28.59 | -0.06 | 0.18 | -0.18 | 28.59 | 0.35 | 28.48 | -9.14 | -0.06 | -27.01 | -0.06 | -28.59 | -0.06 | -78.46 | -0.18 | -46.94 | 0
121.28 | -0.47 | 7.79 | -20.57 | 0 | -30.29 | -0.18 | -32.4 | -0.06 | 21.39 | 0.06 | 21.45 | 0.47 | 21.39 | 0.18 | 11.31 | -23.91 | 0 | 0 | -0.18 | 13.48 | -0.18 | 0 | -0.18 | 28.42 | 0.47 | 28.48 | -6.33 | -0.06 | -24.43 | -0.29 | -41.13 | 0 | -40.72 | 0 | -29.12 | -0.18
121.312 | -0.18 | 13.65 | -18.46 | 0 | -28.83 | -0.06 | -32.76 | -0.06 | 0.29 | 0 | 21.39 | 0.41 | 21.27 | 0.18 | 16.7 | -21.91 | 0.06 | 0 | -0.18 | 0.12 | -0.18 | 0 | -0.18 | 28.65 | 0.35 | 28.59 | -4.1 | -0.18 | -22.74 | 0 | -42.36 | -0.18 | -29.42 | 0 | -37.79 | -0.06
121.344 | -0.23 | 19.22 | -16.7 | 0.06 | -27.19 | -0.18 | -31.47 | -0.18 | 0.06 | 0.12 | 21.21 | 0.29 | 21.39 | 0.18 | 21.74 | -20.74 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.48 | -2.52 | 0.06 | -20.92 | -0.06 | -37.56 | -0.06 | -37.38 | -0.06 | -43.13 | -0.18
121.376 | -0.29 | 21.5 | -15 | -0.06 | -24.61 | -0.06 | -30.24 | -0.18 | 0 | 0.06 | 21.39 | 0.23 | 21.39 | 0.12 | 21.39 | -19.34 | 0.12 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 13.89 | 0.35 | 28.48 | -1.7 | 6.97 | -18.75 | -0.18 | -34.92 | 0 | -42.72 | -0.06 | -40.14 | -0.06
121.408 | -0.23 | 21.5 | -13.54 | -0.06 | -22.56 | -0.06 | -27.6 | -0.18 | 0 | -0.06 | 21.39 | 0.18 | 21.39 | 0.18 | 21.39 | -17.99 | 0 | 0 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0.18 | 0.41 | 28.59 | -0.82 | 14.71 | -16.82 | 0 | -33.69 | 0 | -39.96 | -0.06 | -35.39 | -0.06
121.44 | -0.41 | 21.45 | -11.66 | 0 | -20.45 | -0.06 | -24.79 | -0.06 | 0 | -0.06 | 21.39 | 0.18 | 21.27 | 0.29 | 21.39 | -15.06 | 0 | -4.98 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.29 | 28.59 | -0.35 | 24.79 | -14.24 | -0.18 | -33.22 | -0.18 | -34.45 | -0.18 | -33.05 | -0.06
121.472 | -0.35 | 21.5 | -10.14 | -0.06 | -17.99 | -0.06 | -22.79 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.45 | 0.23 | 21.33 | -13.71 | -0.06 | -71.14 | 0 | 0 | -0.18 | 0 | -0.06 | 0.06 | 0.29 | 28.3 | -0.18 | 28.65 | -12.6 | -0.06 | -31.7 | 0 | -33.11 | 0 | -32.93 | -0.06
121.504 | -0.35 | 21.45 | -8.67 | 0.06 | -16.11 | -0.06 | -20.92 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.45 | 0.23 | 21.39 | -12.25 | -0.18 | -57.48 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.18 | 0.12 | 28.65 | -10.78 | -0.06 | -29.53 | 0 | -32.99 | -0.06 | -32.23 | -0.18
121.536 | -0.29 | 21.45 | -7.44 | -0.06 | -14.77 | -0.18 | -19.04 | -0.06 | 0 | -0.06 | 21.39 | 0.06 | 21.39 | 0.29 | 21.27 | -10.31 | 0.06 | -34.22 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.23 | 28.59 | 0.06 | 28.36 | -9.02 | -0.06 | -27.83 | 0 | -32.4 | 0 | -29.65 | 0
121.568 | -0.18 | 21.5 | -5.74 | -0.06 | -13.3 | -0.06 | -17.05 | -0.18 | 0 | -0.06 | 21.5 | 0.06 | 21.21 | 0.18 | 21.33 | -8.73 | -0.06 | -27.54 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.42 | 0.18 | 28.65 | -6.33 | 0 | -26.31 | -0.06 | -29.94 | 0 | -28.48 | 0
121.6 | -0.18 | 21.45 | -3.93 | -0.06 | -11.07 | -0.06 | -14.82 | -0.06 | 0 | -0.18 | 21.45 | 0 | 21.39 | 0.12 | 21.27 | -6.33 | -0.06 | -34.28 | -0.18 | 0 | -0.06 | 0 | -0.12 | 0 | 0 | 28.59 | 0.35 | 28.65 | -3.87 | -0.06 | -25.02 | -0.06 | -27.6 | 0 | -26.37 | 0
121.632 | -0.23 | 21.45 | -2.99 | 0 | -9.79 | -0.06 | -13.07 | -0.06 | 0 | -0.06 | 21.27 | 0.06 | 21.39 | 0.23 | 21.39 | -4.51 | 0 | -34.92 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 28.65 | 0.41 | 28.48 | -2.29 | 0 | -23.03 | 0 | -26.49 | 0 | -25.55 | 0
121.664 | -0.18 | 21.45 | -1.7 | 1.35 | -8.44 | -0.06 | -10.96 | -0.06 | 0 | 0 | 21.39 | -0.06 | 21.33 | 0.18 | 21.33 | -3.16 | -0.18 | -33.17 | -0.18 | 0 | -0.18 | -40.43 | -0.06 | 0 | 0 | 28.65 | 0.41 | 28.65 | -1.41 | -0.06 | -21.45 | 0 | -25.31 | 0 | -23.32 | -0.06
121.696 | -0.23 | 21.45 | -0.94 | 9.08 | -7.09 | 1.17 | -8.56 | -0.18 | 0 | -0.06 | 17.99 | 0 | 21.15 | 0.18 | 21.39 | -2.05 | -0.06 | -29.65 | -0.18 | 0 | -0.06 | -69.73 | -0.06 | 0 | 0.12 | 28.18 | 0.29 | 28.42 | -1.17 | -0.06 | -19.81 | -0.06 | -23.44 | -0.06 | -22.21 | 0
121.728 | -0.29 | 21.45 | -0.53 | 14.53 | -5.57 | 9.2 | -6.39 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.39 | 0.18 | 21.39 | -1.46 | 0 | -27.6 | -0.06 | -52.09 | -0.06 | -45.41 | -0.18 | 0 | 0 | 28.59 | 0.29 | 28.54 | -0.18 | -0.18 | -18.52 | -0.06 | -22.03 | 0 | -20.45 | -0.06
121.76 | -0.29 | 21.45 | -0.23 | 21.8 | -3.75 | 16.35 | -3.75 | 0.06 | -77.35 | 0 | 0 | -0.06 | 21.39 | 0.12 | 21.27 | -0.59 | -0.18 | -26.6 | -0.06 | -58.77 | -0.18 | -46 | -0.18 | 0 | 0.06 | 28.48 | 0.23 | 28.77 | 0 | 0 | -16.7 | -0.18 | -20.27 | -0.06 | -18.63 | -0.06
121.792 | -0.23 | 21.45 | -0.23 | 21.45 | -2.81 | 21.91 | -2.17 | 0.06 | -52.33 | 0 | 0 | -0.18 | 21.39 | 0.18 | 21.27 | -0.47 | 0.82 | -25.43 | -0.18 | -44.01 | 0 | -44.01 | 0 | -1.23 | 0.06 | 0.23 | 0.23 | 28.42 | 0.12 | 0 | -15 | 0.06 | -18.87 | 0 | -17.23 | 0
121.824 | -0.29 | 21.45 | -0.23 | 21.5 | -1.76 | 21.5 | -1.46 | -0.18 | -42.6 | -0.06 | 0 | 0 | 21.62 | -0.06 | 21.39 | -0.29 | 8.67 | -23.15 | -0.06 | -45.7 | 0 | -40.72 | -0.18 | -76.29 | 0 | 0.06 | 0.18 | 28.59 | 0.23 | 6.27 | -13.48 | -0.06 | -17.17 | 0.06 | -15.65 | -0.18
121.856 | -0.06 | 21.39 | -0.29 | 21.45 | -0.76 | 21.5 | -0.59 | -0.06 | -37.44 | -0.06 | 0 | -0.06 | 21.33 | -0.06 | 21.39 | -0.06 | 13.71 | -21.27 | -0.06 | -45.47 | -0.06 | -40.37 | 0 | -59.83 | 0 | 0 | 0.12 | 28.65 | 0.35 | 14.12 | -11.48 | 4.34 | -15.53 | 0 | -13.89 | 0
121.888 | -0.18 | 21.45 | -0.35 | 21.45 | -0.53 | 21.5 | -0.41 | -0.06 | -33.22 | -0.18 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0.12 | 19.63 | -19.81 | -0.06 | -43.13 | 0.06 | -36.92 | -0.06 | -44.83 | -0.06 | 0.06 | 0.12 | 28.42 | 0.29 | 21.91 | -10.08 | 12.48 | -13.48 | 5.04 | -12.07 | -0.06
121.92 | -0.23 | 21.45 | -0.41 | 21.45 | -0.18 | 21.39 | 0 | -0.06 | -32.46 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0 | 21.39 | -17.99 | -0.06 | -39.67 | 0.06 | -36.15 | 0 | -45.24 | 0 | 0 | 0.18 | 28.59 | 0.29 | 28.83 | -7.79 | 22.85 | -11.54 | 15.7 | -10.37 | 9.14
121.952 | -0.18 | 21.45 | -0.35 | 21.5 | -0.18 | 21.45 | 0 | 0.06 | -32.7 | 0 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.12 | 21.27 | -16.06 | 0 | -38.97 | -0.06 | -33.81 | 0 | -42.54 | -0.06 | 0 | 0.18 | 28.48 | 0.41 | 28.48 | -6.09 | 28.83 | -9.96 | 23.32 | -8.79 | 16.29
121.984 | -0.18 | 21.45 | -0.41 | 21.45 | -0.35 | 21.45 | 0.23 | -0.06 | -31.76 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.27 | -14.18 | 0 | -36.86 | -0.06 | -32.05 | 0 | -40.84 | -0.06 | 0.06 | 0 | 28.42 | 0.35 | 28.59 | -5.16 | 28.65 | -8.38 | 29.01 | -6.33 | 23.97
122.016 | -0.18 | 21.45 | -0.23 | 21.39 | -0.35 | 21.5 | 0.18 | 0 | -30.24 | 0 | 0 | 0 | 0 | 0.06 | 21.39 | 0.18 | 21.33 | -12.77 | -0.06 | -34.75 | -0.06 | -30.24 | 0 | -38.97 | 0 | 0 | 0.06 | 28.65 | 0.29 | 28.42 | -3.16 | 28.48 | -6.86 | 28.59 | -4.75 | 28.71
122.048 | -0.18 | 21.45 | -0.29 | 21.39 | -0.35 | 21.45 | 0.35 | -0.18 | -28.07 | 0.12 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.18 | 21.15 | -11.31 | -0.06 | -31.99 | -0.06 | -27.01 | -0.06 | -37.33 | 0 | 0 | 0.18 | 28.42 | 0.35 | 28.42 | -2.11 | 28.36 | -4.92 | 28.36 | -3.28 | 28.42
122.08 | -0.06 | 21.45 | -0.23 | 21.5 | -0.41 | 21.39 | 0.47 | 4.16 | -25.08 | 0 | 0 | 0.06 | 0.06 | -0.06 | 21.39 | 0.35 | 21.33 | -8.91 | 0.06 | -29.59 | -0.06 | -25.9 | 0 | -34.81 | -0.06 | 0 | 0 | 28.65 | 0.18 | 28.65 | -1.29 | 28.42 | -3.4 | 28.54 | -2.11 | 28.59
122.112 | -0.06 | 21.45 | -0.35 | 21.45 | -0.35 | 21.45 | 0.47 | 10.25 | -22.74 | -0.06 | -0.18 | 0.06 | 0 | 0 | 21.39 | 0.35 | 21.33 | -7.15 | -0.06 | -27.31 | -0.06 | -23.67 | 0 | -33.4 | 0 | 0 | 0 | 28.36 | 0.23 | 28.42 | -0.53 | 28.48 | -2.23 | 28.59 | -1.41 | 28.48
122.144 | -0.06 | 17.64 | -0.18 | 21.45 | -0.23 | 21.5 | 0.47 | 15.59 | -21.09 | -0.06 | -60.76 | -0.06 | 0 | 0.06 | 0.12 | 0.23 | 21.39 | -5.39 | 2.64 | -25.25 | 0 | -21.68 | -0.06 | -31.11 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.59 | -0.41 | 28.83 | -1.29 | 28.59 | -0.82 | 28.48
122.176 | -0.06 | 0.12 | -0.29 | 21.5 | -0.35 | 21.45 | 0.35 | 21.5 | -19.1 | 0 | -61.29 | 0.06 | 0 | -0.06 | 0 | 0.23 | 21.21 | -3.93 | 10.78 | -23.32 | -0.06 | -19.86 | 0 | -29.59 | -0.06 | 0 | -0.06 | 28.59 | 0.18 | 28.65 | -0.23 | 28.42 | -0.94 | 28.42 | -0.53 | 28.48
122.208 | -0.18 | 0 | -0.23 | 21.5 | -0.18 | 21.33 | 0.35 | 21.39 | -16.82 | -0.06 | -44.59 | 0 | 0 | -0.06 | 0 | 0.23 | 21.39 | -2.46 | 18.28 | -21.09 | -0.18 | -17.4 | -0.06 | -27.6 | 0 | -4.63 | 0 | 28.65 | 0.12 | 28.48 | 0 | 28.48 | -0.64 | 28.48 | -0.47 | 28.42
122.24 | -0.06 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | 0.47 | 21.39 | -14.82 | -0.06 | -37.79 | -0.06 | 0 | 0 | 0 | 0.23 | 21.33 | -1.17 | 28.71 | -18.16 | 0 | -15.47 | 0 | -24.38 | 0 | -87.13 | 0 | 26.13 | 0.06 | 28.36 | 0 | 28.48 | -0.29 | 28.42 | -0.06 | 28.48
122.272 | -0.06 | 0 | -0.23 | 21.45 | -0.23 | 21.39 | 0.35 | 21.21 | -12.6 | -0.06 | -32.46 | 0 | 0 | -0.06 | 0 | 0.23 | 21.15 | -1.05 | 28.65 | -16.47 | 0 | -14.53 | 0 | -22.91 | -0.06 | -53.91 | -0.06 | 0.18 | 0 | 28.54 | 0.12 | 28.59 | -0.18 | 28.42 | -0.06 | 28.42
122.304 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.45 | 0.35 | 21.33 | -10.49 | -0.06 | -31.76 | 0 | -31.41 | 0.06 | 0 | 0.29 | 21.39 | -0.47 | 28.65 | -14.53 | 0.06 | -12.95 | 0 | -20.86 | 0 | -49.04 | -0.06 | 0 | 0.06 | 28.54 | 0.23 | 28.48 | -0.18 | 28.42 | 0 | 28.59
122.336 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.29 | 21.39 | -8.14 | 0 | -32.46 | -0.06 | -71.55 | 0.06 | 0 | 0.18 | 21.21 | -0.35 | 28.65 | -13.13 | 9.02 | -11.6 | -0.06 | -18.69 | 0 | -43.54 | 0 | 0 | 0 | 28.59 | 0.12 | 28.48 | 0 | 28.24 | 0 | 28.65
122.368 | -0.06 | 0.06 | -0.18 | 21.39 | -0.29 | 21.5 | 0.18 | 21.39 | -6.21 | -0.06 | -31.58 | -0.18 | -48.22 | -0.06 | 0 | 0.23 | 21.21 | -0.06 | 28.42 | -11.84 | 15.94 | -9.79 | 4.45 | -17.05 | -0.06 | -37.85 | -0.18 | 0 | 0.12 | 28.42 | 0.23 | 28.65 | 0.12 | 28.3 | 0.12 | 28.48
122.4 | 0 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | 0.35 | 21.39 | -3.34 | 0 | -29.53 | -0.06 | -27.77 | -0.06 | 0 | 0.18 | 21.39 | -0.06 | 28.65 | -9.79 | 26.49 | -8.44 | 15.06 | -15 | -0.18 | -36.74 | 0 | 0 | 0.06 | 28.48 | 0.18 | 28.48 | 0.06 | 28.59 | 0.23 | 28.24
122.432 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | 0.12 | 21.39 | -1.93 | -0.06 | -27.31 | -0.06 | -31.88 | -0.06 | 0 | 0.29 | 21.27 | -0.06 | 28.3 | -8.56 | 28.59 | -7.09 | 23.09 | -13.89 | 0 | -38.15 | 0 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | 0.06 | 28.65 | 0.23 | 28.24
122.464 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | 0.12 | 21.33 | -1.41 | 2.29 | -25.55 | -0.06 | -36.21 | 0.06 | 0 | 0.23 | 21.21 | 0.12 | 28.71 | -7.09 | 28.83 | -6.15 | 28.71 | -12.48 | 0 | -37.38 | 0 | 0 | 0 | 28.48 | 0.12 | 28.54 | 0.29 | 28.48 | 0.12 | 28.42
122.496 | -0.06 | 0 | -0.06 | 21.39 | -0.18 | 21.45 | 0.06 | 21.39 | -0.64 | 7.56 | -23.85 | -0.18 | -34.51 | -0.06 | 0 | 0.18 | 21.21 | 0.12 | 28.65 | -6.56 | 28.59 | -4.98 | 28.65 | -10.9 | 0 | -36.1 | 0 | 0 | -0.18 | 28.59 | 0.18 | 28.54 | 0.18 | 28.59 | 0.06 | 28.65
122.528 | -0.18 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.06 | 21.39 | -0.23 | 13.13 | -21.91 | 0 | -31.17 | -0.06 | 0 | 0.18 | 21.33 | 0.18 | 28.3 | -4.98 | 28.59 | -3.46 | 28.65 | -9.79 | 0 | -33.81 | 0 | 0 | 0 | 28.48 | 0.12 | 28.59 | 0.18 | 28.48 | 0.18 | 28.71
122.56 | -0.06 | 0 | -0.18 | 0.12 | -0.18 | 21.45 | 0 | 21.39 | -0.06 | 20.74 | -19.16 | -0.06 | -27.77 | -0.18 | -33.28 | 0.18 | 21.39 | 0.18 | 28.3 | -3.87 | 28.42 | -3.22 | 28.65 | -7.97 | 0.12 | -31.52 | -0.18 | 0 | -0.06 | 28.36 | 0.06 | 28.65 | 0.18 | 28.48 | 0.18 | 28.48
122.592 | -0.18 | -44.01 | -0.18 | 0 | -0.18 | 0.12 | 0 | 21.39 | 0.18 | 21.5 | -17.23 | -0.06 | -26.84 | 0 | -68.03 | 0.23 | 21.33 | 0.12 | 28.48 | -3.11 | 28.48 | -2.23 | 29.06 | -6.8 | 0 | -29.3 | -0.06 | 0 | 0 | 28.48 | 0.06 | 28.48 | 0.12 | 28.48 | 0.12 | 28.89
122.624 | -0.06 | -54.08 | -0.18 | 0 | -0.18 | 0.06 | 0 | 21.39 | 0.35 | 21.39 | -15.23 | -0.06 | -26.43 | -0.06 | -46.7 | 0.18 | 21.39 | 0.12 | 28.59 | -2.46 | 28.3 | -1.76 | 28.65 | -5.51 | 0 | -28.01 | 0.06 | 0 | -0.06 | 28.42 | -0.06 | 28.83 | 0.12 | 28.54 | 0.12 | 28.42
122.656 | -0.18 | -41.25 | -0.18 | 0.06 | -0.18 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | -13.36 | 0 | -24.43 | -0.06 | -29.36 | 0.23 | 21.39 | 0.18 | 28.59 | -1.88 | 28.59 | -1.41 | 28.54 | -4.39 | -0.06 | -26.6 | -0.06 | 0 | -0.18 | 28.48 | 0.06 | 28.48 | 0.12 | 28.42 | 0.12 | 28.59
122.688 | -0.06 | -40.02 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.5 | 0.41 | 21.39 | -11.31 | -0.06 | -22.15 | -0.18 | -29.94 | 0.12 | 13.42 | 0.18 | 28.42 | -1.46 | 28.83 | -0.94 | 28.65 | -3.69 | 0 | -24.55 | -0.06 | -45.47 | -0.06 | 28.42 | 0.12 | 28.48 | 0.12 | 28.48 | 0.12 | 28.42
122.72 | -0.18 | -37.09 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -9.08 | -0.06 | -19.86 | -0.06 | -35.98 | 0.06 | 0.06 | 0.18 | 28.3 | -1.05 | 28.59 | -0.64 | 28.3 | -2.58 | -0.06 | -22.21 | -0.06 | -71.14 | -0.18 | 0.18 | 0 | 28.42 | 0.12 | 28.65 | 0.18 | 28.42
122.752 | 0 | -34.75 | -0.18 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -6.33 | -0.06 | -18.93 | -0.06 | -34.22 | 0.12 | 0 | 0.23 | 29.06 | -0.94 | 28.65 | -0.35 | 28.13 | -2.11 | 0.06 | -19.86 | 0 | -49.1 | -0.06 | 0 | -0.06 | 0.23 | 0.06 | 28.42 | 0.18 | 28.42
122.784 | 0 | -33.17 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.41 | 21.39 | -4.45 | 0 | -17.05 | -0.18 | -31.47 | 0.06 | 0 | 0 | 28.65 | -0.7 | 28.48 | -0.35 | 28.3 | -1.46 | 0 | -18.52 | 0.06 | -46.99 | -0.06 | 0 | -0.06 | 0.06 | 0.06 | 0.29 | 0 | 28.42
122.816 | 0 | -31.47 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -2.29 | 0 | -15.59 | -0.06 | -28.54 | -0.06 | 0 | 0.23 | 28.59 | -0.47 | 28.36 | -0.29 | 28.65 | -1 | 4.51 | -16.06 | -0.06 | -40.55 | -0.06 | 0.06 | 0 | 0 | 0.12 | 0.06 | 0.06 | 0.23
122.848 | -0.06 | -29.06 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.41 | 21.39 | -1.99 | 1.99 | -13.71 | -0.18 | -27.36 | 0 | 0 | 0.18 | 28.48 | -0.41 | 28.36 | -0.29 | 28.54 | -0.64 | 12.89 | -14.65 | -0.18 | -36.45 | -0.18 | 0 | 0 | 0.06 | 0 | 0 | 0.06 | 0.06
122.88 | -0.06 | -26.54 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.29 | 21.39 | -0.64 | 8.79 | -11.13 | -0.06 | -26.08 | 0 | 0 | 0.12 | 28.48 | -0.41 | 28.48 | -0.18 | 28.89 | -0.47 | 22.97 | -11.95 | 0 | -37.03 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0 | 0
122.912 | -0.06 | -24.84 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | -0.29 | 14.24 | -9.9 | -0.06 | -24.32 | -0.06 | 0 | 0.12 | 28.59 | -0.35 | 28.59 | -0.35 | 28.71 | -0.29 | 28.95 | -9.84 | -0.18 | -38.44 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.06 | 0
122.944 | 0 | -22.21 | -0.12 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | -0.06 | 19.69 | -7.91 | 1.88 | -22.15 | 0 | 0 | 0.18 | 28.54 | -0.29 | 28.42 | -0.29 | 28.65 | -0.29 | 28.42 | -7.85 | -0.18 | -36.86 | 0 | 0 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0
122.976 | -0.06 | -20.92 | -0.06 | -2.23 | -0.06 | 0 | 0.06 | 0.23 | 0.18 | 21.45 | 0.12 | 21.5 | -6.39 | 9.26 | -20.39 | 0.12 | 0 | 0.06 | 28.48 | -0.29 | 28.65 | -0.23 | 28.48 | -0.29 | 28.48 | -5.63 | -0.06 | -35.33 | 0 | 0 | -0.23 | 0 | -0.06 | 0 | 0.06 | 0
123.008 | 0 | -18.46 | -0.06 | -61 | -0.18 | -3.52 | -0.18 | 0.06 | 0.18 | 21.39 | 0.23 | 21.21 | -4.45 | 14.47 | -19.28 | 0.06 | 0 | 0.18 | 28.36 | -0.29 | 28.36 | -0.29 | 28.89 | -0.29 | 28.59 | -3.63 | 0 | -33.46 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
123.04 | -0.06 | -15.82 | -0.06 | -45.18 | -0.06 | -62.64 | -0.06 | 0 | 0.12 | 21.45 | 0.29 | 21.39 | -2.46 | 21.8 | -16.93 | 0.12 | 0 | 0.12 | 0.29 | -0.23 | 28.54 | -0.18 | 28.77 | -0.29 | 28.65 | -1.7 | -0.06 | -30.47 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | -0.06 | 0
123.072 | -0.06 | -14.47 | -0.06 | -40.37 | -0.06 | -46.58 | -0.06 | 0 | 0.18 | 21.21 | 0.35 | 21.39 | -2.05 | 21.39 | -15.06 | -0.06 | 0 | 0.12 | 0.12 | -0.18 | 28.42 | -0.18 | 28.36 | -0.29 | 28.54 | -1.23 | -0.06 | -28.48 | 0 | 0 | -0.12 | 0 | 0 | 0 | -0.18 | 0
123.104 | -0.06 | -12.89 | -0.06 | -38.5 | -0.18 | -40.55 | -0.06 | 0 | 0.06 | 21.39 | 0.35 | 21.39 | -1.05 | 21.45 | -13.36 | -0.06 | -0.23 | 0.18 | 0 | -0.23 | 28.59 | -0.06 | 28.59 | -0.18 | 28.71 | -0.64 | -0.06 | -27.01 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0
123.136 | 6.74 | -11.6 | -0.18 | -36.62 | 0.06 | -38.97 | 0 | 0 | 0.06 | 21.45 | 0.47 | 21.39 | -0.94 | 21.15 | -12.13 | 0.12 | -65.98 | 0.12 | 0.06 | -0.23 | 28.36 | -0.18 | 28.42 | -0.29 | 28.07 | -0.18 | -0.06 | -25.25 | -0.06 | -38.03 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0
123.168 | 12.31 | -10.37 | -0.06 | -34.45 | -0.18 | -36.74 | -0.06 | 0 | 0 | 21.33 | 0.53 | 21.33 | -0.35 | 21.33 | -10.31 | -0.06 | -59.12 | 0.12 | 0 | -0.23 | 28.48 | -0.06 | 28.48 | -0.18 | 28.48 | 0 | 0 | -23.67 | -0.06 | -83.56 | -0.18 | -25.25 | -0.06 | 0 | -0.18 | 0
123.2 | 17.93 | -8.32 | -0.18 | -32.23 | -0.18 | -34.04 | -0.06 | 0 | 0.06 | 21.33 | 0.35 | 21.33 | -0.18 | 21.33 | -8.14 | 0.12 | -31.52 | 0.12 | 0 | -0.18 | 14.18 | -0.18 | 28.65 | -0.23 | 28.48 | 0.06 | -0.06 | -21.45 | 0 | -48.69 | -0.06 | -77.76 | -0.06 | -49.22 | -0.18 | 0
123.232 | 21.74 | -7.15 | -0.06 | -30.24 | -0.06 | -31.88 | -0.06 | 0 | 0.06 | 21.5 | 0.47 | 21.45 | -0.06 | 21.33 | -5.86 | 0.06 | -28.54 | 0.12 | 0 | -0.06 | 0.12 | -0.18 | 28.83 | -0.18 | 28.48 | 0.06 | 9.32 | -19.28 | 0 | -46.82 | 0 | -50.45 | -0.12 | -77.41 | -0.18 | -65.22
123.264 | 21.5 | -5.22 | -0.06 | -28.18 | -0.06 | -29.77 | 0 | 0 | 0 | 21.45 | 0.29 | 21.45 | -0.18 | 21.33 | -4.39 | 0 | -34.81 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0.23 | -0.06 | 28.42 | 0.23 | 16.41 | -17.05 | 0 | -41.66 | -0.06 | -30.59 | -0.06 | -51.33 | -0.06 | -75.65
123.296 | 21.5 | -4.1 | -0.06 | -26.25 | -0.06 | -28.18 | -0.18 | 0 | 0 | 21.5 | 0.35 | 21.39 | 0 | 21.27 | -2.93 | -0.06 | -35.51 | 0.12 | 0 | 0 | 0 | -0.18 | 0.06 | -0.18 | 28.59 | 0.41 | 24.26 | -15.06 | -0.06 | -36.74 | -0.06 | -33.17 | -0.06 | -29.42 | -0.18 | -45.24
123.328 | 21.45 | -2.64 | -0.06 | -23.91 | -0.06 | -25.84 | -0.06 | 0 | 0 | 21.45 | 0.29 | 21.45 | 0.12 | 21.39 | -2.52 | 0 | -33.4 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.41 | 28.89 | -13.65 | 0 | -37.21 | -0.06 | -41.6 | -0.06 | -33.46 | -0.06 | -29.71
123.36 | 21.45 | -1.35 | -0.06 | -21.45 | -0.06 | -23.2 | -0.06 | 0 | 0 | 0.59 | 0.18 | 21.27 | 0.12 | 21.33 | -0.94 | 0 | -28.77 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.48 | 0.41 | 28.54 | -11.31 | 0 | -38.85 | -0.06 | -40.2 | -0.06 | -42.95 | -0.06 | -37.97
123.392 | 21.45 | -0.64 | -0.06 | -19.51 | -0.06 | -21.45 | -0.06 | -16.99 | 0 | 0.06 | 0.23 | 21.39 | 0.29 | 21.33 | -0.82 | -0.06 | -27.6 | 0.12 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.29 | 28.59 | -8.79 | 0 | -36.86 | -0.06 | -35.92 | -0.06 | -39.96 | -0.06 | -43.13
123.424 | 21.45 | -0.53 | -0.06 | -17.29 | -0.06 | -19.22 | -0.06 | -77.05 | 0 | 0 | 0.12 | 21.39 | 0.23 | 21.39 | -0.35 | -0.06 | -26.9 | 0.12 | 0 | -0.23 | 0 | -0.18 | 0 | -0.18 | 28.48 | 0.23 | 28.59 | -6.8 | 0 | -35.1 | -0.18 | -33.17 | 0 | -36.1 | -0.06 | -39.67
123.456 | 21.39 | -0.29 | -0.06 | -15.65 | 0 | -17.46 | 0 | -52.38 | -0.06 | 0.06 | 0.12 | 21.33 | 0.23 | 21.39 | -0.47 | 0.06 | -24.9 | 0.12 | -0.76 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.71 | 0.23 | 28.89 | -4.98 | -0.06 | -33.17 | -0.18 | -32.76 | -0.06 | -33.11 | -0.06 | -35.45
123.488 | 21.45 | -0.18 | -0.06 | -14.06 | -0.06 | -15.94 | 0 | -43.13 | 0 | 0 | 0.12 | 21.33 | 0.12 | 21.45 | -0.06 | -0.06 | -22.62 | 0.18 | -67.21 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.35 | 28.48 | -3.11 | -0.06 | -31.58 | -0.18 | -32.7 | 0 | -32.93 | -0.06 | -33.4
123.52 | 21.45 | -0.23 | -0.06 | -12.31 | -0.06 | -13.77 | -0.06 | -34.63 | 0.06 | 0 | 0 | 21.33 | 0.12 | 21.39 | -0.06 | 0.06 | -20.45 | 0.23 | -54.26 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.13 | 0.18 | 28.48 | -1.7 | 10.25 | -29.3 | -0.06 | -30.06 | -0.06 | -32.17 | -0.18 | -32.76
123.552 | 21.45 | -0.35 | -0.06 | -10.96 | -0.06 | -12.31 | -0.06 | -31.99 | -0.18 | 0 | 0.12 | 21.39 | 0.29 | 21.27 | 0.06 | -0.06 | -19.45 | 0.12 | -44.77 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.59 | 0.29 | 28.36 | -0.94 | 17.81 | -26.72 | -0.06 | -28.36 | -0.06 | -29.94 | 0 | -32.23
123.584 | 21.45 | -0.47 | 5.33 | -9.32 | -0.06 | -10.96 | 0 | -32.93 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.27 | -0.06 | 0 | -18.16 | 0.12 | -44.77 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.24 | 0.12 | 28.42 | -0.47 | 25.67 | -25.2 | -0.18 | -26.78 | 0 | -28.18 | -0.06 | -29.94
123.616 | 21.45 | -0.35 | 11.25 | -8.14 | -0.06 | -9.61 | -0.06 | -32.34 | -0.06 | 0 | 0.06 | 21.33 | 0.23 | 21.33 | 0.18 | 2.93 | -15.94 | 0 | -42.89 | -0.06 | 0.06 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.42 | -0.06 | 28.65 | -24.08 | -0.06 | -25.72 | 0 | -27.01 | -0.18 | -27.95
123.648 | 21.45 | -0.35 | 16.7 | -6.68 | -0.06 | -8.14 | -0.06 | -31.47 | 0 | 0 | 0 | 21.39 | 0.12 | 21.45 | 0.18 | 9.2 | -14 | 0.12 | -41.78 | -0.18 | -49.1 | -0.06 | 0 | -0.18 | 28.48 | 0.18 | 28.42 | 0.12 | 28.59 | -21.74 | 0 | -23.91 | -0.18 | -25.43 | -0.06 | -26.84
123.68 | 21.45 | -0.23 | 21.5 | -4.51 | -0.06 | -5.98 | 0 | -29.01 | 0 | 0 | 0.06 | 21.27 | 0.18 | 21.39 | 0.35 | 16.29 | -12.25 | 0.06 | -38.61 | -0.06 | -59.83 | -0.18 | -41.9 | -0.06 | 28.3 | 0 | 28.48 | 0.35 | 28.54 | -19.34 | -0.06 | -22.38 | 0 | -23.79 | 0 | -25.31
123.712 | 21.45 | -0.29 | 21.5 | -3.52 | -0.06 | -4.8 | -0.06 | -26.31 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.39 | 0.29 | 21.74 | -10.55 | 0.12 | -36.45 | -0.18 | -44.83 | -0.18 | -72.31 | -0.06 | 0.29 | 0 | 28.54 | 0.35 | 28.71 | -17.64 | 0 | -20.57 | -0.06 | -21.68 | -0.18 | -23.61
123.744 | 21.39 | -0.18 | 21.45 | -2.11 | 6.21 | -3.63 | -0.06 | -24.96 | 0 | 0 | 0 | 10.61 | 0.12 | 21.33 | 0.29 | 21.39 | -8.85 | 0.06 | -34.86 | -0.06 | -44.94 | 0 | -48.75 | -0.06 | 0.06 | 0.12 | 28.65 | 0.35 | 28.59 | -15.76 | -0.06 | -19.16 | -0.06 | -20.33 | -0.06 | -21.74
123.776 | 21.45 | -0.18 | 21.39 | -1.46 | 12.19 | -2.4 | 0 | -22.79 | -0.06 | -0.06 | 0.06 | 0.12 | 0.18 | 21.39 | 0.23 | 21.33 | -7.15 | -0.06 | -32.58 | -0.06 | -43.48 | 0.06 | -46.23 | -0.06 | 0 | 0 | 28.42 | 0.35 | 28.65 | -13.71 | -0.06 | -17.58 | -0.06 | -18.52 | -0.06 | -20.27
123.808 | 21.45 | -0.18 | 21.45 | -0.76 | 17.75 | -1.64 | 6.27 | -20.68 | 0 | -61.17 | -0.06 | 0 | 0 | 21.27 | 0.18 | 21.33 | -5.39 | 0.12 | -31.52 | -0.06 | -40.55 | 0 | -44.71 | -0.06 | 0.06 | 0 | 28.54 | 0.41 | 28.59 | -11.54 | -0.06 | -15.82 | -0.06 | -17.23 | -0.06 | -19.22
123.84 | 21.45 | -0.35 | 21.39 | -0.18 | 21.5 | -0.64 | 14 | -17.75 | -0.06 | -54.44 | -0.06 | 0 | 0.12 | 0.12 | 0.23 | 21.39 | -3.52 | 0.12 | -28.71 | -0.06 | -39.2 | 0 | -41.43 | 0 | 0 | -0.18 | 28.48 | 0.41 | 28.42 | -8.79 | 9.9 | -14.06 | 0 | -15.65 | 0 | -16.88
123.872 | 21.39 | -0.29 | 21.45 | -0.06 | 21.45 | -0.18 | 19.28 | -16.29 | -0.06 | -43.3 | -0.06 | 0 | 0.06 | 0.06 | 0.23 | 21.39 | -2.46 | 0.06 | -26.54 | 0 | -36.86 | 0 | -40.9 | 0 | 0 | 0 | 28.48 | 0.41 | 28.65 | -7.03 | 17.05 | -12.25 | -0.06 | -13.65 | 8.44 | -15.18
123.904 | 21.45 | -0.23 | 21.45 | -0.18 | 21.5 | -0.18 | 21.56 | -14.18 | -0.18 | -37.38 | -0.06 | 0 | 0 | 0 | 0.29 | 21.33 | -1.52 | 0.18 | -24.84 | 0 | -35.45 | -0.18 | -37.74 | -0.06 | 0 | 0 | 28.3 | 0.35 | 28.36 | -5.16 | 24.84 | -10.31 | 7.73 | -12.07 | 15.59 | -13.48
123.936 | 21.45 | -0.23 | 21.39 | -0.18 | 21.45 | -0.18 | 21.45 | -12.72 | -0.18 | -33.34 | 0 | 0 | 0.12 | 0 | 0.29 | 21.33 | -1 | 0.12 | -22.38 | -0.06 | -34.28 | -0.06 | -36.27 | -0.06 | 0 | 0 | 28.59 | 0.29 | 28.42 | -3.11 | 28.71 | -8.67 | 15.18 | -10.55 | 23.32 | -12.07
123.968 | 21.45 | -0.23 | 21.45 | -0.35 | 21.45 | -0.18 | 21.39 | -10.37 | 0.06 | -32.29 | -0.06 | 0 | 0.12 | 0 | 0.18 | 21.33 | -0.64 | 0 | -20.98 | 0 | -32.23 | -0.06 | -34.4 | 0 | 0 | 0 | 28.24 | 0.18 | 28.65 | -1.76 | 28.54 | -7.03 | 22.85 | -9.08 | 29.01 | -10.37
124 | 21.45 | -0.18 | 21.5 | -0.35 | 21.45 | -0.23 | 21.39 | -7.44 | 0.06 | -32.76 | 0 | 0 | 0 | 0 | 0.18 | 21.33 | -0.18 | 0 | -18.46 | -0.06 | -29.24 | -0.06 | -32.58 | 0 | 0 | -0.06 | 28.42 | 0.18 | 28.65 | -0.82 | 28.65 | -5.04 | 28.54 | -6.86 | 28.54 | -8.38
124.032 | 0.12 | -0.18 | 21.56 | -0.35 | 21.45 | -0.23 | 21.5 | -5.45 | -0.06 | -31.52 | -0.06 | 0 | 0 | 0 | 0.23 | 21.39 | -0.18 | 0.06 | -16.41 | 0.06 | -27.36 | -0.06 | -30.29 | 0 | 0 | -0.06 | 28.42 | 0.29 | 28.54 | -0.59 | 28.48 | -3.93 | 28.48 | -5.22 | 28.83 | -6.68
124.064 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.41 | 21.33 | -3.28 | -0.06 | -29.71 | -0.18 | 0 | 0.06 | 0 | 0.12 | 21.33 | 0 | 0 | -15 | -0.06 | -25.61 | 0 | -28.48 | 0 | 0 | -0.06 | 28.48 | 0.12 | 28.42 | -0.18 | 28.65 | -2.46 | 28.71 | -3.87 | 28.65 | -4.75
124.096 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.35 | 21.39 | -2.05 | 0.06 | -27.48 | -0.18 | 0 | 0 | 0 | 0.12 | 21.27 | 0 | 0.06 | -13.48 | 0 | -23.03 | 0 | -27.07 | 0 | 0 | -0.18 | 28.48 | 0.12 | 28.59 | 0 | 28.42 | -1.93 | 28.65 | -2.52 | 28.65 | -3.57
124.128 | 0 | -0.18 | 21.39 | -0.29 | 21.45 | -0.23 | 21.39 | -1.11 | 0.12 | -25.43 | 0 | 0 | -0.06 | 0 | 0.12 | 21.39 | 0.06 | 1.99 | -12.36 | -0.06 | -21.68 | -0.06 | -24.26 | -0.06 | -20.45 | -0.06 | 0.23 | 0.12 | 28.54 | 0.18 | 28.54 | -1.17 | 28.54 | -1.64 | 28.48 | -2.58
124.16 | 0 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.27 | -0.59 | -0.18 | -22.62 | -0.06 | -3.11 | 0 | 0 | 0.12 | 21.5 | 0.12 | 8.56 | -10.61 | -0.06 | -19.28 | 0 | -22.03 | -0.18 | -79.63 | -0.18 | 0 | 0 | 28.42 | 0.23 | 28.48 | -0.41 | 28.3 | -0.7 | 28.65 | -1.64
124.192 | 0 | -0.18 | 21.45 | -0.18 | 21.45 | -0.18 | 21.39 | 0.06 | -0.18 | -20.68 | 0 | -71.37 | -0.06 | 0 | 0.18 | 21.27 | 0.18 | 13.77 | -9.43 | -0.06 | -17.05 | 0 | -19.98 | -0.06 | -52.68 | -0.06 | 0 | 0 | 28.36 | 0.35 | 28.36 | -0.29 | 28.59 | -0.64 | 28.48 | -0.94
124.224 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | -0.23 | 21.33 | -0.06 | 0.06 | -18.93 | -0.06 | -56.78 | 0 | 0 | 0.06 | 21.33 | 0.12 | 19.4 | -7.91 | 0.06 | -15.47 | -0.18 | -18.69 | -0.06 | -48.46 | -0.06 | 0 | -0.06 | 28.65 | 0.29 | 28.48 | -0.18 | 28.42 | -0.29 | 28.42 | -0.76
124.256 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.29 | 21.39 | 0.18 | -0.06 | -17.46 | 0 | -42.78 | -0.06 | -3.93 | 0.06 | 21.39 | 0.23 | 21.39 | -6.8 | 0 | -14.3 | 0 | -16.58 | -0.06 | -43.3 | 0 | 0 | 0 | 28.48 | 0.29 | 28.48 | -0.06 | 28.48 | -0.18 | 28.48 | -0.29
124.288 | 0 | -0.06 | 21.39 | -0.23 | 21.45 | -0.29 | 21.39 | 0.18 | -0.18 | -15 | -0.18 | -37.97 | -0.06 | -70.55 | 0.12 | 21.33 | 0.18 | 21.39 | -5.92 | 0 | -12.77 | -0.06 | -14.71 | 0 | -37.68 | 0 | 0 | 0 | 28.48 | 0.35 | 28.77 | 0.06 | 28.48 | -0.18 | 28.42 | -0.18
124.32 | 0 | -0.06 | 21.5 | -0.29 | 21.39 | -0.23 | 21.39 | 0.53 | -0.06 | -12.72 | -0.06 | -31.99 | 0.06 | -47.46 | 0 | 21.45 | 0.18 | 21.15 | -4.51 | 0 | -11.13 | 1.99 | -13.18 | -0.06 | -37.03 | -0.06 | 0 | 0 | 28.48 | 0.23 | 28.54 | 0.06 | 28.42 | 0 | 28.65 | 0
124.352 | 0 | -0.18 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | 0.41 | -0.06 | -10.66 | 0 | -32.46 | -0.06 | -29.83 | 0.06 | 21.39 | 0.06 | 21.27 | -3.57 | 0 | -9.55 | 7.85 | -11.54 | 0.06 | -37.68 | -0.06 | 0 | 0 | 28.13 | 0.23 | 28.48 | 0.06 | 28.54 | 0 | 28.59 | 0
124.384 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.18 | 21.45 | 0.53 | -0.06 | -8.44 | -0.06 | -33.11 | 0 | -28.59 | 0 | 21.33 | 0.23 | 21.39 | -2.99 | 1.99 | -8.32 | 15.12 | -9.84 | -0.06 | -36.97 | -0.06 | 0 | 0 | 23.38 | 0.18 | 28.59 | 0.18 | 28.59 | 0.18 | 28.42 | -0.06
124.416 | -0.06 | -0.06 | 21.45 | -0.18 | 21.5 | -0.23 | 21.39 | 0.53 | 0.23 | -6.68 | -0.06 | -31.52 | 0 | -34.4 | 0.06 | 21.33 | 0.12 | 21.39 | -2.4 | 7.73 | -7.03 | 22.91 | -8.91 | 0.06 | -35.57 | 0 | 0 | -0.06 | 0.18 | 0.23 | 28.42 | 0.06 | 28.48 | 0.18 | 28.36 | 0.12
124.448 | -37.97 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.41 | 7.62 | -4.28 | 0 | -30.29 | -0.06 | -35.51 | 0 | 21.33 | 0.23 | 21.27 | -1.82 | 15.12 | -5.92 | 28.65 | -7.44 | 0 | -33.46 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.54 | 0.18 | 28.48 | 0.12 | 28.89 | 0.06
124.48 | -57.89 | -0.06 | 0.12 | -0.18 | 21.45 | -0.23 | 21.39 | 0.47 | 15 | -2.05 | -0.06 | -27.01 | -0.18 | -31 | -0.06 | 21.33 | 0.12 | 21.45 | -1.23 | 25.43 | -4.75 | 28.65 | -5.8 | 0.76 | -30.7 | 0 | 0 | -0.06 | 0 | 0.06 | 28.89 | 0.12 | 28.48 | 0.18 | 28.18 | 0.06
124.512 | -42.13 | -0.06 | 0 | -0.06 | 21.39 | -0.23 | 21.39 | 0.29 | 20.63 | -1.46 | 0 | -25.37 | -0.18 | -28.24 | 0.06 | 0.12 | 0.18 | 21.33 | -0.94 | 28.83 | -3.75 | 28.48 | -4.8 | 10.61 | -29.42 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.42 | 0.18 | 28.65 | 0.23 | 28.42 | 0.12
124.544 | -39.9 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.39 | 0.35 | 21.5 | -0.53 | 0.12 | -23.15 | -0.06 | -27.07 | 0 | 0.06 | 0.35 | 21.39 | -0.76 | 28.59 | -2.87 | 28.59 | -4.04 | 18.16 | -27.77 | -0.06 | -67.21 | 0 | 0 | 0.06 | 28.48 | 0.18 | 28.77 | 0.18 | 28.54 | 0.12
124.576 | -38.15 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.33 | 0.35 | 21.39 | -0.18 | 0 | -21.21 | -0.06 | -26.54 | 0.06 | 0 | 0.12 | 21.27 | -0.64 | 28.59 | -2.23 | 28.42 | -3.16 | 25.72 | -25.78 | 0 | -73.66 | 0 | 0 | 0.06 | 28.71 | 0.12 | 28.42 | 0.29 | 28.65 | 0.12
124.608 | -35.16 | 0 | 0 | -0.18 | 21.45 | -0.18 | 21.56 | 0.35 | 21.39 | 0.06 | -0.06 | -18.93 | -0.06 | -25.14 | 0 | 0 | 0.23 | 21.27 | -0.47 | 28.83 | -1.7 | 28.48 | -2.58 | 28.65 | -24.38 | -0.18 | -50.69 | -0.18 | 0 | 0.06 | 28.77 | 0.18 | 28.48 | 0.12 | 28.48 | 0.18
124.64 | -34.16 | -0.06 | 0 | -0.06 | 0.12 | -0.06 | 21.33 | 0.35 | 21.45 | 0.18 | -0.06 | -16.7 | -0.18 | -22.44 | -0.06 | 0 | 0.18 | 21.33 | -0.53 | 29.06 | -1.29 | 28.18 | -1.7 | 28.59 | -21.68 | -0.18 | -45.7 | -0.18 | 0 | 0 | 28.48 | 0.06 | 28.48 | 0.12 | 28.48 | 0.12
124.672 | -31.23 | -0.18 | 0 | -0.18 | 0 | -0.18 | 14 | 0.18 | 21.45 | 0.47 | -0.06 | -15.18 | -0.06 | -20.68 | 0.06 | 0 | 0.06 | 21.15 | -0.41 | 28.65 | -0.94 | 28.77 | -1.17 | 28.65 | -19.75 | -0.18 | -39.79 | -0.06 | 0 | -0.06 | 28.24 | -0.06 | 28.65 | 0.18 | 28.54 | 0.18
124.704 | -29.42 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0.12 | 0.12 | 21.39 | 0.47 | -0.06 | -12.83 | -0.06 | -19.16 | 0 | 0 | 0.12 | 21.39 | -0.41 | 28.42 | -0.7 | 28.59 | -0.82 | 28.65 | -18.05 | -0.06 | -36.56 | -0.06 | 0 | 0.06 | 26.02 | 0 | 28.48 | 0.06 | 28.54 | 0.18
124.736 | -26.6 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.33 | 0.53 | -0.18 | -11.02 | -0.06 | -17.75 | 0 | 0 | 0.23 | 21.39 | -0.41 | 28.65 | -0.47 | 28.13 | -0.64 | 28.48 | -16 | 0 | -37.38 | -0.18 | 0 | 0 | 0.18 | 0 | 28.48 | 0.18 | 28.48 | 0
124.768 | -25.08 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.33 | 0.47 | 0 | -8.91 | -0.06 | -16.17 | -0.06 | 0 | 0.12 | 21.33 | -0.47 | 28.48 | -0.29 | 28.42 | -0.59 | 28.65 | -14.12 | -0.06 | -38.32 | 0 | 0 | -0.18 | 0 | 0 | 28.48 | 0.12 | 0.23 | 0.12
124.8 | -23.15 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 21.39 | 0.53 | 2.11 | -5.68 | 0 | -13.54 | 0 | 0 | 0.18 | 21.39 | -0.35 | 28.36 | -0.41 | 28.48 | -0.23 | 28.59 | -11.95 | 0 | -36.33 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.18 | 0 | 0 | 0.12
124.832 | -20.22 | -0.06 | 0 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | 7.68 | -3.93 | -0.06 | -12.07 | -0.06 | 0 | 0.23 | 21.39 | -0.29 | 28.65 | -0.29 | 28.36 | -0.35 | 28.95 | -9.61 | 0 | -34.51 | 0 | -81.57 | 0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0.12
124.864 | -18.52 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.39 | 0.53 | 13.18 | -2.23 | -0.06 | -10.37 | -0.06 | 0 | 0.12 | 21.33 | -0.23 | 28.77 | -0.41 | 28.48 | -0.35 | 28.77 | -7.44 | 0 | -32.58 | -0.18 | -69.49 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.06
124.896 | -16.17 | 0 | -25.78 | 0 | 0 | -0.18 | 0 | 0 | 21.39 | 0.41 | 18.75 | -1.64 | 0 | -8.56 | -0.06 | 0 | 0.29 | 21.39 | -0.35 | 28.42 | -0.18 | 28.83 | -0.29 | 28.89 | -5.22 | 0.94 | -30.53 | -0.06 | -48.34 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.18
124.928 | -15 | -0.06 | -61.99 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.62 | -0.59 | -0.18 | -6.68 | 0 | -2.7 | 0.23 | 21.33 | -0.29 | 28.48 | -0.29 | 28.42 | -0.23 | 28.71 | -3.16 | 10.84 | -29.06 | -0.06 | -46.53 | -0.18 | 0 | 0 | 0 | 0.18 | 0 | 0.12
124.96 | -13.54 | -0.06 | -41.37 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.47 | 21.39 | -0.35 | -0.06 | -4.51 | 0 | -71.02 | 0.23 | 21.21 | -0.35 | 28.54 | -0.18 | 28.3 | -0.29 | 28.48 | -1.7 | 20.86 | -27.19 | -0.06 | -39.44 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
124.992 | -12.01 | -0.06 | -39.79 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.23 | 21.39 | 0.12 | 3.98 | -2.99 | -0.18 | -47.35 | 0.18 | 21.39 | -0.35 | 28.3 | -0.23 | 28.48 | -0.18 | 28.3 | -0.94 | 28.65 | -24.96 | -0.06 | -37.21 | -0.06 | 0 | 0 | 0 | 0.06 | 0 | 0.06
125.024 | -10.37 | -0.06 | -38.15 | -0.06 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.29 | 21.39 | 0.18 | 10.14 | -2.4 | -0.06 | -30.06 | 0.18 | 13.95 | -0.23 | 28.89 | -0.29 | 28.36 | -0.29 | 28.42 | -0.76 | 28.65 | -23.32 | -0.06 | -37.91 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
125.056 | -8.56 | -0.06 | -35.16 | -0.06 | -35.16 | -0.18 | 0 | -0.06 | 21.45 | 0.29 | 21.39 | 0.18 | 15.53 | -1.7 | -0.06 | -29.06 | 0.18 | 0.06 | -0.29 | 28.48 | -0.29 | 28.42 | -0.18 | 28.65 | -0.06 | 28.65 | -21.74 | -0.18 | -38.32 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0
125.088 | -7.21 | -0.06 | -33.81 | -0.06 | -59.77 | -0.06 | 0.06 | 0 | 21.39 | 0.23 | 21.33 | 0.41 | 21.15 | -0.76 | 0 | -34.69 | 0.18 | 0 | -0.29 | 28.48 | -0.18 | 28.42 | -0.23 | 28.77 | -0.06 | 28.59 | -19.92 | -0.18 | -36.45 | 0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18
125.12 | -5.63 | -0.06 | -31.23 | -0.06 | -40.37 | -0.18 | -70.73 | -0.18 | 21.39 | 0.06 | 21.39 | 0.41 | 21.45 | -0.59 | -0.06 | -34.69 | 0.18 | 0 | -0.18 | 28.48 | -0.23 | 28.48 | -0.18 | 28.59 | 0.12 | 28.71 | -16.99 | 0 | -33.99 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06
125.152 | -4.16 | -0.06 | -29.24 | -0.06 | -40.02 | -0.06 | -60.12 | -0.06 | 21.39 | 0.12 | 21.33 | 0.47 | 21.33 | -0.29 | 0 | -31.41 | 0.18 | 0 | -0.18 | 28.59 | -0.18 | 28.42 | -0.12 | 28.65 | 0.29 | 28.54 | -15 | -0.06 | -32.17 | -0.06 | -77.05 | -0.18 | 0 | 0 | 0 | 0
125.184 | -3.16 | -0.06 | -27.19 | -0.06 | -37.97 | -0.06 | -43.13 | -0.06 | 21.45 | 0.18 | 21.39 | 0.41 | 21.27 | -0.06 | -0.06 | -28.95 | 0.18 | 0 | -0.18 | 28.36 | -0.18 | 28.48 | -0.06 | 28.42 | 0.18 | 28.48 | -13.36 | -0.06 | -30.7 | -0.06 | -73.25 | -0.18 | 0 | 0 | -57.13 | 0
125.216 | -1.76 | -0.06 | -25.67 | 0 | -35.16 | -0.06 | -39.32 | -0.06 | 21.27 | 0.06 | 21.39 | 0.47 | 21.39 | 0 | -0.06 | -27.6 | 0.06 | 0 | -0.18 | 28.59 | -0.18 | 28.54 | -0.12 | 28.65 | 0.35 | 28.36 | -10.78 | -0.18 | -28.83 | 0 | -41.84 | -0.18 | -48.05 | -0.06 | -78.58 | 0
125.248 | -0.94 | -0.06 | -23.15 | -0.06 | -34.22 | -0.06 | -33.87 | 0 | 21.27 | 0.12 | 21.33 | 0.47 | 21.33 | 0.06 | 2.75 | -26.49 | 0 | 0 | -0.23 | 28.59 | -0.18 | 0.18 | -0.23 | 28.54 | 0.29 | 28.48 | -8.85 | -0.06 | -27.01 | -0.18 | -28.77 | -0.06 | -79.92 | -0.06 | -47.52 | 0.12
125.28 | -0.29 | 6.97 | -20.16 | 0 | -30.53 | -0.18 | -32.58 | -0.06 | 21.39 | 0.06 | 21.39 | 0.47 | 21.39 | 0.12 | 10.66 | -24.84 | 0.06 | 0 | -0.06 | 22.74 | -0.06 | 0 | -0.06 | 28.65 | 0.29 | 28.59 | -6.15 | -0.06 | -24.43 | -0.06 | -39.79 | -0.06 | -41.9 | 0 | -29.53 | -0.06
125.312 | -0.18 | 12.89 | -18.4 | 0.06 | -28.95 | -0.06 | -32.7 | -0.06 | 0.53 | 0 | 21.27 | 0.41 | 21.39 | 0.12 | 16.06 | -22.33 | 0.12 | 0 | -0.18 | 0.18 | -0.18 | 0 | -0.18 | 28.65 | 0.35 | 28.54 | -3.98 | 0 | -23.03 | -0.06 | -42.48 | -0.06 | -28.42 | 0 | -37.74 | -0.18
125.344 | -0.18 | 18.34 | -17.05 | -0.06 | -26.84 | -0.06 | -32.34 | -0.06 | 0.06 | 0 | 21.39 | 0.35 | 21.39 | 0.12 | 21.39 | -20.74 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0 | 28.59 | 0.18 | 28.59 | -2.4 | -0.18 | -21.09 | 0 | -38.38 | -0.18 | -36.8 | 0.06 | -43.13 | -0.18
125.376 | -0.23 | 21.8 | -15.06 | -0.06 | -25.08 | -0.18 | -30.7 | -0.06 | 0 | 0 | 21.21 | 0.41 | 21.39 | 0.12 | 21.45 | -19.63 | 0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.18 | 20.63 | 0.29 | 28.42 | -1.29 | 5.68 | -19.51 | -0.06 | -34.63 | -0.06 | -42.72 | -0.06 | -40.72 | -0.18
125.408 | -0.35 | 21.45 | -13.54 | -0.06 | -23.03 | -0.06 | -28.36 | -0.06 | 0 | 0.06 | 21.15 | 0.23 | 21.15 | 0.18 | 21.39 | -17.75 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0.12 | 0.35 | 28.36 | -1.11 | 13.71 | -17.46 | 0 | -32.93 | -0.06 | -40.55 | -0.06 | -35.74 | -0.06
125.44 | -0.35 | 21.39 | -11.78 | -0.06 | -20.1 | -0.06 | -25.2 | -0.06 | 0 | 0 | 21.33 | 0.29 | 21.39 | 0.18 | 21.33 | -15.47 | 0 | -0.53 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.35 | 28.42 | -0.18 | 23.91 | -14.82 | 0 | -32.93 | -0.18 | -35.16 | 0 | -32.93 | -0.06
125.472 | -0.35 | 21.45 | -10.72 | 0 | -18.34 | -0.18 | -23.15 | -0.06 | 0 | 0.06 | 21.45 | 0.18 | 21.33 | 0.18 | 21.33 | -13.54 | -0.06 | -63.69 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.54 | -0.23 | 28.83 | -12.89 | 0 | -31.52 | -0.18 | -33.34 | 0 | -33.28 | -0.06
125.504 | -0.35 | 21.45 | -8.91 | 0 | -17.17 | -0.06 | -21.62 | -0.06 | 0 | -0.06 | 21.15 | 0.18 | 21.33 | 0.23 | 21.39 | -12.07 | -0.06 | -61.12 | -0.06 | 0 | -0.18 | 0 | -0.18 | 0 | 0.12 | 28.59 | 0.23 | 28.65 | -10.84 | 0 | -30.06 | -0.18 | -32.81 | 0 | -32.05 | -0.18
125.536 | -0.41 | 21.45 | -7.91 | 0.06 | -14.82 | -0.06 | -19.63 | -0.06 | 0 | -0.06 | 21.45 | 0.18 | 21.39 | 0.18 | 21.39 | -10.2 | 0 | -37.62 | -0.18 | 0 | 0 | 0 | -0.18 | 0 | 0.23 | 28.59 | 0.12 | 28.59 | -8.85 | 0 | -28.48 | -0.18 | -31.93 | 0 | -30.12 | -0.06
125.568 | -0.18 | 21.45 | -6.33 | 0.06 | -13.36 | -0.06 | -17.4 | -0.06 | 0 | 0 | 21.21 | 0.12 | 21.27 | 0.23 | 21.39 | -8.91 | 0 | -28.24 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.59 | 0.23 | 28.65 | -6.97 | -0.06 | -26.43 | -0.18 | -30.24 | -0.06 | -27.77 | -0.18
125.6 | -0.18 | 21.45 | -3.93 | 0 | -11.31 | -0.06 | -15.18 | -0.06 | 0 | -0.18 | 21.15 | 0.06 | 21.39 | 0.29 | 21.21 | -6.56 | 0 | -32.99 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0.12 | 28.59 | 0.41 | 28.83 | -4.1 | 0 | -24.9 | -0.18 | -27.77 | 0 | -26.84 | -0.18
125.632 | -0.23 | 21.39 | -2.99 | 0 | -10.37 | -0.06 | -13.07 | -0.06 | 0 | -0.06 | 21.33 | 0.06 | 21.39 | 0.12 | 21.21 | -4.45 | 0 | -35.57 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.12 | 28.65 | 0.47 | 28.42 | -2.4 | 0 | -23.61 | -0.18 | -26.66 | 0.06 | -25.43 | -0.06
125.664 | -0.23 | 21.45 | -1.93 | 0.53 | -8.67 | -0.06 | -11.31 | -0.06 | 0 | 0 | 21.39 | 0 | 21.39 | 0.18 | 21.27 | -3.4 | 0 | -33.63 | -0.18 | 0 | -0.06 | -37.97 | -0.18 | 0 | 0 | 28.65 | 0.29 | 28.48 | -1.46 | 0 | -21.86 | 0 | -25.78 | 0 | -23.67 | -0.18
125.696 | -0.18 | 21.45 | -1.17 | 8.2 | -7.27 | 0.41 | -9.08 | 0 | 0 | 0 | 19.28 | -0.06 | 21.39 | 0.12 | 21.39 | -2.4 | -0.06 | -30 | -0.18 | 0.06 | -0.06 | -70.43 | -0.06 | 0 | 0.06 | 28.59 | 0.35 | 28.42 | -1.11 | -0.18 | -19.98 | -0.18 | -24.14 | 0 | -22.21 | -0.18
125.728 | -0.23 | 21.5 | -0.53 | 13.83 | -5.92 | 8.44 | -6.62 | -0.06 | 0 | -0.18 | 0.12 | -0.06 | 21.56 | 0.12 | 21.27 | -1.64 | 0.06 | -27.6 | -0.06 | -42.07 | -0.06 | -44.77 | -0.06 | 0 | -0.06 | 28.59 | 0.35 | 28.48 | -0.18 | -0.06 | -18.87 | -0.18 | -22.15 | -0.18 | -20.57 | 0
125.76 | -0.18 | 21.45 | -0.18 | 21.45 | -3.98 | 15.59 | -3.93 | -0.06 | -76.7 | -0.06 | 0 | -0.06 | 21.5 | 0.23 | 21.45 | -0.94 | 0 | -26.6 | -0.06 | -61.94 | 0 | -45.7 | -0.18 | 0 | 0.06 | 28.59 | 0.23 | 28.54 | -0.18 | -0.06 | -16.99 | -0.06 | -19.98 | -0.18 | -18.52 | 0
125.792 | -0.23 | 21.5 | -0.18 | 21.45 | -2.64 | 21.5 | -2.46 | 0.12 | -54.32 | -0.06 | 0 | -0.18 | 21.27 | 0.12 | 21.33 | -0.47 | 2.11 | -25.43 | -0.06 | -45.29 | 0 | -43.89 | -0.06 | -0.18 | 0 | 0.29 | 0.35 | 28.48 | 0.18 | -0.06 | -14.71 | -0.06 | -18.75 | 0 | -17.4 | -0.18
125.824 | -0.23 | 21.39 | -0.23 | 21.45 | -1.7 | 21.5 | -1.46 | -0.06 | -42.89 | -0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.33 | -0.35 | 7.62 | -23.32 | -0.06 | -45.06 | 0 | -41.49 | -0.06 | -72.42 | 0 | 0.12 | 0.18 | 28.36 | 0.23 | 5.16 | -13.18 | -0.18 | -17.23 | -0.06 | -15.7 | -0.18
125.856 | -0.18 | 21.45 | -0.23 | 21.39 | -1.11 | 21.45 | -1 | 0 | -37.85 | 0 | 0 | -0.06 | 21.45 | 0.12 | 21.33 | -0.18 | 13.24 | -21.74 | -0.06 | -43.71 | 0 | -39.08 | 0 | -62.64 | 0.06 | 0 | 0.18 | 28.36 | 0.23 | 13.24 | -11.31 | 3.69 | -15.29 | -0.06 | -14 | -0.18
125.888 | -0.18 | 21.39 | -0.35 | 21.45 | -0.64 | 21.45 | -0.35 | -0.06 | -33.52 | -0.06 | 0 | -0.06 | 0.23 | 0 | 21.39 | -0.06 | 18.63 | -20.39 | -0.18 | -40.61 | -0.06 | -37.68 | 0 | -45.76 | -0.06 | 0.06 | 0.12 | 28.59 | 0.41 | 20.92 | -9.79 | 11.84 | -13.77 | 4.22 | -12.13 | -0.18
125.92 | 0 | 21.45 | -0.35 | 21.45 | -0.23 | 21.45 | 0 | -0.06 | -32.17 | -0.18 | 0 | -0.18 | 0 | 0 | 21.21 | -0.06 | 21.39 | -18.28 | 0 | -39.44 | -0.06 | -35.92 | 0 | -46 | -0.06 | 0 | 0.12 | 28.65 | 0.29 | 28.89 | -7.97 | 22.33 | -11.6 | 15 | -10.08 | 7.85
125.952 | -0.06 | 21.39 | -0.23 | 21.45 | -0.18 | 21.5 | 0.23 | -0.06 | -32.76 | 0 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.06 | 21.45 | -16.17 | -0.06 | -36.86 | -0.06 | -34.4 | -0.06 | -42.78 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.48 | -5.92 | 29.01 | -10.02 | 22.79 | -8.56 | 15.12
125.984 | -0.06 | 21.39 | -0.18 | 21.45 | -0.23 | 21.5 | 0.29 | 0.06 | -31.58 | 0 | 0 | -0.06 | 0 | 0 | 21.39 | 0.06 | 21.39 | -14.24 | -0.06 | -35.57 | -0.06 | -32.4 | 0 | -41.66 | -0.06 | 0 | 0.06 | 28.42 | 0.29 | 28.42 | -4.39 | 28.18 | -8.32 | 29.06 | -6.86 | 22.85
126.016 | -0.18 | 21.45 | -0.23 | 21.45 | -0.29 | 21.45 | 0.35 | -0.06 | -29.77 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.39 | -13.01 | -0.06 | -33.52 | 0 | -30.7 | -0.06 | -39.44 | 0 | 0 | 0.12 | 28.42 | 0.29 | 28.42 | -3.4 | 28.65 | -7.03 | 28.54 | -4.92 | 29.01
126.048 | -0.06 | 21.45 | -0.29 | 21.5 | -0.35 | 21.45 | 0.47 | -0.06 | -27.31 | -0.06 | 0 | 0 | 0 | 0 | 21.33 | 0.12 | 21.39 | -11.31 | -0.06 | -31.82 | 0 | -28.07 | -0.06 | -37.38 | 0 | 0 | 0 | 28.54 | 0.18 | 28.42 | -2.23 | 28.48 | -5.57 | 28.42 | -3.75 | 28.36
126.08 | -0.18 | 21.45 | -0.23 | 21.39 | -0.41 | 21.45 | 0.53 | 3.34 | -24.9 | -0.06 | 0 | 0 | 0 | 0 | 21.33 | 0.18 | 21.09 | -9.32 | -0.06 | -29.36 | 0 | -26.13 | -0.18 | -35.57 | -0.06 | 0 | 0 | 28.95 | 0.18 | 28.59 | -1.29 | 28.65 | -3.34 | 28.59 | -2.29 | 28.48
126.112 | -0.18 | 21.45 | -0.23 | 21.5 | -0.35 | 21.39 | 0.41 | 9.49 | -23.03 | 0 | -0.06 | -0.06 | 0 | 0 | 21.39 | 0.12 | 21.33 | -7.03 | 0 | -27.42 | 0 | -23.91 | -0.06 | -33.52 | 0 | 0 | 0 | 28.65 | 0.18 | 28.42 | -0.59 | 28.18 | -2.46 | 28.42 | -1.46 | 28.48
126.144 | -0.18 | 21.45 | -0.41 | 21.45 | -0.29 | 21.45 | 0.41 | 14.82 | -20.98 | -0.06 | -56.31 | -0.18 | 0 | 0 | 0.23 | 0.23 | 21.21 | -5.39 | 2.29 | -26.02 | 0 | -21.97 | 0 | -31.47 | -0.18 | 0 | 0 | 28.65 | 0.12 | 28.59 | -0.59 | 28.59 | -1.41 | 28.42 | -0.76 | 28.54
126.176 | -0.18 | 0.12 | -0.29 | 21.45 | -0.23 | 21.45 | 0.35 | 20.45 | -19.04 | 0 | -61.17 | -0.06 | 0 | 0.06 | 0.06 | 0.23 | 21.27 | -4.04 | 9.96 | -23.5 | 0 | -19.98 | 0 | -29.71 | -0.06 | 0 | -0.06 | 28.65 | 0.12 | 28.42 | -0.23 | 28.36 | -0.94 | 28.59 | -0.53 | 28.54
126.208 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.53 | 21.56 | -17.4 | 0 | -44.83 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.39 | -2.99 | 17.11 | -21.86 | 0 | -17.4 | 0 | -27.31 | -0.18 | -0.06 | 0 | 28.42 | 0.12 | 28.59 | -0.06 | 28.48 | -0.64 | 28.59 | -0.35 | 28.65
126.24 | 0 | 0 | -0.29 | 21.45 | -0.23 | 21.39 | 0.41 | 21.39 | -14.82 | 0.06 | -37.62 | -0.06 | 0 | 0.06 | 0.06 | 0.18 | 21.21 | -1.76 | 27.54 | -18.87 | -0.06 | -15.7 | 0 | -25.25 | -0.06 | -93.34 | 0 | 28.65 | -0.06 | 28.48 | 0 | 28.42 | -0.53 | 28.59 | -0.29 | 28.59
126.272 | -0.18 | 0 | -0.29 | 21.5 | -0.18 | 21.39 | 0.41 | 21.39 | -12.72 | 0 | -32.7 | -0.06 | 0 | 0 | 0 | 0.23 | 21.33 | -1.05 | 28.59 | -17.23 | -0.06 | -14.53 | 0 | -23.2 | 0 | -57.25 | 0 | 0.23 | -0.06 | 28.59 | 0 | 28.48 | -0.18 | 28.48 | 0 | 28.48
126.304 | -0.06 | 0 | -0.29 | 21.45 | -0.29 | 21.45 | 0.47 | 21.39 | -10.61 | -0.06 | -32.17 | 0 | -24.79 | 0 | 0 | 0.18 | 21.45 | -0.7 | 28.59 | -15.35 | 0 | -13.07 | 0 | -21.27 | -0.06 | -48.52 | 0 | 0 | -0.06 | 28.13 | 0.18 | 28.71 | 0 | 28.3 | 0 | 28.42
126.336 | -0.06 | 0 | -0.29 | 21.45 | -0.29 | 21.5 | 0.35 | 21.39 | -8.44 | 0.06 | -32.7 | -0.18 | -74.07 | 0 | 0 | 0.29 | 21.39 | -0.23 | 28.65 | -14.24 | 8.2 | -11.6 | -0.06 | -19.45 | -0.06 | -44.42 | 0 | 0 | 0 | 28.59 | 0.12 | 28.42 | -0.06 | 28.48 | 0 | 28.65
126.368 | -0.06 | 0 | -0.18 | 21.45 | -0.35 | 21.5 | 0.29 | 21.39 | -6.39 | -0.06 | -32.34 | -0.06 | -49.46 | -0.06 | 0 | 0.12 | 21.39 | -0.06 | 28.54 | -12.83 | 15.29 | -9.67 | 3.4 | -17.64 | -0.06 | -39.38 | -0.06 | 0.06 | 0 | 28.54 | 0.18 | 28.48 | 0.12 | 28.59 | 0.06 | 28.36
126.4 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | 0.23 | 21.21 | -3.63 | 0.06 | -30.29 | 0 | -27.95 | -0.06 | 0 | 0.12 | 21.33 | -0.06 | 28.54 | -10.96 | 25.9 | -8.44 | 14 | -15.18 | -0.18 | -36.97 | -0.06 | 0 | 0 | 28.48 | 0.18 | 28.48 | 0.12 | 28.36 | 0.06 | 28.65
126.432 | -0.18 | 0 | -0.23 | 21.45 | -0.23 | 21.45 | 0.23 | 21.39 | -1.93 | 0.06 | -27.77 | -0.06 | -30.82 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 28.48 | -9.67 | 28.77 | -7.21 | 22.03 | -14.24 | 0 | -38.15 | -0.06 | 0 | -0.06 | 28.59 | 0.12 | 28.18 | 0.12 | 28.42 | 0.18 | 28.36
126.464 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.06 | 21.39 | -1.7 | 1.99 | -25.72 | -0.06 | -35.68 | -0.06 | 0 | 0.18 | 21.39 | 0.06 | 28.3 | -8.38 | 28.54 | -5.98 | 29.01 | -12.36 | -0.06 | -38.09 | 0 | 0 | 0 | 28.59 | 0.18 | 28.42 | 0.12 | 28.48 | 0.18 | 28.71
126.496 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.12 | 21.45 | -0.41 | 7.03 | -23.5 | -0.06 | -34.69 | -0.06 | 0 | 0.18 | 21.33 | 0 | 28.71 | -7.27 | 28.65 | -4.69 | 28.48 | -10.96 | -0.06 | -36.27 | -0.06 | 0 | -0.06 | 28.48 | 0.18 | 28.36 | 0.06 | 28.54 | 0.23 | 28.77
126.528 | -0.18 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.12 | 21.27 | -0.35 | 12.72 | -21.68 | 0 | -31.47 | -0.06 | 0 | 0.12 | 21.39 | 0.12 | 28.89 | -6.15 | 28.42 | -3.46 | 28.77 | -9.61 | -0.06 | -34.4 | 0 | 0 | -0.06 | 28.42 | 0.12 | 28.24 | 0.06 | 28.48 | 0.29 | 28.48
126.56 | -0.06 | -0.06 | -0.06 | 0.23 | -0.18 | 21.45 | 0 | 21.39 | -0.06 | 20.16 | -19.63 | -0.06 | -28.36 | 0 | -24.32 | 0.18 | 21.39 | 0.18 | 28.65 | -4.69 | 28.48 | -2.93 | 28.59 | -8.03 | 0 | -31.93 | 0 | 0 | -0.06 | 28.42 | 0.12 | 28.83 | 0.06 | 28.18 | 0.18 | 28.65
126.592 | -0.06 | -35.8 | -0.06 | 0.06 | -0.18 | 0.23 | 0.06 | 21.39 | 0.23 | 21.39 | -17.4 | -0.06 | -27.83 | -0.06 | -70.02 | 0.23 | 21.39 | 0.12 | 28.59 | -3.75 | 28.65 | -2.17 | 28.54 | -6.21 | -0.06 | -30.18 | -0.06 | 0 | 0 | 28.42 | 0.06 | 28.48 | 0.18 | 28.48 | 0.18 | 28.42
126.624 | 0 | -56.78 | -0.06 | 0 | -0.18 | 0.06 | 0 | 21.39 | 0.29 | 21.39 | -15.06 | 0 | -26.78 | -0.06 | -49.98 | 0.12 | 21.21 | 0.06 | 28.54 | -2.99 | 28.42 | -1.82 | 28.42 | -5.22 | 0 | -28.48 | -0.06 | 0 | -0.06 | 28.42 | 0.06 | 28.59 | 0.06 | 28.48 | 0.23 | 28.48
126.656 | -0.06 | -43.36 | -0.18 | 0.06 | -0.18 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -13.54 | 0.06 | -24.14 | -0.06 | -30.18 | 0.23 | 21.21 | 0.18 | 28.77 | -2.52 | 28.77 | -1.29 | 28.54 | -4.45 | 0 | -26.6 | 0 | 0 | -0.06 | 28.59 | 0 | 28.48 | 0.12 | 28.59 | 0.18 | 28.42
126.688 | -0.06 | -39.85 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.47 | 21.39 | -11.13 | -0.06 | -22.33 | -0.06 | -29.01 | 0.12 | 19.57 | 0.23 | 28.65 | -1.82 | 28.42 | -0.82 | 28.3 | -3.63 | 0 | -24.9 | -0.06 | -36.1 | -0.06 | 28.42 | -0.06 | 28.48 | 0.06 | 28.54 | 0.06 | 28.48
126.72 | -0.06 | -36.74 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.35 | 21.33 | -8.85 | 0 | -20.16 | -0.06 | -34.92 | 0.18 | 0.06 | 0.18 | 28.59 | -1.29 | 28.65 | -0.59 | 28.18 | -2.58 | 0 | -22.56 | 0 | -74.53 | -0.18 | 0.23 | 0 | 28.59 | 0.06 | 28.42 | 0.12 | 28.48
126.752 | 0 | -35.92 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.35 | 21.39 | -6.68 | -0.18 | -18.87 | -0.06 | -34.81 | 0.18 | 0 | 0.18 | 29.06 | -1.05 | 28.48 | -0.47 | 28.42 | -2.11 | -0.06 | -20.51 | -0.06 | -50.16 | -0.18 | 0 | 0 | 0.18 | 0.12 | 28.48 | 0.06 | 28.59
126.784 | 0 | -33.17 | -0.18 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.33 | -4.16 | -0.06 | -17.05 | 0.06 | -32.17 | 0.18 | 0 | 0.18 | 28.48 | -0.94 | 28.54 | -0.23 | 28.18 | -1.64 | 0 | -18.63 | -0.06 | -47.93 | -0.06 | 0 | 0 | 0.06 | 0.12 | 0.29 | 0.06 | 28.59
126.816 | -0.06 | -31.29 | -0.18 | 0 | -0.18 | 0 | -0.18 | 21.39 | 0.35 | 21.39 | -2.4 | -0.06 | -15.47 | -0.18 | -28.48 | 0.12 | 0 | 0.12 | 28.59 | -0.7 | 28.48 | -0.29 | 28.59 | -0.82 | 2.81 | -16.47 | 0.06 | -41.25 | -0.18 | 0 | 0 | 0 | 0.12 | 0.06 | 0.06 | 0.23
126.848 | 0 | -29.3 | -0.06 | 0 | -0.06 | 0 | 0 | 21.56 | 0.41 | 21.39 | -1.7 | -0.06 | -13.77 | -0.18 | -27.07 | 0.12 | 0 | 0.18 | 28.54 | -0.41 | 29.06 | -0.29 | 28.48 | -0.82 | 11.78 | -14.82 | 0 | -36.74 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.06
126.88 | -0.06 | -26.54 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.62 | 0.41 | 21.39 | -0.59 | 8.03 | -11.48 | -0.18 | -26.84 | 0.18 | 0 | 0.12 | 28.54 | -0.47 | 28.65 | -0.23 | 28.48 | -0.41 | 22.09 | -12.31 | -0.06 | -37.21 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0
126.912 | -0.06 | -24.61 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.39 | -0.41 | 13.54 | -9.67 | -0.18 | -24.2 | -0.06 | 0 | 0.29 | 28.48 | -0.41 | 28.48 | -0.18 | 28.48 | -0.35 | 29.12 | -10.25 | -0.06 | -37.62 | -0.18 | 0 | 0 | 0 | 0 | 0 | 0.12 | 0
126.944 | -0.06 | -22.85 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.5 | -0.06 | 19.04 | -8.14 | 1.17 | -21.86 | 0 | 0 | 0.18 | 28.59 | -0.23 | 28.83 | -0.23 | 28.48 | -0.18 | 28.65 | -8.2 | 0 | -37.15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
126.976 | -0.06 | -20.92 | -0.06 | -0.06 | -0.06 | 0 | 0.06 | 0.47 | 0.29 | 21.27 | 0.29 | 21.45 | -6.56 | 8.97 | -20.39 | -0.06 | 0 | 0.12 | 28.48 | -0.35 | 28.3 | -0.29 | 28.48 | -0.18 | 28.36 | -5.68 | -0.06 | -34.75 | -0.06 | 0 | -0.06 | 0 | 0.12 | 0 | 0 | 0
127.008 | 0 | -18.69 | 0 | -55.9 | -0.18 | -0.29 | 0 | 0.06 | 0.35 | 21.39 | 0.29 | 21.39 | -4.69 | 14.18 | -19.45 | 0.06 | 0 | 0.18 | 28.48 | -0.18 | 28.54 | -0.23 | 28.65 | -0.18 | 28.48 | -3.98 | 0 | -33.52 | 0 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0
127.04 | 0 | -16.17 | -0.06 | -46.47 | -0.06 | -64.46 | 0 | 0 | 0.06 | 21.39 | 0.47 | 21.39 | -2.93 | 21.5 | -17.23 | 0 | 0 | 0.18 | 0.64 | -0.29 | 28.59 | -0.18 | 28.42 | -0.29 | 28.59 | -1.93 | -0.18 | -31.11 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
127.072 | -0.06 | -14.82 | -0.18 | -39.67 | -0.18 | -48.4 | -0.06 | 0 | 0.12 | 21.45 | 0.53 | 21.39 | -1.99 | 21.39 | -15.41 | 0 | 0 | 0.12 | 0.12 | -0.35 | 28.65 | -0.23 | 28.65 | -0.29 | 28.42 | -1.17 | -0.06 | -28.77 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
127.104 | 0 | -13.59 | -0.06 | -39.14 | 0 | -40.9 | 0 | 0 | 0.12 | 21.39 | 0.53 | 21.39 | -1.05 | 21.45 | -13.54 | 0 | -0.06 | 0.18 | 0 | -0.23 | 28.59 | -0.18 | 28.42 | -0.29 | 28.42 | -0.64 | 0 | -27.01 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
127.136 | 5.92 | -11.95 | -0.06 | -37.44 | -0.06 | -38.97 | 0 | 0 | 0.06 | 21.33 | 0.53 | 21.27 | -0.94 | 21.33 | -11.89 | -0.06 | -57.48 | 0.18 | 0 | -0.18 | 28.65 | -0.18 | 28.42 | -0.29 | 28.65 | -0.23 | -0.06 | -25.61 | -0.18 | -24.55 | -0.06 | 0 | 0 | 0 | -0.18 | 0
127.168 | 11.78 | -10.37 | -0.06 | -34.69 | -0.06 | -37.09 | -0.06 | 0 | 0.23 | 21.45 | 0.47 | 21.33 | -0.35 | 21.39 | -10.37 | -0.06 | -61.94 | 0.23 | 0 | -0.18 | 28.48 | -0.18 | 28.42 | -0.18 | 28.48 | -0.18 | -0.06 | -24.2 | 0 | -90.41 | -0.06 | -15.82 | -0.06 | 0 | -0.18 | 0
127.2 | 17.17 | -8.44 | 0 | -32.4 | -0.06 | -34.4 | 0.12 | 0 | 0 | 21.5 | 0.53 | 21.33 | -0.18 | 21.39 | -8.09 | 0.06 | -33.52 | 0.12 | 0 | -0.23 | 20.74 | -0.18 | 28.42 | -0.23 | 28.48 | 0.18 | 0 | -21.74 | -0.06 | -51.45 | -0.06 | -81.21 | 0 | -44.83 | 0 | 0
127.232 | 21.8 | -7.03 | -0.06 | -30.7 | -0.06 | -32.34 | -0.18 | 0 | -0.06 | 21.45 | 0.47 | 21.21 | -0.06 | 21.27 | -6.15 | 0 | -27.95 | 0.23 | 0 | -0.18 | 0.18 | -0.18 | 28.36 | -0.18 | 28.42 | 0.23 | 8.97 | -19.75 | -0.06 | -48.69 | -0.18 | -53.85 | 0 | -78.75 | -0.06 | -52.68
127.264 | 21.5 | -5.45 | -0.06 | -27.89 | -0.06 | -30.24 | -0.18 | 0 | 0.06 | 21.39 | 0.47 | 21.15 | -0.06 | 21.39 | -4.51 | 0.06 | -33.52 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0.23 | -0.18 | 28.59 | 0.29 | 16.17 | -17.4 | -0.18 | -41.9 | -0.06 | -30.7 | -0.06 | -52.44 | -0.06 | -78.58
127.296 | 21.45 | -4.16 | -0.06 | -26.31 | -0.06 | -28.48 | -0.06 | 0 | 0 | 21.56 | 0.47 | 21.33 | 0.06 | 21.39 | -3.16 | -0.06 | -36.04 | 0.18 | 0 | -0.18 | 0 | -0.06 | 0.06 | -0.18 | 28.59 | 0.41 | 24.02 | -15.7 | -0.18 | -36.97 | -0.18 | -32.52 | -0.06 | -30.94 | -0.06 | -49.1
127.328 | 21.45 | -2.81 | -0.06 | -24.08 | -0.06 | -26.13 | -0.06 | 0 | -0.06 | 21.39 | 0.35 | 21.27 | 0 | 21.27 | -2.23 | 0.06 | -33.28 | 0.18 | 0 | -0.06 | 0.06 | -0.18 | 0 | -0.18 | 28.89 | 0.35 | 28.83 | -14.12 | -0.06 | -37.74 | -0.06 | -42.78 | 0 | -32.64 | -0.06 | -31.06
127.36 | 21.39 | -1.76 | -0.06 | -21.33 | -0.06 | -23.61 | -0.06 | 0 | 0.06 | 6.91 | 0.35 | 21.33 | 0.06 | 21.39 | -1.41 | 0 | -29.24 | 0.23 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 28.54 | 0.18 | 28.48 | -11.54 | -0.18 | -38.44 | -0.18 | -40.14 | -0.06 | -42.83 | 0 | -36.62
127.392 | 21.45 | -0.94 | -0.06 | -19.4 | -0.06 | -21.91 | -0.06 | -6.91 | -0.06 | 0.12 | 0.18 | 21.33 | 0.12 | 21.39 | -0.7 | -0.06 | -27.36 | 0.18 | 0 | -0.18 | 0 | -0.18 | 0 | -0.18 | 28.83 | 0.35 | 28.54 | -9.43 | -0.06 | -37.56 | -0.06 | -36.21 | -0.06 | -39.79 | 0 | -42.6
127.424 | 21.45 | -0.29 | -0.06 | -17.29 | -0.06 | -19.28 | 0 | -77.23 | -0.18 | 0 | 0.12 | 21.33 | 0.23 | 21.33 | -0.35 | -0.06 | -26.66 | 0.18 | 0 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.65 | 0.41 | 28.71 | -7.44 | 0 | -35.39 | -0.18 | -32.99 | -0.06 | -36.1 | 0 | -40.72
127.456 | 21.45 | -0.18 | 0 | -16.11 | -0.18 | -18.16 | -0.06 | -54.67 | -0.06 | 0 | 0 | 21.39 | 0.23 | 21.15 | -0.23 | 0 | -25.55 | 0.18 | -0.06 | -0.18 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.54 | -4.92 | 0 | -33.87 | 0 | -32.7 | 0 | -33.22 | 0 | -36.21
127.488 | 21.45 | -0.18 | -0.06 | -14.65 | -0.18 | -15.88 | -0.06 | -42.78 | -0.06 | 0 | 0.06 | 21.39 | 0.18 | 21.39 | -0.18 | -0.06 | -23.38 | 0.12 | -56.31 | -0.06 | 0 | -0.18 | 0 | -0.18 | 28.59 | 0.23 | 28.48 | -3.16 | -0.06 | -31.58 | -0.06 | -32.81 | -0.06 | -32.64 | -0.06 | -33.28
127.52 | 21.5 | -0.29 | -0.06 | -12.6 | 0 | -14.12 | -0.06 | -36.33 | -0.18 | 0 | 0.12 | 21.39 | 0.18 | 21.33 | -0.06 | -0.06 | -20.98 | 0.12 | -57.25 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.42 | -1.93 | 9.32 | -29.01 | -0.06 | -30.47 | -0.06 | -32.29 | 0 | -33.11
127.552 | 21.45 | -0.35 | -0.06 | -11.13 | -0.06 | -12.48 | 0 | -31.93 | -0.18 | 0 | 0.06 | 21.33 | 0.18 | 21.33 | 0.12 | 0 | -19.92 | 0.12 | -45.35 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.23 | 28.65 | -0.94 | 16.52 | -27.01 | -0.18 | -28.48 | -0.06 | -30.53 | -0.06 | -32.34
127.584 | 21.45 | -0.35 | 5.04 | -9.32 | -0.18 | -10.84 | 0 | -32.52 | 0 | 0 | 0 | 21.27 | 0.06 | 21.21 | 0.12 | -0.06 | -18.46 | 0.06 | -45.35 | -0.18 | 0 | 0 | 0 | -0.18 | 28.42 | 0.18 | 28.48 | -0.41 | 24.49 | -25.43 | 0 | -27.25 | -0.18 | -28.54 | 0 | -30.76
127.616 | 21.45 | -0.41 | 11.13 | -8.09 | -0.06 | -9.61 | -0.06 | -32.4 | -0.06 | 0 | 0 | 21.33 | 0.06 | 21.27 | 0.12 | 2.7 | -16.41 | 0.12 | -42.48 | -0.06 | 0 | -0.06 | 0 | -0.06 | 28.48 | 0.12 | 28.48 | -0.23 | 28.54 | -24.26 | -0.06 | -26.25 | -0.06 | -26.72 | -0.06 | -28.18
127.648 | 21.45 | -0.35 | 16.47 | -6.68 | -0.06 | -8.2 | 0 | -31.29 | -0.18 | 0 | 0 | 21.45 | 0.23 | 21.33 | 0.23 | 9.08 | -15.06 | 0.06 | -40.72 | 0 | -43.48 | -0.06 | 0 | -0.06 | 28.48 | 0.06 | 28.18 | 0 | 28.3 | -22.56 | -0.06 | -24.2 | -0.06 | -25.9 | -0.06 | -26.6
127.68 | 21.45 | -0.35 | 21.62 | -4.63 | -0.06 | -6.68 | 0.12 | -29.12 | -0.06 | 0 | 0 | 21.27 | 0.18 | 21.39 | 0.29 | 16.11 | -12.48 | 0.18 | -37.97 | -0.06 | -61.88 | -0.18 | -36.15 | -0.06 | 28.42 | 0.06 | 28.59 | 0.12 | 28.48 | -19.98 | -0.06 | -22.33 | -0.06 | -24.14 | 0 | -25.08
127.712 | 21.45 | -0.23 | 21.5 | -3.34 | -0.06 | -4.92 | 0.06 | -27.07 | -0.06 | 0 | -0.18 | 21.39 | 0.18 | 21.39 | 0.18 | 21.56 | -11.13 | 0.18 | -36.92 | -0.06 | -45.59 | -0.06 | -75.82 | -0.06 | 0.29 | 0.12 | 28.54 | 0.35 | 28.59 | -17.34 | -0.06 | -20.39 | -0.06 | -22.38 | -0.18 | -23.85
127.744 | 21.39 | -0.18 | 21.5 | -2.29 | 5.68 | -3.4 | -0.06 | -24.67 | -0.06 | 0 | -0.06 | 17.7 | 0.23 | 21.39 | 0.06 | 21.39 | -9.14 | 0.18 | -35.22 | 0 | -46 | -0.06 | -50.63 | 0 | 0.06 | 0 | 28.59 | 0.29 | 28.65 | -15.65 | -0.18 | -19.16 | -0.06 | -20.68 | 0 | -22.38
127.776 | 21.39 | -0.23 | 21.45 | -1.41 | 11.72 | -2.46 | 0.06 | -22.68 | -0.06 | 0 | -0.06 | 0.12 | 0.18 | 21.27 | 0.18 | 21.33 | -7.56 | 0.18 | -33.4 | 0 | -43.24 | 0 | -45.94 | -0.06 | 0 | 0 | 28.71 | 0.41 | 28.42 | -14.24 | -0.06 | -17.58 | -0.06 | -19.04 | -0.18 | -20.51
127.808 | 21.45 | -0.23 | 21.5 | -0.53 | 17.11 | -1.35 | 5.39 | -21.09 | 0 | -54.08 | 0 | 0 | 0.18 | 21.21 | 0.23 | 21.39 | -5.8 | 0.12 | -31.23 | -0.18 | -40.9 | 0 | -44.83 | -0.06 | 0 | 0 | 28.71 | 0.47 | 28.65 | -12.25 | 0 | -16.06 | -0.18 | -17.87 | -0.06 | -19.1
127.84 | 21.45 | -0.23 | 21.45 | -0.29 | 21.62 | -0.64 | 12.95 | -18.28 | 0 | -56.31 | 0 | 0 | 0.06 | 0.12 | 0.23 | 21.39 | -3.69 | 0.12 | -29.01 | -0.18 | -38.85 | -0.06 | -41.25 | 0 | 0 | 0 | 28.65 | 0.41 | 28.54 | -9.32 | 8.97 | -13.54 | -0.06 | -15.29 | -0.06 | -16.93
127.872 | 21.39 | -0.29 | 21.45 | -0.18 | 21.45 | -0.29 | 18.63 | -16.47 | -0.06 | -43.89 | 0 | 0 | 0 | 0.06 | 0.29 | 21.39 | -2.4 | 0 | -26.72 | 0 | -37.56 | -0.06 | -39.55 | 0.06 | 0 | 0 | 28.65 | 0.29 | 28.48 | -7.56 | 16 | -12.42 | -0.06 | -13.77 | 7.09 | -15.23
127.904 | 21.39 | -0.29 | 21.5 | -0.18 | 21.5 | -0.06 | 21.56 | -14.77 | -0.06 | -38.85 | 0 | 0 | 0.06 | 0 | 0.23 | 21.33 | -1 | 0.06 | -25.08 | -0.06 | -34.98 | -0.06 | -38.09 | 0 | 0 | 0 | 28.48 | 0.23 | 28.48 | -5.39 | 23.67 | -10.72 | 6.33 | -12.13 | 14.77 | -13.83
127.936 | 21.45 | -0.23 | 21.39 | -0.29 | 21.45 | -0.18 | 21.39 | -12.77 | -0.06 | -33.4 | -0.06 | 0 | 0 | 0 | 0.29 | 21.33 | -0.76 | 0.06 | -23.15 | -0.06 | -34.1 | -0.06 | -36.97 | -0.06 | 0 | 0 | 28.42 | 0.29 | 28.77 | -3.4 | 28.95 | -8.85 | 14.18 | -10.2 | 22.44 | -12.36
127.968 | 21.45 | -0.18 | 21.45 | -0.35 | 21.45 | -0.29 | 21.39 | -10.9 | -0.06 | -31.93 | -0.18 | 0 | 0 | 0 | 0.29 | 21.33 | -0.76 | 0 | -21.21 | -0.06 | -32.64 | -0.06 | -34.98 | 0 | 0 | -0.06 | 28.89 | 0.23 | 28.65 | -2.23 | 28.48 | -7.27 | 21.8 | -9.2 | 28.83 | -10.72
128 | 21.45 | -0.23 | 21.39 | -0.35 | 21.5 | -0.35 | 21.39 | -7.79 | -0.06 | -32.4 | -0.06 | 0 | 0 | 0 | 0.23 | 21.33 | -0.29 | 0.06 | -18.69 | -0.18 | -29.12 | -0.06 | -32.4 | 0 | 0 | -0.06 | 28.42 | 0.23 | 28.48 | -1 | 28.42 | -5.22 | 28.71 | -7.09 | 28.65 | -8.79
128.032 | 0.23 | -0.23 | 21.45 | -0.29 | 21.5 | -0.35 | 21.39 | -5.57 | -0.06 | -31.99 | -0.06 | 0 | 0.06 | 0 | 0.29 | 21.39 | -0.29 | 0.12 | -16.7 | -0.06 | -27.6 | -0.06 | -30.59 | -0.06 | 0 | -0.06 | 28.42 | 0.18 | 28.48 | -0.35 | 28.48 | -3.93 | 28.42 | -5.04 | 28.48 | -6.86
128.064 | 0 | -0.23 | 21.39 | -0.23 | 21.39 | -0.41 | 21.45 | -3.46 | -0.18 | -30.35 | 0 | 0 | 0.12 | 0 | 0.12 | 21.33 | -0.18 | 0 | -15.65 | -0.06 | -25.43 | 0 | -28.3 | 0.06 | 0 | 0 | 28.54 | 0.12 | 28.48 | -0.23 | 28.59 | -2.64 | 28.54 | -3.75 | 28.71 | -5.22
128.096 | 0 | -0.06 | 21.39 | -0.35 | 21.45 | -0.23 | 21.33 | -1.82 | 0 | -28.48 | -0.06 | 0 | -0.06 | 0 | 0.12 | 21.33 | -0.18 | 0.06 | -13.95 | 0 | -23.15 | 0.06 | -26.84 | -0.06 | 0 | -0.06 | 28.89 | 0.18 | 28.42 | 0 | 28.3 | -1.99 | 28.65 | -2.64 | 28.42 | -3.69
128.128 | 0 | -0.18 | 21.39 | -0.23 | 21.45 | -0.29 | 21.39 | -1.64 | 0.06 | -25.61 | -0.06 | 0 | -0.06 | 0 | 0.18 | 21.27 | -0.06 | 0.06 | -12.25 | -0.06 | -21.86 | 0 | -24.67 | -0.06 | -6.91 | -0.06 | 0.18 | 0 | 28.48 | 0.12 | 28.36 | -0.94 | 28.48 | -1.88 | 28.24 | -2.4
128.16 | 0 | -0.18 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | -0.59 | 0 | -23.38 | -0.06 | -0.18 | -0.06 | 0 | 0.18 | 21.33 | -0.06 | 7.56 | -10.66 | -0.06 | -19.28 | 0 | -21.86 | 0 | -84.85 | -0.18 | 0 | 0 | 28.36 | 0.18 | 28.3 | -0.64 | 28.48 | -0.82 | 28.83 | -1.23
128.192 | 0 | -0.18 | 21.5 | -0.18 | 21.45 | -0.23 | 21.39 | -0.06 | -0.06 | -21.39 | -0.06 | -61.23 | 0 | 0 | 0.23 | 21.15 | 0 | 13.01 | -9.38 | -0.06 | -17.64 | 0 | -20.51 | -0.06 | -54.44 | 0 | 0 | 0 | 28.36 | 0.41 | 28.36 | -0.29 | 28.42 | -0.64 | 28.65 | -1
128.224 | 0.06 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.06 | -0.06 | -19.16 | -0.06 | -60.3 | 0 | 0 | 0.23 | 21.21 | 0.06 | 18.63 | -8.09 | -0.06 | -16 | 0 | -18.22 | -0.06 | -48.81 | -0.06 | 0 | 0.06 | 28.42 | 0.29 | 28.36 | -0.18 | 28.65 | -0.35 | 28.42 | -0.53
128.256 | 0 | -0.06 | 21.5 | -0.18 | 21.45 | -0.29 | 21.39 | 0.06 | 0 | -17.52 | -0.06 | -45 | 0 | -1.29 | 0.18 | 21.39 | 0 | 21.56 | -6.86 | 0.06 | -14.41 | 0.06 | -16.47 | -0.06 | -43.89 | 0.06 | 0 | 0 | 28.65 | 0.29 | 28.54 | -0.06 | 28.48 | -0.29 | 28.42 | -0.59
128.288 | 0 | -0.06 | 21.45 | -0.23 | 21.45 | -0.29 | 21.39 | 0.23 | -0.06 | -15.59 | 0 | -38.91 | -0.18 | -68.56 | 0.18 | 21.39 | 0 | 21.45 | -5.86 | 0.06 | -12.83 | 0 | -14.53 | 0 | -37.44 | 0 | 0 | 0.06 | 28.77 | 0.35 | 28.59 | 0 | 28.48 | -0.18 | 28.48 | 0
128.32 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | -0.29 | 21.33 | 0.35 | 0 | -12.6 | -0.06 | -32.7 | -0.12 | -48.75 | 0.12 | 21.39 | 0.12 | 21.39 | -4.63 | 0 | -10.78 | 0 | -13.13 | 0 | -36.86 | 0 | 0 | 0 | 28.71 | 0.41 | 28.42 | 0 | 28.42 | 0.12 | 28.59 | 0
128.352 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.45 | 0.35 | -0.06 | -10.66 | -0.06 | -32.64 | -0.06 | -30.94 | 0.18 | 21.39 | 0.06 | 21.39 | -3.52 | 0 | -9.79 | 6.33 | -11.66 | 0 | -37.97 | -0.18 | 0 | -0.06 | 28.48 | 0.29 | 28.36 | 0.06 | 28.48 | 0.06 | 28.48 | 0
128.384 | 0 | -0.06 | 21.45 | -0.18 | 21.5 | -0.23 | 21.39 | 0.41 | 0 | -8.67 | -0.06 | -32.93 | -0.06 | -29.06 | 0.06 | 21.27 | 0.06 | 21.39 | -2.87 | 0 | -8.14 | 14.06 | -10.31 | 0.06 | -37.33 | -0.06 | 0 | 0 | 28.42 | 0.23 | 28.42 | 0.06 | 28.18 | -0.06 | 28.48 | 0
128.416 | 0 | 0 | 21.39 | -0.18 | 21.45 | -0.23 | 21.39 | 0.41 | -0.06 | -6.56 | -0.06 | -31.82 | -0.06 | -34.16 | 0.06 | 21.27 | 0.06 | 21.45 | -2.52 | 6.27 | -7.32 | 21.74 | -9.32 | -0.06 | -35.92 | 0 | 0 | -0.06 | 0.18 | 0.23 | 28.42 | 0.18 | 28.36 | 0.06 | 28.48 | 0
128.448 | -32.52 | -0.06 | 21.45 | -0.18 | 21.45 | -0.23 | 21.39 | 0.41 | 6.74 | -4.22 | -0.06 | -29.83 | -0.18 | -35.92 | 0.12 | 21.33 | 0.12 | 21.27 | -2.05 | 14 | -5.8 | 28.95 | -7.68 | -0.06 | -34.1 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.59 | 0.18 | 28.65 | 0.06 | 28.89 | 0.18
128.48 | -60.82 | -0.06 | 0.12 | -0.18 | 21.5 | -0.18 | 21.39 | 0.35 | 14.41 | -2.11 | -0.06 | -27.42 | 0 | -31.41 | 0.06 | 21.39 | 0.06 | 21.27 | -1.46 | 24.26 | -4.63 | 28.65 | -6.15 | 2.46 | -31.06 | -0.06 | 0 | 0 | 0 | 0.12 | 28.48 | 0.06 | 28.24 | 0 | 28.89 | 0.12
128.512 | -44.24 | -0.06 | 0 | 0 | 21.39 | -0.23 | 21.45 | 0.41 | 19.63 | -1.29 | -0.06 | -25.25 | -0.06 | -28.89 | 0 | 0.12 | 0 | 21.39 | -1.05 | 28.77 | -3.75 | 28.65 | -4.8 | 10.37 | -29.77 | -0.18 | 0 | 0.06 | 0 | 0 | 28.48 | 0.12 | 28.65 | 0.06 | 28.42 | 0.06
128.544 | -41.02 | -0.06 | 0 | -0.06 | 21.45 | -0.18 | 21.45 | 0.29 | 21.39 | -0.76 | 0.06 | -23.32 | 0 | -27.01 | -0.06 | 0.06 | 0.06 | 21.45 | -0.76 | 28.54 | -3.11 | 28.71 | -4.1 | 17.87 | -27.77 | -0.06 | -61.64 | -0.18 | 0 | 0.18 | 28.89 | 0.12 | 28.59 | 0.12 | 28.3 | 0.18
128.576 | -38.09 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.39 | 0.35 | 21.39 | 0 | -0.18 | -21.56 | 0 | -26.49 | -0.06 | 0 | 0.06 | 21.27 | -0.64 | 28.48 | -2.58 | 28.48 | -3.46 | 25.49 | -25.67 | 0 | -74.71 | -0.06 | 0 | 0.06 | 28.24 | 0.06 | 28.42 | 0.12 | 28.42 | 0.18
128.608 | -35.51 | -0.06 | 0 | -0.18 | 21.39 | -0.18 | 21.39 | 0.23 | 21.39 | -0.18 | -0.23 | -19.98 | -0.06 | -25.02 | -0.06 | 0 | 0.12 | 21.33 | -0.76 | 28.48 | -1.88 | 28.59 | -2.7 | 28.95 | -24.14 | 0 | -51.27 | -0.06 | 0 | 0.12 | 28.42 | 0 | 28.48 | 0.06 | 28.42 | 0.12
128.64 | -33.52 | -0.06 | 0 | -0.06 | 0.12 | -0.18 | 21.39 | 0.18 | 21.39 | 0.35 | -0.18 | -17.29 | -0.18 | -22.21 | -0.06 | 0 | 0.06 | 21.33 | -0.53 | 28.54 | -1.17 | 28.48 | -1.88 | 28.54 | -22.44 | -0.06 | -46.64 | -0.06 | 0 | 0 | 28.59 | 0 | 28.42 | 0.18 | 28.42 | 0.12
128.672 | -31.64 | -0.06 | 0 | -0.06 | 0 | -0.18 | 19.98 | 0.12 | 21.27 | 0.35 | -0.06 | -15.41 | -0.06 | -20.68 | -0.06 | 0 | 0.18 | 21.45 | -0.35 | 28.48 | -0.94 | 28.54 | -1.29 | 28.59 | -20.1 | -0.06 | -40.61 | -0.06 | 0 | 0 | 28.36 | 0 | 28.59 | 0.18 | 28.42 | 0.06
128.704 | -30 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0.12 | 0.18 | 21.39 | 0.35 | -0.23 | -13.24 | -0.06 | -19.63 | -0.06 | 0 | -0.06 | 21.33 | -0.35 | 28.18 | -0.94 | 28.83 | -0.94 | 28.65 | -18.16 | -0.06 | -37.21 | -0.06 | 0 | 0 | 28.42 | -0.06 | 28.48 | 0.12 | 28.42 | 0.12
128.736 | -27.54 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.33 | 0.41 | -0.06 | -11.13 | -0.06 | -17.87 | 0 | 0 | 0.06 | 21.39 | -0.41 | 28.48 | -0.53 | 28.13 | -0.64 | 28.48 | -15.94 | -0.06 | -37.68 | 0 | 0 | 0 | 0.23 | 0 | 27.89 | 0.12 | 28.42 | 0.23
128.768 | -25.14 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 21.27 | 0.47 | -0.06 | -9.32 | 0 | -16.06 | 0 | 0 | 0.06 | 21.15 | -0.41 | 28.42 | -0.59 | 28.65 | -0.41 | 28.54 | -14.18 | 0 | -37.74 | 0 | 0 | 0 | 0 | 0 | 28.48 | 0.12 | 0.23 | 0.12
128.8 | -23.15 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.47 | -0.18 | -6.15 | 0 | -13.83 | -0.06 | 0 | 0 | 21.39 | -0.29 | 28.65 | -0.35 | 28.65 | -0.23 | 28.54 | -11.78 | 0 | -36.39 | 0 | 0 | 0 | 0 | 0 | 0.18 | 0.06 | 0 | 0.18
128.832 | -20.92 | -0.18 | 0 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.53 | 6.68 | -4.28 | -0.06 | -12.13 | -0.18 | 0 | 0 | 21.39 | -0.29 | 28.54 | -0.23 | 28.36 | -0.41 | 28.59 | -9.61 | 0 | -34.45 | -0.18 | -70.73 | 0.06 | 0 | 0 | 0 | 0 | 0 | 0.12
128.864 | -19.34 | -0.06 | 0 | -0.18 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.41 | 12.31 | -2.23 | -0.06 | -10.37 | 0 | 0 | 0.06 | 21.33 | -0.35 | 28.59 | -0.29 | 28.42 | -0.29 | 28.65 | -7.68 | 0 | -32.76 | 0 | -72.48 | 0 | 0 | 0 | 0 | 0.06 | 0 | 0.06
128.896 | -16.88 | -0.06 | -23.67 | -0.06 | 0 | -0.06 | 0 | 0 | 21.33 | 0.41 | 18.22 | -1.88 | -0.06 | -8.91 | -0.06 | 0 | 0.06 | 21.39 | -0.41 | 28.59 | -0.23 | 28.48 | -0.29 | 28.59 | -5.51 | 2.34 | -30.35 | -0.18 | -50.04 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0 | 0.12
128.928 | -15.18 | -0.18 | -62.46 | -0.06 | 0 | -0.06 | 0 | 0 | 21.39 | 0.47 | 21.62 | -0.53 | -0.18 | -6.8 | 0.06 | -0.94 | 0 | 21.33 | -0.23 | 28.65 | -0.29 | 28.48 | -0.23 | 28.42 | -3.63 | 10.25 | -29.06 | -0.18 | -48.28 | 0 | 0 | -0.06 | 0 | 0.12 | 0 | 0.06
128.96 | -13.71 | -0.06 | -41.72 | 0 | 0 | -0.06 | 0 | 0 | 21.45 | 0.35 | 21.39 | -0.29 | -0.06 | -4.63 | 0 | -71.14 | 0.12 | 21.33 | -0.23 | 28.36 | -0.29 | 28.24 | -0.23 | 28.77 | -2.05 | 20.45 | -26.54 | 0 | -40.14 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.12
128.992 | -12.07 | -0.06 | -39.2 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.39 | 0.06 | 3.75 | -3.16 | 0 | -48.58 | 0.12 | 21.33 | -0.29 | 28.54 | -0.29 | 28.42 | -0.23 | 28.42 | -1 | 27.77 | -25.14 | -0.18 | -36.92 | -0.06 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06
129.024 | -10.55 | -0.06 | -38.32 | -0.06 | -0.12 | -0.18 | 0 | -0.06 | 21.27 | 0.23 | 21.45 | 0.23 | 10.02 | -2.23 | 0 | -31.11 | 0.23 | 19.92 | -0.23 | 28.42 | -0.29 | 28.48 | -0.23 | 28.89 | -0.76 | 28.77 | -23.44 | 0 | -37.85 | -0.06 | 0 | 0 | 0 | 0 | 0 | 0.12
129.056 | -8.61 | 0 | -35.45 | -0.06 | -31.35 | -0.06 | 0 | 0.06 | 21.45 | 0.35 | 21.39 | 0.12 | 15.35 | -1.7 | -0.06 | -29.36 | 0 | 0.12 | -0.18 | 28.48 | -0.18 | 28.48 | -0.23 | 28.48 | -0.35 | 28.65 | -21.27 | 0 | -38.61 | 0.12 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06
129.088 | -7.56 | -0.06 | -34.75 | -0.06 | -60.47 | -0.06 | 0 | 0 | 21.39 | 0.18 | 21.39 | 0.35 | 20.98 | -0.76 | -0.06 | -34.28 | 0.12 | 0 | -0.23 | 28.48 | -0.18 | 28.42 | -0.18 | 28.42 | -0.18 | 28.48 | -19.63 | 0 | -37.09 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.12
129.12 | -6.09 | 0 | -31.41 | 0 | -42.07 | -0.06 | -62.81 | -0.06 | 21.33 | 0.12 | 21.39 | 0.47 | 21.39 | -0.53 | 0 | -34.45 | 0.06 | 0 | -0.18 | 28.54 | -0.18 | 28.54 | -0.18 | 28.71 | 0.18 | 28.48 | -16.93 | -0.18 | -34.22 | 0.06 | 0 | -0.06 | 0 | 0 | 0 | 0.18
129.152 | -4.51 | -0.06 | -29.3 | -0.06 | -39.26 | -0.06 | -62.93 | -0.06 | 21.39 | 0.23 | 21.39 | 0.41 | 21.39 | -0.06 | 0 | -31.99 | 0.12 | 0 | -0.18 | 28.3 | -0.18 | 28.3 | -0.18 | 28.42 | 0.06 | 28.77 | -14.77 | 0 | -32.76 | 0 | -61.53 | -0.06 | 0 | -0.18 | 0.06 | 0
129.184 | -3.16 | -0.06 | -27.31 | 0 | -38.5 | -0.06 | -45.47 | -0.18 | 21.39 | 0.23 | 21.45 | 0.41 | 21.27 | -0.18 | -0.06 | -29.12 | 0.12 | 0 | -0.18 | 28.71 | -0.18 | 28.24 | -0.23 | 28.36 | 0.35 | 28.65 | -13.36 | -0.06 | -30.35 | 0.06 | -77.52 | -0.18 | 0 | 0 | -45.35 | 0.12
129.216 | -1.7 | -0.06 | -25.43 | 0 | -34.98 | -0.06 | -39.96 | -0.06 | 21.39 | 0.12 | 21.39 | 0.47 | 21.39 | -0.06 | -0.06 | -27.48 | 0.12 | 0 | -0.18 | 28.48 | -0.18 | 28.65 | -0.06 | 28.36 | 0.29 | 28.48 | -11.07 | 0 | -28.13 | 0 | -46.41 | -0.06 | -34.63 | -0.06 | -83.09 | 0.06
129.248 | -1.11 | -0.06 | -23.15 | -0.06 | -33.75 | -0.06 | -35.1 | -0.06 | 21.39 | 0.12 | 21.27 | 0.29 | 21.39 | 0 | 2.4 | -26.37 | 0.12 | 0 | -0.06 | 28.54 | -0.18 | 0.18 | -0.18 | 28.83 | 0.35 | 28.36 | -9.08 | 0 | -26.9 | -0.18 | -29.77 | 0 | -83.26 | 0.06 | -51.68 | 0
129.28 | -0.53 | 6.45 | -20.92 | 0 | -30.7 | -0.06 | -32.05 | -0.06 | 21.27 | 0.06 | 21.39 | 0.41 | 21.39 | 0.18 | 10.14 | -24.38 | 0.12 | 0 | -0.18 | 28.42 | -0.18 | 0 | -0.06 | 28.54 | 0.18 | 28.65 | -6.56 | -0.06 | -24.61 | 0 | -37.91 | -0.18 | -45.12 | 0.12 | -29.47 | 0
129.312 | -0.35 | 12.42 | -18.93 | 0 | -28.71 | -0.06 | -32.46 | -0.06 | 10.72 | 0 | 21.45 | 0.41 | 21.33 | 0.06 | 15.47 | -22.44 | 0.06 | 0 | -0.18 | 0.18 | -0.18 | 0 | -0.18 | 28.42 | 0.35 | 28.42 | -4.16 | 0.06 | -23.38 | -0.18 | -43.24 | 0 | -29.3 | 0 | -35.92 | 0
129.344 | -0.23 | 17.7 | -16.93 | -0.06 | -27.01 | -0.06 | -32.4 | -0.18 | 0.12 | 0 | 21.39 | 0.41 | 21.39 | 0.12 | 21.33 | -21.33 | 0.18 | 0 | -0.06 | 0 | -0.06 | 0 | -0.18 | 28.59 | 0.18 | 28.36 | -2.52 | 0 | -21.86 | 0 | -39.67 | 0 | -35.74 | 0 | -42.42 | 0
129.376 | -0.29 | 21.74 | -15.53 | -0.06 | -25.61 | -0.18 | -30.18 | -0.06 | 0 | -0.06 | 21.39 | 0.29 | 21.27 | 0.12 | 21.45 | -19.34 | 0.12 | 0 | -0.06 | 0 | -0.06 | 0 | -0.06 | 22.74 | 0.29 | 28.42 | -1.76 | 4.39 | -19.75 | 0 | -35.39 | -0.23 | -42.95 | -0.18 | -40.84 | 0.06
129.408 | -0.23 | 21.5 | -13.54 | -0.06 | -23.2 | -0.06 | -28.65 | -0.18 | 0 | -0.06 | 21.45 | 0.35 | 21.33 | 0.23 | 21.39 | -17.99 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | -0.12 | 0.18 | 0.29 | 28.3 | -0.82 | 12.72 | -17.46 | 0 | -32.7 | -0.06 | -40.84 | 0 | -36.62 | 0
129.44 | -0.35 | 21.45 | -11.89 | 0.06 | -20.63 | -0.06 | -25.72 | -0.06 | 0 | -0.18 | 21.39 | 0.18 | 21.27 | 0.23 | 21.33 | -15.82 | 0.06 | 0 | -0.18 | 0 | -0.18 | 0 | -0.06 | 0 | 0.41 | 28.59 | -0.35 | 22.85 | -14.82 | 0 | -32.64 | 0 | -35.33 | 0 | -33.11 | 0
129.472 | -0.29 | 21.45 | -10.78 | -0.06 | -17.99 | -0.06 | -24.08 | -0.06 | 0 | -0.06 | 21.45 | 0.18 | 21.39 | 0.18 | 21.33 | -13.71 | 0.12 | -55.08 | 0 | 0 | 0 | 0 | -0.06 | 0.06 | 0.35 | 28.65 | -0.18 | 28.95 | -13.24 | -0.18 | -31.82 | 0 | -33.05 | 0 | -32.58 | 0
129.504 | -0.41 | 21.45 | -9.08 | 0 | -16.76 | -0.06 | -21.68 | 0 | 0 | -0.06 | 21.39 | 0.12 | 21.33 | 0.23 | 21.39 | -12.25 | 0.06 | -63.05 | -0.06 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 28.54 | 0.06 | 28.54 | -11.02 | 0 | -30.24 | 0 | -32.52 | 0 | -32.58 | -0.06
129.536 | -0.18 | 21.45 | -7.85 | -0.06 | -15.41 | -0.18 | -19.75 | 0.06 | 0 | -0.06 | 21.39 | 0.12 | 21.39 | 0.18 | 21.39 | -10.55 | 0.06 | -38.91 | -0.18 | 0 | -0.06 | 0 | 0 | 0 | 0.23 | 28.65 | 0.12 | 28.3 | -9.14 | 0.06 | -28.07 | -0.06 | -32.4 | 0 | -30.94 | 0
129.568 | -0.23 | 21.5 | -6.62 | 0 | -13.77 | -0.06 | -17.58 | 0 | 0 | -0.06 | 20.98 | 0.06 | 21.45 | 0.23 | 21.39 | -8.91 | 0.06 | -27.83 | -0.18 | 0 | -0.06 | 0 | -0.18 | 0 | 0.18 | 28.48 | 0.35 | 28.48 | -7.03 | 0 | -26.6 | -0.06 | -31 | 0 | -28.48 | 0
129.6 | -0.35 | 21.45 | -4.34 | -0.06 | -11.6 | -0.06 | -15.41 | -0.06 | 0 | -0.06 | 21.39 | 0 | 21.33 | 0.23 | 21.39 | -6.5 | 0.06 | -33.05 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0.18 | 28.89 | 0.47 | 28.42 | -4.28 | 0 | -25.55 | -0.06 | -27.95 | 0.06 | -26.31 | 0
129.632 | -0.18 | 21.45 | -3.11 | 0 | -10.02 | -0.06 | -13.54 | 0 | 0 | -0.06 | 21.39 | 0 | 21.27 | 0.18 | 21.27 | -4.98 | 0 | -35.51 | -0.06 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 28.65 | 0.35 | 28.42 | -2.52 | 0 | -23.38 | -0.06 | -26.84 | -0.06 | -25.67 | 0
129.664 | -0.29 | 21.45 | -2.05 | 0.18 | -8.85 | -0.06 | -11.07 | -0.06 | 0 | 0.06 | 21.39 | 0.12 | 21.39 | 0.12 | 21.39 | -3.52 | 0.18 | -33.99 | -0.18 | 0 | -0.06 | -31 | -0.06 | 0 | 0.06 | 28.65 | 0.41 | 28.42 | -1.82 | -0.06 | -21.68 | 0 | -25.43 | -0.06 | -23.91 | -0.06
129.696 | -0.23 | 21.5 | -1.35 | 7.38 | -7.44 | 0.06 | -9.2 | 0 | 0 | 0.06 | 21.45 | 0.06 | 21.39 | 0.18 | 21.39 | -2.23 | 0.06 | -30.53 | -0.18 | 0 | 0 | -74.18 | -0.18 | 0 | 0 | 28.42 | 0.29 | 28.36 | -0.94 | 0 | -20.27 | -0.06 | -24.08 | -0.06 | -21.97 | 0
129.728 | -0.29 | 21.45 | -0.64 | 13.18 | -6.15 | 7.5 | -6.86 | -0.06 | 0 | -0.06 | 0.12 | 0.06 | 21.39 | 0 | 21.39 | -1.41 | 0 | -27.89 | -0.18 | -36.51 | 0 | -46.47 | -0.06 | 0 | -0.06 | 28.48 | 0.29 | 28.48 | -0.7 | 0.06 | -19.22 | -0.06 | -22.44 | 0 | -20.74 | 0.06
129.76 | -0.29 | 21.45 | -0.29 | 20.63 | -4.1 | 14.82 | -4.16 | -0.06 | -72.83 | 0.06 | 0 | 0 | 21.33 | 0.18 | 21.27 | -0.94 | -0.06 | -27.42 | 0 | -64.51 | -0.06 | -45.76 | -0.18 | 0 | 0.06 | 28.89 | 0.47 | 28.3 | 0 | 0 | -17.05 | 0 | -20.16 | 0.06 | -18.63 | 0.06
129.792 | -0.23 | 21.39 | -0.18 | 21.5 | -2.81 | 20.74 | -2.29 | -0.06 | -57.19 | -0.06 | 0 | 0.06 | 21.27 | 0.06 | 21.39 | -0.41 | 1.99 | -25.43 | -0.06 | -46.35 | 0 | -44.3 | -0.06 | -0.12 | -0.06 | 0.29 | 0.29 | 28.36 | 0 | 0.06 | -14.71 | 0 | -18.87 | -0.06 | -17.4 | -0.06
129.824 | -0.18 | 21.45 | -0.29 | 21.45 | -1.64 | 21.45 | -1.7 | -0.06 | -43.89 | -0.06 | 0 | 0 | 21.45 | 0.12 | 21.21 | -0.29 | 6.62 | -23.44 | -0.06 | -45.53 | -0.06 | -41.25 | -0.06 | -69.96 | 0.06 | 0.12 | 0.23 | 28.54 | 0.06 | 3.87 | -13.54 | -0.06 | -17.58 | 0 | -15.47 | -0.18
129.856 | -0.18 | 21.45 | -0.23 | 21.39 | -1.11 | 21.39 | -0.64 | 0.12 | -39.14 | 0 | 0 | -0.06 | 21.15 | -0.06 | 21.39 | -0.06 | 12.48 | -21.5 | -0.06 | -43.48 | -0.06 | -40.78 | 0 | -63.64 | 0 | 0 | 0.18 | 28.65 | 0.18 | 12.13 | -12.01 | 2.17 | -15.94 | -0.06 | -13.95 | -0.18
129.888 | -0.18 | 21.45 | -0.23 | 21.45 | -0.47 | 21.39 | -0.41 | -0.06 | -34.04 | -0.18 | 0 | 0 | 0.18 | -0.06 | 21.39 | -0.06 | 17.99 | -20.45 | -0.06 | -41.08 | -0.06 | -37.79 | 0 | -45.47 | -0.06 | 0 | 0.12 | 28.59 | 0.29 | 19.92 | -10.14 | 11.72 | -14.24 | 3.81 | -12.07 | -0.18
129.92 | -0.18 | 21.45 | -0.47 | 21.45 | -0.18 | 21.39 | -0.06 | 0 | -32.46 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.27 | 0.06 | 21.39 | -18.52 | -0.06 | -39.2 | -0.06 | -36.45 | -0.06 | -45.76 | 0 | 0 | 0.12 | 28.48 | 0.29 | 29.18 | -8.2 | 21.91 | -11.95 | 14.59 | -10.2 | 6.33
129.952 | -0.18 | 21.45 | -0.29 | 21.5 | -0.18 | 21.45 | -0.06 | 0 | -32.4 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.12 | 21.39 | -16.41 | -0.06 | -37.74 | -0.06 | -33.69 | -0.12 | -42.78 | 0.06 | 0 | 0.18 | 28.59 | 0.18 | 28.65 | -6.68 | 29.18 | -10.66 | 22.27 | -8.91 | 14.12
129.984 | -0.18 | 21.45 | -0.29 | 21.39 | -0.18 | 21.45 | 0.35 | 0.06 | -31.93 | 0 | 0 | -0.06 | 0.06 | 0 | 21.39 | 0.35 | 21.39 | -14.59 | -0.06 | -35.74 | 0.06 | -32.34 | -0.06 | -41.9 | -0.18 | 0 | 0.12 | 28.54 | 0.35 | 28.65 | -4.69 | 28.65 | -8.79 | 29.01 | -7.09 | 21.91
130.016 | -0.18 | 21.39 | -0.35 | 21.45 | -0.23 | 21.39 | 0.23 | -0.06 | -30.35 | 0 | 0 | -0.06 | 0 | -0.06 | 21.21 | 0.35 | 21.33 | -13.13 | 0 | -34.45 | -0.06 | -30.53 | -0.06 | -39.61 | -0.06 | 0 | 0.06 | 28.48 | 0.29 | 28.59 | -3.4 | 28.54 | -7.21 | 29.12 | -5.22 | 28.83
130.048 | -0.06 | 21.45 | -0.18 | 21.39 | -0.29 | 21.33 | 0.41 | 0 | -27.95 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.39 | 0.47 | 21.45 | -11.6 | -0.06 | -32.23 | 0.06 | -27.31 | 0 | -37.03 | 0 | 0 | 0.06 | 28.48 | 0.35 | 28.48 | -2.17 | 28.42 | -5.22 | 28.48 | -3.63 | 28.42
130.08 | -0.18 | 21.45 | -0.23 | 21.39 | -0.29 | 21.39 | 0.35 | 1.46 | -24.96 | 0 | 0 | 0 | 0.06 | -0.06 | 21.39 | 0.47 | 21.21 | -9.49 | 0 | -29.83 | 0 | -26.54 | 0 | -35.39 | 0 | 0 | 0.06 | 28.36 | 0.23 | 28.42 | -1.17 | 28.54 | -3.57 | 28.18 | -2.23 | 28.48
130.112 | -0.06 | 21.45 | -0.29 | 21.45 | -0.23 | 21.39 | 0.41 | 8.79 | -22.79 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.21 | -7.44 | 0 | -27.6 | 0 | -23.67 | 0 | -33.52 | 0 | 0 | 0 | 28.59 | 0.18 | 28.48 | -0.94 | 28.65 | -2.52 | 28.54 | -1.76 | 28.65
130.144 | -0.06 | 21.45 | -0.18 | 21.45 | -0.35 | 21.39 | 0.41 | 14.24 | -21.09 | 0 | -46.88 | -0.06 | 0 | -0.06 | 0.23 | 0.35 | 21.45 | -5.86 | -0.06 | -25.61 | 0 | -21.68 | 0 | -31.29 | 0 | 0 | -0.06 | 28.65 | 0.18 | 28.59 | -0.41 | 28.54 | -1.17 | 28.54 | -0.7 | 28.42
130.176 | -0.18 | 0.12 | -0.29 | 21.5 | -0.18 | 21.39 | 0.41 | 19.98 | -19.51 | -0.06 | -64.28 | 0 | 0 | 0 | 0.06 | 0.29 | 21.21 | -4.22 | 8.97 | -24.38 | 0 | -20.1 | -0.06 | -29.77 | 0 | 0 | -0.06 | 28.54 | 0.06 | 28.36 | -0.29 | 28.54 | -0.94 | 28.59 | -0.7 | 28.48
130.208 | -0.06 | 0 | -0.23 | 21.39 | -0.23 | 21.33 | 0.41 | 21.45 | -17.4 | -0.06 | -47.46 | -0.06 | 0 | -0.06 | 0 | 0.35 | 21.33 | -3.16 | 16 | -21.97 | 0 | -17.52 | 0 | -28.01 | -0.18 | -0.06 | 0 | 28.42 | 0.06 | 28.59 | -0.18 | 28.48 | -1 | 28.42 | -0.35 | 28.48
130.24 | -0.18 | 0 | -0.23 | 21.5 | -0.23 | 21.39 | 0.47 | 21.39 | -14.47 | -0.06 | -38.03 | -0.06 | 0 | -0.06 | 0 | 0.29 | 21.33 | -1.7 | 26.19 | -19.45 | 0 | -16.17 | -0.06 | -24.67 | -0.18 | -92.99 | 0 | 28.36 | 0.06 | 28.59 | 0.06 | 28.42 | -0.18 | 28.42 | -0.41 | 28.42
130.272 | -0.06 | 0 | -0.23 | 21.45 | -0.29 | 21.39 | 0.29 | 21.39 | -12.6 | -0.06 | -32.81 | -0.06 | 0 | 0.06 | 0 | 0.35 | 21.39 | -0.76 | 28.71 | -17.46 | 0 | -14.82 | 0 | -23.15 | -0.06 | -59.18 | 0 | 0.29 | 0.06 | 28.48 | 0 | 28.48 | -0.29 | 28.48 | 0.06 | 28.42
130.304 | -0.06 | 0 | -0.18 | 21.39 | -0.35 | 21.33 | 0.41 | 21.39 | -11.07 | -0.06 | -31.99 | -0.06 | -18.93 | 0 | 0 | 0.35 | 21.39 | -0.76 | 28.71 | -15.7 | 1.99 | -13.54 | 0 | -21.56 | 0 | -48.17 | 0 | 0.06 | 0 | 28.48 | 0.06 | 28.42 | -0.06 | 28.65 | 0.06 | 28.42
130.336 | -0.06 | 0 | -0.18 | 21.39 | -0.29 | 21.33 | 0.35 | 21.33 | -8.91 | -0.06 | -32.64 | 0.06 | -73.89 | 0.06 | 0 | 0.29 | 21.15 | -0.29 | 28.36 | -14.24 | 7.73 | -11.54 | 0.12 | -19.16 | -0.06 | -45 | 0 | 0 | 0 | 28.48 | 0.23 | 28.59 | -0.06 | 28.65 | 0 | 28.59
130.368 | -0.06 | 0 | -0.18 | 21.45 | -0.29 | 21.39 | 0.29 | 21.39 | -6.62 | 0.06 | -32.58 | 0 | -51.39 | -0.06 | 0 | 0.29 | 21.39 | -0.23 | 28.71 | -12.95 | 14.94 | -10.02 | 2.93 | -17.58 | -0.18 | -38.85 | -0.18 | 0.06 | 0 | 28.42 | 0.18 | 28.48 | 0 | 28.48 | 0 | 28.54
130.4 | -0.06 | 0 | -0.06 | 21.39 | -0.35 | 21.39 | 0.12 | 21.33 | -3.93 | -0.06 | -30.29 | 0.06 | -28.18 | -0.06 | 0 | 0.35 | 21.33 | -0.06 | 28.3 | -11.25 | 25.72 | -8.32 | 13.3 | -14.82 | -0.06 | -36.97 | 0.06 | 0 | -0.06 | 28.42 | 0.12 | 28.54 | 0 | 28.59 | 0 | 28.48
130.432 | -0.06 | 0 | -0.18 | 21.39 | -0.29 | 21.33 | 0.18 | 21.33 | -2.29 | 0 | -28.48 | -0.06 | -30.41 | 0.06 | 0 | 0.23 | 21.39 | 0.06 | 28.54 | -9.84 | 28.77 | -7.09 | 21.15 | -14.12 | -0.06 | -38.15 | -0.06 | 0 | 0.06 | 28.48 | 0.12 | 28.59 | 0.12 | 28.59 | 0 | 28.59
130.464 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.33 | 0.18 | 21.39 | -1.64 | -0.06 | -26.02 | -0.06 | -35.74 | 0 | 0 | 0.18 | 21.39 | 0 | 28.59 | -8.2 | 28.59 | -5.98 | 29.06 | -12.42 | -0.18 | -37.85 | -0.06 | 0.06 | -0.06 | 28.65 | 0.12 | 28.65 | 0.12 | 28.48 | 0.18 | 28.3
130.496 | -0.18 | 0 | -0.18 | 21.39 | -0.23 | 21.39 | 0.06 | 21.39 | -0.41 | 6.86 | -23.61 | 0 | -35.27 | -0.06 | 0 | 0.23 | 21.33 | 0.18 | 28.59 | -7.38 | 28.65 | -4.92 | 28.71 | -11.13 | -0.06 | -36.45 | -0.06 | 0 | 0.06 | 28.77 | 0.18 | 28.42 | 0.12 | 29.06 | 0.18 | 28.48
130.528 | -0.06 | 0 | -0.18 | 21.45 | -0.18 | 21.5 | 0.06 | 21.33 | -0.35 | 12.31 | -21.86 | -0.06 | -31.7 | -0.06 | 0 | 0.18 | 21.39 | 0.12 | 28.65 | -6.15 | 28.42 | -3.69 | 28.65 | -9.32 | -0.06 | -34.63 | 0 | 0 | -0.06 | 28.65 | 0.06 | 28.48 | 0.06 | 28.59 | 0.18 | 28.48
130.56 | -0.06 | 0 | -0.06 | 0.64 | -0.18 | 21.45 | 0.12 | 21.21 | -0.06 | 19.69 | -19.45 | -0.06 | -27.83 | -0.06 | -16.23 | 0.23 | 21.21 | 0.23 | 28.59 | -4.86 | 28.59 | -3.16 | 28.77 | -8.03 | -0.06 | -31.23 | 0 | 0 | -0.06 | 28.42 | 0.06 | 28.48 | 0.23 | 28.54 | 0.18 | 28.48
130.592 | -0.06 | -33.75 | -0.06 | 0.06 | -0.18 | 0.59 | 0.06 | 21.39 | 0.12 | 21.5 | -17.23 | 0 | -27.07 | -0.18 | -74.07 | 0.12 | 21.33 | 0.18 | 28.36 | -3.87 | 28.65 | -2.23 | 28.65 | -6.8 | -0.18 | -29.83 | 0 | 0 | -0.18 | 28.3 | 0.06 | 28.24 | 0.18 | 28.42 | 0.06 | 28.42
130.624 | -0.06 | -58.07 | -0.18 | 0 | -0.18 | 0.06 | 0.06 | 21.39 | 0.29 | 21.39 | -15.29 | -0.18 | -25.9 | 0 | -51.92 | 0.06 | 21.39 | 0.12 | 28.42 | -3.16 | 28.71 | -1.99 | 28.83 | -5.39 | -0.06 | -28.54 | 0.06 | 0 | -0.06 | 28.42 | -0.06 | 28.48 | 0.18 | 28.48 | 0.12 | 28.42
130.656 | -0.06 | -44.01 | -0.06 | 0 | -0.06 | 0 | 0.06 | 21.39 | 0.29 | 21.39 | -13.36 | -0.06 | -24.32 | -0.06 | -30.94 | 0 | 21.39 | 0.12 | 28.65 | -2.58 | 28.59 | -1.64 | 28.24 | -4.51 | -0.06 | -26.54 | 0 | 0 | 0 | 28.42 | 0 | 28.48 | 0.12 | 28.48 | 0.12 | 28.42
130.688 | -0.18 | -40.37 | -0.06 | 0 | -0.18 | 0 | 0.06 | 21.39 | 0.53 | 21.33 | -11.6 | 0 | -22.33 | -0.06 | -29.01 | 0.06 | 21.39 | 0.12 | 28.65 | -1.99 | 28.42 | -0.94 | 28.54 | -3.75 | -0.06 | -24.96 | 0 | -20.1 | 0 | 28.48 | -0.06 | 28.42 | 0.12 | 28.48 | 0.18 | 28.48
130.72 | 0 | -37.38 | -0.06 | 0 | -0.18 | 0 | 0 | 21.33 | 0.53 | 21.39 | -8.73 | 0.06 | -20.51 | -0.18 | -35.1 | 0 | 0.06 | 0.23 | 28.65 | -1.46 | 28.95 | -0.7 | 28.42 | -2.7 | 0 | -22.38 | 0 | -79.28 | -0.06 | 0.29 | 0 | 28.48 | 0.18 | 28.83 | 0.12 | 28.48
130.752 | -0.06 | -34.92 | -0.06 | 0 | -0.18 | 0 | 0 | 21.39 | 0.47 | 21.33 | -6.33 | -0.06 | -18.99 | -0.06 | -34.69 | 0 | 0 | 0.12 | 28.42 | -1.17 | 28.48 | -0.47 | 28.42 | -2.11 | -0.18 | -20.63 | -0.06 | -51.51 | 0 | 0.06 | 0 | 0.23 | 0 | 28.48 | 0.12 | 28.54
130.784 | -0.06 | -33.46 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.21 | 0.53 | 21.33 | -4.45 | -0.06 | -17.4 | -0.06 | -31.41 | 0 | 0 | 0.18 | 28.54 | -0.76 | 28.65 | -0.29 | 28.83 | -1.41 | 0 | -18.87 | 0 | -48.28 | 0 | 0 | -0.18 | 0.06 | 0.12 | 0.23 | 0.12 | 28.48
130.816 | -0.06 | -31.47 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.45 | 0.41 | 21.33 | -2.64 | -0.06 | -15.47 | -0.06 | -29.3 | 0.06 | 0 | 0.29 | 28.54 | -0.7 | 28.42 | -0.23 | 28.42 | -0.94 | 0.94 | -16.76 | 0 | -42.13 | 0 | 0.06 | -0.18 | 0 | 0 | 0.06 | 0 | 0.23
130.848 | -0.06 | -29.06 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.41 | 21.33 | -1.99 | -0.06 | -13.83 | -0.06 | -27.72 | 0.06 | 0 | 0.23 | 28.54 | -0.47 | 28.42 | -0.35 | 28.65 | -0.94 | 10.96 | -15.29 | 0 | -37.09 | 0.06 | 0 | -0.06 | 0.06 | 0 | 0 | 0.06 | 0.06
130.88 | -0.06 | -27.01 | -0.18 | 0 | -0.18 | 0 | 0.06 | 21.45 | 0.47 | 21.39 | -0.7 | 7.09 | -11.48 | -0.06 | -26.37 | 0 | 0 | 0.23 | 28.48 | -0.41 | 28.48 | -0.18 | 28.83 | -0.53 | 20.86 | -12.36 | 0 | -37.91 | 0 | 0 | -0.06 | 0 | 0.06 | 0 | 0.06 | 0
130.912 | -0.06 | -24.84 | -0.06 | 0 | -0.06 | 0 | -0.06 | 21.33 | 0.47 | 21.39 | -0.06 | 12.77 | -9.84 | -0.06 | -24.55 | 0 | 0 | 0.12 | 28.59 | -0.35 | 28.65 | -0.23 | 28.42 | -0.41 | 28.95 | -10.25 | 0 | -38.32 | 0 | 0 | -0.18 | 0 | 0.12 | 0 | 0.06 | 0
130.944 | -0.06 | -22.74 | -0.06 | 0 | -0.18 | 0 | -0.06 | 21.33 | 0.29 | 21.45 | 0.06 | 18.28 | -8.32 | 0.35 | -22.85 | -0.06 | 0 | 0.23 | 28.77 | -0.41 | 28.36 | -0.29 | 28.42 | -0.29 | 28.65 | -8.14 | -0.06 | -36.68 | 0 | 0 | -0.18 | 0 | 0 | 0 | 0 | 0
130.976 | -0.06 | -20.92 | -0.18 | 0 | -0.06 | 0 | -0.06 | 10.55 | 0.23 | 21.39 | 0.06 | 21.5 | -6.39 | 7.97 | -20.98 | 0 | 0 | 0.18 | 28.54 | -0.29 | 28.65 | -0.23 | 28.42 | -0.23 | 28.65 | -5.98 | 0 | -35.63 | 0 | 0 | -0.06 | 0 | -0.06 | 0 | 0 | 0
131.008 | -0.06 | -18.87 | -0.06 | -47.05 | -0.06 | -0.12 | -0.06 | 0.06 | 0.18 | 21.27 | 0.41 | 21.39 | -5.04 | 13.54 | -19.51 | -0.06 | 0 | 0.18 | 28.36 | -0.29 | 28.65 | -0.18 | 28.65 | -0.29 | 28.48 | -4.04 | 0 | -33.4 | -0.06 | 0 | -0.18 | 0 | 0 | 0 | 0.12 | 0
131.04 | -0.06 | -16.76 | -0.18 | -48.87 | -0.06 | -61.99 | 0 | 0 | 0.06 | 21.45 | 0.23 | 21.39 | -3.16 | 20.98 | -17.52 | 0 | 0 | 0.23 | 14.18 | -0.29 | 28.65 | -0.35 | 28.89 | -0.18 | 28.48 | -1.99 | 0 | -31 | -0.06 | 0 | -0.06 | 0 | 0 | 0 | 0 | 0

## Measured Pressure
Airway Pressure cmH2O
Epoch No | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36
Time (s)
0 | 11.72 | 0.1 | 4.49 | 0.1 | 0 | 0.2 | 0.1 | 4.88 | 4.88 | 18.07 | 4.79 | 20.7 | 4.88 | 22.95 | 9.67 | 13.18 | 9.67 | 9.86 | 9.96 | 0.1 | 2.25 | 0 | 5.08 | 0.1 | 17.97 | 4.79 | 19.82 | 4.59 | 13.48 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
0.032 | 12.21 | 0.1 | 5.08 | 0 | 0.68 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.79 | 20.22 | 4.88 | 23.34 | 9.67 | 14.06 | 9.67 | 9.86 | 10.16 | 0.1 | 2.05 | 0 | 2.15 | 0.1 | 16.8 | 4.79 | 20.41 | 4.59 | 14.26 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
0.064 | 12.79 | 0.1 | 5.57 | 0 | 1.37 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 18.56 | 4.88 | 23.73 | 9.77 | 14.75 | 9.67 | 9.86 | 10.16 | 0.1 | 1.76 | 0 | 2.25 | 0.1 | 17.48 | 4.79 | 21 | 4.69 | 15.14 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
0.096 | 13.38 | 0 | 6.15 | 0.1 | 2.54 | 0 | 0.59 | 4.79 | 4.88 | 9.96 | 4.88 | 18.46 | 4.88 | 24.12 | 9.67 | 15.23 | 9.67 | 9.86 | 9.86 | 0.1 | 1.86 | 0 | 1.95 | 0 | 17.68 | 4.79 | 21.68 | 4.69 | 15.82 | 4.59 | 4.88 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
0.128 | 13.87 | 0.1 | 6.64 | 0.1 | 3.81 | 0 | 1.27 | 4.79 | 4.88 | 3.32 | 4.88 | 18.46 | 4.88 | 24.41 | 9.67 | 15.72 | 9.67 | 9.86 | 9.86 | 0.1 | 1.46 | 0.1 | 1.86 | 0 | 17.68 | 4.79 | 22.17 | 4.69 | 16.5 | 4.59 | 5.86 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77
0.16 | 14.36 | 0.1 | 7.32 | 0.1 | 4.59 | 0.1 | 2.44 | 4.79 | 4.88 | 3.22 | 4.79 | 18.26 | 4.88 | 22.27 | 9.77 | 16.31 | 9.67 | 9.86 | 10.06 | 0.1 | 1.46 | 0 | 1.56 | 0.1 | 17.68 | 4.88 | 22.75 | 4.69 | 17.29 | 4.49 | 6.93 | 9.77 | 9.77 | 9.77 | 10.74 | 9.77
0.192 | 14.84 | 0.1 | 8.01 | 0 | 5.08 | 0.1 | 3.81 | 4.69 | 4.88 | 3.91 | 4.79 | 18.16 | 4.88 | 22.07 | 9.77 | 16.8 | 9.57 | 9.86 | 9.96 | 0.1 | 1.27 | 0 | 1.56 | 0.1 | 17.77 | 4.88 | 23.34 | 4.69 | 17.77 | 4.49 | 8.69 | 9.67 | 10.64 | 9.67 | 11.62 | 9.77
0.224 | 15.43 | 0.1 | 8.5 | 0 | 5.66 | 0.1 | 4.49 | 4.59 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 21.78 | 9.77 | 17.19 | 9.67 | 9.96 | 9.67 | 0 | 1.37 | 0.1 | 1.56 | 0 | 17.77 | 4.79 | 23.93 | 4.79 | 18.46 | 4.59 | 10.45 | 9.67 | 11.43 | 9.67 | 13.28 | 9.77
0.256 | 15.92 | 0.1 | 9.08 | 0.1 | 6.25 | 0 | 5.08 | 4.59 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 21.58 | 9.77 | 17.68 | 9.57 | 9.86 | 9.67 | 0.1 | 1.07 | 0 | 1.27 | 0 | 17.68 | 4.79 | 24.41 | 4.79 | 18.95 | 4.49 | 11.72 | 9.67 | 12.99 | 9.67 | 15.14 | 9.67
0.288 | 16.41 | 0.1 | 9.67 | 0.1 | 6.84 | 0.1 | 5.57 | 4.49 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 18.16 | 9.57 | 9.96 | 9.77 | 0.1 | 1.07 | 0.1 | 1.17 | 0 | 17.77 | 4.79 | 25 | 4.79 | 19.73 | 4.49 | 12.79 | 9.67 | 15.04 | 9.67 | 16.41 | 9.67
0.32 | 16.9 | 0.1 | 10.35 | 0.1 | 7.42 | 0.1 | 6.15 | 4.59 | 4.88 | 5.27 | 4.79 | 18.07 | 4.88 | 21.58 | 9.77 | 18.65 | 9.67 | 9.86 | 9.77 | 0.1 | 0.88 | 0 | 1.07 | 0.1 | 17.68 | 4.88 | 25.49 | 4.79 | 20.31 | 4.49 | 13.67 | 9.67 | 16.31 | 9.67 | 17.29 | 9.67
0.352 | 17.38 | 0.1 | 10.84 | 0 | 8.01 | 0.1 | 6.74 | 4.59 | 4.88 | 4.98 | 4.79 | 18.07 | 4.88 | 21.49 | 9.77 | 19.14 | 9.67 | 9.86 | 9.67 | 0.1 | 0.88 | 0 | 0.88 | 0.1 | 17.68 | 4.79 | 25.98 | 4.79 | 20.9 | 4.49 | 14.36 | 9.67 | 17.19 | 9.67 | 17.97 | 9.67
0.384 | 15.43 | 0.1 | 11.43 | 0 | 8.69 | 0.1 | 7.42 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 19.43 | 9.67 | 9.96 | 9.67 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 17.68 | 4.79 | 26.37 | 4.79 | 21.49 | 4.59 | 15.14 | 9.67 | 17.97 | 9.57 | 18.75 | 9.67
0.416 | 15.04 | 0 | 12.11 | 0.1 | 9.28 | 0 | 8.01 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 19.92 | 9.67 | 9.86 | 9.67 | 0 | 0.68 | 0.1 | 0.78 | 0.1 | 17.58 | 4.88 | 22.75 | 4.79 | 22.07 | 4.59 | 15.82 | 9.67 | 18.75 | 9.57 | 19.43 | 9.67
0.448 | 14.84 | 0.1 | 12.5 | 0.1 | 9.86 | 0 | 8.5 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.31 | 9.67 | 10.55 | 9.57 | 0.1 | 0.68 | 0 | 0.68 | 0 | 5.76 | 4.79 | 22.27 | 4.79 | 22.66 | 4.59 | 16.5 | 9.57 | 19.34 | 9.67 | 20.12 | 9.67
0.48 | 14.75 | 0.1 | 13.09 | 0.1 | 10.45 | 0.1 | 9.28 | 4.49 | 4.79 | 4.98 | 4.79 | 8.98 | 4.88 | 21.39 | 9.86 | 20.8 | 9.67 | 11.33 | 9.67 | 0.1 | 0.49 | 0 | 0.68 | 0.1 | 1.76 | 4.88 | 21.78 | 4.79 | 23.24 | 4.59 | 17.19 | 9.67 | 20.02 | 9.67 | 20.7 | 9.67
0.512 | 14.75 | 0.1 | 13.67 | 0 | 11.04 | 0.1 | 9.86 | 4.49 | 4.79 | 4.69 | 4.79 | 2.73 | 4.88 | 21.39 | 9.86 | 21.19 | 9.67 | 12.4 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.25 | 4.79 | 21.58 | 4.79 | 23.73 | 4.59 | 17.87 | 9.67 | 20.61 | 9.67 | 21.29 | 9.57
0.544 | 14.55 | 0.1 | 14.16 | 0 | 11.62 | 0.1 | 10.35 | 4.59 | 4.88 | 4.59 | 4.88 | 3.22 | 4.79 | 21.39 | 9.86 | 21.58 | 9.67 | 13.57 | 9.67 | 0 | 0.49 | 0.1 | 0.59 | 0.1 | 2.05 | 4.79 | 21.58 | 4.79 | 24.22 | 4.69 | 18.56 | 9.67 | 21.19 | 9.67 | 21.88 | 9.67
0.576 | 14.45 | 0.1 | 14.65 | 0.1 | 12.21 | 0.1 | 11.04 | 4.59 | 4.88 | 4.69 | 4.88 | 4.1 | 4.88 | 12.5 | 9.77 | 21.97 | 9.67 | 14.36 | 9.57 | 0.1 | 0.39 | 0 | 0.49 | 0 | 1.66 | 4.79 | 21.39 | 4.79 | 24.81 | 4.69 | 19.24 | 9.67 | 21.88 | 9.57 | 22.36 | 9.67
0.608 | 14.45 | 0.1 | 15.14 | 0.1 | 12.79 | 0.1 | 11.52 | 4.59 | 4.88 | 4.69 | 4.88 | 5.27 | 4.88 | 6.93 | 9.77 | 22.36 | 9.67 | 14.94 | 9.57 | 0.1 | 0.29 | 0 | 0.49 | 0 | 1.56 | 4.79 | 21.29 | 4.79 | 25.29 | 4.69 | 19.82 | 9.67 | 22.36 | 9.67 | 22.95 | 9.67
0.64 | 14.45 | 0.1 | 15.72 | 0.1 | 13.28 | 0.1 | 12.11 | 4.69 | 4.79 | 4.59 | 4.88 | 5.96 | 4.88 | 8.11 | 9.86 | 22.75 | 9.77 | 15.43 | 9.57 | 0.1 | 0.2 | 0.88 | 0.39 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 25.68 | 4.69 | 20.31 | 9.67 | 22.95 | 9.67 | 23.54 | 9.67
0.672 | 14.45 | 0.1 | 16.21 | 0 | 13.77 | 0.1 | 12.6 | 4.69 | 4.79 | 4.49 | 4.79 | 5.66 | 4.88 | 11.23 | 9.86 | 23.15 | 9.77 | 15.92 | 9.57 | 0.1 | 0.2 | 1.76 | 0.29 | 0.1 | 1.37 | 4.79 | 21.29 | 4.79 | 26.27 | 4.69 | 21 | 9.67 | 23.44 | 9.67 | 24.02 | 9.67
0.704 | 14.45 | 0.1 | 16.7 | 0 | 14.26 | 0.1 | 13.18 | 4.69 | 4.88 | 4.49 | 4.88 | 5.18 | 4.88 | 11.23 | 9.86 | 23.54 | 9.77 | 16.5 | 9.67 | 0.98 | 0.2 | 3.42 | 0.29 | 0.1 | 1.27 | 4.79 | 21.29 | 4.88 | 22.85 | 4.79 | 21.49 | 9.67 | 23.93 | 9.67 | 24.51 | 9.67
0.736 | 14.36 | 0.1 | 17.19 | 0.1 | 14.84 | 0 | 13.77 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 4.88 | 10.16 | 9.77 | 23.83 | 9.77 | 17.09 | 9.57 | 1.86 | 0.2 | 5.27 | 0.2 | 0.1 | 1.17 | 4.79 | 21.19 | 4.88 | 22.27 | 4.79 | 22.17 | 9.67 | 24.51 | 9.67 | 25.1 | 9.67
0.768 | 14.36 | 0 | 13.77 | 0.1 | 15.43 | 0 | 14.26 | 4.69 | 6.35 | 4.49 | 4.88 | 4.98 | 4.88 | 9.77 | 9.77 | 24.22 | 9.77 | 17.38 | 9.57 | 3.52 | 0.2 | 6.35 | 0.2 | 0 | 1.07 | 4.79 | 21.19 | 4.88 | 21.78 | 4.79 | 22.66 | 9.67 | 24.9 | 9.67 | 25.59 | 9.67
0.8 | 8.5 | 0.1 | 14.16 | 0.1 | 15.82 | 0.1 | 14.75 | 4.79 | 7.72 | 4.49 | 4.79 | 4.98 | 4.88 | 10.06 | 9.86 | 20.8 | 9.77 | 17.87 | 9.67 | 5.18 | 0.1 | 7.23 | 0.2 | 0.88 | 0.98 | 4.79 | 21.19 | 4.79 | 21.58 | 4.79 | 23.24 | 9.77 | 25.59 | 9.67 | 26.08 | 9.67
0.832 | 2.44 | 0.1 | 14.26 | 0 | 16.31 | 0.1 | 15.33 | 4.79 | 8.69 | 4.49 | 4.79 | 4.88 | 4.88 | 10.16 | 9.86 | 20.9 | 9.77 | 18.36 | 9.67 | 6.45 | 0.1 | 8.11 | 0.1 | 1.76 | 1.07 | 4.79 | 10.64 | 4.79 | 21.58 | 4.79 | 23.73 | 9.77 | 26.08 | 9.67 | 26.47 | 9.67
0.864 | 1.95 | 0.1 | 14.45 | 0 | 16.8 | 0.1 | 15.72 | 4.79 | 9.67 | 4.49 | 4.79 | 4.79 | 4.88 | 10.16 | 9.86 | 21.19 | 9.77 | 18.75 | 9.67 | 7.42 | 0.1 | 8.98 | 0.1 | 3.42 | 0.78 | 4.79 | 3.03 | 4.88 | 21.39 | 4.79 | 24.32 | 9.77 | 26.56 | 9.77 | 27.05 | 9.67
0.896 | 1.56 | 0.1 | 14.45 | 0.1 | 17.29 | 0.1 | 16.31 | 4.79 | 10.25 | 4.49 | 4.79 | 4.79 | 4.88 | 9.96 | 9.77 | 21.49 | 9.77 | 19.24 | 9.67 | 8.11 | 0.1 | 9.67 | 0.1 | 5.37 | 0.68 | 4.79 | 3.22 | 4.88 | 21.29 | 4.79 | 24.71 | 9.77 | 27.05 | 9.77 | 27.54 | 9.77
0.928 | 1.56 | 0.1 | 14.45 | 0.1 | 15.23 | 0 | 16.8 | 4.79 | 10.94 | 4.59 | 4.79 | 4.79 | 4.88 | 9.86 | 9.77 | 21.39 | 9.77 | 19.73 | 9.67 | 8.79 | 0.1 | 10.45 | 0.1 | 6.35 | 0.59 | 4.79 | 3.61 | 4.88 | 21.39 | 4.79 | 25.29 | 9.77 | 27.54 | 9.67 | 28.03 | 9.67
0.96 | 1.46 | 0.1 | 14.55 | 0.1 | 15.23 | 0.1 | 17.29 | 4.79 | 11.52 | 4.59 | 4.79 | 4.69 | 4.88 | 9.86 | 9.86 | 21.39 | 9.86 | 20.22 | 9.67 | 9.67 | 0.1 | 11.23 | 0.1 | 7.42 | 0.59 | 4.79 | 4.69 | 4.88 | 21.29 | 4.79 | 25.78 | 9.77 | 28.03 | 9.77 | 28.52 | 9.77
0.992 | 1.27 | 0.1 | 14.45 | 0 | 14.84 | 0.1 | 14.55 | 4.88 | 12.01 | 4.59 | 4.79 | 4.59 | 4.88 | 9.86 | 9.86 | 21.49 | 9.86 | 20.51 | 9.67 | 10.45 | 0.1 | 12.01 | 0.1 | 8.11 | 0.59 | 4.79 | 6.15 | 4.79 | 21.29 | 4.79 | 26.27 | 9.77 | 28.42 | 9.77 | 29.01 | 9.77
1.024 | 1.17 | 0.1 | 14.45 | 0 | 14.75 | 0.1 | 15.14 | 4.88 | 12.6 | 4.59 | 4.79 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.86 | 20.9 | 9.67 | 11.33 | 0.1 | 12.79 | 0.1 | 8.98 | 0.49 | 4.79 | 6.15 | 4.88 | 21.19 | 4.79 | 23.24 | 9.77 | 28.91 | 9.77 | 29.49 | 9.77
1.056 | 1.07 | 0.1 | 14.45 | 0.1 | 14.75 | 0 | 14.75 | 4.79 | 13.18 | 4.69 | 4.79 | 4.59 | 4.88 | 9.77 | 9.77 | 21.49 | 9.86 | 21.29 | 9.67 | 12.11 | 0.1 | 13.48 | 0.1 | 9.77 | 0.29 | 4.79 | 5.66 | 4.88 | 21.19 | 4.79 | 22.36 | 9.77 | 29.49 | 9.77 | 25.59 | 9.77
1.088 | 0.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0 | 14.55 | 4.79 | 13.77 | 4.69 | 4.88 | 4.49 | 4.88 | 9.77 | 9.86 | 21.49 | 9.86 | 21.78 | 9.67 | 12.7 | 0.1 | 14.26 | 0.1 | 10.55 | 0.29 | 4.79 | 5.27 | 4.88 | 21.19 | 4.79 | 21.88 | 9.77 | 25.2 | 9.77 | 24.81 | 9.77
1.12 | 0.88 | 0.1 | 14.45 | 0 | 14.55 | 0.1 | 14.65 | 4.88 | 14.26 | 4.69 | 5.57 | 4.59 | 4.88 | 9.77 | 9.86 | 21.39 | 9.86 | 22.17 | 9.77 | 13.48 | 0.1 | 15.04 | 0.1 | 11.33 | 0.29 | 4.79 | 5.37 | 4.88 | 11.72 | 4.79 | 21.68 | 9.77 | 24.41 | 9.77 | 24.71 | 9.77
1.152 | 0.78 | 0.1 | 14.36 | 0 | 14.55 | 0.1 | 14.55 | 4.88 | 14.75 | 4.69 | 6.25 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 22.46 | 9.77 | 14.26 | 0.1 | 15.72 | 0.1 | 12.01 | 0.29 | 4.79 | 5.18 | 4.88 | 2.93 | 4.79 | 21.58 | 9.77 | 24.51 | 9.77 | 24.41 | 9.77
1.184 | 0.59 | 0.1 | 14.16 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 15.33 | 4.69 | 7.42 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.95 | 9.77 | 14.94 | 0.1 | 16.5 | 0.1 | 12.79 | 0.2 | 4.79 | 5.18 | 4.88 | 3.13 | 4.79 | 21.49 | 9.77 | 24.32 | 9.77 | 24.51 | 9.86
1.216 | 0.59 | 0 | 3.81 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 15.72 | 4.79 | 8.59 | 4.49 | 4.88 | 9.67 | 9.77 | 21.39 | 9.86 | 23.34 | 9.67 | 15.63 | 0.1 | 16.99 | 0.1 | 13.77 | 0.2 | 5.86 | 4.98 | 4.88 | 3.52 | 4.79 | 21.39 | 9.77 | 24.41 | 9.77 | 24.41 | 9.86
1.248 | 0.49 | 0.1 | 1.86 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 16.21 | 4.79 | 9.47 | 4.49 | 4.88 | 9.67 | 9.77 | 10.64 | 9.86 | 23.73 | 9.77 | 16.31 | 0.1 | 17.68 | 0.1 | 14.26 | 0.1 | 7.13 | 4.98 | 4.88 | 4.49 | 4.79 | 21.39 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77
1.28 | 0.49 | 0.1 | 1.86 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 16.7 | 4.79 | 10.16 | 4.49 | 4.88 | 9.67 | 9.86 | 6.74 | 9.86 | 24.02 | 9.77 | 16.9 | 0.1 | 18.36 | 0.1 | 15.04 | 0.2 | 9.08 | 4.98 | 4.88 | 6.05 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.41 | 9.86
1.312 | 0.39 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.19 | 4.79 | 10.64 | 4.49 | 5.57 | 9.67 | 9.86 | 9.47 | 9.86 | 24.41 | 9.77 | 17.68 | 0.1 | 18.95 | 0.1 | 15.63 | 0.1 | 10.74 | 4.88 | 4.79 | 6.15 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
1.344 | 0.39 | 0.1 | 1.46 | 0 | 10.35 | 0.1 | 14.45 | 4.88 | 17.68 | 4.79 | 11.33 | 4.59 | 6.64 | 9.67 | 9.86 | 11.52 | 9.86 | 22.27 | 9.77 | 18.36 | 0.1 | 19.63 | 0.1 | 16.5 | 0.1 | 11.91 | 4.79 | 4.88 | 5.66 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
1.376 | 0.29 | 0 | 1.27 | 0.1 | 2.93 | 0 | 14.36 | 4.88 | 18.07 | 4.79 | 11.91 | 4.59 | 7.91 | 9.67 | 9.77 | 10.94 | 9.86 | 22.27 | 9.77 | 18.95 | 0.1 | 20.31 | 0.1 | 17.09 | 0.1 | 12.89 | 4.69 | 4.88 | 5.27 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
1.408 | 0.2 | 0.1 | 1.17 | 0.1 | 2.15 | 0 | 11.91 | 4.88 | 18.56 | 4.79 | 12.6 | 4.59 | 8.98 | 9.57 | 9.86 | 10.16 | 9.86 | 21.88 | 9.77 | 19.73 | 0.1 | 20.9 | 0.1 | 17.87 | 0.1 | 13.87 | 4.69 | 4.88 | 5.47 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
1.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.66 | 0.1 | 3.52 | 4.88 | 18.95 | 4.79 | 13.09 | 4.69 | 9.77 | 9.57 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 20.22 | 0.1 | 21.49 | 0.1 | 18.46 | 0.1 | 14.45 | 4.79 | 4.88 | 5.18 | 4.79 | 11.72 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
1.472 | 0.2 | 0.1 | 0.88 | 0 | 1.46 | 0.1 | 2.15 | 4.88 | 19.43 | 4.79 | 13.57 | 4.69 | 10.45 | 9.67 | 9.86 | 10.16 | 9.86 | 21.78 | 9.77 | 20.9 | 0.1 | 21.97 | 0.1 | 19.14 | 0.1 | 15.23 | 4.69 | 4.88 | 5.08 | 4.79 | 2.93 | 9.77 | 24.22 | 9.77 | 19.43 | 9.86
1.504 | 0.1 | 0.1 | 0.88 | 0 | 1.56 | 0.1 | 1.86 | 4.88 | 19.92 | 4.79 | 14.06 | 4.69 | 10.94 | 9.67 | 9.86 | 10.35 | 9.86 | 21.58 | 9.86 | 21.49 | 0.1 | 22.66 | 0.1 | 19.82 | 0.1 | 15.92 | 4.59 | 4.88 | 5.08 | 4.79 | 3.13 | 9.77 | 19.53 | 9.77 | 7.62 | 9.86
1.536 | 0 | 0 | 0.68 | 0.1 | 1.37 | 0 | 1.37 | 4.88 | 20.22 | 4.79 | 14.65 | 4.69 | 11.62 | 9.67 | 9.86 | 10.06 | 9.86 | 21.58 | 9.77 | 21.97 | 0.1 | 19.24 | 0.1 | 20.31 | 0 | 16.7 | 4.59 | 4.88 | 4.98 | 4.79 | 3.52 | 9.77 | 7.81 | 9.77 | 6.74 | 9.86
1.568 | 0.1 | 0.68 | 0.68 | 0.1 | 1.27 | 0 | 1.46 | 4.88 | 20.7 | 4.79 | 15.14 | 4.79 | 12.21 | 9.67 | 10.45 | 9.86 | 9.86 | 21.49 | 9.77 | 22.46 | 0.1 | 18.46 | 0.1 | 20.8 | 0 | 17.19 | 4.59 | 4.88 | 4.88 | 4.79 | 4.69 | 9.77 | 6.64 | 9.77 | 10.74 | 9.86
1.6 | 0.1 | 1.27 | 0.68 | 0.1 | 1.07 | 0.1 | 1.27 | 4.88 | 20.22 | 4.88 | 15.63 | 4.79 | 12.79 | 9.67 | 11.43 | 9.86 | 9.86 | 21.49 | 9.86 | 19.43 | 0 | 18.16 | 0.1 | 21.49 | 0.1 | 17.97 | 4.59 | 4.79 | 4.88 | 4.79 | 6.15 | 9.77 | 10.45 | 9.77 | 12.21 | 9.86
1.632 | 0.1 | 2.44 | 0.59 | 0 | 0.88 | 0.1 | 1.17 | 4.88 | 18.56 | 4.88 | 16.11 | 4.79 | 13.28 | 9.67 | 12.7 | 9.86 | 9.96 | 21.49 | 9.86 | 18.56 | 0.1 | 17.97 | 0.1 | 22.07 | 0.1 | 18.65 | 4.59 | 4.79 | 4.79 | 4.79 | 6.15 | 9.77 | 12.11 | 9.86 | 10.74 | 9.86
1.664 | 0.1 | 3.71 | 0.59 | 0 | 0.88 | 0 | 1.07 | 4.88 | 18.56 | 4.79 | 16.6 | 4.79 | 13.87 | 9.67 | 13.67 | 9.86 | 9.86 | 21.49 | 9.86 | 18.07 | 0.1 | 17.97 | 0.1 | 22.66 | 0.1 | 19.24 | 4.59 | 5.37 | 4.79 | 4.79 | 5.76 | 9.77 | 10.74 | 9.77 | 10.16 | 9.86
1.696 | 0 | 4.59 | 0.39 | 0.1 | 0.68 | 0 | 0.98 | 4.88 | 18.46 | 4.88 | 17.19 | 4.79 | 14.36 | 9.67 | 14.45 | 9.77 | 9.86 | 21.39 | 9.77 | 17.87 | 0.1 | 17.87 | 0.1 | 19.53 | 0.1 | 19.92 | 4.49 | 6.35 | 4.69 | 4.79 | 5.37 | 9.77 | 10.25 | 9.77 | 10.06 | 9.86
1.728 | 0.1 | 5.08 | 0.39 | 0.1 | 0.68 | 0.1 | 0.78 | 4.88 | 18.26 | 4.88 | 17.58 | 4.79 | 14.84 | 9.67 | 14.94 | 9.77 | 9.86 | 21.49 | 9.86 | 17.87 | 0.1 | 17.77 | 0.1 | 18.56 | 0.1 | 20.7 | 4.49 | 8.01 | 4.69 | 4.79 | 5.08 | 9.77 | 10.16 | 9.77 | 10.55 | 9.86
1.76 | 0.1 | 5.57 | 0.29 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 18.16 | 4.88 | 17.97 | 4.79 | 15.43 | 9.67 | 15.63 | 9.67 | 9.86 | 15.14 | 9.86 | 17.77 | 0.1 | 17.77 | 0.1 | 18.16 | 0.1 | 21.09 | 4.59 | 10.16 | 4.69 | 4.79 | 5.18 | 9.77 | 10.45 | 9.77 | 10.55 | 9.86
1.792 | 0.1 | 6.15 | 0.29 | 0 | 0.59 | 0.1 | 0.68 | 4.88 | 18.16 | 4.79 | 18.46 | 4.79 | 15.82 | 9.67 | 16.11 | 9.67 | 9.86 | 7.62 | 9.86 | 17.68 | 0 | 17.68 | 0.1 | 17.97 | 0.1 | 21.68 | 4.59 | 11.43 | 4.69 | 4.79 | 5.27 | 9.77 | 10.45 | 9.77 | 10.16 | 9.86
1.824 | 0.1 | 6.74 | 0.2 | 0 | 0.49 | 0.1 | 0.59 | 4.88 | 18.16 | 4.88 | 18.95 | 4.79 | 16.31 | 9.77 | 16.6 | 9.67 | 9.86 | 7.72 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 22.17 | 4.69 | 12.4 | 4.59 | 4.79 | 5.08 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
1.856 | 0.1 | 7.42 | 0.1 | 0.1 | 0.49 | 0 | 0.49 | 4.88 | 18.16 | 4.88 | 19.34 | 4.88 | 16.8 | 9.77 | 17.09 | 9.67 | 9.86 | 10.55 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 22.75 | 4.59 | 13.28 | 4.59 | 4.79 | 4.98 | 9.77 | 10.25 | 9.77 | 10.16 | 9.86
1.888 | 0.1 | 8.01 | 0.1 | 0.1 | 0.29 | 0 | 0.49 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 17.29 | 9.77 | 17.58 | 9.67 | 9.86 | 11.43 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 23.34 | 4.59 | 14.16 | 4.59 | 4.79 | 4.98 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
1.92 | 0.1 | 8.59 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 17.77 | 9.77 | 18.07 | 9.67 | 9.86 | 10.55 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 23.83 | 4.69 | 14.94 | 4.59 | 4.79 | 4.98 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
1.952 | 0.1 | 9.18 | 0.1 | 0.2 | 0.29 | 0.1 | 0.29 | 4.88 | 18.07 | 4.79 | 20.61 | 4.79 | 18.16 | 9.77 | 18.46 | 9.67 | 9.86 | 10.16 | 9.86 | 17.58 | 0 | 12.89 | 0.1 | 17.68 | 0.1 | 24.41 | 4.79 | 15.63 | 4.49 | 4.79 | 4.88 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
1.984 | 0.1 | 9.77 | 0.1 | 1.07 | 0.2 | 0.1 | 0.29 | 4.88 | 18.07 | 4.79 | 21 | 4.88 | 18.56 | 9.77 | 18.95 | 9.67 | 9.86 | 9.77 | 9.86 | 17.58 | 0.1 | 3.13 | 0.1 | 17.77 | 0.1 | 24.9 | 4.79 | 16.31 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
2.016 | 0.1 | 10.45 | 0.1 | 1.86 | 0.2 | 0.78 | 0.2 | 4.88 | 18.07 | 4.88 | 19.14 | 4.88 | 19.14 | 9.77 | 19.34 | 9.67 | 9.86 | 10.16 | 9.86 | 10.35 | 0.1 | 2.25 | 0.1 | 17.68 | 0 | 25.59 | 4.69 | 16.99 | 4.59 | 4.79 | 4.79 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
2.048 | 0.1 | 11.04 | 0 | 3.22 | 0.1 | 1.56 | 0.2 | 4.88 | 8.89 | 4.88 | 18.85 | 4.88 | 19.43 | 9.77 | 19.82 | 9.57 | 9.86 | 10.06 | 9.86 | 2.64 | 0.1 | 1.95 | 0.1 | 17.68 | 0.1 | 25.98 | 4.79 | 17.68 | 4.59 | 4.79 | 4.79 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
2.08 | 0.1 | 11.62 | 0.1 | 4.1 | 0.1 | 2.83 | 0.2 | 4.88 | 2.83 | 4.88 | 18.56 | 4.88 | 19.92 | 9.77 | 20.22 | 9.57 | 9.86 | 9.86 | 9.86 | 2.15 | 0.1 | 1.66 | 0.1 | 17.68 | 0.1 | 24.51 | 4.79 | 18.36 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
2.112 | 0.1 | 12.11 | 0.1 | 4.79 | 0 | 3.91 | 0.1 | 4.88 | 3.22 | 4.79 | 18.36 | 4.79 | 20.31 | 9.77 | 20.7 | 9.57 | 10.74 | 9.77 | 9.96 | 2.05 | 0.1 | 1.76 | 0.1 | 12.89 | 0.1 | 21.09 | 4.79 | 18.85 | 4.59 | 5.76 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
2.144 | 0.1 | 12.7 | 0.1 | 5.37 | 0.1 | 4.59 | 0.1 | 4.88 | 4 | 4.88 | 18.46 | 4.88 | 20.8 | 9.77 | 21 | 9.57 | 11.13 | 9.86 | 9.96 | 1.66 | 0.1 | 1.46 | 0.1 | 3.03 | 0.1 | 21.49 | 4.79 | 19.53 | 4.59 | 6.35 | 4.59 | 9.96 | 9.77 | 9.86 | 9.86 | 9.86
2.176 | 0.1 | 13.18 | 0.1 | 5.86 | 0.1 | 5.18 | 0 | 4.88 | 5.18 | 4.88 | 18.26 | 4.88 | 19.14 | 9.77 | 21.39 | 9.67 | 12.21 | 9.86 | 9.86 | 1.86 | 0.1 | 1.37 | 0.1 | 2.34 | 0.1 | 21.29 | 4.79 | 20.12 | 4.59 | 7.81 | 4.59 | 10.94 | 9.77 | 10.64 | 9.77 | 9.86
2.208 | 0.1 | 13.77 | 0.1 | 6.45 | 0.1 | 5.76 | 0 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 17.68 | 9.77 | 21.78 | 9.67 | 13.38 | 9.77 | 9.86 | 1.46 | 0.1 | 1.27 | 0.1 | 2.05 | 0 | 21.29 | 4.79 | 20.8 | 4.69 | 10.06 | 4.59 | 12.4 | 9.77 | 11.33 | 9.77 | 10.74
2.24 | 0.1 | 14.26 | 0.1 | 7.13 | 0.1 | 6.35 | 0.1 | 4.88 | 5.57 | 4.88 | 18.16 | 4.88 | 17.97 | 9.77 | 22.27 | 9.67 | 14.26 | 9.67 | 9.96 | 1.46 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 21.29 | 4.88 | 21.49 | 4.69 | 11.43 | 4.59 | 14.55 | 9.77 | 12.79 | 9.67 | 11.52
2.272 | 0.1 | 14.84 | 0.1 | 7.62 | 0 | 6.93 | 0.1 | 4.88 | 5.27 | 4.79 | 18.07 | 4.88 | 18.16 | 9.77 | 22.56 | 9.67 | 14.84 | 9.67 | 9.96 | 1.46 | 0.1 | 1.07 | 0.1 | 1.76 | 0.1 | 21.29 | 4.88 | 21.97 | 4.69 | 12.4 | 4.49 | 15.92 | 9.67 | 14.94 | 9.67 | 12.99
2.304 | 0.1 | 15.33 | 0.1 | 8.2 | 0.1 | 7.52 | 0.1 | 4.88 | 4.98 | 4.79 | 18.07 | 4.88 | 17.97 | 9.77 | 23.05 | 9.67 | 15.33 | 9.67 | 9.96 | 1.07 | 0.1 | 1.07 | 0.1 | 1.56 | 0.1 | 21.29 | 4.88 | 22.56 | 4.69 | 13.28 | 4.59 | 16.99 | 9.67 | 16.21 | 9.67 | 15.04
2.336 | 0 | 15.82 | 0.1 | 8.89 | 0.1 | 8.11 | 0 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 23.34 | 9.67 | 15.92 | 9.67 | 9.86 | 1.17 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 21.19 | 4.79 | 23.05 | 4.69 | 14.16 | 4.59 | 17.87 | 9.67 | 17.29 | 9.67 | 16.31
2.368 | 0.1 | 16.21 | 0 | 9.57 | 0.1 | 8.69 | 0 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 23.73 | 9.67 | 16.41 | 9.67 | 9.86 | 0.98 | 0.1 | 0.88 | 0.1 | 1.27 | 0.1 | 21.19 | 4.88 | 23.54 | 4.79 | 14.94 | 4.59 | 18.46 | 9.67 | 17.87 | 9.67 | 17.19
2.4 | 0.1 | 16.8 | 0.1 | 10.16 | 0.1 | 9.38 | 0.1 | 5.47 | 4.79 | 4.88 | 18.07 | 4.88 | 18.07 | 9.86 | 24.12 | 9.67 | 16.9 | 9.67 | 9.96 | 1.07 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 21.19 | 4.88 | 24.22 | 4.79 | 15.53 | 4.59 | 19.14 | 9.67 | 18.46 | 9.67 | 18.07
2.432 | 0.1 | 17.19 | 0.1 | 10.64 | 0.1 | 9.96 | 0.1 | 6.45 | 4.79 | 4.88 | 10.64 | 4.88 | 18.07 | 9.86 | 23.73 | 9.67 | 17.38 | 9.67 | 9.96 | 0.88 | 0.1 | 0.59 | 0.1 | 1.07 | 0.1 | 21.19 | 4.88 | 24.71 | 4.79 | 16.21 | 4.59 | 19.82 | 9.67 | 19.24 | 9.67 | 18.65
2.464 | 0.1 | 14.75 | 0.1 | 11.33 | 0.1 | 10.55 | 0.1 | 7.72 | 4.79 | 4.79 | 3.32 | 4.88 | 18.07 | 9.77 | 21.58 | 9.67 | 17.77 | 9.67 | 10.64 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 0.1 | 21.19 | 4.88 | 25.2 | 4.79 | 16.99 | 4.59 | 20.41 | 9.57 | 19.92 | 9.67 | 19.24
2.496 | 0.1 | 15.23 | 0.1 | 11.82 | 0.1 | 11.13 | 0.1 | 8.89 | 4.79 | 4.88 | 3.13 | 4.88 | 18.07 | 9.77 | 21.49 | 9.67 | 18.26 | 9.57 | 11.33 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 0.1 | 21.19 | 4.88 | 25.68 | 4.79 | 17.58 | 4.69 | 21 | 9.67 | 20.51 | 9.67 | 20.02
2.528 | 0.1 | 14.75 | 0 | 12.4 | 0.1 | 11.72 | 0 | 9.67 | 4.69 | 4.88 | 3.81 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 18.75 | 9.57 | 12.6 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 0.1 | 8.01 | 4.88 | 26.17 | 4.79 | 18.16 | 4.69 | 21.58 | 9.67 | 21.09 | 9.67 | 20.61
2.56 | 0.1 | 14.65 | 0.1 | 12.99 | 0.1 | 12.31 | 0.1 | 10.35 | 4.59 | 4.79 | 5.08 | 4.88 | 18.07 | 9.77 | 21.58 | 9.67 | 19.24 | 9.57 | 13.67 | 0.49 | 0 | 0.39 | 0.1 | 0.78 | 0.1 | 2.25 | 4.88 | 22.27 | 4.79 | 18.85 | 4.69 | 22.17 | 9.67 | 21.78 | 9.67 | 21.19
2.592 | 0.1 | 14.65 | 0.1 | 13.48 | 0.1 | 12.89 | 0.1 | 10.94 | 4.59 | 4.79 | 5.86 | 4.88 | 17.97 | 9.77 | 21.58 | 9.77 | 19.63 | 9.57 | 14.45 | 0.49 | 0.1 | 0.29 | 0.1 | 0.68 | 0.1 | 3.22 | 4.88 | 20.8 | 4.79 | 19.53 | 4.69 | 22.75 | 9.67 | 22.27 | 9.67 | 21.78
2.624 | 0.1 | 14.55 | 0.1 | 13.97 | 0.1 | 13.38 | 0.1 | 11.62 | 4.59 | 4.79 | 5.76 | 4.88 | 6.93 | 9.77 | 21.49 | 9.77 | 20.12 | 9.57 | 15.04 | 0.39 | 0.88 | 0.29 | 0.1 | 0.59 | 0.1 | 3.91 | 4.88 | 21.29 | 4.79 | 20.12 | 4.79 | 23.24 | 9.67 | 22.85 | 9.67 | 22.36
2.656 | 0 | 14.55 | 0.1 | 14.55 | 0.1 | 13.87 | 0 | 12.21 | 4.59 | 4.88 | 5.27 | 4.88 | 2.64 | 9.77 | 21.39 | 9.77 | 20.51 | 9.57 | 15.53 | 0.39 | 1.76 | 0.2 | 0.2 | 0.59 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 20.8 | 4.79 | 23.83 | 9.67 | 23.44 | 9.67 | 22.95
2.688 | 0.1 | 14.45 | 0 | 15.04 | 0.1 | 14.36 | 0.1 | 12.7 | 4.59 | 4.88 | 5.08 | 4.88 | 3.32 | 9.77 | 21.49 | 9.77 | 20.9 | 9.57 | 16.02 | 0.29 | 3.42 | 0.2 | 1.56 | 0.49 | 0.1 | 6.15 | 4.88 | 21.29 | 4.79 | 21.29 | 4.79 | 24.41 | 9.67 | 23.93 | 9.67 | 23.34
2.72 | 0.1 | 14.45 | 0.1 | 15.53 | 0 | 14.94 | 0.1 | 13.28 | 4.49 | 4.88 | 4.88 | 4.88 | 4.39 | 9.86 | 21.49 | 9.77 | 21.29 | 9.57 | 16.6 | 0.29 | 5.27 | 0.1 | 2.73 | 0.49 | 0.1 | 6.05 | 4.88 | 21.39 | 4.79 | 21.88 | 4.79 | 24.9 | 9.67 | 24.51 | 9.77 | 23.83
2.752 | 0.1 | 14.45 | 0.1 | 16.11 | 0 | 15.43 | 0.1 | 13.77 | 4.49 | 4.79 | 4.88 | 4.88 | 5.47 | 9.86 | 21.39 | 9.77 | 21.68 | 9.67 | 17.09 | 0.2 | 6.45 | 0.1 | 4.88 | 0.39 | 0.1 | 5.57 | 4.88 | 21.29 | 4.88 | 22.46 | 4.79 | 25.39 | 9.67 | 25 | 9.67 | 24.32
2.784 | 0.1 | 14.36 | 0.1 | 16.5 | 0.1 | 15.92 | 0.1 | 14.36 | 4.49 | 5.47 | 4.88 | 4.88 | 5.96 | 9.86 | 21.39 | 9.77 | 22.17 | 9.67 | 17.58 | 0.2 | 7.23 | 0.1 | 6.25 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.79 | 23.05 | 4.79 | 25.88 | 9.67 | 25.49 | 9.77 | 24.9
2.816 | 0.1 | 14.36 | 0.1 | 16.99 | 0.1 | 16.41 | 0 | 14.94 | 4.49 | 6.25 | 4.79 | 4.88 | 5.57 | 9.77 | 21.39 | 9.77 | 22.46 | 9.67 | 17.97 | 0.2 | 8.11 | 0.1 | 7.13 | 0.29 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 23.54 | 4.88 | 26.37 | 9.77 | 25.88 | 9.77 | 25.49
2.848 | 0.1 | 14.36 | 0.1 | 15.53 | 0.1 | 16.9 | 0.1 | 15.33 | 4.59 | 7.32 | 4.88 | 4.88 | 5.18 | 9.86 | 21.39 | 9.77 | 22.85 | 9.67 | 18.46 | 0.2 | 8.89 | 0.1 | 7.91 | 0.2 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 24.12 | 4.88 | 26.86 | 9.77 | 26.47 | 9.77 | 25.78
2.88 | 0.1 | 11.33 | 0.1 | 14.06 | 0.1 | 17.38 | 0.1 | 15.92 | 4.49 | 8.59 | 4.69 | 4.88 | 4.98 | 9.86 | 12.6 | 9.77 | 23.24 | 9.67 | 18.85 | 0.1 | 9.67 | 0 | 8.59 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.88 | 24.61 | 4.88 | 27.34 | 9.77 | 26.86 | 9.77 | 26.37
2.912 | 0.1 | 3.03 | 0.1 | 14.16 | 0.1 | 15.23 | 0.1 | 16.41 | 4.59 | 9.47 | 4.59 | 4.88 | 4.88 | 9.77 | 6.84 | 9.77 | 23.63 | 9.67 | 19.34 | 0.1 | 10.45 | 0 | 9.47 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.88 | 25.29 | 4.79 | 28.03 | 9.77 | 27.44 | 9.77 | 26.86
2.944 | 0.1 | 1.95 | 0.1 | 14.55 | 0.1 | 15.04 | 0.1 | 16.8 | 4.59 | 10.16 | 4.69 | 4.88 | 4.88 | 9.77 | 8.4 | 9.86 | 23.93 | 9.67 | 19.73 | 0.1 | 11.23 | 0.1 | 10.25 | 0.1 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 25.68 | 4.79 | 28.42 | 9.77 | 27.93 | 9.77 | 27.34
2.976 | 0 | 1.76 | 0.1 | 14.45 | 0.1 | 14.75 | 0 | 17.38 | 4.69 | 10.74 | 4.59 | 4.88 | 4.79 | 9.77 | 11.33 | 9.86 | 24.32 | 9.67 | 20.22 | 0.1 | 11.91 | 0.1 | 11.04 | 0.1 | 0.1 | 4.88 | 4.88 | 19.82 | 4.79 | 26.17 | 4.88 | 28.91 | 9.77 | 28.42 | 9.77 | 28.03
3.008 | 0.1 | 1.56 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 17.77 | 4.69 | 11.43 | 4.59 | 4.88 | 4.79 | 9.77 | 11.13 | 9.86 | 21.29 | 9.67 | 20.61 | 0.1 | 12.7 | 0 | 11.82 | 0.1 | 0.1 | 4.79 | 4.88 | 4.49 | 4.88 | 21.78 | 4.88 | 29.3 | 9.77 | 28.91 | 9.86 | 28.42
3.04 | 0.1 | 1.56 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 18.26 | 4.69 | 12.01 | 4.49 | 4.88 | 4.79 | 9.77 | 10.35 | 9.86 | 22.17 | 9.67 | 21 | 0.1 | 13.48 | 0 | 12.6 | 0.1 | 0.1 | 4.79 | 4.88 | 3.03 | 4.79 | 21.29 | 4.88 | 24.61 | 9.77 | 29.3 | 9.86 | 28.91
3.072 | 0.1 | 1.27 | 0.1 | 14.36 | 0 | 14.55 | 0.1 | 18.65 | 4.69 | 12.6 | 4.49 | 4.88 | 4.79 | 9.86 | 9.86 | 9.86 | 21.49 | 9.67 | 21.39 | 0.1 | 14.06 | 0 | 13.38 | 0.1 | 0.1 | 4.69 | 4.88 | 3.52 | 4.88 | 21.29 | 4.88 | 23.34 | 9.77 | 24.81 | 9.77 | 29.3
3.104 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.14 | 4.69 | 12.99 | 4.59 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 21.78 | 0.1 | 14.84 | 0 | 14.16 | 0.1 | 0.2 | 4.69 | 4.88 | 4.39 | 4.88 | 21.19 | 4.88 | 24.02 | 9.77 | 23.73 | 9.86 | 25.1
3.136 | 0 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.53 | 4.79 | 13.57 | 4.59 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.68 | 9.77 | 22.17 | 0.1 | 15.63 | 0.1 | 14.84 | 0.1 | 1.46 | 4.59 | 4.88 | 5.66 | 4.79 | 21.29 | 4.88 | 24.22 | 9.77 | 24.41 | 9.86 | 24.81
3.168 | 0.1 | 0.88 | 0 | 14.36 | 0.1 | 14.45 | 0.1 | 19.92 | 4.79 | 14.16 | 4.49 | 5.66 | 4.59 | 9.77 | 10.06 | 9.86 | 21.49 | 9.77 | 22.56 | 0.1 | 16.41 | 0.1 | 15.53 | 0 | 2.64 | 4.59 | 4.88 | 6.25 | 4.88 | 21.29 | 4.88 | 24.22 | 9.77 | 24.12 | 9.86 | 24.32
3.2 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 20.41 | 4.69 | 14.65 | 4.49 | 6.25 | 4.59 | 9.77 | 9.96 | 9.86 | 21.58 | 9.67 | 22.95 | 0.1 | 17.09 | 0 | 16.31 | 0.1 | 4.49 | 4.59 | 4.88 | 5.76 | 4.88 | 21.29 | 4.88 | 24.22 | 9.77 | 24.32 | 9.86 | 24.32
3.232 | 0.1 | 0.78 | 0.1 | 14.36 | 0 | 14.36 | 0.1 | 20.8 | 4.79 | 15.23 | 4.49 | 7.62 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 23.34 | 0.1 | 17.68 | 0 | 16.99 | 0.1 | 5.96 | 4.59 | 4.88 | 5.47 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.32
3.264 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 20.31 | 4.79 | 15.63 | 4.59 | 8.69 | 4.49 | 10.45 | 9.86 | 9.86 | 21.49 | 9.77 | 23.73 | 0.1 | 18.36 | 0.1 | 17.48 | 0.1 | 7.03 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.32
3.296 | 0 | 0.59 | 0.1 | 6.15 | 0.1 | 14.36 | 0 | 17.87 | 4.79 | 16.11 | 4.59 | 9.57 | 4.49 | 11.23 | 9.86 | 9.86 | 21.49 | 9.77 | 24.12 | 0.1 | 19.04 | 0 | 18.26 | 0.1 | 7.81 | 4.49 | 4.88 | 5.08 | 4.79 | 21.29 | 4.88 | 24.22 | 9.86 | 24.32 | 9.86 | 24.22
3.328 | 0.1 | 0.49 | 0 | 1.95 | 0.1 | 8.01 | 0 | 18.36 | 4.79 | 16.6 | 4.59 | 10.25 | 4.59 | 12.31 | 9.77 | 9.86 | 21.49 | 9.77 | 24.41 | 0.1 | 19.53 | 0 | 18.95 | 0 | 8.4 | 4.49 | 4.88 | 5.27 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.32
3.36 | 0.1 | 0.49 | 0.1 | 1.86 | 0.1 | 2.25 | 0.1 | 18.46 | 4.79 | 17.09 | 4.59 | 10.84 | 4.49 | 13.48 | 9.77 | 9.86 | 21.49 | 9.77 | 22.46 | 0.1 | 20.31 | 0 | 19.43 | 0.1 | 9.47 | 4.49 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
3.392 | 0.1 | 0.39 | 0.1 | 1.76 | 0.1 | 1.95 | 0.1 | 18.26 | 4.79 | 17.58 | 4.59 | 11.43 | 4.49 | 14.26 | 9.77 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 20.8 | 0 | 20.12 | 0.1 | 10.25 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
3.424 | 0.1 | 0.29 | 0.1 | 1.37 | 0.1 | 1.66 | 0.1 | 18.26 | 4.88 | 18.07 | 4.69 | 12.01 | 4.49 | 14.84 | 9.77 | 9.86 | 19.82 | 9.77 | 21.88 | 0.1 | 21.39 | 0 | 20.8 | 0.1 | 10.94 | 4.49 | 4.88 | 4.88 | 4.79 | 19.53 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
3.456 | 0.1 | 0.29 | 0.1 | 1.46 | 0.1 | 1.37 | 0 | 18.16 | 4.88 | 18.46 | 4.69 | 12.6 | 4.59 | 15.43 | 9.67 | 9.86 | 8.2 | 9.77 | 21.68 | 0.1 | 21.97 | 0 | 21.39 | 0 | 11.72 | 4.49 | 4.88 | 4.88 | 4.79 | 4.3 | 4.88 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
3.488 | 0.68 | 0.2 | 0 | 1.17 | 0.1 | 1.46 | 0 | 18.16 | 4.88 | 18.95 | 4.69 | 13.09 | 4.59 | 15.92 | 9.67 | 9.86 | 7.13 | 9.77 | 21.68 | 0.1 | 22.56 | 0.1 | 21.88 | 0 | 12.4 | 4.49 | 4.88 | 4.79 | 4.88 | 3.03 | 4.88 | 10.45 | 9.86 | 23.63 | 9.86 | 24.12
3.52 | 1.37 | 0.2 | 0.1 | 1.17 | 0 | 1.37 | 0.1 | 18.16 | 4.79 | 19.34 | 4.69 | 13.67 | 4.59 | 16.41 | 9.67 | 9.86 | 10.25 | 9.77 | 21.58 | 0.1 | 19.34 | 0 | 22.46 | 0.1 | 13.28 | 4.59 | 5.76 | 4.79 | 4.79 | 3.52 | 4.88 | 5.86 | 9.77 | 8.11 | 9.86 | 20.7
3.552 | 2.64 | 0.1 | 0.1 | 0.98 | 0 | 1.07 | 0.1 | 18.16 | 4.88 | 19.82 | 4.69 | 14.16 | 4.59 | 16.9 | 9.67 | 9.86 | 11.52 | 9.77 | 21.58 | 0.1 | 18.46 | 0 | 17.77 | 0.1 | 13.97 | 4.59 | 6.84 | 4.69 | 4.79 | 4.2 | 4.88 | 9.28 | 9.77 | 6.35 | 9.86 | 7.72
3.584 | 3.81 | 0.1 | 0.1 | 0.88 | 0.1 | 0.98 | 0.1 | 18.07 | 4.88 | 20.12 | 4.79 | 14.75 | 4.59 | 17.38 | 9.67 | 9.86 | 10.64 | 9.86 | 21.58 | 0 | 18.07 | 0.1 | 16.8 | 0.1 | 14.75 | 4.59 | 8.5 | 4.69 | 4.79 | 5.66 | 4.88 | 12.31 | 9.77 | 10.06 | 9.86 | 6.54
3.616 | 4.59 | 0 | 0.1 | 0.78 | 0.1 | 0.88 | 0 | 18.16 | 4.88 | 20.61 | 4.79 | 15.23 | 4.69 | 17.77 | 9.57 | 9.86 | 9.96 | 9.86 | 21.49 | 0.1 | 17.87 | 0 | 17.48 | 0.1 | 15.53 | 4.59 | 10.45 | 4.69 | 4.79 | 6.25 | 4.88 | 11.04 | 9.86 | 12.21 | 9.86 | 10.64
3.648 | 5.08 | 0.1 | 0 | 0.68 | 0.1 | 0.78 | 0.1 | 18.07 | 4.88 | 21 | 4.79 | 15.72 | 4.69 | 18.36 | 9.57 | 9.86 | 9.96 | 9.86 | 21.49 | 0.1 | 17.97 | 0 | 17.68 | 0 | 16.11 | 4.59 | 11.72 | 4.59 | 4.79 | 5.76 | 4.88 | 10.16 | 9.77 | 10.94 | 9.86 | 12.11
3.68 | 5.66 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 18.07 | 4.88 | 19.04 | 4.79 | 16.21 | 4.69 | 18.75 | 9.57 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 17.77 | 0 | 17.68 | 0.1 | 16.8 | 4.69 | 12.79 | 4.59 | 4.79 | 5.27 | 4.88 | 10.06 | 9.77 | 10.06 | 9.86 | 11.04
3.712 | 6.25 | 0.1 | 0.1 | 0.49 | 0 | 0.68 | 0.1 | 8.4 | 4.88 | 18.85 | 4.79 | 16.7 | 4.79 | 19.24 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 17.58 | 4.69 | 13.57 | 4.59 | 4.79 | 5.08 | 4.88 | 10.25 | 9.77 | 10.16 | 9.86 | 10.16
3.744 | 6.84 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.73 | 4.88 | 18.56 | 4.88 | 17.19 | 4.69 | 19.73 | 9.67 | 9.86 | 10.06 | 9.86 | 21.49 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 18.16 | 4.69 | 14.36 | 4.59 | 4.79 | 5.08 | 4.88 | 10.55 | 9.77 | 10.45 | 9.86 | 10.16
3.776 | 7.42 | 0 | 0.1 | 0.49 | 0.1 | 0.49 | 0 | 3.22 | 4.88 | 18.46 | 4.88 | 17.58 | 4.79 | 20.12 | 9.67 | 9.86 | 9.86 | 9.86 | 13.28 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 18.85 | 4.69 | 15.14 | 4.59 | 4.79 | 5.08 | 4.88 | 10.35 | 9.86 | 10.45 | 9.86 | 10.45
3.808 | 8.01 | 0.1 | 0 | 0.39 | 0.1 | 0.39 | 0.1 | 4 | 4.88 | 18.46 | 4.88 | 18.16 | 4.79 | 20.51 | 9.57 | 9.86 | 9.96 | 9.86 | 7.23 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 19.43 | 4.69 | 15.82 | 4.59 | 5.66 | 5.18 | 4.88 | 10.16 | 9.86 | 10.25 | 9.86 | 10.45
3.84 | 8.69 | 0.1 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 5.27 | 4.79 | 18.26 | 4.79 | 18.46 | 4.79 | 20.9 | 9.67 | 9.86 | 9.86 | 9.86 | 8.01 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 20.12 | 4.79 | 16.5 | 4.49 | 6.64 | 4.98 | 4.88 | 10.25 | 9.77 | 10.06 | 9.86 | 10.25
3.872 | 9.28 | 0.1 | 0.2 | 0.2 | 0 | 0.29 | 0.1 | 5.96 | 4.88 | 18.16 | 4.79 | 18.95 | 4.79 | 21.29 | 9.67 | 9.86 | 9.86 | 9.86 | 11.04 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 20.61 | 4.79 | 17.29 | 4.59 | 8.3 | 4.88 | 4.88 | 10.16 | 9.86 | 10.16 | 9.86 | 10.06
3.904 | 9.86 | 0.1 | 1.17 | 0.2 | 0.1 | 0.29 | 0.1 | 5.57 | 4.88 | 18.16 | 4.88 | 19.34 | 4.79 | 21.78 | 9.67 | 9.96 | 9.77 | 9.86 | 11.43 | 0 | 17.58 | 0 | 17.68 | 0.1 | 21.29 | 4.79 | 17.97 | 4.59 | 10.45 | 4.88 | 4.88 | 10.06 | 9.77 | 10.16 | 9.86 | 10.16
3.936 | 10.45 | 0.1 | 2.15 | 0.2 | 0.1 | 0.2 | 0 | 5.27 | 4.88 | 18.07 | 4.88 | 19.82 | 4.79 | 22.17 | 9.67 | 10.55 | 9.77 | 9.86 | 10.35 | 0.1 | 13.18 | 0 | 17.68 | 0 | 21.88 | 4.79 | 18.46 | 4.59 | 11.62 | 4.79 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
3.968 | 11.04 | 0.1 | 3.42 | 0.1 | 0.1 | 0.2 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.31 | 4.79 | 22.56 | 9.67 | 11.62 | 9.77 | 9.86 | 9.96 | 0.1 | 3.13 | 0 | 17.68 | 0 | 22.46 | 4.79 | 19.24 | 4.59 | 12.6 | 4.69 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
4 | 11.62 | 0.1 | 4.3 | 0.1 | 0 | 0.1 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.61 | 4.88 | 22.95 | 9.67 | 13.09 | 9.67 | 9.86 | 9.96 | 0.1 | 2.34 | 0 | 6.05 | 0.1 | 19.82 | 4.79 | 19.82 | 4.59 | 13.48 | 4.69 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
4.032 | 12.21 | 0.1 | 4.98 | 0 | 0.68 | 0.1 | 0.1 | 4.98 | 4.88 | 18.07 | 4.79 | 21 | 4.88 | 23.34 | 9.67 | 13.87 | 9.67 | 9.86 | 10.16 | 0.1 | 2.05 | 0 | 1.76 | 0.1 | 17.09 | 4.79 | 20.41 | 4.59 | 14.26 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
4.064 | 12.7 | 0.1 | 5.47 | 0 | 1.37 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 19.04 | 4.88 | 23.63 | 9.67 | 14.65 | 9.67 | 9.86 | 10.16 | 0.1 | 1.76 | 0 | 2.25 | 0.1 | 17.68 | 4.79 | 21 | 4.69 | 15.04 | 4.59 | 4.88 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86
4.096 | 13.28 | 0 | 6.05 | 0.1 | 2.54 | 0 | 0.59 | 4.79 | 4.88 | 11.13 | 4.88 | 18.85 | 4.79 | 24.12 | 9.67 | 15.14 | 9.67 | 9.86 | 9.86 | 0.1 | 1.76 | 0 | 2.15 | 0 | 17.68 | 4.79 | 21.58 | 4.69 | 15.72 | 4.59 | 4.88 | 9.77 | 9.86 | 9.67 | 9.86 | 9.77
4.128 | 13.77 | 0.1 | 6.54 | 0.1 | 3.81 | 0 | 1.37 | 4.79 | 4.88 | 3.42 | 4.88 | 18.56 | 4.88 | 24.41 | 9.67 | 15.63 | 9.67 | 9.86 | 9.96 | 0.1 | 1.56 | 0 | 1.66 | 0 | 17.77 | 4.79 | 22.17 | 4.69 | 16.41 | 4.59 | 5.76 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
4.16 | 14.36 | 0.1 | 7.23 | 0.1 | 4.59 | 0.1 | 2.34 | 4.69 | 4.88 | 3.22 | 4.79 | 18.36 | 4.88 | 22.46 | 9.77 | 16.21 | 9.67 | 9.86 | 9.96 | 0.1 | 1.46 | 0 | 1.66 | 0.1 | 17.77 | 4.88 | 22.75 | 4.69 | 16.99 | 4.49 | 6.84 | 9.77 | 9.86 | 9.77 | 10.74 | 9.77
4.192 | 14.84 | 0.1 | 7.81 | 0 | 5.08 | 0.1 | 3.81 | 4.69 | 4.88 | 3.71 | 4.88 | 18.36 | 4.88 | 22.27 | 9.77 | 16.7 | 9.67 | 9.86 | 9.86 | 0.1 | 1.27 | 0 | 1.46 | 0.1 | 17.77 | 4.79 | 23.24 | 4.69 | 17.77 | 4.49 | 8.5 | 9.67 | 10.74 | 9.77 | 11.43 | 9.77
4.224 | 15.33 | 0.1 | 8.4 | 0.1 | 5.66 | 0.1 | 4.49 | 4.59 | 4.88 | 4.88 | 4.88 | 18.16 | 4.79 | 21.88 | 9.77 | 17.09 | 9.67 | 9.86 | 9.86 | 0.1 | 1.27 | 0 | 1.46 | 0.1 | 17.77 | 4.88 | 23.83 | 4.79 | 18.46 | 4.59 | 10.35 | 9.67 | 11.33 | 9.67 | 12.99 | 9.77
4.256 | 15.82 | 0 | 9.08 | 0.1 | 6.25 | 0 | 5.08 | 4.59 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 17.68 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 1.37 | 0 | 17.68 | 4.79 | 24.32 | 4.79 | 19.04 | 4.49 | 11.72 | 9.67 | 12.79 | 9.67 | 15.04 | 9.67
4.288 | 16.31 | 0.1 | 9.67 | 0.1 | 6.84 | 0 | 5.57 | 4.49 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 18.07 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 0.98 | 0 | 17.77 | 4.79 | 24.81 | 4.79 | 19.82 | 4.59 | 12.7 | 9.67 | 14.94 | 9.67 | 16.31 | 9.67
4.32 | 16.8 | 0.1 | 10.25 | 0.1 | 7.42 | 0.1 | 6.15 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.58 | 9.77 | 18.56 | 9.67 | 9.86 | 9.77 | 0.1 | 0.88 | 0 | 1.07 | 0.1 | 17.68 | 4.79 | 25.39 | 4.79 | 20.22 | 4.49 | 13.57 | 9.67 | 16.21 | 9.67 | 17.38 | 9.67
4.352 | 17.19 | 0 | 10.74 | 0 | 8.01 | 0.1 | 6.74 | 4.49 | 4.88 | 4.98 | 4.79 | 18.07 | 4.88 | 21.49 | 9.77 | 18.95 | 9.67 | 9.86 | 9.67 | 0.1 | 0.78 | 0 | 0.88 | 0.1 | 17.77 | 4.88 | 25.88 | 4.79 | 20.9 | 4.49 | 14.36 | 9.67 | 17.19 | 9.67 | 17.97 | 9.67
4.384 | 15.23 | 0.1 | 11.43 | 0 | 8.59 | 0.1 | 7.32 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 19.43 | 9.67 | 9.86 | 9.67 | 0 | 0.78 | 0 | 0.98 | 0.1 | 17.68 | 4.88 | 26.37 | 4.79 | 21.49 | 4.59 | 15.04 | 9.67 | 17.87 | 9.67 | 18.75 | 9.67
4.416 | 15.23 | 0 | 12.01 | 0.1 | 9.28 | 0 | 8.01 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.92 | 9.67 | 9.86 | 9.67 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 17.68 | 4.79 | 23.24 | 4.79 | 21.97 | 4.59 | 15.72 | 9.57 | 18.65 | 9.57 | 19.34 | 9.67
4.448 | 14.84 | 0 | 12.5 | 0.1 | 9.77 | 0 | 8.59 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.31 | 9.67 | 10.64 | 9.67 | 0.1 | 0.68 | 0 | 0.78 | 0 | 6.93 | 4.79 | 22.27 | 4.79 | 22.56 | 4.59 | 16.5 | 9.67 | 19.34 | 9.57 | 20.02 | 9.57
4.48 | 14.75 | 0.1 | 13.09 | 0.1 | 10.45 | 0.1 | 9.18 | 4.49 | 4.88 | 4.88 | 4.88 | 10.45 | 4.88 | 21.39 | 9.86 | 20.8 | 9.67 | 11.23 | 9.67 | 0.1 | 0.49 | 0 | 0.68 | 0.1 | 1.66 | 4.88 | 21.88 | 4.79 | 23.15 | 4.59 | 17.19 | 9.67 | 19.92 | 9.67 | 20.61 | 9.67
4.512 | 14.75 | 0.1 | 13.57 | 0 | 11.04 | 0.1 | 9.77 | 4.49 | 4.88 | 4.79 | 4.88 | 3.03 | 4.88 | 21.39 | 9.86 | 21.09 | 9.67 | 12.21 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.25 | 4.88 | 21.68 | 4.79 | 23.63 | 4.59 | 17.77 | 9.67 | 20.51 | 9.67 | 21.19 | 9.67
4.544 | 14.55 | 0.1 | 14.16 | 0 | 11.62 | 0.1 | 10.35 | 4.59 | 4.88 | 4.69 | 4.88 | 3.32 | 4.88 | 21.39 | 9.86 | 21.49 | 9.67 | 13.48 | 9.57 | 0 | 0.49 | 0 | 0.49 | 0.1 | 2.05 | 4.88 | 21.68 | 4.79 | 24.12 | 4.69 | 18.46 | 9.67 | 21.19 | 9.67 | 21.78 | 9.67
4.576 | 14.55 | 0 | 14.65 | 0.1 | 12.11 | 0 | 10.94 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 4.88 | 13.09 | 9.77 | 21.97 | 9.67 | 14.26 | 9.57 | 0 | 0.39 | 0 | 0.49 | 0 | 1.95 | 4.79 | 21.49 | 4.79 | 24.61 | 4.69 | 19.14 | 9.67 | 21.68 | 9.57 | 22.46 | 9.67
4.608 | 14.55 | 0.1 | 15.04 | 0.1 | 12.79 | 0 | 11.52 | 4.59 | 4.88 | 4.59 | 4.88 | 4.98 | 4.88 | 7.13 | 9.77 | 22.27 | 9.67 | 14.84 | 9.67 | 0.1 | 0.29 | 0 | 0.49 | 0 | 1.66 | 4.79 | 21.39 | 4.88 | 25.2 | 4.69 | 19.73 | 9.67 | 22.27 | 9.67 | 22.95 | 9.67
4.64 | 14.45 | 0.1 | 15.63 | 0.1 | 13.28 | 0.1 | 12.11 | 4.59 | 4.88 | 4.59 | 4.88 | 5.96 | 4.88 | 7.81 | 9.86 | 22.75 | 9.67 | 15.33 | 9.67 | 0.1 | 0.2 | 0.68 | 0.39 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 25.78 | 4.69 | 20.31 | 9.67 | 22.85 | 9.67 | 23.44 | 9.67
4.672 | 14.45 | 0.1 | 16.11 | 0 | 13.77 | 0.1 | 12.6 | 4.69 | 4.88 | 4.49 | 4.88 | 5.66 | 4.88 | 10.94 | 9.86 | 23.05 | 9.77 | 15.92 | 9.67 | 0.1 | 0.2 | 1.66 | 0.29 | 0.1 | 1.37 | 4.88 | 21.29 | 4.88 | 26.27 | 4.69 | 20.9 | 9.67 | 23.54 | 9.67 | 23.83 | 9.67
4.704 | 14.45 | 0 | 16.6 | 0 | 14.26 | 0.1 | 13.18 | 4.69 | 4.88 | 4.49 | 4.88 | 5.27 | 4.79 | 11.33 | 9.86 | 23.44 | 9.77 | 16.41 | 9.67 | 0.88 | 0.2 | 3.22 | 0.29 | 0.1 | 1.37 | 4.88 | 21.39 | 4.88 | 23.24 | 4.79 | 21.49 | 9.67 | 23.83 | 9.67 | 24.51 | 9.67
4.736 | 14.36 | 0 | 17.09 | 0.1 | 14.84 | 0 | 13.67 | 4.69 | 5.66 | 4.49 | 4.88 | 5.08 | 4.88 | 10.35 | 9.77 | 23.83 | 9.77 | 16.8 | 9.67 | 1.76 | 0.2 | 5.18 | 0.29 | 0.1 | 1.17 | 4.79 | 21.29 | 4.88 | 22.07 | 4.69 | 22.17 | 9.77 | 24.41 | 9.67 | 25 | 9.67
4.768 | 14.36 | 0 | 13.77 | 0.1 | 15.33 | 0 | 14.26 | 4.69 | 6.35 | 4.49 | 4.88 | 5.08 | 4.88 | 9.77 | 9.77 | 24.12 | 9.77 | 17.38 | 9.67 | 3.42 | 0.2 | 6.35 | 0.2 | 0 | 1.17 | 4.79 | 21.29 | 4.88 | 21.68 | 4.79 | 22.66 | 9.67 | 25.1 | 9.67 | 25.49 | 9.67
4.8 | 10.06 | 0.1 | 13.97 | 0.1 | 15.82 | 0.1 | 14.84 | 4.79 | 7.62 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 21.09 | 9.77 | 17.77 | 9.67 | 5.27 | 0.1 | 7.23 | 0.2 | 0.68 | 1.07 | 4.88 | 21.29 | 4.88 | 21.58 | 4.79 | 23.15 | 9.77 | 25.39 | 9.67 | 25.98 | 9.67
4.832 | 2.93 | 0 | 14.16 | 0 | 16.21 | 0.1 | 15.33 | 4.79 | 8.69 | 4.49 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 21 | 9.77 | 18.26 | 9.67 | 6.45 | 0.1 | 8.01 | 0.1 | 1.66 | 0.88 | 4.88 | 13.28 | 4.79 | 21.58 | 4.79 | 23.73 | 9.77 | 25.98 | 9.67 | 26.47 | 9.77
4.864 | 1.95 | 0.1 | 14.45 | 0 | 16.8 | 0.1 | 15.82 | 4.79 | 9.57 | 4.49 | 4.88 | 4.79 | 4.88 | 10.06 | 9.86 | 21.29 | 9.77 | 18.75 | 9.67 | 7.32 | 0.1 | 8.89 | 0.2 | 3.22 | 0.78 | 4.88 | 3.32 | 4.88 | 21.49 | 4.79 | 24.12 | 9.77 | 26.47 | 9.67 | 26.95 | 9.67
4.896 | 1.66 | 0 | 14.36 | 0.1 | 17.29 | 0 | 16.31 | 4.79 | 10.25 | 4.49 | 4.88 | 4.79 | 4.88 | 9.96 | 9.86 | 21.49 | 9.77 | 19.14 | 9.67 | 8.01 | 0.1 | 9.57 | 0.2 | 5.27 | 0.68 | 4.79 | 3.13 | 4.88 | 21.29 | 4.79 | 24.71 | 9.77 | 26.95 | 9.67 | 27.44 | 9.77
4.928 | 1.46 | 0 | 14.36 | 0.1 | 14.94 | 0.1 | 16.7 | 4.79 | 10.84 | 4.59 | 4.88 | 4.69 | 4.88 | 9.96 | 9.77 | 21.39 | 9.77 | 19.53 | 9.67 | 8.89 | 0.1 | 10.25 | 0.2 | 6.45 | 0.59 | 4.79 | 3.42 | 4.88 | 21.39 | 4.79 | 25.2 | 9.77 | 27.44 | 9.67 | 27.83 | 9.67
4.96 | 1.46 | 0.1 | 14.45 | 0.1 | 15.33 | 0.1 | 17.19 | 4.79 | 11.52 | 4.59 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.39 | 9.77 | 20.02 | 9.77 | 9.57 | 0.1 | 11.23 | 0.1 | 7.42 | 0.59 | 4.88 | 4.59 | 4.88 | 21.29 | 4.79 | 25.68 | 9.77 | 28.03 | 9.77 | 28.42 | 9.77
4.992 | 1.27 | 0.1 | 14.45 | 0 | 14.75 | 0.1 | 14.26 | 4.88 | 12.01 | 4.59 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.49 | 9.77 | 20.41 | 9.67 | 10.45 | 0.1 | 12.01 | 0.1 | 8.01 | 0.49 | 4.88 | 6.05 | 4.79 | 21.29 | 4.79 | 26.37 | 9.77 | 28.32 | 9.77 | 28.91 | 9.77
5.024 | 1.17 | 0 | 14.45 | 0 | 14.75 | 0.1 | 14.75 | 4.88 | 12.6 | 4.59 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.49 | 9.86 | 20.8 | 9.77 | 11.33 | 0.1 | 12.79 | 0.1 | 8.79 | 0.49 | 4.88 | 6.25 | 4.88 | 21.19 | 4.79 | 23.05 | 9.77 | 28.91 | 9.77 | 29.4 | 9.77
5.056 | 1.07 | 0 | 14.45 | 0.1 | 14.65 | 0 | 14.65 | 4.79 | 13.09 | 4.59 | 4.88 | 4.59 | 4.88 | 9.67 | 9.77 | 21.49 | 9.77 | 21.19 | 9.67 | 12.01 | 0.1 | 13.48 | 0.1 | 9.77 | 0.29 | 4.79 | 5.66 | 4.88 | 21.19 | 4.79 | 22.27 | 9.77 | 29.4 | 9.77 | 24.9 | 9.77
5.088 | 0.88 | 0 | 14.36 | 0.1 | 14.55 | 0 | 14.55 | 4.79 | 13.67 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 21.68 | 9.67 | 12.79 | 0.1 | 14.16 | 0.1 | 10.45 | 0.39 | 4.79 | 5.47 | 4.88 | 21.19 | 4.79 | 21.78 | 9.77 | 24.81 | 9.77 | 23.83 | 9.77
5.12 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 14.26 | 4.69 | 5.47 | 4.59 | 4.88 | 9.77 | 9.86 | 21.39 | 9.86 | 22.07 | 9.77 | 13.48 | 0.1 | 14.94 | 0.1 | 11.23 | 0.29 | 4.88 | 5.18 | 4.88 | 14.06 | 4.79 | 21.58 | 9.77 | 23.73 | 9.77 | 24.41 | 9.77
5.152 | 0.78 | 0.1 | 14.36 | 0 | 14.55 | 0.1 | 14.55 | 4.88 | 14.65 | 4.69 | 6.15 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 22.46 | 9.77 | 14.26 | 0.1 | 15.53 | 0.1 | 11.91 | 0.29 | 4.79 | 5.18 | 4.79 | 3.32 | 4.79 | 21.58 | 9.77 | 24.41 | 9.77 | 24.22 | 9.77
5.184 | 0.68 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 15.23 | 4.69 | 7.32 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 9.77 | 15.04 | 0.1 | 16.31 | 0.1 | 12.79 | 0.2 | 4.88 | 5.08 | 4.88 | 3.13 | 4.79 | 21.49 | 9.77 | 24.22 | 9.77 | 24.32 | 9.77
5.216 | 0.59 | 0 | 4.2 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 15.72 | 4.79 | 8.59 | 4.49 | 4.88 | 9.67 | 9.77 | 21.39 | 9.86 | 23.24 | 9.77 | 15.72 | 0.1 | 16.99 | 0.1 | 13.67 | 0.2 | 5.86 | 5.08 | 4.88 | 3.42 | 4.79 | 21.39 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86
5.248 | 0.49 | 0 | 1.56 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 16.21 | 4.79 | 9.47 | 4.49 | 4.88 | 9.67 | 9.77 | 11.43 | 9.86 | 23.54 | 9.77 | 16.31 | 0.1 | 17.58 | 0.1 | 14.26 | 0.1 | 7.13 | 4.98 | 4.88 | 4.49 | 4.79 | 21.39 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77
5.28 | 0.49 | 0.1 | 1.95 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 16.7 | 4.79 | 10.16 | 4.59 | 4.88 | 9.67 | 9.86 | 6.54 | 9.86 | 23.93 | 9.77 | 16.99 | 0.1 | 18.26 | 0.1 | 14.84 | 0.2 | 8.98 | 4.98 | 4.88 | 5.86 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86
5.312 | 0.49 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.09 | 4.79 | 10.74 | 4.49 | 5.57 | 9.67 | 9.86 | 8.89 | 9.86 | 24.32 | 9.77 | 17.58 | 0.1 | 18.85 | 0.1 | 15.72 | 0.1 | 10.74 | 4.88 | 4.79 | 6.35 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
5.344 | 0.29 | 0.1 | 1.56 | 0 | 11.04 | 0.1 | 14.45 | 4.88 | 17.58 | 4.79 | 11.43 | 4.59 | 6.54 | 9.67 | 9.86 | 11.52 | 9.86 | 21.58 | 9.86 | 18.26 | 0.1 | 19.53 | 0.1 | 16.41 | 0.1 | 12.11 | 4.88 | 4.88 | 5.66 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
5.376 | 0.29 | 0 | 1.37 | 0.1 | 3.03 | 0 | 14.36 | 4.88 | 18.07 | 4.79 | 11.91 | 4.59 | 7.91 | 9.57 | 9.77 | 11.13 | 9.86 | 22.17 | 9.77 | 18.95 | 0.1 | 20.12 | 0.1 | 17.19 | 0.1 | 12.79 | 4.79 | 4.88 | 5.27 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
5.408 | 0.2 | 0.1 | 1.17 | 0.1 | 2.05 | 0 | 12.99 | 4.88 | 18.56 | 4.79 | 12.5 | 4.59 | 8.89 | 9.57 | 9.77 | 10.25 | 9.86 | 21.58 | 9.77 | 19.63 | 0.1 | 20.8 | 0.1 | 17.87 | 0.1 | 13.67 | 4.69 | 4.88 | 5.18 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
5.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.76 | 0.1 | 3.91 | 4.88 | 19.04 | 4.79 | 13.09 | 4.69 | 9.67 | 9.67 | 9.86 | 9.77 | 9.86 | 21.68 | 9.86 | 20.22 | 0 | 21.39 | 0.1 | 18.46 | 0.1 | 14.65 | 4.79 | 4.88 | 5.18 | 4.79 | 13.48 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
5.472 | 0.2 | 0 | 0.98 | 0 | 1.56 | 0.1 | 1.95 | 4.88 | 19.34 | 4.79 | 13.57 | 4.69 | 10.35 | 9.67 | 9.86 | 10.06 | 9.86 | 21.58 | 9.86 | 20.8 | 0.1 | 21.88 | 0.1 | 18.95 | 0.1 | 15.33 | 4.69 | 4.88 | 5.18 | 4.79 | 3.22 | 9.77 | 24.22 | 9.77 | 23.44 | 9.86
5.504 | 0.1 | 0.1 | 0.88 | 0 | 1.56 | 0.1 | 1.86 | 4.88 | 19.92 | 4.79 | 14.16 | 4.69 | 11.04 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 9.86 | 21.29 | 0.1 | 22.56 | 0.1 | 19.53 | 0.1 | 15.92 | 4.69 | 4.88 | 4.98 | 4.79 | 3.22 | 9.77 | 23.54 | 9.77 | 8.4 | 9.86
5.536 | 0.1 | 0 | 0.78 | 0.1 | 1.27 | 0 | 1.46 | 4.88 | 20.22 | 4.88 | 14.65 | 4.69 | 11.52 | 9.67 | 9.86 | 10.06 | 9.86 | 21.49 | 9.86 | 21.97 | 0.1 | 18.46 | 0.1 | 20.22 | 0.1 | 16.7 | 4.59 | 4.88 | 5.08 | 4.79 | 3.52 | 9.77 | 8.2 | 9.77 | 6.45 | 9.86
5.568 | 0.1 | 0.59 | 0.68 | 0.1 | 1.17 | 0 | 1.46 | 4.88 | 20.7 | 4.88 | 15.14 | 4.69 | 12.11 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 9.86 | 22.56 | 0.1 | 17.77 | 0.1 | 20.8 | 0.1 | 17.29 | 4.59 | 4.88 | 4.98 | 4.79 | 4.59 | 9.77 | 6.35 | 9.77 | 10.16 | 9.77
5.6 | 0.1 | 1.37 | 0.59 | 0.1 | 1.07 | 0.1 | 1.27 | 4.88 | 21.09 | 4.88 | 15.63 | 4.79 | 12.7 | 9.67 | 11.33 | 9.86 | 9.86 | 21.39 | 9.86 | 19.24 | 0 | 17.97 | 0.1 | 21.49 | 0.1 | 17.97 | 4.59 | 4.88 | 4.88 | 4.79 | 6.05 | 9.77 | 10.25 | 9.77 | 12.31 | 9.86
5.632 | 0.1 | 2.44 | 0.59 | 0 | 0.88 | 0.1 | 1.17 | 4.88 | 19.14 | 4.79 | 16.11 | 4.79 | 13.18 | 9.67 | 12.6 | 9.96 | 9.96 | 21.49 | 9.86 | 18.46 | 0.1 | 17.77 | 0.1 | 22.07 | 0.1 | 18.65 | 4.59 | 4.79 | 4.88 | 4.79 | 6.15 | 9.77 | 12.21 | 9.77 | 10.94 | 9.86
5.664 | 0.1 | 3.71 | 0.49 | 0 | 0.88 | 0.1 | 1.17 | 4.88 | 18.95 | 4.88 | 16.6 | 4.79 | 13.77 | 9.67 | 13.57 | 9.86 | 9.86 | 21.39 | 9.86 | 18.07 | 0.1 | 17.87 | 0.1 | 22.46 | 0.1 | 19.34 | 4.59 | 5.08 | 4.79 | 4.79 | 5.66 | 9.77 | 11.04 | 9.77 | 10.06 | 9.86
5.696 | 0 | 4.49 | 0.49 | 0.1 | 0.78 | 0.1 | 0.98 | 4.88 | 18.56 | 4.88 | 17.09 | 4.79 | 14.26 | 9.67 | 14.36 | 9.77 | 9.86 | 21.39 | 9.86 | 17.87 | 0.1 | 17.77 | 0.1 | 18.75 | 0 | 19.92 | 4.59 | 6.25 | 4.69 | 4.79 | 5.47 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86
5.728 | 0.1 | 5.08 | 0.29 | 0.1 | 0.59 | 0 | 0.78 | 4.88 | 18.46 | 4.88 | 17.58 | 4.79 | 14.84 | 9.67 | 14.94 | 9.67 | 9.86 | 21.39 | 9.86 | 17.87 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 20.51 | 4.59 | 7.72 | 4.79 | 4.79 | 5.18 | 9.77 | 10.06 | 9.77 | 10.35 | 9.86
5.76 | 0.1 | 5.57 | 0.29 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 18.36 | 4.88 | 17.97 | 4.79 | 15.33 | 9.67 | 15.53 | 9.77 | 9.86 | 16.8 | 9.96 | 17.77 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 21.09 | 4.59 | 9.96 | 4.69 | 4.79 | 5.08 | 9.77 | 10.35 | 9.77 | 10.55 | 9.86
5.792 | 0.1 | 6.15 | 0.29 | 0 | 0.49 | 0.1 | 0.68 | 4.88 | 18.26 | 4.79 | 18.46 | 4.79 | 15.82 | 9.67 | 16.02 | 9.67 | 9.86 | 8.2 | 9.96 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 21.68 | 4.59 | 11.33 | 4.59 | 4.79 | 5.08 | 9.77 | 10.45 | 9.77 | 10.25 | 9.86
5.824 | 0.1 | 6.74 | 0.29 | 0 | 0.49 | 0.1 | 0.68 | 4.88 | 18.26 | 4.79 | 18.95 | 4.88 | 16.31 | 9.77 | 16.6 | 9.67 | 9.86 | 7.42 | 9.96 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.27 | 4.59 | 12.5 | 4.59 | 4.79 | 5.08 | 9.77 | 10.25 | 9.77 | 10.16 | 9.86
5.856 | 0.1 | 7.42 | 0.2 | 0.1 | 0.39 | 0 | 0.59 | 4.88 | 18.16 | 4.88 | 19.34 | 4.88 | 16.8 | 9.77 | 16.99 | 9.67 | 9.86 | 10.35 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.85 | 4.59 | 13.28 | 4.59 | 4.79 | 5.08 | 9.77 | 10.16 | 9.77 | 10.25 | 9.86
5.888 | 0.1 | 8.01 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.88 | 18.16 | 4.88 | 19.82 | 4.88 | 17.19 | 9.77 | 17.48 | 9.67 | 9.86 | 11.43 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 23.34 | 4.59 | 14.06 | 4.59 | 4.79 | 5.08 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
5.92 | 0.1 | 8.5 | 0.1 | 0 | 0.29 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 17.68 | 9.77 | 17.97 | 9.67 | 9.86 | 10.64 | 9.96 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 0.1 | 23.93 | 4.69 | 14.94 | 4.59 | 4.79 | 4.88 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86
5.952 | 0.1 | 9.08 | 0.1 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 18.07 | 4.79 | 20.61 | 4.79 | 18.16 | 9.77 | 18.46 | 9.67 | 9.86 | 9.86 | 9.96 | 17.58 | 0.1 | 15.72 | 0.1 | 17.68 | 0.1 | 24.51 | 4.69 | 15.53 | 4.59 | 4.79 | 4.88 | 9.77 | 10.06 | 9.77 | 10.06 | 9.86
5.984 | 0.1 | 9.77 | 0.1 | 0.98 | 0.2 | 0 | 0.29 | 4.88 | 18.07 | 4.79 | 20.9 | 4.88 | 18.65 | 9.77 | 18.85 | 9.67 | 9.86 | 9.86 | 9.96 | 17.58 | 0.1 | 4 | 0.1 | 17.77 | 0.1 | 25 | 4.69 | 16.21 | 4.49 | 4.79 | 4.79 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86
6.016 | 0.1 | 10.45 | 0.1 | 1.76 | 0.2 | 0.68 | 0.2 | 4.88 | 18.07 | 4.88 | 18.65 | 4.88 | 19.14 | 9.77 | 19.34 | 9.67 | 9.86 | 10.16 | 9.86 | 11.72 | 0.1 | 2.15 | 0.1 | 17.68 | 0 | 25.49 | 4.69 | 16.99 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
6.048 | 0.1 | 10.94 | 0 | 3.13 | 0.1 | 1.37 | 0.2 | 4.88 | 10.45 | 4.88 | 18.95 | 4.88 | 19.43 | 9.77 | 19.73 | 9.67 | 9.86 | 10.16 | 9.86 | 2.93 | 0.1 | 2.15 | 0.1 | 17.68 | 0.1 | 25.98 | 4.69 | 17.68 | 4.59 | 4.79 | 4.79 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
6.08 | 0.1 | 11.52 | 0.1 | 4.1 | 0.1 | 2.64 | 0.2 | 4.88 | 3.13 | 4.88 | 18.56 | 4.88 | 19.92 | 9.77 | 20.12 | 9.57 | 9.86 | 9.86 | 9.96 | 2.15 | 0 | 1.56 | 0.1 | 17.68 | 0.1 | 25.1 | 4.79 | 18.26 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
6.112 | 0.1 | 12.11 | 0.1 | 4.79 | 0 | 3.81 | 0.2 | 4.88 | 3.22 | 4.88 | 18.36 | 4.88 | 20.22 | 9.77 | 20.61 | 9.57 | 10.25 | 9.86 | 9.96 | 2.05 | 0.1 | 1.76 | 0.1 | 14.84 | 0.1 | 21.58 | 4.79 | 18.85 | 4.59 | 5.08 | 4.59 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
6.144 | 0.1 | 12.7 | 0.1 | 5.27 | 0.1 | 4.59 | 0.1 | 4.88 | 3.91 | 4.79 | 18.46 | 4.88 | 20.7 | 9.77 | 21 | 9.67 | 11.04 | 9.96 | 9.96 | 1.66 | 0.1 | 1.56 | 0.1 | 3.71 | 0.1 | 21.58 | 4.79 | 19.53 | 4.59 | 6.05 | 4.59 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86
6.176 | 0.1 | 13.18 | 0.1 | 5.86 | 0.1 | 5.18 | 0.1 | 4.88 | 4.98 | 4.88 | 18.26 | 4.88 | 19.53 | 9.77 | 21.39 | 9.67 | 12.01 | 9.86 | 9.86 | 1.66 | 0.1 | 1.46 | 0.1 | 2.25 | 0.1 | 21.39 | 4.79 | 20.22 | 4.59 | 7.62 | 4.59 | 10.74 | 9.77 | 10.35 | 9.77 | 9.86
6.208 | 0.1 | 13.77 | 0 | 6.45 | 0.1 | 5.66 | 0 | 4.88 | 5.96 | 4.88 | 18.26 | 4.88 | 17.77 | 9.77 | 21.78 | 9.67 | 13.28 | 9.77 | 9.96 | 1.56 | 0.1 | 1.37 | 0.1 | 2.15 | 0.1 | 21.39 | 4.79 | 20.8 | 4.69 | 9.77 | 4.59 | 12.11 | 9.77 | 11.13 | 9.67 | 10.74
6.24 | 0.1 | 14.26 | 0.1 | 7.03 | 0.1 | 6.25 | 0.1 | 4.88 | 5.57 | 4.88 | 18.26 | 4.88 | 18.07 | 9.77 | 22.17 | 9.67 | 14.16 | 9.77 | 9.96 | 1.37 | 0 | 1.27 | 0.1 | 1.86 | 0.1 | 21.39 | 4.88 | 21.39 | 4.69 | 11.43 | 4.59 | 14.16 | 9.67 | 12.7 | 9.67 | 11.43
6.272 | 0.1 | 14.75 | 0.1 | 7.72 | 0.1 | 6.84 | 0.1 | 4.88 | 5.27 | 4.88 | 18.07 | 4.79 | 18.26 | 9.77 | 22.56 | 9.67 | 14.75 | 9.77 | 9.96 | 1.37 | 0.1 | 1.07 | 0.1 | 1.66 | 0.1 | 21.29 | 4.88 | 21.97 | 4.69 | 12.31 | 4.49 | 15.72 | 9.67 | 14.75 | 9.67 | 12.79
6.304 | 0.1 | 15.23 | 0.1 | 8.3 | 0.1 | 7.42 | 0.1 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 17.97 | 9.77 | 22.95 | 9.67 | 15.33 | 9.77 | 9.96 | 1.27 | 0.1 | 1.07 | 0.1 | 1.56 | 0.1 | 21.29 | 4.88 | 22.46 | 4.69 | 13.18 | 4.59 | 16.9 | 9.67 | 16.21 | 9.67 | 14.94
6.336 | 0 | 15.72 | 0.1 | 8.89 | 0.1 | 8.11 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 23.44 | 9.67 | 15.82 | 9.67 | 9.96 | 1.17 | 0.1 | 0.88 | 0.1 | 1.46 | 0 | 21.19 | 4.79 | 23.05 | 4.69 | 13.97 | 4.59 | 17.77 | 9.67 | 17.09 | 9.67 | 16.31
6.368 | 0.1 | 16.21 | 0.1 | 9.47 | 0.1 | 8.59 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 23.73 | 9.67 | 16.41 | 9.67 | 9.96 | 0.98 | 0.1 | 0.78 | 0.1 | 1.46 | 0.1 | 21.19 | 4.88 | 23.54 | 4.79 | 14.75 | 4.59 | 18.36 | 9.67 | 17.77 | 9.67 | 17.19
6.4 | 0.1 | 16.7 | 0.1 | 10.06 | 0.1 | 9.28 | 0.1 | 5.66 | 4.88 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 24.12 | 9.67 | 16.9 | 9.67 | 9.96 | 0.88 | 0 | 0.78 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.12 | 4.79 | 15.53 | 4.59 | 19.14 | 9.67 | 18.56 | 9.67 | 17.87
6.432 | 0.1 | 17.19 | 0.1 | 10.64 | 0.1 | 9.86 | 0.1 | 6.25 | 4.79 | 4.88 | 13.67 | 4.88 | 18.07 | 9.77 | 24.12 | 9.67 | 17.38 | 9.67 | 9.96 | 0.88 | 0.1 | 0.59 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.61 | 4.79 | 16.11 | 4.59 | 19.73 | 9.57 | 19.24 | 9.67 | 18.75
6.464 | 0.1 | 14.55 | 0.1 | 11.23 | 0.1 | 10.45 | 0.1 | 7.62 | 4.79 | 4.79 | 3.91 | 4.88 | 18.07 | 9.77 | 21.97 | 9.67 | 17.77 | 9.67 | 10.64 | 0.78 | 0.1 | 0.59 | 0.1 | 0.98 | 0.1 | 21.19 | 4.88 | 25.2 | 4.79 | 16.9 | 4.59 | 20.31 | 9.57 | 19.82 | 9.67 | 19.24
6.496 | 0.1 | 15.23 | 0.1 | 11.82 | 0.1 | 11.04 | 0.1 | 8.79 | 4.79 | 4.88 | 3.32 | 4.88 | 17.97 | 9.77 | 21.68 | 9.77 | 18.26 | 9.57 | 11.33 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 0.1 | 21.19 | 4.88 | 25.68 | 4.79 | 17.58 | 4.69 | 21 | 9.67 | 20.51 | 9.67 | 20.02
6.528 | 0.1 | 14.65 | 0.1 | 12.4 | 0.1 | 11.62 | 0.1 | 9.57 | 4.69 | 4.88 | 3.81 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 18.65 | 9.57 | 12.5 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 0.1 | 8.5 | 4.88 | 26.27 | 4.79 | 18.16 | 4.69 | 21.49 | 9.67 | 21.09 | 9.67 | 20.51
6.56 | 0.1 | 14.65 | 0.1 | 12.89 | 0.1 | 12.21 | 0.1 | 10.25 | 4.59 | 4.88 | 4.88 | 4.88 | 17.97 | 9.77 | 21.68 | 9.77 | 19.24 | 9.57 | 13.67 | 0.49 | 0 | 0.39 | 0.1 | 0.78 | 0.1 | 2.25 | 4.88 | 22.07 | 4.79 | 18.85 | 4.69 | 22.17 | 9.67 | 21.68 | 9.67 | 21.09
6.592 | 0.1 | 14.65 | 0.1 | 13.48 | 0.1 | 12.79 | 0.1 | 10.84 | 4.59 | 4.88 | 5.96 | 4.88 | 17.97 | 9.77 | 21.58 | 9.77 | 19.53 | 9.57 | 14.45 | 0.59 | 0.1 | 0.29 | 0.1 | 0.68 | 0.1 | 3.22 | 4.88 | 20.41 | 4.79 | 19.43 | 4.69 | 22.66 | 9.67 | 22.27 | 9.67 | 21.78
6.624 | 0.1 | 14.55 | 0.1 | 13.87 | 0.1 | 13.28 | 0.1 | 11.43 | 4.59 | 4.79 | 5.76 | 4.88 | 7.32 | 9.77 | 21.49 | 9.77 | 20.02 | 9.57 | 15.04 | 0.39 | 0.68 | 0.29 | 0.1 | 0.59 | 0.1 | 3.81 | 4.88 | 21.09 | 4.79 | 20.12 | 4.79 | 23.24 | 9.67 | 22.85 | 9.67 | 22.27
6.656 | 0 | 14.55 | 0.1 | 14.55 | 0.1 | 13.87 | 0 | 12.11 | 4.59 | 4.88 | 5.37 | 4.88 | 2.64 | 9.77 | 21.39 | 9.77 | 20.41 | 9.57 | 15.53 | 0.39 | 1.66 | 0.29 | 0.2 | 0.49 | 0 | 4.98 | 4.88 | 21.19 | 4.79 | 20.7 | 4.79 | 23.73 | 9.67 | 23.34 | 9.67 | 22.85
6.688 | 0.1 | 14.45 | 0.1 | 15.04 | 0.1 | 14.36 | 0.1 | 12.6 | 4.59 | 4.88 | 5.08 | 4.88 | 3.13 | 9.77 | 21.49 | 9.77 | 20.8 | 9.57 | 16.02 | 0.29 | 3.22 | 0.2 | 1.46 | 0.49 | 0 | 6.15 | 4.88 | 21.29 | 4.79 | 21.29 | 4.79 | 24.22 | 9.67 | 23.83 | 9.67 | 23.34
6.72 | 0.1 | 14.45 | 0.1 | 15.53 | 0 | 14.84 | 0.1 | 13.18 | 4.49 | 4.88 | 4.88 | 4.88 | 4.39 | 9.77 | 21.49 | 9.77 | 21.29 | 9.57 | 16.6 | 0.29 | 5.18 | 0.1 | 2.73 | 0.49 | 0.1 | 6.15 | 4.88 | 21.29 | 4.79 | 21.88 | 4.79 | 24.81 | 9.67 | 24.41 | 9.77 | 23.93
6.752 | 0.1 | 14.45 | 0.1 | 16.02 | 0.1 | 15.33 | 0.1 | 13.67 | 4.49 | 4.88 | 4.88 | 4.88 | 5.47 | 9.77 | 21.39 | 9.77 | 21.68 | 9.57 | 17.09 | 0.2 | 6.35 | 0.1 | 4.79 | 0.39 | 0.1 | 5.57 | 4.88 | 21.29 | 4.79 | 22.36 | 4.79 | 25.29 | 9.67 | 24.9 | 9.67 | 24.51
6.784 | 0.1 | 14.36 | 0.1 | 16.41 | 0.1 | 15.82 | 0.1 | 14.26 | 4.59 | 5.47 | 4.88 | 4.88 | 5.96 | 9.77 | 21.39 | 9.86 | 22.07 | 9.67 | 17.58 | 0.2 | 7.23 | 0.1 | 6.25 | 0.29 | 0.1 | 5.18 | 4.88 | 21.39 | 4.79 | 23.05 | 4.79 | 25.68 | 9.67 | 25.49 | 9.77 | 25
6.816 | 0 | 14.36 | 0.1 | 16.9 | 0.1 | 16.41 | 0.1 | 14.84 | 4.49 | 6.25 | 4.88 | 4.88 | 5.57 | 9.77 | 21.39 | 9.86 | 22.36 | 9.67 | 17.87 | 0.2 | 8.2 | 0.1 | 7.13 | 0.29 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 23.54 | 4.79 | 26.27 | 9.77 | 25.88 | 9.77 | 25.49
6.848 | 0.1 | 14.36 | 0.1 | 15.82 | 0.1 | 16.8 | 0.1 | 15.23 | 4.59 | 7.42 | 4.79 | 4.88 | 5.27 | 9.77 | 21.39 | 9.86 | 22.75 | 9.67 | 18.36 | 0.2 | 8.79 | 0.1 | 7.81 | 0.2 | 0 | 5.08 | 4.88 | 21.29 | 4.79 | 24.02 | 4.79 | 26.76 | 9.77 | 26.47 | 9.77 | 25.88
6.88 | 0.1 | 11.62 | 0.1 | 14.16 | 0 | 17.29 | 0.1 | 15.72 | 4.59 | 8.59 | 4.79 | 4.88 | 4.98 | 9.77 | 12.89 | 9.86 | 23.15 | 9.67 | 18.85 | 0.1 | 9.67 | 0 | 8.69 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.88 | 24.61 | 4.79 | 27.25 | 9.77 | 26.86 | 9.77 | 26.37
6.912 | 0.1 | 3.22 | 0.1 | 14.26 | 0.1 | 15.43 | 0.1 | 16.21 | 4.59 | 9.47 | 4.69 | 4.88 | 4.88 | 9.77 | 7.03 | 9.86 | 23.54 | 9.67 | 19.24 | 0.1 | 10.35 | 0 | 9.47 | 0.2 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 25 | 4.79 | 27.83 | 9.77 | 27.34 | 9.77 | 26.86
6.944 | 0.1 | 1.95 | 0.1 | 14.55 | 0.1 | 15.23 | 0.1 | 16.7 | 4.59 | 10.16 | 4.69 | 4.88 | 4.88 | 9.77 | 8.3 | 9.86 | 23.93 | 9.67 | 19.82 | 0.1 | 11.13 | 0.1 | 10.25 | 0.2 | 0.1 | 4.98 | 4.88 | 21.29 | 4.79 | 25.59 | 4.88 | 28.22 | 9.77 | 27.83 | 9.86 | 27.44
6.976 | 0.1 | 1.66 | 0.1 | 14.45 | 0.1 | 14.84 | 0.1 | 17.19 | 4.59 | 10.74 | 4.59 | 4.88 | 4.79 | 9.77 | 11.23 | 9.86 | 24.32 | 9.67 | 20.12 | 0.1 | 11.91 | 0.1 | 10.94 | 0.1 | 0.1 | 4.79 | 4.88 | 20.51 | 4.79 | 25.98 | 4.88 | 28.71 | 9.77 | 28.32 | 9.77 | 28.03
7.008 | 0.1 | 1.46 | 0 | 14.45 | 0.1 | 14.65 | 0.1 | 17.68 | 4.69 | 11.43 | 4.59 | 4.88 | 4.79 | 9.77 | 11.23 | 9.86 | 21.29 | 9.67 | 20.51 | 0.1 | 12.79 | 0 | 11.72 | 0.1 | 0.1 | 4.79 | 4.88 | 4.98 | 4.79 | 21.49 | 4.88 | 29.1 | 9.77 | 28.91 | 9.77 | 28.32
7.04 | 0.1 | 1.56 | 0.1 | 14.45 | 0.1 | 14.75 | 0.1 | 18.07 | 4.69 | 11.91 | 4.59 | 4.88 | 4.79 | 9.77 | 10.35 | 9.86 | 21.29 | 9.67 | 21 | 0.1 | 13.48 | 0 | 12.6 | 0.1 | 0.1 | 4.79 | 4.88 | 3.13 | 4.88 | 20.51 | 4.88 | 26.47 | 9.77 | 29.3 | 9.77 | 28.81
7.072 | 0.1 | 1.27 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.56 | 4.69 | 12.6 | 4.49 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 21.39 | 9.67 | 21.39 | 0.1 | 14.16 | 0 | 13.28 | 0.1 | 0.1 | 4.79 | 4.88 | 3.22 | 4.79 | 20.9 | 4.88 | 23.63 | 9.77 | 24.61 | 9.77 | 29.3
7.104 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 19.04 | 4.79 | 13.09 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.49 | 9.67 | 21.78 | 0.1 | 15.04 | 0.1 | 13.97 | 0.1 | 0.1 | 4.69 | 4.88 | 4.2 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 23.54 | 9.86 | 24.81
7.136 | 0 | 0.98 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 19.43 | 4.79 | 13.67 | 4.59 | 4.88 | 4.59 | 9.77 | 10.35 | 9.86 | 21.49 | 9.67 | 22.17 | 0.1 | 15.63 | 0 | 14.75 | 0.1 | 1.17 | 4.59 | 4.88 | 5.57 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
7.168 | 0.1 | 0.98 | 0 | 14.36 | 0.1 | 14.45 | 0.1 | 19.82 | 4.79 | 14.06 | 4.59 | 5.57 | 4.59 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 22.56 | 0.1 | 16.31 | 0.1 | 15.33 | 0 | 2.54 | 4.59 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
7.2 | 0.1 | 0.88 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 20.31 | 4.79 | 14.65 | 4.49 | 6.25 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.67 | 22.95 | 0.1 | 16.99 | 0 | 16.21 | 0.1 | 4.3 | 4.59 | 4.88 | 5.86 | 4.79 | 21.29 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
7.232 | 0.1 | 0.88 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 20.7 | 4.79 | 15.23 | 4.49 | 7.42 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 23.34 | 0.1 | 17.58 | 0 | 16.99 | 0.1 | 5.86 | 4.59 | 4.88 | 5.37 | 4.88 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.32
7.264 | 0.1 | 0.59 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 21.09 | 4.79 | 15.63 | 4.59 | 8.59 | 4.49 | 10.45 | 9.86 | 9.86 | 21.49 | 9.77 | 23.73 | 0.1 | 18.26 | 0.1 | 17.48 | 0.1 | 6.84 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
7.296 | 0.1 | 0.59 | 0.1 | 6.54 | 0.1 | 14.36 | 0.1 | 18.65 | 4.79 | 16.11 | 4.59 | 9.57 | 4.49 | 11.23 | 9.77 | 9.86 | 21.49 | 9.77 | 24.12 | 0.1 | 18.95 | 0 | 18.16 | 0.1 | 7.72 | 4.49 | 4.88 | 5.18 | 4.79 | 21.29 | 4.88 | 24.22 | 9.86 | 24.32 | 9.86 | 24.22
7.328 | 0.1 | 0.49 | 0.1 | 1.86 | 0.1 | 9.67 | 0.1 | 18.75 | 4.79 | 16.6 | 4.59 | 10.16 | 4.49 | 12.31 | 9.86 | 9.86 | 21.39 | 9.77 | 24.51 | 0.1 | 19.53 | 0 | 18.85 | 0 | 8.4 | 4.49 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
7.36 | 0.1 | 0.49 | 0.1 | 1.95 | 0.1 | 2.54 | 0.1 | 18.56 | 4.79 | 17.09 | 4.59 | 10.74 | 4.59 | 13.38 | 9.77 | 9.86 | 21.49 | 9.77 | 22.36 | 0.1 | 20.12 | 0 | 19.43 | 0.1 | 9.38 | 4.49 | 4.88 | 5.18 | 4.88 | 21.19 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
7.392 | 0.1 | 0.39 | 0.1 | 1.76 | 0 | 1.95 | 0.1 | 18.36 | 4.79 | 17.58 | 4.59 | 11.43 | 4.49 | 14.16 | 9.77 | 9.86 | 21.39 | 9.77 | 22.36 | 0.1 | 20.7 | 0 | 20.12 | 0.1 | 10.16 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
7.424 | 0.1 | 0.29 | 0.1 | 1.37 | 0.1 | 1.86 | 0.1 | 18.36 | 4.88 | 18.07 | 4.69 | 11.91 | 4.49 | 14.75 | 9.77 | 9.86 | 21 | 9.77 | 21.97 | 0.1 | 21.39 | 0.1 | 20.7 | 0.1 | 10.94 | 4.49 | 4.88 | 4.98 | 4.79 | 21.09 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
7.456 | 0 | 0.29 | 0.1 | 1.37 | 0.1 | 1.37 | 0 | 18.16 | 4.88 | 18.46 | 4.69 | 12.5 | 4.59 | 15.33 | 9.67 | 9.86 | 8.69 | 9.77 | 21.78 | 0.1 | 21.88 | 0 | 21.29 | 0 | 11.72 | 4.49 | 4.88 | 4.98 | 4.79 | 5.47 | 4.88 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
7.488 | 0.59 | 0.2 | 0.1 | 1.27 | 0.1 | 1.56 | 0.1 | 18.16 | 4.88 | 18.85 | 4.69 | 13.09 | 4.59 | 15.92 | 9.67 | 9.86 | 7.13 | 9.77 | 21.78 | 0.1 | 22.56 | 0 | 21.88 | 0 | 12.5 | 4.49 | 4.88 | 4.79 | 4.79 | 2.64 | 4.88 | 11.43 | 9.86 | 23.93 | 9.86 | 24.12
7.52 | 1.37 | 0.2 | 0.1 | 1.17 | 0.1 | 1.27 | 0.1 | 18.16 | 4.79 | 19.34 | 4.69 | 13.57 | 4.59 | 16.41 | 9.67 | 9.86 | 9.96 | 9.77 | 21.68 | 0.1 | 18.65 | 0 | 22.46 | 0.1 | 13.28 | 4.59 | 5.76 | 4.79 | 4.79 | 3.42 | 4.88 | 5.57 | 9.86 | 8.5 | 9.86 | 23.63
7.552 | 2.44 | 0.2 | 0.1 | 0.88 | 0 | 1.17 | 0.1 | 18.16 | 4.79 | 19.82 | 4.69 | 14.16 | 4.69 | 16.9 | 9.67 | 9.86 | 11.52 | 9.77 | 21.58 | 0.1 | 17.77 | 0 | 17.87 | 0.1 | 14.06 | 4.59 | 6.54 | 4.69 | 4.79 | 4.2 | 4.79 | 8.59 | 9.77 | 6.25 | 9.86 | 8.5
7.584 | 3.81 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 18.07 | 4.88 | 20.22 | 4.79 | 14.65 | 4.69 | 17.38 | 9.67 | 9.86 | 10.74 | 9.86 | 21.58 | 0.1 | 17.87 | 0 | 16.8 | 0.1 | 14.65 | 4.59 | 8.2 | 4.69 | 4.79 | 5.47 | 4.88 | 12.21 | 9.77 | 9.96 | 9.86 | 6.54
7.616 | 4.49 | 0 | 0.1 | 0.78 | 0.1 | 0.88 | 0 | 18.07 | 4.88 | 20.61 | 4.79 | 15.14 | 4.69 | 17.77 | 9.67 | 9.86 | 9.86 | 9.77 | 21.49 | 0.1 | 17.77 | 0 | 17.58 | 0.1 | 15.43 | 4.69 | 10.25 | 4.69 | 4.79 | 6.25 | 4.88 | 11.23 | 9.86 | 12.21 | 9.86 | 10.45
7.648 | 4.98 | 0 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 18.07 | 4.88 | 21 | 4.79 | 15.63 | 4.69 | 18.26 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 0.1 | 17.77 | 0 | 17.68 | 0 | 16.11 | 4.59 | 11.62 | 4.69 | 4.79 | 5.86 | 4.88 | 10.16 | 9.86 | 10.94 | 9.86 | 12.4
7.68 | 5.57 | 0 | 0.1 | 0.59 | 0.1 | 0.78 | 0.1 | 18.07 | 4.79 | 18.95 | 4.79 | 16.21 | 4.79 | 18.75 | 9.67 | 9.86 | 10.16 | 9.77 | 21.49 | 0.1 | 17.77 | 0 | 17.68 | 0.1 | 16.8 | 4.69 | 12.6 | 4.59 | 4.79 | 5.47 | 4.88 | 10.16 | 9.86 | 10.16 | 9.86 | 10.94
7.712 | 6.15 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 9.57 | 4.88 | 18.95 | 4.79 | 16.6 | 4.79 | 19.24 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 17.38 | 4.69 | 13.38 | 4.59 | 4.79 | 5.18 | 4.88 | 10.35 | 9.77 | 10.25 | 9.86 | 9.96
7.744 | 6.84 | 0 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.22 | 4.88 | 18.46 | 4.88 | 17.09 | 4.79 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 18.07 | 4.69 | 14.26 | 4.59 | 4.79 | 5.08 | 4.88 | 10.55 | 9.77 | 10.35 | 9.86 | 9.96
7.776 | 7.32 | 0 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 3.32 | 4.88 | 18.36 | 4.88 | 17.58 | 4.79 | 20.12 | 9.67 | 9.86 | 9.86 | 9.86 | 13.87 | 0.1 | 17.68 | 0 | 17.77 | 0 | 18.85 | 4.69 | 15.04 | 4.59 | 4.79 | 5.18 | 4.88 | 10.45 | 9.86 | 10.45 | 9.86 | 10.45
7.808 | 8.01 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 3.91 | 4.88 | 18.46 | 4.88 | 17.97 | 4.79 | 20.51 | 9.67 | 9.86 | 9.96 | 9.86 | 7.62 | 0.1 | 17.68 | 0 | 17.77 | 0 | 19.34 | 4.69 | 15.63 | 4.49 | 5.76 | 4.98 | 4.88 | 9.96 | 9.77 | 10.25 | 9.86 | 10.45
7.84 | 8.59 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 4.98 | 4.88 | 18.26 | 4.88 | 18.46 | 4.88 | 20.9 | 9.67 | 9.86 | 9.86 | 9.86 | 7.72 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 20.12 | 4.79 | 16.41 | 4.49 | 6.45 | 4.98 | 4.88 | 10.06 | 9.86 | 10.16 | 9.86 | 10.25
7.872 | 9.28 | 0.1 | 0.1 | 0.2 | 0 | 0.29 | 0.1 | 5.96 | 4.88 | 18.16 | 4.79 | 18.95 | 4.88 | 21.29 | 9.67 | 9.86 | 9.86 | 9.86 | 10.64 | 0.1 | 17.58 | 0 | 17.68 | 0 | 20.61 | 4.79 | 17.09 | 4.59 | 8.01 | 4.88 | 4.88 | 10.25 | 9.86 | 10.06 | 9.86 | 10.16
7.904 | 9.77 | 0.1 | 1.17 | 0.2 | 0.1 | 0.29 | 0.1 | 5.66 | 4.88 | 18.16 | 4.88 | 19.34 | 4.79 | 21.78 | 9.67 | 9.96 | 9.77 | 9.86 | 11.52 | 0 | 17.58 | 0 | 17.77 | 0.1 | 21.19 | 4.79 | 17.77 | 4.59 | 10.25 | 4.98 | 4.88 | 10.16 | 9.77 | 10.06 | 9.86 | 10.25
7.936 | 10.45 | 0 | 1.95 | 0.2 | 0.1 | 0.2 | 0.1 | 5.27 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 22.17 | 9.67 | 10.45 | 9.77 | 9.86 | 10.45 | 0.1 | 15.14 | 0 | 17.68 | 0 | 21.88 | 4.79 | 18.46 | 4.59 | 11.43 | 4.79 | 4.88 | 10.16 | 9.86 | 9.96 | 9.86 | 10.06
7.968 | 10.94 | 0 | 3.22 | 0.1 | 0.1 | 0.1 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.22 | 4.88 | 22.46 | 9.67 | 11.52 | 9.77 | 9.86 | 9.86 | 0.1 | 3.81 | 0 | 17.68 | 0 | 22.36 | 4.79 | 19.04 | 4.59 | 12.5 | 4.79 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
8 | 11.52 | 0.1 | 4.2 | 0.1 | 0.1 | 0.1 | 0.1 | 4.98 | 4.79 | 18.07 | 4.88 | 20.61 | 4.88 | 22.85 | 9.77 | 12.89 | 9.77 | 9.86 | 9.96 | 0.1 | 2.05 | 0 | 6.35 | 0.1 | 20.9 | 4.79 | 19.73 | 4.59 | 13.28 | 4.69 | 4.88 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
8.032 | 12.11 | 0.1 | 4.88 | 0 | 0.49 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.79 | 21 | 4.88 | 23.24 | 9.77 | 13.77 | 9.67 | 9.86 | 10.16 | 0.1 | 2.15 | 0 | 1.76 | 0.1 | 17.68 | 4.79 | 20.31 | 4.59 | 14.16 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
8.064 | 12.7 | 0.1 | 5.37 | 0 | 1.37 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 18.95 | 4.88 | 23.63 | 9.77 | 14.55 | 9.67 | 9.86 | 10.16 | 0.1 | 1.86 | 0 | 2.15 | 0.1 | 17.97 | 4.79 | 20.9 | 4.69 | 14.94 | 4.69 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
8.096 | 13.18 | 0 | 5.96 | 0.1 | 2.34 | 0 | 0.68 | 4.88 | 4.88 | 12.11 | 4.88 | 18.95 | 4.88 | 24.02 | 9.77 | 15.04 | 9.67 | 9.86 | 9.86 | 0.1 | 1.76 | 0 | 2.15 | 0 | 17.77 | 4.79 | 21.58 | 4.69 | 15.53 | 4.59 | 4.88 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
8.128 | 13.77 | 0.1 | 6.45 | 0.1 | 3.81 | 0 | 1.37 | 4.69 | 4.88 | 3.52 | 4.88 | 18.56 | 4.88 | 24.32 | 9.77 | 15.63 | 9.67 | 9.86 | 9.96 | 0.1 | 1.66 | 0 | 1.76 | 0 | 17.87 | 4.79 | 22.07 | 4.69 | 16.41 | 4.59 | 5.76 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77
8.16 | 14.26 | 0.1 | 7.13 | 0 | 4.59 | 0.1 | 2.25 | 4.69 | 4.79 | 3.13 | 4.88 | 18.36 | 4.88 | 22.17 | 9.77 | 16.11 | 9.67 | 9.86 | 9.96 | 0.1 | 1.46 | 0 | 1.66 | 0.1 | 17.87 | 4.79 | 22.66 | 4.69 | 16.99 | 4.59 | 6.64 | 9.77 | 9.86 | 9.77 | 10.45 | 9.77
8.192 | 14.75 | 0.1 | 7.72 | 0 | 4.98 | 0.1 | 3.71 | 4.69 | 4.88 | 3.81 | 4.79 | 18.46 | 4.88 | 22.27 | 9.77 | 16.6 | 9.57 | 9.86 | 9.96 | 0.1 | 1.37 | 0 | 1.56 | 0.1 | 17.77 | 4.79 | 23.15 | 4.69 | 17.58 | 4.49 | 8.5 | 9.77 | 10.74 | 9.67 | 11.04 | 9.77
8.224 | 15.23 | 0 | 8.4 | 0 | 5.57 | 0.1 | 4.49 | 4.59 | 4.88 | 4.79 | 4.88 | 18.26 | 4.88 | 21.78 | 9.77 | 17.19 | 9.57 | 9.86 | 9.77 | 0.1 | 1.27 | 0 | 1.46 | 0.1 | 17.77 | 4.79 | 23.73 | 4.79 | 18.26 | 4.59 | 10.25 | 9.67 | 11.23 | 9.67 | 12.7 | 9.77
8.256 | 15.72 | 0 | 8.98 | 0.1 | 6.15 | 0 | 5.08 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 17.58 | 9.57 | 9.86 | 9.77 | 0.1 | 1.27 | 0 | 1.37 | 0 | 17.77 | 4.79 | 24.22 | 4.79 | 18.85 | 4.49 | 11.62 | 9.67 | 12.7 | 9.67 | 14.75 | 9.67
8.288 | 16.21 | 0 | 9.57 | 0.1 | 6.74 | 0 | 5.57 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 18.07 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 1.17 | 0 | 17.77 | 4.79 | 24.71 | 4.79 | 19.63 | 4.49 | 12.7 | 9.67 | 14.84 | 9.67 | 16.11 | 9.67
8.32 | 16.8 | 0.1 | 10.16 | 0.1 | 7.32 | 0.1 | 6.15 | 4.59 | 4.79 | 5.37 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 18.46 | 9.67 | 9.86 | 9.77 | 0.1 | 0.88 | 0 | 1.17 | 0.1 | 17.77 | 4.88 | 25.29 | 4.79 | 20.12 | 4.49 | 13.48 | 9.67 | 16.11 | 9.67 | 17.09 | 9.67
8.352 | 17.19 | 0 | 10.74 | 0 | 8.01 | 0.1 | 6.74 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 18.95 | 9.67 | 9.86 | 9.67 | 0.1 | 0.88 | 0 | 0.98 | 0.1 | 17.77 | 4.79 | 25.78 | 4.79 | 20.8 | 4.49 | 14.26 | 9.67 | 17.09 | 9.67 | 17.87 | 9.67
8.384 | 14.55 | 0.1 | 11.33 | 0 | 8.59 | 0.1 | 7.32 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 19.34 | 9.67 | 9.86 | 9.67 | 0 | 0.78 | 0 | 0.88 | 0.1 | 17.68 | 4.79 | 26.27 | 4.79 | 21.39 | 4.59 | 15.04 | 9.67 | 17.87 | 9.67 | 18.56 | 9.67
8.416 | 15.23 | 0 | 12.01 | 0.1 | 9.08 | 0 | 8.01 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 19.82 | 9.67 | 9.86 | 9.67 | 0 | 0.78 | 0 | 0.88 | 0.1 | 17.68 | 4.88 | 22.75 | 4.79 | 21.97 | 4.59 | 15.72 | 9.57 | 18.56 | 9.57 | 19.24 | 9.67
8.448 | 14.65 | 0 | 12.4 | 0.1 | 9.77 | 0 | 8.5 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 20.22 | 9.67 | 10.35 | 9.67 | 0.1 | 0.59 | 0 | 0.78 | 0 | 7.42 | 4.79 | 21.88 | 4.79 | 22.56 | 4.59 | 16.41 | 9.67 | 19.24 | 9.57 | 19.82 | 9.57
8.48 | 14.65 | 0.1 | 12.99 | 0.1 | 10.35 | 0.1 | 9.08 | 4.59 | 4.79 | 4.79 | 4.88 | 12.31 | 4.88 | 21.39 | 9.86 | 20.7 | 9.67 | 11.04 | 9.67 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 1.95 | 4.88 | 21.58 | 4.79 | 23.15 | 4.59 | 17.09 | 9.67 | 19.92 | 9.67 | 20.51 | 9.67
8.512 | 14.65 | 0.1 | 13.57 | 0 | 10.94 | 0.1 | 9.77 | 4.59 | 4.79 | 4.69 | 4.79 | 3.32 | 4.88 | 21.39 | 9.86 | 21 | 9.67 | 12.01 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.25 | 4.88 | 21.49 | 4.79 | 23.73 | 4.59 | 17.68 | 9.67 | 20.51 | 9.67 | 21.09 | 9.67
8.544 | 14.55 | 0.1 | 14.06 | 0 | 11.52 | 0.1 | 10.35 | 4.59 | 4.88 | 4.69 | 4.88 | 3.32 | 4.88 | 21.39 | 9.86 | 21.49 | 9.67 | 13.38 | 9.67 | 0 | 0.49 | 0.1 | 0.68 | 0.1 | 2.15 | 4.88 | 21.49 | 4.79 | 24.12 | 4.69 | 18.46 | 9.67 | 21 | 9.67 | 21.68 | 9.67
8.576 | 14.55 | 0 | 14.55 | 0.1 | 12.11 | 0 | 10.94 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 4.88 | 13.97 | 9.86 | 21.88 | 9.67 | 14.16 | 9.57 | 0.1 | 0.39 | 0 | 0.59 | 0.1 | 1.66 | 4.79 | 21.49 | 4.79 | 24.81 | 4.69 | 19.14 | 9.67 | 21.78 | 9.57 | 22.17 | 9.67
8.608 | 14.45 | 0.1 | 15.04 | 0.1 | 12.6 | 0 | 11.43 | 4.59 | 4.88 | 4.59 | 4.88 | 4.88 | 4.88 | 7.52 | 9.86 | 22.27 | 9.67 | 14.84 | 9.57 | 0.1 | 0.29 | 0 | 0.49 | 0 | 1.56 | 4.79 | 21.39 | 4.79 | 25.2 | 4.69 | 19.73 | 9.67 | 22.27 | 9.67 | 22.75 | 9.67
8.64 | 14.45 | 0.1 | 15.63 | 0.1 | 13.18 | 0.1 | 12.11 | 4.69 | 4.79 | 4.59 | 4.88 | 5.86 | 4.88 | 7.72 | 9.86 | 22.66 | 9.77 | 15.33 | 9.67 | 0.1 | 0.29 | 0.59 | 0.39 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 25.68 | 4.69 | 20.22 | 9.67 | 22.85 | 9.67 | 23.44 | 9.67
8.672 | 14.45 | 0.1 | 16.11 | 0 | 13.67 | 0.1 | 12.6 | 4.69 | 4.88 | 4.49 | 4.79 | 5.66 | 4.88 | 10.74 | 9.86 | 23.05 | 9.77 | 15.92 | 9.67 | 0.1 | 0.2 | 1.56 | 0.29 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 26.17 | 4.69 | 20.9 | 9.67 | 23.34 | 9.67 | 23.83 | 9.67
8.704 | 14.45 | 0 | 16.6 | 0 | 14.16 | 0.1 | 13.18 | 4.69 | 4.88 | 4.49 | 4.88 | 5.27 | 4.88 | 11.33 | 9.86 | 23.44 | 9.77 | 16.41 | 9.67 | 0.88 | 0.2 | 3.13 | 0.29 | 0.1 | 1.27 | 4.88 | 21.39 | 4.88 | 22.75 | 4.79 | 21.49 | 9.67 | 23.83 | 9.67 | 24.32 | 9.67
8.736 | 14.36 | 0 | 17.09 | 0.1 | 14.75 | 0 | 13.67 | 4.69 | 5.57 | 4.49 | 4.88 | 5.08 | 4.88 | 10.35 | 9.86 | 23.73 | 9.77 | 16.8 | 9.67 | 1.76 | 0.2 | 4.98 | 0.29 | 0.1 | 1.17 | 4.79 | 21.29 | 4.79 | 21.39 | 4.69 | 22.07 | 9.77 | 24.51 | 9.67 | 24.9 | 9.67
8.768 | 14.36 | 0 | 15.23 | 0.1 | 15.33 | 0 | 14.16 | 4.69 | 6.35 | 4.49 | 4.88 | 4.88 | 4.98 | 9.86 | 9.86 | 24.12 | 9.77 | 17.29 | 9.67 | 3.42 | 0.2 | 6.25 | 0.2 | 0 | 1.07 | 4.79 | 21.29 | 4.79 | 21.39 | 4.79 | 22.56 | 9.67 | 25 | 9.67 | 25.39 | 9.67
8.8 | 12.01 | 0.1 | 14.06 | 0.1 | 15.72 | 0.1 | 14.75 | 4.69 | 7.42 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 22.95 | 9.77 | 17.77 | 9.67 | 5.18 | 0.1 | 7.23 | 0.2 | 0.49 | 0.88 | 4.88 | 21.19 | 4.79 | 21.29 | 4.79 | 23.15 | 9.77 | 25.39 | 9.67 | 25.98 | 9.67
8.832 | 3.42 | 0 | 14.36 | 0 | 16.21 | 0.1 | 15.33 | 4.79 | 8.59 | 4.49 | 4.79 | 4.88 | 4.88 | 10.25 | 9.86 | 21.58 | 9.77 | 18.26 | 9.67 | 6.45 | 0.1 | 7.91 | 0.2 | 1.56 | 0.88 | 4.88 | 16.6 | 4.79 | 21.39 | 4.79 | 23.73 | 9.77 | 25.78 | 9.67 | 26.47 | 9.67
8.864 | 1.95 | 0.1 | 14.55 | 0 | 16.7 | 0 | 15.72 | 4.79 | 9.47 | 4.49 | 4.88 | 4.79 | 4.88 | 10.16 | 9.86 | 21.68 | 9.77 | 18.65 | 9.67 | 7.32 | 0.1 | 8.69 | 0.2 | 3.03 | 0.88 | 4.88 | 4.1 | 4.88 | 21.39 | 4.79 | 24.22 | 9.77 | 26.27 | 9.77 | 26.95 | 9.67
8.896 | 1.86 | 0 | 14.36 | 0.1 | 17.19 | 0 | 16.31 | 4.79 | 10.16 | 4.59 | 4.88 | 4.79 | 4.88 | 9.86 | 9.86 | 21.68 | 9.77 | 19.14 | 9.67 | 8.01 | 0.1 | 9.57 | 0.2 | 5.18 | 0.78 | 4.79 | 3.42 | 4.88 | 21.29 | 4.79 | 24.71 | 9.77 | 26.95 | 9.67 | 27.44 | 9.77
8.928 | 1.56 | 0 | 14.45 | 0.1 | 14.26 | 0 | 16.7 | 4.79 | 10.74 | 4.49 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.58 | 9.77 | 19.53 | 9.67 | 8.79 | 0.1 | 10.35 | 0.2 | 6.35 | 0.59 | 4.79 | 3.42 | 4.79 | 21.39 | 4.79 | 25.2 | 9.77 | 27.44 | 9.77 | 27.93 | 9.67
8.96 | 1.46 | 0.1 | 14.45 | 0.1 | 15.33 | 0.1 | 17.19 | 4.79 | 11.43 | 4.59 | 4.88 | 4.59 | 4.88 | 9.96 | 9.86 | 21.49 | 9.86 | 20.02 | 9.77 | 9.57 | 0.1 | 11.23 | 0.1 | 7.23 | 0.59 | 4.88 | 4.3 | 4.88 | 21.19 | 4.79 | 25.59 | 9.77 | 27.83 | 9.77 | 28.42 | 9.77
8.992 | 1.27 | 0.1 | 14.45 | 0 | 14.55 | 0.1 | 14.06 | 4.79 | 11.91 | 4.59 | 4.79 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.86 | 20.31 | 9.67 | 10.35 | 0.1 | 11.91 | 0.1 | 8.01 | 0.59 | 4.88 | 5.86 | 4.79 | 21.29 | 4.79 | 26.27 | 9.77 | 28.32 | 9.77 | 28.71 | 9.77
9.024 | 1.17 | 0.1 | 14.45 | 0 | 14.55 | 0.1 | 14.45 | 4.79 | 12.5 | 4.59 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.86 | 20.8 | 9.77 | 11.13 | 0.1 | 12.7 | 0.1 | 8.79 | 0.49 | 4.88 | 6.25 | 4.88 | 21.19 | 4.79 | 22.85 | 9.77 | 28.81 | 9.77 | 29.3 | 9.77
9.056 | 1.07 | 0 | 14.45 | 0.1 | 14.55 | 0 | 14.55 | 4.79 | 13.09 | 4.59 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 9.67 | 11.82 | 0.1 | 13.48 | 0.1 | 9.67 | 0.39 | 4.79 | 5.76 | 4.88 | 21.19 | 4.79 | 22.07 | 9.77 | 29.4 | 9.77 | 24.51 | 9.77
9.088 | 0.88 | 0.1 | 14.36 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 13.67 | 4.69 | 4.88 | 4.59 | 4.88 | 9.67 | 9.86 | 21.49 | 9.86 | 21.68 | 9.77 | 12.7 | 0.1 | 14.16 | 0.1 | 10.35 | 0.29 | 4.79 | 5.37 | 4.88 | 21.19 | 4.79 | 21.68 | 9.77 | 24.61 | 9.77 | 23.34 | 9.77
9.12 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 14.16 | 4.69 | 5.66 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 21.97 | 9.77 | 13.48 | 0.1 | 14.94 | 0.1 | 11.04 | 0.29 | 4.88 | 5.27 | 4.79 | 18.46 | 4.79 | 21.58 | 9.77 | 23.54 | 9.77 | 24.12 | 9.77
9.152 | 0.78 | 0 | 14.45 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 14.65 | 4.69 | 6.05 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 22.36 | 9.77 | 14.26 | 0.1 | 15.53 | 0.1 | 11.91 | 0.29 | 4.88 | 5.18 | 4.79 | 4.2 | 4.79 | 21.58 | 9.77 | 24.32 | 9.77 | 24.12 | 9.77
9.184 | 0.68 | 0.1 | 14.45 | 0 | 14.45 | 0 | 14.55 | 4.88 | 15.14 | 4.69 | 7.03 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 9.77 | 14.84 | 0.1 | 16.21 | 0.1 | 12.79 | 0.29 | 4.88 | 5.08 | 4.88 | 3.13 | 4.79 | 21.49 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
9.216 | 0.59 | 0 | 5.27 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 15.63 | 4.69 | 8.5 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 9.77 | 15.63 | 0.1 | 16.9 | 0.1 | 13.48 | 0.2 | 5.76 | 5.08 | 4.79 | 3.52 | 4.79 | 21.39 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
9.248 | 0.59 | 0 | 1.66 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 16.21 | 4.79 | 9.38 | 4.49 | 4.88 | 9.67 | 9.86 | 12.4 | 9.86 | 23.54 | 9.77 | 16.21 | 0.1 | 17.48 | 0.1 | 14.16 | 0.2 | 6.84 | 4.98 | 4.88 | 4.3 | 4.79 | 21.39 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
9.28 | 0.49 | 0.1 | 2.05 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 16.7 | 4.79 | 10.06 | 4.49 | 4.88 | 9.67 | 9.86 | 6.54 | 9.86 | 24.02 | 9.86 | 16.99 | 0 | 18.16 | 0.1 | 14.94 | 0.2 | 8.79 | 4.98 | 4.88 | 5.76 | 4.79 | 21.29 | 9.86 | 24.32 | 9.77 | 24.32 | 9.86
9.312 | 0.49 | 0 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.09 | 4.79 | 10.64 | 4.49 | 5.57 | 9.67 | 9.86 | 8.5 | 9.86 | 24.32 | 9.86 | 17.58 | 0.1 | 18.85 | 0.1 | 15.63 | 0.1 | 10.55 | 4.88 | 4.79 | 6.25 | 4.79 | 21.29 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
9.344 | 0.39 | 0.1 | 1.56 | 0 | 13.18 | 0 | 14.45 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 6.54 | 9.57 | 9.86 | 11.23 | 9.86 | 21.09 | 9.86 | 18.26 | 0.1 | 19.43 | 0.1 | 16.41 | 0.1 | 11.91 | 4.88 | 4.88 | 5.76 | 4.79 | 21.19 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86
9.376 | 0.29 | 0 | 1.37 | 0.1 | 4 | 0 | 14.36 | 4.88 | 17.97 | 4.79 | 11.82 | 4.59 | 7.72 | 9.57 | 9.86 | 11.13 | 9.86 | 21.39 | 9.86 | 18.85 | 0.1 | 20.12 | 0.1 | 16.99 | 0.1 | 12.79 | 4.79 | 4.88 | 5.27 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
9.408 | 0.29 | 0.1 | 1.17 | 0.1 | 1.95 | 0 | 14.06 | 4.88 | 18.46 | 4.79 | 12.4 | 4.59 | 8.79 | 9.67 | 9.86 | 10.35 | 9.86 | 21.39 | 9.77 | 19.43 | 0.1 | 20.7 | 0.1 | 17.58 | 0.1 | 13.67 | 4.69 | 4.88 | 5.18 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
9.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 3.91 | 4.88 | 18.95 | 4.79 | 13.09 | 4.69 | 9.67 | 9.57 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.22 | 0.1 | 21.39 | 0.1 | 18.26 | 0.1 | 14.36 | 4.69 | 4.79 | 5.08 | 4.79 | 14.45 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
9.472 | 0.2 | 0.1 | 1.07 | 0 | 1.56 | 0.1 | 1.95 | 4.88 | 19.34 | 4.79 | 13.48 | 4.69 | 10.35 | 9.57 | 9.86 | 10.06 | 9.86 | 21.58 | 9.86 | 20.8 | 0.1 | 21.88 | 0.1 | 18.95 | 0.1 | 15.23 | 4.69 | 4.88 | 5.18 | 4.79 | 3.13 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
9.504 | 0.1 | 0 | 0.98 | 0 | 1.46 | 0.1 | 1.95 | 4.88 | 19.92 | 4.88 | 14.06 | 4.69 | 11.04 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 9.86 | 21.39 | 0.1 | 22.46 | 0.1 | 19.43 | 0.1 | 16.02 | 4.59 | 4.88 | 4.98 | 4.79 | 3.32 | 9.86 | 23.93 | 9.86 | 9.57 | 9.86
9.536 | 0.1 | 0 | 0.88 | 0.1 | 1.37 | 0 | 1.56 | 4.88 | 20.22 | 4.79 | 14.55 | 4.69 | 11.52 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 9.86 | 21.88 | 0.1 | 18.07 | 0.1 | 20.22 | 0.1 | 16.6 | 4.59 | 4.88 | 5.08 | 4.79 | 3.61 | 9.86 | 8.69 | 9.77 | 6.05 | 9.86
9.568 | 0 | 0.49 | 0.68 | 0.1 | 1.17 | 0 | 1.56 | 4.79 | 20.7 | 4.79 | 15.14 | 4.69 | 12.11 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 9.86 | 22.46 | 0.1 | 17.09 | 0.1 | 20.8 | 0.1 | 17.29 | 4.59 | 4.79 | 4.98 | 4.79 | 4.49 | 9.77 | 6.25 | 9.86 | 9.67 | 9.86
9.6 | 0.1 | 1.27 | 0.68 | 0 | 1.07 | 0.1 | 1.37 | 4.88 | 21 | 4.88 | 15.63 | 4.79 | 12.7 | 9.67 | 11.33 | 10.06 | 9.86 | 21.39 | 9.86 | 19.14 | 0.1 | 17.77 | 0.1 | 21.29 | 0.1 | 17.97 | 4.59 | 4.79 | 4.98 | 4.79 | 5.96 | 9.86 | 9.86 | 9.86 | 12.31 | 9.86
9.632 | 0.1 | 2.34 | 0.68 | 0 | 0.98 | 0.1 | 1.27 | 4.88 | 19.14 | 4.79 | 16.02 | 4.79 | 13.28 | 9.67 | 12.6 | 9.86 | 9.86 | 21.49 | 9.86 | 18.36 | 0.1 | 17.58 | 0.1 | 21.88 | 0.1 | 18.56 | 4.59 | 4.79 | 4.79 | 4.79 | 6.05 | 9.86 | 12.11 | 9.86 | 11.13 | 9.86
9.664 | 0.1 | 3.71 | 0.59 | 0 | 0.78 | 0.1 | 1.07 | 4.88 | 18.95 | 4.79 | 16.6 | 4.79 | 13.77 | 9.67 | 13.57 | 9.86 | 9.86 | 21.39 | 9.86 | 18.07 | 0.1 | 17.68 | 0.1 | 22.56 | 0.1 | 19.24 | 4.59 | 5.08 | 4.79 | 4.79 | 5.66 | 9.86 | 11.13 | 9.86 | 10.16 | 9.86
9.696 | 0 | 4.39 | 0.49 | 0.1 | 0.78 | 0 | 0.98 | 4.88 | 18.56 | 4.88 | 17.09 | 4.79 | 14.26 | 9.67 | 14.45 | 9.86 | 9.86 | 21.39 | 9.86 | 17.87 | 0.1 | 17.68 | 0.1 | 18.36 | 0.1 | 19.92 | 4.59 | 6.05 | 4.69 | 4.79 | 5.27 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
9.728 | 0.1 | 4.98 | 0.39 | 0.1 | 0.68 | 0 | 0.88 | 4.88 | 18.46 | 4.88 | 17.48 | 4.79 | 14.84 | 9.67 | 14.94 | 9.77 | 9.86 | 21.39 | 9.86 | 17.87 | 0.1 | 17.68 | 0.1 | 17.19 | 0.1 | 20.41 | 4.59 | 7.72 | 4.69 | 4.79 | 5.18 | 9.77 | 9.96 | 9.86 | 10.35 | 9.86
9.76 | 0.1 | 5.47 | 0.39 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 18.46 | 4.88 | 17.97 | 4.79 | 15.33 | 9.67 | 15.53 | 9.67 | 9.86 | 18.65 | 9.96 | 17.77 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 21 | 4.59 | 9.86 | 4.69 | 4.79 | 5.18 | 9.86 | 10.35 | 9.86 | 10.45 | 9.86
9.792 | 0.1 | 6.15 | 0.29 | 0 | 0.59 | 0.1 | 0.78 | 4.88 | 18.26 | 4.79 | 18.46 | 4.79 | 15.82 | 9.67 | 16.02 | 9.67 | 9.96 | 8.79 | 9.96 | 17.68 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.68 | 4.59 | 11.33 | 4.59 | 4.79 | 5.18 | 9.86 | 10.45 | 9.86 | 10.35 | 9.86
9.824 | 0.1 | 6.64 | 0.29 | 0 | 0.49 | 0.1 | 0.68 | 4.88 | 18.26 | 4.79 | 18.85 | 4.79 | 16.31 | 9.67 | 16.5 | 9.67 | 9.86 | 7.13 | 9.96 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.17 | 4.59 | 12.4 | 4.49 | 4.79 | 5.08 | 9.86 | 10.35 | 9.86 | 10.16 | 9.86
9.856 | 0 | 7.23 | 0.2 | 0.1 | 0.49 | 0 | 0.49 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 16.7 | 9.67 | 16.99 | 9.67 | 9.86 | 10.16 | 9.86 | 17.58 | 0.1 | 17.58 | 0.2 | 17.77 | 0 | 22.75 | 4.59 | 13.28 | 4.59 | 4.79 | 4.98 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
9.888 | 0.1 | 7.91 | 0.1 | 0.1 | 0.39 | 0 | 0.59 | 4.88 | 18.16 | 4.88 | 19.73 | 4.88 | 17.29 | 9.67 | 17.48 | 9.67 | 9.86 | 11.52 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 23.15 | 4.59 | 14.06 | 4.59 | 4.79 | 4.98 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
9.92 | 0.1 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 17.68 | 9.77 | 17.97 | 9.67 | 9.96 | 10.64 | 9.96 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 23.83 | 4.69 | 14.84 | 4.59 | 4.79 | 4.88 | 9.86 | 10.25 | 9.86 | 10.16 | 9.86
9.952 | 0.1 | 9.08 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 18.07 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 9.96 | 17.58 | 0.1 | 15.43 | 0.1 | 17.68 | 0.1 | 24.32 | 4.69 | 15.53 | 4.49 | 4.79 | 4.88 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86
9.984 | 0.1 | 9.77 | 0.1 | 0.78 | 0.2 | 0.1 | 0.29 | 4.88 | 18.07 | 4.79 | 20.9 | 4.88 | 18.56 | 9.77 | 18.95 | 9.67 | 9.86 | 9.86 | 9.96 | 17.58 | 0.1 | 4.1 | 0.1 | 17.77 | 0.1 | 24.81 | 4.69 | 16.21 | 4.49 | 4.79 | 4.79 | 9.86 | 9.86 | 9.86 | 9.96 | 9.86
10.016 | 0 | 10.25 | 0.1 | 1.66 | 0.2 | 0.59 | 0.29 | 4.88 | 18.07 | 4.88 | 17.97 | 4.88 | 19.04 | 9.77 | 19.24 | 9.57 | 9.86 | 10.06 | 9.96 | 12.7 | 0.1 | 1.95 | 0.1 | 17.68 | 0.1 | 25.49 | 4.69 | 16.9 | 4.59 | 4.79 | 4.79 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
10.048 | 0 | 10.84 | 0.1 | 2.93 | 0.1 | 1.37 | 0.2 | 4.88 | 11.13 | 4.88 | 18.46 | 4.88 | 19.43 | 9.77 | 19.73 | 9.57 | 9.86 | 10.16 | 9.96 | 3.13 | 0.1 | 2.25 | 0.1 | 17.68 | 0.1 | 25.88 | 4.79 | 17.58 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
10.08 | 0.1 | 11.43 | 0.1 | 3.91 | 0.1 | 2.64 | 0.2 | 4.88 | 3.22 | 4.88 | 18.26 | 4.88 | 19.82 | 9.77 | 20.22 | 9.57 | 9.96 | 10.06 | 9.96 | 2.15 | 0.1 | 1.76 | 0.1 | 17.68 | 0.1 | 26.37 | 4.79 | 18.26 | 4.59 | 4.79 | 4.69 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
10.112 | 0.1 | 12.01 | 0.1 | 4.59 | 0 | 3.81 | 0.2 | 4.88 | 3.03 | 4.88 | 18.16 | 4.88 | 20.22 | 9.77 | 20.61 | 9.57 | 10.16 | 9.86 | 9.96 | 2.25 | 0 | 1.66 | 0.1 | 17.09 | 0.1 | 22.75 | 4.79 | 18.85 | 4.59 | 4.98 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
10.144 | 0.1 | 12.6 | 0.1 | 5.27 | 0.1 | 4.59 | 0.1 | 4.88 | 3.81 | 4.79 | 18.26 | 4.88 | 20.7 | 9.77 | 21 | 9.67 | 11.04 | 9.86 | 9.96 | 1.56 | 0.1 | 1.56 | 0.1 | 4.1 | 0.1 | 22.07 | 4.79 | 19.53 | 4.59 | 5.96 | 4.59 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
10.176 | 0 | 13.09 | 0.1 | 5.86 | 0.1 | 5.08 | 0.1 | 4.88 | 4.88 | 4.88 | 18.16 | 4.88 | 20.61 | 9.77 | 21.49 | 9.67 | 12.01 | 9.86 | 9.96 | 1.76 | 0.1 | 1.46 | 0.1 | 2.25 | 0.1 | 21.58 | 4.79 | 20.12 | 4.59 | 7.42 | 4.59 | 10.74 | 9.77 | 9.96 | 9.77 | 9.86
10.208 | 0 | 13.77 | 0.1 | 6.45 | 0.1 | 5.66 | 0.1 | 4.88 | 5.86 | 4.88 | 18.26 | 4.88 | 18.56 | 9.77 | 21.88 | 9.67 | 13.28 | 9.77 | 9.96 | 1.56 | 0.1 | 1.27 | 0.1 | 1.95 | 0.1 | 21.58 | 4.79 | 20.7 | 4.69 | 9.57 | 4.59 | 12.01 | 9.77 | 10.94 | 9.67 | 10.16
10.24 | 0.1 | 14.16 | 0.1 | 7.03 | 0.1 | 6.15 | 0.1 | 4.88 | 5.66 | 4.88 | 18.26 | 4.88 | 18.46 | 9.77 | 22.17 | 9.67 | 14.16 | 9.67 | 9.96 | 1.37 | 0.1 | 1.27 | 0.1 | 1.95 | 0.1 | 21.49 | 4.88 | 21.29 | 4.69 | 11.23 | 4.59 | 13.97 | 9.77 | 12.5 | 9.67 | 11.23
10.272 | 0.1 | 14.65 | 0.1 | 7.62 | 0.1 | 6.84 | 0.1 | 4.88 | 5.37 | 4.88 | 18.07 | 4.88 | 18.46 | 9.77 | 22.66 | 9.67 | 14.75 | 9.77 | 9.96 | 1.46 | 0.1 | 1.17 | 0.1 | 1.66 | 0.1 | 21.39 | 4.88 | 21.88 | 4.69 | 12.31 | 4.49 | 15.53 | 9.67 | 14.55 | 9.67 | 12.79
10.304 | 0.1 | 15.23 | 0.1 | 8.2 | 0.1 | 7.42 | 0.1 | 4.88 | 5.08 | 4.88 | 18.16 | 4.88 | 18.16 | 9.77 | 22.95 | 9.67 | 15.33 | 9.67 | 9.96 | 1.17 | 0.1 | 1.07 | 0.1 | 1.56 | 0.1 | 21.29 | 4.88 | 22.56 | 4.69 | 13.09 | 4.59 | 16.7 | 9.67 | 16.02 | 9.67 | 14.84
10.336 | 0 | 15.72 | 0.1 | 8.89 | 0.1 | 8.01 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 9.77 | 23.34 | 9.67 | 15.82 | 9.67 | 9.96 | 1.17 | 0.1 | 0.98 | 0.1 | 1.56 | 0.1 | 21.29 | 4.88 | 23.05 | 4.69 | 13.97 | 4.59 | 17.68 | 9.67 | 17.09 | 9.67 | 16.31
10.368 | 0 | 16.11 | 0.1 | 9.47 | 0.1 | 8.59 | 0.1 | 4.79 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 23.73 | 9.67 | 16.31 | 9.67 | 9.96 | 1.07 | 0.1 | 0.88 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 23.54 | 4.79 | 14.75 | 4.59 | 18.26 | 9.67 | 17.77 | 9.67 | 17.19
10.4 | 0.1 | 16.7 | 0.1 | 9.96 | 0.1 | 9.28 | 0.1 | 5.57 | 4.79 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 24.12 | 9.67 | 16.8 | 9.67 | 9.96 | 0.88 | 0 | 0.78 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 24.02 | 4.79 | 15.53 | 4.59 | 18.95 | 9.67 | 18.46 | 9.67 | 17.97
10.432 | 0.1 | 17.09 | 0.1 | 10.64 | 0.1 | 9.77 | 0.1 | 6.35 | 4.79 | 4.79 | 15.33 | 4.88 | 18.16 | 9.77 | 24.41 | 9.67 | 17.29 | 9.57 | 9.96 | 0.88 | 0 | 0.68 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.61 | 4.79 | 16.21 | 4.59 | 19.63 | 9.67 | 19.14 | 9.67 | 18.56
10.464 | 0.1 | 13.97 | 0.1 | 11.13 | 0.1 | 10.45 | 0.1 | 7.52 | 4.79 | 4.88 | 4.49 | 4.88 | 18.07 | 9.77 | 22.36 | 9.77 | 17.77 | 9.57 | 10.55 | 0.68 | 0.1 | 0.59 | 0.1 | 0.98 | 0.1 | 21.19 | 4.88 | 25.1 | 4.79 | 16.8 | 4.69 | 20.12 | 9.57 | 19.82 | 9.67 | 19.24
10.496 | 0 | 14.75 | 0.1 | 11.82 | 0.1 | 11.04 | 0.1 | 8.69 | 4.69 | 4.88 | 3.22 | 4.88 | 18.07 | 9.77 | 21.88 | 9.77 | 18.26 | 9.57 | 11.33 | 0.68 | 0.1 | 0.59 | 0.1 | 0.98 | 0 | 21.19 | 4.88 | 25.59 | 4.79 | 17.48 | 4.69 | 20.9 | 9.67 | 20.51 | 9.67 | 20.02
10.528 | 0.1 | 14.36 | 0.1 | 12.31 | 0.1 | 11.52 | 0.1 | 9.57 | 4.69 | 4.88 | 3.71 | 4.88 | 18.07 | 9.77 | 21.78 | 9.77 | 18.65 | 9.57 | 12.4 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 0 | 9.77 | 4.88 | 26.08 | 4.79 | 18.07 | 4.69 | 21.49 | 9.67 | 20.9 | 9.67 | 20.51
10.56 | 0.1 | 14.45 | 0.1 | 12.89 | 0.1 | 12.11 | 0.1 | 10.16 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 19.14 | 9.57 | 13.57 | 0.59 | 0.1 | 0.39 | 0.1 | 0.78 | 0.1 | 2.64 | 4.88 | 21.49 | 4.79 | 18.85 | 4.69 | 22.07 | 9.67 | 21.58 | 9.67 | 21
10.592 | 0.1 | 14.45 | 0.1 | 13.38 | 0.1 | 12.7 | 0.1 | 10.84 | 4.59 | 4.79 | 5.76 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 19.53 | 9.57 | 14.36 | 0.49 | 0.1 | 0.29 | 0.1 | 0.78 | 0.1 | 3.32 | 4.88 | 20.31 | 4.79 | 19.34 | 4.69 | 22.56 | 9.67 | 22.17 | 9.67 | 21.68
10.624 | 0.1 | 14.45 | 0.1 | 13.97 | 0.1 | 13.28 | 0.1 | 11.43 | 4.59 | 4.79 | 5.96 | 4.88 | 8.3 | 9.77 | 21.58 | 9.77 | 19.92 | 9.57 | 14.94 | 0.39 | 0.39 | 0.29 | 0.1 | 0.59 | 0.1 | 3.71 | 4.88 | 21 | 4.79 | 20.02 | 4.79 | 23.15 | 9.67 | 22.75 | 9.67 | 22.27
10.656 | 0 | 14.45 | 0.1 | 14.45 | 0.1 | 13.77 | 0.1 | 12.01 | 4.59 | 4.88 | 5.37 | 4.88 | 2.83 | 9.77 | 21.49 | 9.77 | 20.51 | 9.57 | 15.43 | 0.39 | 1.46 | 0.29 | 0.2 | 0.59 | 0 | 4.69 | 4.88 | 21.19 | 4.79 | 20.7 | 4.79 | 23.63 | 9.67 | 23.34 | 9.67 | 22.85
10.688 | 0.1 | 14.45 | 0.1 | 14.94 | 0.1 | 14.26 | 0.1 | 12.6 | 4.59 | 4.88 | 5.08 | 4.88 | 3.32 | 9.77 | 21.49 | 9.77 | 20.9 | 9.57 | 16.02 | 0.29 | 3.03 | 0.2 | 1.17 | 0.49 | 0 | 5.96 | 4.88 | 21.19 | 4.88 | 21.29 | 4.79 | 24.32 | 9.67 | 23.83 | 9.67 | 23.34
10.72 | 0.1 | 14.45 | 0.1 | 15.43 | 0.1 | 14.84 | 0.1 | 13.18 | 4.49 | 4.88 | 4.98 | 4.88 | 4.1 | 9.77 | 21.49 | 9.77 | 21.19 | 9.57 | 16.5 | 0.29 | 5.08 | 0.1 | 2.54 | 0.49 | 0.1 | 6.05 | 4.88 | 21.29 | 4.79 | 21.88 | 4.79 | 24.81 | 9.67 | 24.32 | 9.67 | 23.93
10.752 | 0.1 | 14.45 | 0.1 | 15.92 | 0.1 | 15.33 | 0.1 | 13.67 | 4.49 | 4.79 | 4.98 | 4.88 | 5.37 | 9.77 | 21.49 | 9.77 | 21.49 | 9.57 | 16.99 | 0.2 | 6.25 | 0.1 | 4.49 | 0.39 | 0.1 | 5.66 | 4.88 | 21.29 | 4.79 | 22.36 | 4.79 | 25.29 | 9.67 | 24.9 | 9.67 | 24.51
10.784 | 0.1 | 14.36 | 0.1 | 16.5 | 0.1 | 15.82 | 0.1 | 14.16 | 4.59 | 5.57 | 4.98 | 4.88 | 5.96 | 9.77 | 21.49 | 9.86 | 21.97 | 9.67 | 17.48 | 0.2 | 7.13 | 0.1 | 5.96 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.79 | 22.95 | 4.79 | 25.78 | 9.67 | 25.39 | 9.77 | 24.9
10.816 | 0 | 14.36 | 0.1 | 16.99 | 0.1 | 16.31 | 0.1 | 14.84 | 4.49 | 6.15 | 4.98 | 4.88 | 5.66 | 9.77 | 21.39 | 9.86 | 22.36 | 9.67 | 17.87 | 0.2 | 8.01 | 0.1 | 6.84 | 0.29 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 23.44 | 4.79 | 26.37 | 9.77 | 25.88 | 9.77 | 25.49
10.848 | 0 | 14.36 | 0.1 | 16.9 | 0.1 | 16.8 | 0.1 | 15.14 | 4.59 | 7.23 | 4.88 | 4.88 | 5.27 | 9.77 | 21.39 | 9.86 | 22.75 | 9.67 | 18.26 | 0.2 | 8.69 | 0.1 | 7.72 | 0.2 | 0.1 | 4.98 | 4.88 | 21.29 | 4.79 | 24.02 | 4.79 | 26.76 | 9.77 | 26.37 | 9.77 | 25.88
10.88 | 0.1 | 13.77 | 0.1 | 14.94 | 0.1 | 17.19 | 0.1 | 15.72 | 4.49 | 8.59 | 4.79 | 4.88 | 4.98 | 9.77 | 13.18 | 9.86 | 23.15 | 9.67 | 18.75 | 0.1 | 9.57 | 0 | 8.4 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 24.51 | 4.79 | 27.25 | 9.77 | 26.76 | 9.77 | 26.47
10.912 | 0.1 | 3.81 | 0.1 | 14.55 | 0.1 | 15.23 | 0.1 | 16.21 | 4.59 | 9.47 | 4.69 | 4.88 | 4.98 | 9.77 | 7.13 | 9.86 | 23.54 | 9.67 | 19.24 | 0.1 | 10.25 | 0 | 9.47 | 0.2 | 0.1 | 4.98 | 4.88 | 21.19 | 4.88 | 25.1 | 4.79 | 27.74 | 9.77 | 27.34 | 9.77 | 26.86
10.944 | 0.1 | 1.76 | 0.1 | 14.75 | 0.1 | 15.23 | 0.1 | 16.7 | 4.59 | 10.06 | 4.69 | 4.88 | 4.88 | 9.77 | 8.3 | 9.86 | 23.93 | 9.67 | 19.73 | 0.1 | 11.13 | 0 | 10.16 | 0.2 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 25.49 | 4.79 | 28.22 | 9.77 | 27.83 | 9.77 | 27.34
10.976 | 0 | 1.76 | 0.1 | 14.65 | 0.1 | 14.84 | 0.1 | 17.19 | 4.59 | 10.64 | 4.69 | 4.88 | 4.88 | 9.77 | 11.23 | 9.86 | 24.32 | 9.67 | 20.12 | 0.1 | 11.82 | 0 | 10.94 | 0.1 | 0.1 | 4.88 | 4.88 | 21.19 | 4.79 | 26.08 | 4.88 | 28.71 | 9.77 | 28.32 | 9.77 | 28.03
11.008 | 0 | 1.56 | 0.1 | 14.65 | 0.1 | 14.65 | 0.1 | 17.58 | 4.69 | 11.33 | 4.59 | 4.88 | 4.88 | 9.77 | 11.13 | 9.86 | 20.9 | 9.67 | 20.51 | 0.1 | 12.6 | 0.1 | 11.72 | 0.1 | 0 | 4.79 | 4.88 | 5.86 | 4.88 | 21.78 | 4.79 | 29.01 | 9.77 | 28.71 | 9.77 | 28.32
11.04 | 0.1 | 1.46 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 18.07 | 4.69 | 11.82 | 4.59 | 4.88 | 4.69 | 9.77 | 10.35 | 9.86 | 21.09 | 9.67 | 20.9 | 0.1 | 13.38 | 0 | 12.5 | 0.1 | 0.1 | 4.79 | 4.88 | 2.83 | 4.88 | 20.22 | 4.88 | 28.52 | 9.77 | 29.2 | 9.77 | 28.81
11.072 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.56 | 4.69 | 12.5 | 4.49 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.29 | 9.67 | 21.39 | 0.1 | 14.16 | 0 | 13.28 | 0.1 | 0.1 | 4.69 | 4.88 | 3.42 | 4.88 | 20.9 | 4.79 | 24.81 | 9.77 | 24.61 | 9.77 | 29.3
11.104 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 19.04 | 4.79 | 12.99 | 4.59 | 4.88 | 4.69 | 9.77 | 10.16 | 9.86 | 21.39 | 9.67 | 21.68 | 0.1 | 14.84 | 0 | 13.87 | 0.1 | 0.1 | 4.69 | 4.88 | 4.1 | 4.88 | 21.19 | 4.88 | 24.61 | 9.77 | 23.34 | 9.77 | 24.61
11.136 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.43 | 4.79 | 13.57 | 4.49 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 22.17 | 0.1 | 15.53 | 0.1 | 14.65 | 0.1 | 1.07 | 4.59 | 4.88 | 5.37 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 23.93 | 9.77 | 23.83
11.168 | 0 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.82 | 4.79 | 14.06 | 4.59 | 5.66 | 4.59 | 9.77 | 10.25 | 9.86 | 21.49 | 9.67 | 22.46 | 0.1 | 16.21 | 0 | 15.33 | 0 | 2.34 | 4.59 | 4.88 | 6.15 | 4.79 | 21.19 | 4.79 | 24.41 | 9.77 | 24.12 | 9.86 | 24.12
11.2 | 0.1 | 0.78 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 20.31 | 4.79 | 14.65 | 4.49 | 6.15 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.67 | 22.95 | 0.1 | 16.99 | 0 | 16.02 | 0.1 | 4.2 | 4.59 | 4.88 | 5.86 | 4.88 | 21.19 | 4.88 | 24.41 | 9.77 | 24.12 | 9.77 | 24.12
11.232 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 20.7 | 4.79 | 15.14 | 4.49 | 7.32 | 4.59 | 9.77 | 10.06 | 9.86 | 21.49 | 9.77 | 23.24 | 0.1 | 17.58 | 0 | 16.9 | 0.1 | 5.66 | 4.59 | 4.88 | 5.47 | 4.88 | 21.19 | 4.79 | 24.41 | 9.77 | 24.12 | 9.86 | 24.22
11.264 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 21.09 | 4.88 | 15.53 | 4.59 | 8.5 | 4.49 | 10.45 | 9.96 | 9.86 | 21.49 | 9.77 | 23.73 | 0.1 | 18.16 | 0.1 | 17.48 | 0.1 | 6.74 | 4.59 | 4.88 | 5.27 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22
11.296 | 0 | 0.59 | 0.1 | 7.62 | 0.1 | 14.36 | 0.1 | 19.04 | 4.88 | 16.11 | 4.59 | 9.47 | 4.59 | 11.13 | 9.86 | 9.86 | 21.49 | 9.77 | 24.02 | 0.1 | 18.85 | 0.1 | 18.16 | 0.1 | 7.62 | 4.49 | 4.88 | 5.18 | 4.79 | 21.29 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
11.328 | 0.1 | 0.49 | 0.1 | 2.05 | 0.1 | 10.35 | 0.1 | 18.95 | 4.79 | 16.6 | 4.69 | 10.16 | 4.59 | 12.21 | 9.77 | 9.86 | 21.39 | 9.77 | 24.41 | 0.1 | 19.53 | 0 | 18.75 | 0 | 8.4 | 4.49 | 4.88 | 5.18 | 4.88 | 21.19 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
11.36 | 0.1 | 0.49 | 0.1 | 1.95 | 0.1 | 2.83 | 0.1 | 18.56 | 4.79 | 16.99 | 4.59 | 10.64 | 4.49 | 13.38 | 9.86 | 9.86 | 21.49 | 9.77 | 21.88 | 0.1 | 20.12 | 0 | 19.43 | 0.1 | 9.28 | 4.49 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
11.392 | 0.1 | 0.39 | 0.1 | 1.76 | 0.1 | 2.05 | 0.1 | 18.46 | 4.79 | 17.58 | 4.69 | 11.33 | 4.59 | 14.16 | 9.86 | 9.86 | 21.39 | 9.77 | 22.27 | 0.1 | 20.8 | 0 | 20.02 | 0.1 | 9.96 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
11.424 | 0.1 | 0.29 | 0.1 | 1.46 | 0.1 | 1.95 | 0.1 | 18.46 | 4.88 | 17.97 | 4.69 | 11.91 | 4.49 | 14.84 | 9.77 | 9.86 | 21.19 | 9.77 | 21.78 | 0.1 | 21.29 | 0 | 20.61 | 0.1 | 10.84 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
11.456 | 0 | 0.2 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 18.26 | 4.88 | 18.46 | 4.69 | 12.4 | 4.59 | 15.33 | 9.67 | 9.86 | 8.79 | 9.77 | 21.78 | 0.1 | 21.88 | 0 | 21.19 | 0 | 11.72 | 4.49 | 4.88 | 5.08 | 4.79 | 6.64 | 4.88 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
11.488 | 0.59 | 0.2 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 18.16 | 4.88 | 18.85 | 4.79 | 12.99 | 4.59 | 15.82 | 9.67 | 9.86 | 7.03 | 9.77 | 21.68 | 0.1 | 22.46 | 0 | 21.78 | 0 | 12.5 | 4.49 | 4.88 | 4.79 | 4.88 | 2.44 | 4.79 | 12.5 | 9.86 | 24.12 | 9.86 | 24.12
11.52 | 1.37 | 0.2 | 0.1 | 1.17 | 0 | 1.27 | 0.1 | 18.26 | 4.79 | 19.34 | 4.79 | 13.48 | 4.59 | 16.41 | 9.67 | 9.86 | 9.86 | 9.77 | 21.68 | 0.1 | 18.36 | 0 | 22.46 | 0.1 | 13.18 | 4.59 | 5.37 | 4.79 | 4.88 | 3.42 | 4.79 | 5.86 | 9.86 | 10.45 | 9.86 | 23.93
11.552 | 2.25 | 0.1 | 0.1 | 1.07 | 0.1 | 1.17 | 0.1 | 18.16 | 4.88 | 19.73 | 4.69 | 14.06 | 4.59 | 16.9 | 9.67 | 9.86 | 11.62 | 9.77 | 21.58 | 0.1 | 16.9 | 0 | 19.73 | 0.1 | 13.97 | 4.59 | 6.45 | 4.69 | 4.79 | 4.2 | 4.79 | 8.01 | 9.86 | 5.86 | 9.86 | 8.79
11.584 | 3.71 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 18.16 | 4.88 | 20.12 | 4.79 | 14.65 | 4.69 | 17.38 | 9.67 | 9.86 | 10.94 | 9.77 | 21.58 | 0.1 | 17.58 | 0 | 17.09 | 0.1 | 14.65 | 4.59 | 8.01 | 4.69 | 4.79 | 5.37 | 4.88 | 11.82 | 9.77 | 9.08 | 9.86 | 6.25
11.616 | 4.49 | 0 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 18.07 | 4.88 | 20.61 | 4.79 | 15.14 | 4.69 | 17.77 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 0.1 | 17.58 | 0 | 17.87 | 0.1 | 15.43 | 4.69 | 10.16 | 4.69 | 4.79 | 6.25 | 4.88 | 11.43 | 9.86 | 12.31 | 9.86 | 10.35
11.648 | 4.98 | 0 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 18.07 | 4.88 | 21 | 4.88 | 15.53 | 4.69 | 18.26 | 9.67 | 9.86 | 9.77 | 9.86 | 21.49 | 0.1 | 17.68 | 0 | 17.87 | 0 | 16.11 | 4.59 | 11.52 | 4.59 | 4.88 | 5.86 | 4.79 | 10.25 | 9.86 | 11.23 | 9.86 | 12.31
11.68 | 5.47 | 0 | 0.1 | 0.68 | 0 | 0.68 | 0.1 | 18.07 | 4.88 | 18.16 | 4.79 | 16.11 | 4.69 | 18.65 | 9.67 | 9.86 | 10.06 | 9.77 | 21.49 | 0.1 | 17.68 | 0 | 17.87 | 0.1 | 16.8 | 4.69 | 12.6 | 4.59 | 4.79 | 5.47 | 4.88 | 9.96 | 9.86 | 10.25 | 9.86 | 11.13
11.712 | 6.15 | 0 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 10.16 | 4.88 | 18.95 | 4.79 | 16.6 | 4.69 | 19.24 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 17.48 | 4.69 | 13.28 | 4.59 | 4.79 | 5.27 | 4.88 | 10.35 | 9.77 | 9.96 | 9.86 | 10.25
11.744 | 6.64 | 0 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.13 | 4.88 | 18.26 | 4.88 | 16.99 | 4.69 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 18.16 | 4.69 | 14.16 | 4.59 | 4.79 | 5.27 | 4.88 | 10.45 | 9.77 | 10.35 | 9.86 | 9.96
11.776 | 7.23 | 0 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.13 | 4.88 | 18.26 | 4.88 | 17.48 | 4.79 | 20.12 | 9.67 | 9.86 | 9.86 | 9.86 | 14.94 | 0.1 | 17.68 | 0 | 17.77 | 0 | 18.75 | 4.69 | 14.94 | 4.59 | 4.79 | 5.18 | 4.88 | 10.35 | 9.86 | 10.45 | 9.86 | 10.45
11.808 | 7.91 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4 | 4.88 | 18.26 | 4.88 | 17.97 | 4.79 | 20.51 | 9.67 | 9.86 | 9.96 | 9.86 | 7.81 | 0.1 | 17.68 | 0 | 17.77 | 0 | 19.34 | 4.69 | 15.63 | 4.59 | 5.08 | 4.98 | 4.88 | 10.16 | 9.86 | 10.35 | 9.86 | 10.55
11.84 | 8.5 | 0.1 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 4.98 | 4.88 | 18.16 | 4.88 | 18.46 | 4.79 | 20.9 | 9.67 | 9.86 | 9.96 | 9.86 | 7.62 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 20.02 | 4.79 | 16.31 | 4.59 | 6.25 | 4.98 | 4.79 | 10.16 | 9.86 | 9.96 | 9.86 | 10.35
11.872 | 9.08 | 0.1 | 0.1 | 0.2 | 0.1 | 0.29 | 0.1 | 5.76 | 4.88 | 18.16 | 4.88 | 18.85 | 4.79 | 21.29 | 9.67 | 9.86 | 9.86 | 9.86 | 10.55 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 20.61 | 4.79 | 16.99 | 4.59 | 7.81 | 4.88 | 4.79 | 10.16 | 9.86 | 10.06 | 9.77 | 10.16
11.904 | 9.67 | 0 | 1.07 | 0.2 | 0.1 | 0.29 | 0.1 | 5.66 | 4.88 | 18.16 | 4.88 | 19.34 | 4.79 | 21.68 | 9.67 | 9.86 | 9.77 | 9.86 | 11.62 | 0 | 17.58 | 0 | 17.68 | 0 | 21.19 | 4.79 | 17.68 | 4.59 | 10.16 | 4.88 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.06
11.936 | 10.25 | 0 | 1.95 | 0.2 | 0.1 | 0.2 | 0.1 | 5.37 | 4.88 | 18.07 | 4.88 | 19.73 | 4.79 | 22.07 | 9.67 | 10.64 | 9.67 | 9.86 | 10.64 | 0.1 | 17.38 | 0 | 17.68 | 0 | 21.68 | 4.79 | 18.36 | 4.59 | 11.23 | 4.79 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
11.968 | 10.84 | 0 | 3.13 | 0.1 | 0.1 | 0.2 | 0.1 | 4.98 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 22.46 | 9.67 | 11.33 | 9.77 | 9.86 | 10.06 | 0.1 | 4.39 | 0 | 17.68 | 0 | 22.36 | 4.79 | 18.95 | 4.59 | 12.4 | 4.79 | 4.79 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
12 | 11.52 | 0.1 | 4.1 | 0.1 | 0.1 | 0.1 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.51 | 4.88 | 22.85 | 9.67 | 12.7 | 9.67 | 9.86 | 9.96 | 0.1 | 2.05 | 0 | 7.32 | 0.1 | 21.58 | 4.79 | 19.53 | 4.59 | 13.28 | 4.69 | 4.88 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
12.032 | 12.11 | 0.1 | 4.79 | 0 | 0.29 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 23.24 | 9.77 | 13.67 | 9.67 | 9.86 | 10.16 | 0.1 | 2.15 | 0 | 1.76 | 0.1 | 18.07 | 4.79 | 20.22 | 4.59 | 14.06 | 4.69 | 4.79 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
12.064 | 12.6 | 0.1 | 5.37 | 0 | 1.27 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 4.79 | 23.54 | 9.77 | 14.45 | 9.67 | 9.86 | 10.25 | 0 | 2.05 | 0 | 2.25 | 0.1 | 18.07 | 4.79 | 20.8 | 4.69 | 14.84 | 4.69 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
12.096 | 13.18 | 0 | 5.96 | 0.1 | 2.25 | 0 | 0.29 | 4.88 | 4.88 | 14.16 | 4.88 | 18.75 | 4.79 | 24.02 | 9.77 | 14.94 | 9.67 | 9.86 | 9.96 | 0.1 | 2.05 | 0.1 | 2.15 | 0 | 17.87 | 4.79 | 21.39 | 4.69 | 15.43 | 4.69 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
12.128 | 13.67 | 0 | 6.45 | 0.1 | 3.71 | 0 | 1.27 | 4.69 | 4.88 | 4 | 4.88 | 18.26 | 4.88 | 24.32 | 9.77 | 15.53 | 9.67 | 9.86 | 9.86 | 0.1 | 1.56 | 0.1 | 1.66 | 0 | 17.87 | 4.79 | 21.97 | 4.69 | 16.21 | 4.59 | 5.76 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
12.16 | 14.16 | 0.1 | 7.13 | 0.1 | 4.49 | 0.1 | 2.25 | 4.79 | 4.88 | 2.93 | 4.88 | 18.26 | 4.88 | 21.97 | 9.77 | 16.11 | 9.67 | 9.86 | 9.96 | 0.1 | 1.37 | 0 | 1.66 | 0.1 | 17.87 | 4.79 | 22.46 | 4.69 | 16.9 | 4.59 | 6.45 | 9.77 | 9.86 | 9.77 | 10.06 | 9.77
12.192 | 14.65 | 0 | 7.72 | 0 | 4.98 | 0.1 | 3.61 | 4.69 | 4.88 | 3.61 | 4.88 | 18.26 | 4.88 | 22.17 | 9.77 | 16.6 | 9.67 | 9.86 | 9.96 | 0.1 | 1.37 | 0 | 1.56 | 0.1 | 17.87 | 4.79 | 23.15 | 4.69 | 17.58 | 4.49 | 8.2 | 9.67 | 9.96 | 9.77 | 10.94 | 9.77
12.224 | 15.14 | 0 | 8.3 | 0 | 5.57 | 0.1 | 4.49 | 4.59 | 4.88 | 4.69 | 4.88 | 18.16 | 4.79 | 21.68 | 9.77 | 16.99 | 9.57 | 9.86 | 9.86 | 0.1 | 1.27 | 0 | 1.37 | 0.1 | 17.77 | 4.88 | 23.63 | 4.79 | 18.26 | 4.59 | 10.16 | 9.67 | 10.94 | 9.67 | 12.4 | 9.77
12.256 | 15.72 | 0 | 8.98 | 0.1 | 6.05 | 0 | 4.98 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 17.58 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 1.27 | 0 | 17.77 | 4.79 | 24.22 | 4.79 | 18.75 | 4.49 | 11.43 | 9.67 | 12.5 | 9.67 | 14.45 | 9.67
12.288 | 16.21 | 0.1 | 9.47 | 0.1 | 6.64 | 0 | 5.47 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 21.58 | 9.77 | 17.97 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 1.17 | 0 | 17.68 | 4.79 | 24.71 | 4.79 | 19.53 | 4.59 | 12.5 | 9.67 | 14.55 | 9.67 | 15.92 | 9.67
12.32 | 16.7 | 0.1 | 10.16 | 0.1 | 7.32 | 0.1 | 6.15 | 4.59 | 4.88 | 5.37 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 18.46 | 9.67 | 9.86 | 9.77 | 0.1 | 0.98 | 0 | 1.17 | 0.1 | 17.77 | 4.88 | 25.2 | 4.79 | 20.12 | 4.49 | 13.48 | 9.67 | 16.02 | 9.67 | 17.09 | 9.67
12.352 | 17.19 | 0.1 | 10.74 | 0 | 7.81 | 0.1 | 6.74 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 18.85 | 9.67 | 9.86 | 9.77 | 0.1 | 0.68 | 0 | 1.07 | 0.1 | 17.68 | 4.79 | 25.68 | 4.79 | 20.7 | 4.49 | 14.16 | 9.67 | 16.99 | 9.67 | 17.77 | 9.67
12.384 | 14.06 | 0.1 | 11.33 | 0 | 8.4 | 0.1 | 7.32 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 21.49 | 9.86 | 19.34 | 9.67 | 9.86 | 9.67 | 0 | 0.78 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 26.27 | 4.79 | 21.39 | 4.49 | 14.75 | 9.67 | 17.77 | 9.67 | 18.46 | 9.67
12.416 | 14.75 | 0 | 11.91 | 0.1 | 9.08 | 0 | 8.01 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.82 | 9.67 | 9.86 | 9.67 | 0.1 | 0.68 | 0 | 0.88 | 0.1 | 17.68 | 4.88 | 22.27 | 4.79 | 21.88 | 4.59 | 15.63 | 9.57 | 18.56 | 9.67 | 19.24 | 9.67
12.448 | 14.36 | 0 | 12.4 | 0.1 | 9.67 | 0 | 8.4 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 20.22 | 9.67 | 10.06 | 9.67 | 0.1 | 0.68 | 0 | 0.78 | 0 | 8.01 | 4.79 | 20.8 | 4.79 | 22.46 | 4.59 | 16.31 | 9.67 | 19.24 | 9.57 | 19.82 | 9.57
12.48 | 14.45 | 0.1 | 12.99 | 0 | 10.35 | 0 | 9.08 | 4.59 | 4.88 | 4.79 | 4.88 | 14.55 | 4.88 | 21.39 | 9.86 | 20.61 | 9.67 | 10.94 | 9.67 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 1.95 | 4.88 | 21.29 | 4.79 | 23.05 | 4.59 | 16.9 | 9.67 | 19.82 | 9.67 | 20.51 | 9.57
12.512 | 14.45 | 0.1 | 13.48 | 0 | 10.74 | 0.1 | 9.67 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 4.88 | 21.39 | 9.86 | 21 | 9.67 | 11.82 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.25 | 4.79 | 21.29 | 4.79 | 23.54 | 4.59 | 17.77 | 9.67 | 20.31 | 9.67 | 21 | 9.67
12.544 | 14.45 | 0.1 | 14.06 | 0 | 11.43 | 0.1 | 10.35 | 4.59 | 4.88 | 4.69 | 4.98 | 3.22 | 4.79 | 21.39 | 9.86 | 21.39 | 9.67 | 13.28 | 9.67 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.25 | 4.79 | 21.29 | 4.79 | 24.12 | 4.69 | 18.46 | 9.67 | 21.09 | 9.67 | 21.58 | 9.67
12.576 | 14.55 | 0 | 14.55 | 0.1 | 12.01 | 0 | 10.84 | 4.59 | 4.88 | 4.69 | 4.88 | 3.71 | 4.88 | 14.65 | 9.86 | 21.78 | 9.67 | 14.06 | 9.57 | 0.1 | 0.39 | 0 | 0.59 | 0 | 1.76 | 4.79 | 21.39 | 4.79 | 24.61 | 4.69 | 19.04 | 9.67 | 21.58 | 9.67 | 22.17 | 9.67
12.608 | 14.45 | 0.1 | 15.04 | 0.1 | 12.5 | 0 | 11.43 | 4.59 | 4.88 | 4.69 | 4.88 | 4.88 | 4.88 | 7.72 | 9.86 | 22.27 | 9.67 | 14.75 | 9.67 | 0.1 | 0.29 | 0 | 0.49 | 0 | 1.76 | 4.79 | 21.29 | 4.79 | 25.1 | 4.69 | 19.53 | 9.67 | 22.27 | 9.67 | 22.75 | 9.67
12.64 | 14.45 | 0 | 15.53 | 0.1 | 13.09 | 0.1 | 12.11 | 4.59 | 4.88 | 4.59 | 4.88 | 5.76 | 4.88 | 7.62 | 9.86 | 22.66 | 9.67 | 15.23 | 9.67 | 0.1 | 0.29 | 0.68 | 0.49 | 0.1 | 1.56 | 4.88 | 21.39 | 4.79 | 25.59 | 4.69 | 20.22 | 9.67 | 22.75 | 9.67 | 23.24 | 9.67
12.672 | 14.45 | 0 | 16.02 | 0 | 13.67 | 0.1 | 12.6 | 4.69 | 4.88 | 4.49 | 4.88 | 5.76 | 4.88 | 10.64 | 9.86 | 22.95 | 9.67 | 15.82 | 9.67 | 0.1 | 0.2 | 1.46 | 0.29 | 0.1 | 1.56 | 4.79 | 21.29 | 4.79 | 26.08 | 4.69 | 20.8 | 9.67 | 23.44 | 9.67 | 23.73 | 9.67
12.704 | 14.45 | 0 | 16.5 | 0 | 14.16 | 0.1 | 13.09 | 4.69 | 4.88 | 4.49 | 4.88 | 5.47 | 4.79 | 11.33 | 9.86 | 23.44 | 9.77 | 16.31 | 9.67 | 0.78 | 0.2 | 3.03 | 0.29 | 0.1 | 1.37 | 4.79 | 21.29 | 4.79 | 22.36 | 4.79 | 21.39 | 9.67 | 23.83 | 9.67 | 24.32 | 9.67
12.736 | 14.36 | 0 | 17.09 | 0.1 | 14.65 | 0 | 13.67 | 4.69 | 5.57 | 4.49 | 4.88 | 5.08 | 4.79 | 10.45 | 9.86 | 23.73 | 9.77 | 16.8 | 9.67 | 1.66 | 0.2 | 4.98 | 0.29 | 0.1 | 1.17 | 4.79 | 21.29 | 4.88 | 20.51 | 4.79 | 21.97 | 9.67 | 24.32 | 9.67 | 24.81 | 9.67
12.768 | 14.36 | 0 | 16.02 | 0 | 15.23 | 0 | 14.16 | 4.69 | 6.25 | 4.49 | 4.88 | 4.98 | 4.88 | 9.86 | 9.86 | 24.12 | 9.77 | 17.19 | 9.67 | 3.22 | 0.2 | 6.35 | 0.2 | 0 | 1.07 | 4.79 | 21.19 | 4.88 | 21.09 | 4.79 | 22.56 | 9.67 | 24.81 | 9.67 | 25.29 | 9.67
12.8 | 13.97 | 0 | 14.26 | 0.1 | 15.72 | 0.1 | 14.75 | 4.79 | 7.42 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 24.02 | 9.77 | 17.68 | 9.67 | 4.98 | 0.1 | 7.13 | 0.2 | 0.39 | 1.07 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 23.05 | 9.77 | 25.39 | 9.67 | 25.88 | 9.67
12.832 | 3.91 | 0.1 | 14.55 | 0 | 16.11 | 0 | 15.23 | 4.79 | 8.59 | 4.39 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 22.27 | 9.77 | 18.16 | 9.67 | 6.25 | 0.1 | 8.01 | 0.2 | 1.46 | 0.88 | 4.79 | 20.51 | 4.79 | 21.19 | 4.79 | 23.63 | 9.77 | 25.88 | 9.77 | 26.27 | 9.67
12.864 | 1.86 | 0.1 | 14.65 | 0 | 16.7 | 0.1 | 15.63 | 4.79 | 9.57 | 4.49 | 4.88 | 4.88 | 4.79 | 10.16 | 9.86 | 22.07 | 9.77 | 18.56 | 9.67 | 7.13 | 0.1 | 8.69 | 0.2 | 3.03 | 0.88 | 4.79 | 4.88 | 4.79 | 21.29 | 4.79 | 24.22 | 9.77 | 26.37 | 9.77 | 26.86 | 9.77
12.896 | 1.86 | 0 | 14.45 | 0.1 | 17.19 | 0 | 16.21 | 4.79 | 10.16 | 4.49 | 4.88 | 4.79 | 4.79 | 9.86 | 9.86 | 21.78 | 9.77 | 19.14 | 9.67 | 7.91 | 0.1 | 9.47 | 0.1 | 5.18 | 0.68 | 4.79 | 3.03 | 4.88 | 21.29 | 4.79 | 24.61 | 9.77 | 26.86 | 9.77 | 27.34 | 9.77
12.928 | 1.56 | 0 | 14.55 | 0.1 | 13.77 | 0 | 16.7 | 4.79 | 10.74 | 4.49 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.68 | 9.77 | 19.53 | 9.67 | 8.69 | 0.1 | 10.35 | 0.1 | 6.35 | 0.68 | 4.79 | 3.32 | 4.88 | 21.29 | 4.79 | 25.2 | 9.77 | 27.34 | 9.77 | 27.83 | 9.67
12.96 | 1.46 | 0.1 | 14.55 | 0.1 | 14.84 | 0.1 | 17.19 | 4.79 | 11.43 | 4.49 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.58 | 9.77 | 20.02 | 9.77 | 9.47 | 0.1 | 11.13 | 0.1 | 7.23 | 0.59 | 4.79 | 4.2 | 4.79 | 21.19 | 4.79 | 25.68 | 9.77 | 27.83 | 9.77 | 28.42 | 9.77
12.992 | 1.27 | 0.1 | 14.45 | 0 | 14.26 | 0.1 | 13.77 | 4.79 | 11.91 | 4.49 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.58 | 9.86 | 20.31 | 9.77 | 10.35 | 0.1 | 11.82 | 0.1 | 8.01 | 0.59 | 4.79 | 5.57 | 4.79 | 21.29 | 4.79 | 26.17 | 9.77 | 28.22 | 9.77 | 28.81 | 9.77
13.024 | 1.17 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 14.06 | 4.79 | 12.6 | 4.59 | 4.88 | 4.59 | 4.79 | 9.77 | 9.86 | 21.49 | 9.86 | 20.8 | 9.77 | 11.23 | 0.1 | 12.6 | 0.1 | 8.79 | 0.49 | 4.79 | 6.25 | 4.88 | 21.19 | 4.79 | 22.36 | 9.77 | 28.71 | 9.77 | 29.3 | 9.77
13.056 | 0.98 | 0 | 14.45 | 0.1 | 14.45 | 0 | 14.26 | 4.79 | 13.09 | 4.59 | 4.88 | 4.59 | 4.79 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 9.67 | 11.82 | 0.1 | 13.28 | 0.1 | 9.47 | 0.39 | 4.79 | 5.86 | 4.88 | 21.19 | 4.79 | 21.39 | 9.77 | 29.3 | 9.77 | 25 | 9.77
13.088 | 0.98 | 0 | 14.45 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 13.67 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.86 | 21.58 | 9.77 | 12.7 | 0.1 | 14.16 | 0.1 | 10.35 | 0.29 | 4.79 | 5.37 | 4.79 | 21.19 | 4.79 | 21.39 | 9.77 | 24.51 | 9.77 | 23.24 | 9.77
13.12 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 14.16 | 4.69 | 5.66 | 4.49 | 4.88 | 9.67 | 9.96 | 21.39 | 9.86 | 21.97 | 9.77 | 13.48 | 0.1 | 14.94 | 0.1 | 11.13 | 0.29 | 4.79 | 5.18 | 4.79 | 21 | 4.79 | 21.29 | 9.77 | 23.34 | 9.86 | 23.93 | 9.77
13.152 | 0.78 | 0 | 14.45 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 14.65 | 4.69 | 6.05 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.36 | 9.77 | 14.16 | 0.1 | 15.53 | 0.1 | 11.91 | 0.29 | 4.79 | 5.18 | 4.79 | 5.18 | 4.79 | 21.49 | 9.77 | 24.02 | 9.86 | 24.12 | 9.77
13.184 | 0.68 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 15.23 | 4.69 | 7.03 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.75 | 9.77 | 14.75 | 0.1 | 16.31 | 0.1 | 12.79 | 0.2 | 4.79 | 5.27 | 4.88 | 2.93 | 4.79 | 21.39 | 9.77 | 24.12 | 9.86 | 24.22 | 9.77
13.216 | 0.59 | 0 | 5.86 | 0 | 14.45 | 0 | 14.45 | 4.88 | 15.72 | 4.79 | 8.4 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 9.77 | 15.53 | 0.1 | 16.99 | 0.1 | 13.48 | 0.2 | 5.96 | 5.08 | 4.79 | 3.32 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
13.248 | 0.49 | 0 | 1.95 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 16.21 | 4.79 | 9.38 | 4.49 | 4.79 | 9.67 | 9.86 | 13.18 | 9.86 | 23.54 | 9.77 | 16.21 | 0.1 | 17.58 | 0.1 | 14.26 | 0.1 | 6.74 | 4.98 | 4.79 | 4.3 | 4.79 | 21.29 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77
13.28 | 0.49 | 0.1 | 2.15 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 16.7 | 4.79 | 10.06 | 4.49 | 4.88 | 9.67 | 9.86 | 6.93 | 9.86 | 23.93 | 9.77 | 16.8 | 0.1 | 18.26 | 0.1 | 15.04 | 0.2 | 8.5 | 4.98 | 4.79 | 5.57 | 4.79 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
13.312 | 0.39 | 0 | 1.66 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.19 | 4.79 | 10.64 | 4.49 | 5.57 | 9.57 | 9.86 | 8.2 | 9.86 | 24.22 | 9.86 | 17.48 | 0.1 | 18.85 | 0.1 | 15.63 | 0.1 | 10.35 | 4.88 | 4.79 | 6.25 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
13.344 | 0.39 | 0.1 | 1.46 | 0 | 14.26 | 0.1 | 14.45 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 6.35 | 9.57 | 9.86 | 11.04 | 9.86 | 20.8 | 9.86 | 18.26 | 0.1 | 19.43 | 0.1 | 16.31 | 0.1 | 11.72 | 4.79 | 4.88 | 5.86 | 4.79 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
13.376 | 0.29 | 0 | 1.37 | 0.1 | 4.59 | 0 | 14.45 | 4.88 | 18.07 | 4.79 | 11.82 | 4.59 | 7.62 | 9.57 | 9.86 | 11.23 | 9.86 | 20.9 | 9.77 | 18.85 | 0.1 | 20.12 | 0.1 | 16.99 | 0 | 12.7 | 4.79 | 4.88 | 5.47 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
13.408 | 0.2 | 0.1 | 1.27 | 0.1 | 1.76 | 0 | 14.36 | 4.88 | 18.46 | 4.79 | 12.4 | 4.59 | 8.69 | 9.57 | 9.86 | 10.45 | 9.86 | 21.19 | 9.86 | 19.43 | 0.1 | 20.7 | 0.1 | 17.68 | 0.1 | 13.48 | 4.79 | 4.79 | 5.37 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
13.44 | 0.2 | 0.1 | 1.27 | 0.1 | 1.86 | 0.1 | 4.79 | 4.88 | 18.95 | 4.88 | 13.09 | 4.69 | 9.57 | 9.67 | 9.86 | 9.86 | 9.86 | 21.39 | 9.86 | 20.12 | 0 | 21.39 | 0.1 | 18.36 | 0.1 | 14.26 | 4.69 | 4.79 | 5.08 | 4.79 | 17.77 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
13.472 | 0.2 | 0.1 | 1.07 | 0 | 1.76 | 0 | 1.76 | 4.88 | 19.34 | 4.79 | 13.48 | 4.59 | 10.25 | 9.67 | 9.86 | 9.96 | 9.86 | 21.39 | 9.86 | 20.7 | 0.1 | 21.97 | 0.1 | 18.95 | 0.1 | 15.04 | 4.69 | 4.79 | 5.18 | 4.79 | 3.81 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
13.504 | 0.1 | 0 | 0.88 | 0 | 1.46 | 0 | 2.05 | 4.88 | 19.92 | 4.79 | 14.06 | 4.69 | 10.74 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 9.86 | 21.19 | 0.1 | 22.46 | 0.1 | 19.53 | 0.1 | 15.82 | 4.59 | 4.88 | 5.08 | 4.79 | 3.03 | 9.86 | 24.22 | 9.86 | 10.55 | 9.86
13.536 | 0 | 0 | 0.78 | 0.1 | 1.37 | 0 | 1.66 | 4.88 | 20.22 | 4.88 | 14.65 | 4.69 | 11.43 | 9.57 | 9.86 | 10.16 | 9.86 | 21.49 | 9.86 | 21.88 | 0.1 | 17.87 | 0.1 | 20.31 | 0 | 16.5 | 4.59 | 4.88 | 5.08 | 4.88 | 3.52 | 9.86 | 10.45 | 9.86 | 5.86 | 9.86
13.568 | 0 | 0.2 | 0.68 | 0.1 | 1.17 | 0 | 1.46 | 4.79 | 20.7 | 4.88 | 15.04 | 4.69 | 12.01 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 9.86 | 22.46 | 0.1 | 16.99 | 0.1 | 20.8 | 0.1 | 17.19 | 4.59 | 4.88 | 4.98 | 4.79 | 4.39 | 9.77 | 5.76 | 9.86 | 9.18 | 9.86
13.6 | 0.1 | 1.17 | 0.78 | 0.1 | 1.07 | 0 | 1.37 | 4.88 | 21.09 | 4.88 | 15.53 | 4.79 | 12.6 | 9.67 | 11.33 | 9.86 | 9.86 | 21.39 | 9.86 | 18.36 | 0.1 | 17.68 | 0.1 | 21.39 | 0.1 | 17.77 | 4.59 | 4.79 | 4.88 | 4.79 | 5.76 | 9.86 | 9.28 | 9.86 | 12.31 | 9.86
13.632 | 0.1 | 2.15 | 0.59 | 0 | 0.98 | 0.1 | 1.17 | 4.88 | 18.95 | 4.79 | 16.02 | 4.69 | 13.09 | 9.67 | 12.4 | 9.96 | 9.86 | 21.49 | 9.86 | 17.68 | 0.1 | 17.58 | 0.1 | 21.97 | 0.1 | 18.46 | 4.59 | 4.79 | 4.98 | 4.79 | 6.15 | 9.86 | 12.4 | 9.86 | 11.13 | 9.86
13.664 | 0.1 | 3.52 | 0.59 | 0 | 0.88 | 0 | 1.07 | 4.88 | 19.04 | 4.88 | 16.5 | 4.79 | 13.67 | 9.67 | 13.57 | 9.86 | 9.86 | 21.39 | 9.86 | 17.77 | 0.1 | 17.68 | 0.1 | 22.56 | 0.1 | 19.14 | 4.59 | 4.98 | 4.69 | 4.79 | 5.86 | 9.86 | 11.33 | 9.86 | 10.16 | 9.86
13.696 | 0 | 4.39 | 0.49 | 0.1 | 0.88 | 0 | 1.07 | 4.88 | 18.56 | 4.88 | 16.99 | 4.79 | 14.16 | 9.67 | 14.36 | 9.77 | 9.86 | 21.39 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 18.46 | 0.1 | 19.82 | 4.59 | 5.96 | 4.69 | 4.79 | 5.37 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
13.728 | 0.1 | 4.98 | 0.39 | 0.1 | 0.68 | 0 | 0.88 | 4.88 | 18.46 | 4.88 | 17.38 | 4.79 | 14.65 | 9.67 | 14.94 | 9.77 | 9.86 | 21.39 | 9.86 | 17.68 | 0.1 | 17.68 | 0.1 | 16.9 | 0.1 | 20.31 | 4.49 | 7.32 | 4.69 | 4.79 | 5.18 | 9.77 | 10.16 | 9.86 | 10.35 | 9.86
13.76 | 0.1 | 5.47 | 0.39 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 18.46 | 4.88 | 17.87 | 4.79 | 15.14 | 9.67 | 15.53 | 9.77 | 9.96 | 21.09 | 9.96 | 17.68 | 0.1 | 17.68 | 0.1 | 17.58 | 0.1 | 21 | 4.59 | 9.67 | 4.69 | 4.79 | 5.08 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86
13.792 | 0.1 | 6.05 | 0.29 | 0 | 0.49 | 0.1 | 0.78 | 4.88 | 18.26 | 4.88 | 18.26 | 4.79 | 15.72 | 9.67 | 15.92 | 9.67 | 9.96 | 9.86 | 9.96 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 21.49 | 4.59 | 11.13 | 4.69 | 4.79 | 5.18 | 9.86 | 10.45 | 9.86 | 10.35 | 9.86
13.824 | 0.1 | 6.54 | 0.29 | 0 | 0.49 | 0 | 0.68 | 4.88 | 18.26 | 4.88 | 18.75 | 4.79 | 16.21 | 9.67 | 16.5 | 9.67 | 9.86 | 6.84 | 9.96 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 22.17 | 4.59 | 12.21 | 4.59 | 4.79 | 5.08 | 9.86 | 10.35 | 9.86 | 10.16 | 9.86
13.856 | 0.1 | 7.23 | 0.2 | 0.1 | 0.49 | 0 | 0.59 | 4.88 | 18.26 | 4.88 | 19.24 | 4.88 | 16.7 | 9.67 | 17.09 | 9.67 | 9.86 | 9.57 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.66 | 4.59 | 13.09 | 4.49 | 4.79 | 5.08 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
13.888 | 0 | 7.72 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.79 | 18.16 | 4.88 | 19.73 | 4.88 | 17.09 | 9.67 | 17.38 | 9.67 | 9.86 | 11.62 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 23.24 | 4.59 | 13.97 | 4.59 | 4.79 | 4.98 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86
13.92 | 0.1 | 8.4 | 0.2 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.16 | 4.88 | 20.12 | 4.88 | 17.58 | 9.77 | 17.87 | 9.67 | 9.86 | 10.74 | 9.96 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 0.1 | 23.73 | 4.69 | 14.75 | 4.59 | 4.79 | 4.98 | 9.86 | 10.16 | 9.86 | 10.06 | 9.86
13.952 | 0.1 | 8.98 | 0.1 | 0.1 | 0.2 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 17.97 | 9.77 | 18.46 | 9.67 | 9.86 | 10.16 | 9.96 | 17.58 | 0.1 | 17.38 | 0.1 | 17.77 | 0.1 | 24.22 | 4.69 | 15.33 | 4.59 | 4.79 | 4.79 | 9.86 | 10.16 | 9.86 | 9.86 | 9.86
13.984 | 0.1 | 9.57 | 0.1 | 0.68 | 0.2 | 0 | 0.39 | 4.88 | 18.16 | 4.88 | 20.9 | 4.88 | 18.46 | 9.77 | 18.85 | 9.67 | 9.86 | 9.77 | 9.96 | 17.58 | 0.1 | 4.59 | 0.1 | 17.77 | 0.1 | 24.81 | 4.69 | 16.11 | 4.49 | 4.79 | 4.69 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
14.016 | 0 | 10.25 | 0.1 | 1.56 | 0.2 | 0.68 | 0.29 | 4.88 | 18.07 | 4.88 | 17.68 | 4.88 | 18.85 | 9.77 | 19.34 | 9.57 | 9.86 | 10.06 | 9.96 | 15.04 | 0.1 | 1.86 | 0.1 | 17.68 | 0 | 25.39 | 4.69 | 16.8 | 4.49 | 4.79 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
14.048 | 0.1 | 10.74 | 0.1 | 2.73 | 0.1 | 1.27 | 0.2 | 4.79 | 12.11 | 4.88 | 18.07 | 4.88 | 19.34 | 9.77 | 19.73 | 9.57 | 9.86 | 10.25 | 9.96 | 3.71 | 0.1 | 2.25 | 0.1 | 17.68 | 0.1 | 25.78 | 4.69 | 17.58 | 4.59 | 4.79 | 4.69 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
14.08 | 0.1 | 11.43 | 0.1 | 3.81 | 0.1 | 2.44 | 0.2 | 4.88 | 3.52 | 4.88 | 18.16 | 4.88 | 19.82 | 9.77 | 20.12 | 9.57 | 9.86 | 9.96 | 9.96 | 2.05 | 0 | 1.76 | 0.1 | 17.68 | 0.1 | 26.27 | 4.79 | 18.16 | 4.59 | 4.79 | 4.69 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
14.112 | 0.1 | 11.91 | 0.1 | 4.59 | 0 | 3.81 | 0.2 | 4.88 | 3.13 | 4.88 | 18.07 | 4.88 | 20.12 | 9.77 | 20.61 | 9.57 | 9.96 | 9.86 | 9.96 | 2.15 | 0 | 1.76 | 0.1 | 17.48 | 0.1 | 23.24 | 4.79 | 18.75 | 4.59 | 4.98 | 4.59 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
14.144 | 0.1 | 12.5 | 0.1 | 5.18 | 0.1 | 4.49 | 0.2 | 4.88 | 3.81 | 4.88 | 18.16 | 4.88 | 20.61 | 9.77 | 21 | 9.67 | 10.74 | 9.96 | 9.96 | 1.76 | 0.1 | 1.66 | 0.1 | 4.49 | 0.1 | 22.27 | 4.79 | 19.43 | 4.59 | 5.96 | 4.59 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
14.176 | 0 | 13.09 | 0.1 | 5.76 | 0.1 | 5.18 | 0.1 | 4.88 | 4.79 | 4.88 | 18.16 | 4.88 | 21 | 9.77 | 21.29 | 9.67 | 11.91 | 9.86 | 9.96 | 1.76 | 0.1 | 1.46 | 0.1 | 1.95 | 0 | 21.78 | 4.79 | 20.12 | 4.59 | 7.42 | 4.59 | 10.84 | 9.77 | 9.86 | 9.67 | 9.86
14.208 | 0 | 13.67 | 0.1 | 6.35 | 0.1 | 5.66 | 0.1 | 4.79 | 5.96 | 4.88 | 18.16 | 4.88 | 18.95 | 9.77 | 21.78 | 9.67 | 13.09 | 9.86 | 9.96 | 1.56 | 0.1 | 1.37 | 0.1 | 2.15 | 0.1 | 21.58 | 4.79 | 20.61 | 4.59 | 9.47 | 4.59 | 11.82 | 9.77 | 10.84 | 9.77 | 9.96
14.24 | 0.1 | 14.16 | 0.1 | 6.84 | 0.1 | 6.15 | 0.1 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 18.75 | 9.77 | 22.17 | 9.67 | 13.97 | 9.67 | 9.96 | 1.46 | 0 | 1.27 | 0.1 | 1.95 | 0.1 | 21.58 | 4.79 | 21.19 | 4.69 | 11.13 | 4.59 | 13.67 | 9.77 | 12.21 | 9.67 | 11.04
14.272 | 0.1 | 14.65 | 0.1 | 7.52 | 0.1 | 6.74 | 0.1 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 18.46 | 9.77 | 22.56 | 9.67 | 14.65 | 9.67 | 9.96 | 1.37 | 0 | 1.07 | 0.1 | 1.86 | 0.1 | 21.49 | 4.79 | 21.88 | 4.69 | 12.11 | 4.49 | 15.43 | 9.67 | 14.16 | 9.67 | 12.5
14.304 | 0.1 | 15.14 | 0.1 | 8.2 | 0.1 | 7.32 | 0.1 | 4.88 | 5.08 | 4.79 | 18.16 | 4.88 | 18.26 | 9.77 | 22.95 | 9.67 | 15.23 | 9.77 | 9.96 | 1.17 | 0.1 | 1.07 | 0.1 | 1.56 | 0.1 | 21.39 | 4.79 | 22.36 | 4.69 | 13.09 | 4.59 | 16.6 | 9.67 | 15.72 | 9.67 | 14.55
14.336 | 0 | 15.63 | 0.1 | 8.69 | 0.1 | 8.01 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 18.26 | 9.77 | 23.34 | 9.67 | 15.72 | 9.67 | 9.96 | 1.17 | 0.1 | 0.98 | 0.1 | 1.27 | 0.1 | 21.29 | 4.79 | 22.95 | 4.69 | 13.87 | 4.59 | 17.58 | 9.67 | 16.9 | 9.67 | 16.11
14.368 | 0 | 16.11 | 0.1 | 9.38 | 0.1 | 8.59 | 0.1 | 4.79 | 4.98 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 23.63 | 9.67 | 16.31 | 9.67 | 9.96 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 0 | 21.29 | 4.79 | 23.44 | 4.79 | 14.65 | 4.59 | 18.16 | 9.67 | 17.58 | 9.67 | 16.99
14.4 | 0.1 | 16.6 | 0.1 | 9.86 | 0.1 | 9.28 | 0.2 | 5.57 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 24.12 | 9.67 | 16.8 | 9.67 | 9.96 | 0.88 | 0.1 | 0.78 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.02 | 4.79 | 15.43 | 4.59 | 18.85 | 9.67 | 18.36 | 9.67 | 17.77
14.432 | 0.1 | 16.99 | 0.1 | 10.55 | 0.1 | 9.77 | 0.1 | 6.35 | 4.79 | 4.79 | 17.68 | 4.88 | 18.16 | 9.77 | 24.41 | 9.67 | 17.29 | 9.67 | 9.96 | 0.88 | 0.1 | 0.68 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 24.51 | 4.79 | 16.11 | 4.59 | 19.53 | 9.57 | 19.14 | 9.67 | 18.56
14.464 | 0.1 | 13.67 | 0.1 | 11.13 | 0.1 | 10.45 | 0.1 | 7.42 | 4.79 | 4.79 | 4.69 | 4.88 | 18.07 | 9.77 | 22.46 | 9.67 | 17.68 | 9.67 | 10.74 | 0.88 | 0.1 | 0.59 | 0.1 | 1.07 | 0.1 | 21.19 | 4.88 | 25 | 4.69 | 16.8 | 4.59 | 20.31 | 9.67 | 19.73 | 9.67 | 19.24
14.496 | 0 | 14.06 | 0.1 | 11.72 | 0.1 | 10.94 | 0.1 | 8.69 | 4.79 | 4.88 | 2.93 | 4.88 | 18.07 | 9.77 | 22.17 | 9.67 | 18.16 | 9.57 | 11.33 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 0 | 21.19 | 4.79 | 25.59 | 4.79 | 17.48 | 4.69 | 20.8 | 9.67 | 20.22 | 9.67 | 20.02
14.528 | 0 | 14.16 | 0.1 | 12.21 | 0.1 | 11.52 | 0.1 | 9.47 | 4.69 | 4.88 | 3.71 | 4.88 | 18.07 | 9.77 | 21.88 | 9.77 | 18.56 | 9.57 | 12.21 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 0.1 | 12.21 | 4.88 | 26.08 | 4.79 | 18.07 | 4.69 | 21.29 | 9.67 | 20.9 | 9.67 | 20.41
14.56 | 0.1 | 14.36 | 0.1 | 12.79 | 0.1 | 12.11 | 0.1 | 10.25 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.78 | 9.67 | 19.14 | 9.57 | 13.48 | 0.59 | 0.1 | 0.39 | 0.1 | 0.68 | 0.1 | 3.13 | 4.88 | 21.49 | 4.79 | 18.75 | 4.69 | 21.97 | 9.67 | 21.58 | 9.67 | 21.09
14.592 | 0.1 | 14.36 | 0.1 | 13.38 | 0.1 | 12.7 | 0.1 | 10.84 | 4.59 | 4.79 | 5.66 | 4.88 | 17.97 | 9.77 | 21.78 | 9.77 | 19.43 | 9.57 | 14.26 | 0.49 | 0.1 | 0.29 | 0.1 | 0.68 | 0.1 | 3.22 | 4.88 | 20.31 | 4.79 | 19.34 | 4.69 | 22.56 | 9.67 | 22.17 | 9.67 | 21.68
14.624 | 0.1 | 14.36 | 0.1 | 13.87 | 0.1 | 13.18 | 0.1 | 11.43 | 4.59 | 4.79 | 5.76 | 4.88 | 8.89 | 9.77 | 21.58 | 9.77 | 19.92 | 9.57 | 14.94 | 0.49 | 0.29 | 0.29 | 0.1 | 0.59 | 0.1 | 3.52 | 4.88 | 21 | 4.79 | 20.02 | 4.69 | 23.15 | 9.67 | 22.66 | 9.67 | 22.27
14.656 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 13.77 | 0.1 | 12.01 | 4.59 | 4.88 | 5.47 | 4.88 | 3.03 | 9.77 | 21.58 | 9.77 | 20.31 | 9.57 | 15.43 | 0.39 | 1.56 | 0.29 | 0.1 | 0.49 | 0.1 | 4.59 | 4.79 | 21.19 | 4.79 | 20.61 | 4.79 | 23.63 | 9.67 | 23.24 | 9.67 | 22.75
14.688 | 0 | 14.36 | 0.1 | 14.84 | 0.1 | 14.26 | 0.1 | 12.6 | 4.59 | 4.88 | 5.27 | 4.88 | 3.32 | 9.77 | 21.58 | 9.77 | 20.8 | 9.57 | 15.92 | 0.29 | 2.83 | 0.2 | 1.17 | 0.49 | 0 | 5.86 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 24.12 | 9.67 | 23.73 | 9.67 | 23.24
14.72 | 0.1 | 14.45 | 0.1 | 15.33 | 0.1 | 14.84 | 0.1 | 13.09 | 4.49 | 4.79 | 4.88 | 4.88 | 4 | 9.77 | 21.49 | 9.77 | 21.09 | 9.57 | 16.5 | 0.29 | 4.88 | 0.2 | 2.34 | 0.49 | 0.1 | 5.96 | 4.88 | 21.19 | 4.79 | 21.78 | 4.79 | 24.81 | 9.67 | 24.32 | 9.67 | 23.83
14.752 | 0.1 | 14.45 | 0.1 | 15.92 | 0.1 | 15.33 | 0.1 | 13.67 | 4.49 | 4.79 | 4.88 | 4.88 | 5.18 | 9.77 | 21.49 | 9.77 | 21.58 | 9.57 | 16.99 | 0.29 | 6.15 | 0.2 | 4.3 | 0.49 | 0.1 | 5.86 | 4.88 | 21.19 | 4.79 | 22.36 | 4.79 | 25.29 | 9.67 | 24.71 | 9.67 | 24.32
14.784 | 0.1 | 14.36 | 0.1 | 16.31 | 0.1 | 15.82 | 0.1 | 14.16 | 4.59 | 5.66 | 4.98 | 4.88 | 5.86 | 9.77 | 21.49 | 9.77 | 21.97 | 9.67 | 17.48 | 0.2 | 7.13 | 0.1 | 5.86 | 0.29 | 0.1 | 5.37 | 4.88 | 21.29 | 4.79 | 22.95 | 4.79 | 25.78 | 9.67 | 25.29 | 9.77 | 24.81
14.816 | 0 | 14.36 | 0.1 | 16.9 | 0.1 | 16.31 | 0.1 | 14.75 | 4.49 | 6.05 | 4.88 | 4.88 | 5.66 | 9.77 | 21.39 | 9.77 | 22.36 | 9.67 | 17.97 | 0.2 | 7.72 | 0.1 | 6.74 | 0.29 | 0 | 5.08 | 4.88 | 21.19 | 4.79 | 23.44 | 4.79 | 26.27 | 9.77 | 25.88 | 9.77 | 25.39
14.848 | 0 | 14.36 | 0.1 | 17.29 | 0.1 | 16.7 | 0.1 | 15.14 | 4.59 | 6.93 | 4.79 | 4.88 | 5.27 | 9.77 | 21.39 | 9.86 | 22.75 | 9.67 | 18.46 | 0.2 | 8.69 | 0.1 | 7.62 | 0.2 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 23.93 | 4.79 | 26.66 | 9.77 | 26.27 | 9.77 | 25.78
14.88 | 0.1 | 14.36 | 0.1 | 15.33 | 0 | 17.29 | 0.1 | 15.72 | 4.49 | 8.4 | 4.79 | 4.88 | 4.98 | 9.77 | 15.04 | 9.77 | 23.15 | 9.67 | 18.85 | 0.2 | 9.47 | 0 | 8.4 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.79 | 24.51 | 4.79 | 27.25 | 9.77 | 26.86 | 9.77 | 26.37
14.912 | 0.1 | 4.79 | 0.1 | 15.04 | 0 | 14.94 | 0.1 | 16.21 | 4.59 | 9.38 | 4.69 | 4.88 | 4.98 | 9.77 | 7.42 | 9.77 | 23.54 | 9.67 | 19.24 | 0.1 | 10.06 | 0.1 | 9.28 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 25.1 | 4.79 | 27.54 | 9.77 | 27.34 | 9.77 | 26.86
14.944 | 0.1 | 1.76 | 0.1 | 14.84 | 0.1 | 15.23 | 0.1 | 16.7 | 4.59 | 9.96 | 4.69 | 4.88 | 4.98 | 9.77 | 7.91 | 9.77 | 23.83 | 9.67 | 19.73 | 0.1 | 10.94 | 0.1 | 10.16 | 0.1 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 25.59 | 4.79 | 28.13 | 9.67 | 27.83 | 9.77 | 27.44
14.976 | 0 | 1.86 | 0.1 | 14.75 | 0.1 | 14.84 | 0.1 | 17.19 | 4.69 | 10.55 | 4.59 | 4.88 | 4.79 | 9.77 | 11.04 | 9.77 | 24.22 | 9.67 | 20.12 | 0.1 | 11.72 | 0.1 | 10.94 | 0.1 | 0.1 | 4.88 | 4.88 | 21.19 | 4.79 | 26.08 | 4.79 | 28.61 | 9.77 | 28.32 | 9.77 | 27.93
15.008 | 0 | 1.66 | 0.1 | 14.75 | 0.1 | 14.65 | 0.1 | 17.58 | 4.69 | 11.23 | 4.59 | 4.88 | 4.88 | 9.77 | 11.33 | 9.86 | 20.8 | 9.67 | 20.51 | 0.1 | 12.5 | 0.1 | 11.72 | 0.1 | 0.1 | 4.88 | 4.88 | 7.03 | 4.79 | 21.88 | 4.79 | 29.1 | 9.77 | 28.71 | 9.77 | 28.22
15.04 | 0.1 | 1.27 | 0.1 | 14.55 | 0.1 | 14.65 | 0.1 | 18.07 | 4.69 | 11.82 | 4.59 | 4.88 | 4.79 | 9.77 | 10.45 | 9.77 | 20.9 | 9.67 | 21 | 0.1 | 13.38 | 0 | 12.5 | 0.1 | 0.1 | 4.79 | 4.88 | 2.34 | 4.79 | 20.22 | 4.79 | 29.49 | 9.77 | 29.2 | 9.77 | 28.71
15.072 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.56 | 4.69 | 12.4 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.09 | 9.67 | 21.39 | 0.1 | 14.06 | 0 | 13.18 | 0.1 | 0.1 | 4.69 | 4.88 | 3.42 | 4.79 | 20.9 | 4.79 | 25.88 | 9.77 | 26.56 | 9.77 | 29.2
15.104 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.95 | 4.69 | 12.89 | 4.59 | 4.88 | 4.69 | 9.77 | 10.25 | 9.86 | 21.39 | 9.77 | 21.68 | 0.1 | 14.75 | 0.1 | 13.87 | 0.1 | 0.1 | 4.69 | 4.88 | 4.1 | 4.79 | 21.19 | 4.79 | 25.1 | 9.77 | 23.63 | 9.77 | 24.61
15.136 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 19.43 | 4.79 | 13.48 | 4.59 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.39 | 9.67 | 22.17 | 0.1 | 15.43 | 0.1 | 14.65 | 0.1 | 1.07 | 4.69 | 4.88 | 5.37 | 4.79 | 21.19 | 4.88 | 24.51 | 9.77 | 24.02 | 9.77 | 23.34
15.168 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 19.82 | 4.79 | 13.97 | 4.59 | 5.27 | 4.59 | 9.77 | 10.06 | 9.86 | 21.39 | 9.77 | 22.46 | 0.1 | 16.21 | 0.1 | 15.43 | 0 | 2.15 | 4.59 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 24.61 | 9.77 | 24.22 | 9.77 | 23.93
15.2 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 20.22 | 4.79 | 14.55 | 4.59 | 6.05 | 4.59 | 9.77 | 10.06 | 9.86 | 21.49 | 9.67 | 22.95 | 0.1 | 16.99 | 0 | 16.11 | 0.1 | 3.91 | 4.59 | 4.88 | 6.05 | 4.79 | 21.19 | 4.79 | 24.61 | 9.77 | 24.22 | 9.77 | 24.12
15.232 | 0.1 | 0.68 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 20.7 | 4.79 | 15.04 | 4.59 | 7.03 | 4.59 | 9.77 | 10.06 | 9.86 | 21.49 | 9.67 | 23.24 | 0.1 | 17.58 | 0 | 16.8 | 0.1 | 5.57 | 4.59 | 4.88 | 5.57 | 4.79 | 21.19 | 4.79 | 24.51 | 9.77 | 24.22 | 9.77 | 24.22
15.264 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 21 | 4.79 | 15.43 | 4.59 | 8.4 | 4.49 | 10.64 | 9.96 | 9.86 | 21.49 | 9.77 | 23.73 | 0.1 | 18.26 | 0.1 | 17.38 | 0.1 | 6.74 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.12
15.296 | 0 | 0.59 | 0.1 | 8.5 | 0.1 | 14.36 | 0.1 | 19.14 | 4.79 | 16.02 | 4.59 | 9.38 | 4.59 | 11.04 | 9.86 | 9.86 | 21.49 | 9.77 | 24.02 | 0.1 | 18.95 | 0.1 | 18.16 | 0.1 | 7.52 | 4.49 | 4.88 | 4.98 | 4.79 | 21.29 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
15.328 | 0 | 0.49 | 0.1 | 2.44 | 0.1 | 10.74 | 0.1 | 18.95 | 4.88 | 16.5 | 4.59 | 10.06 | 4.49 | 12.01 | 9.86 | 9.86 | 21.39 | 9.77 | 24.41 | 0.1 | 19.53 | 0.1 | 18.75 | 0 | 8.4 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
15.36 | 0.1 | 0.49 | 0.1 | 2.05 | 0.1 | 2.93 | 0.1 | 18.56 | 4.79 | 16.99 | 4.59 | 10.64 | 4.59 | 13.38 | 9.86 | 9.86 | 21.39 | 9.77 | 21.58 | 0.1 | 20.12 | 0 | 19.34 | 0.1 | 9.18 | 4.49 | 4.88 | 5.27 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
15.392 | 0.1 | 0.39 | 0.1 | 1.66 | 0 | 1.95 | 0.1 | 18.46 | 4.79 | 17.38 | 4.69 | 11.23 | 4.49 | 14.16 | 9.77 | 9.86 | 21.39 | 9.77 | 22.17 | 0.1 | 20.7 | 0 | 20.02 | 0.1 | 9.86 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22
15.424 | 0.1 | 0.29 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 18.46 | 4.88 | 17.87 | 4.79 | 11.82 | 4.49 | 14.75 | 9.77 | 9.86 | 21.39 | 9.77 | 21.68 | 0.1 | 21.29 | 0.1 | 20.61 | 0.1 | 10.74 | 4.59 | 4.88 | 4.88 | 4.79 | 21.19 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
15.456 | 0 | 0.29 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 18.26 | 4.88 | 18.36 | 4.79 | 12.4 | 4.59 | 15.23 | 9.67 | 9.86 | 10.06 | 9.77 | 21.68 | 0.1 | 21.78 | 0.1 | 21.19 | 0.1 | 11.52 | 4.49 | 4.88 | 4.88 | 4.79 | 6.93 | 4.79 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
15.488 | 0.2 | 0.2 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 18.26 | 4.88 | 18.75 | 4.79 | 12.99 | 4.59 | 15.82 | 9.67 | 9.86 | 6.54 | 9.77 | 21.68 | 0.1 | 22.36 | 0.1 | 21.78 | 0 | 12.31 | 4.49 | 4.88 | 4.88 | 4.79 | 2.25 | 4.88 | 13.87 | 9.86 | 24.12 | 9.86 | 24.12
15.52 | 1.17 | 0.2 | 0.1 | 1.07 | 0 | 1.37 | 0.1 | 18.26 | 4.88 | 19.24 | 4.79 | 13.48 | 4.59 | 16.31 | 9.67 | 9.86 | 9.47 | 9.77 | 21.58 | 0.1 | 17.77 | 0 | 22.36 | 0.1 | 12.99 | 4.59 | 5.08 | 4.88 | 4.88 | 3.32 | 4.88 | 6.05 | 9.77 | 11.43 | 9.86 | 24.12
15.552 | 2.15 | 0.1 | 0.1 | 0.98 | 0 | 1.27 | 0.1 | 18.16 | 4.88 | 19.73 | 4.79 | 13.97 | 4.69 | 16.8 | 9.67 | 9.86 | 11.52 | 9.77 | 21.58 | 0.1 | 16.7 | 0 | 20.8 | 0.1 | 13.77 | 4.59 | 6.05 | 4.69 | 4.79 | 4.1 | 4.79 | 7.62 | 9.77 | 5.57 | 9.86 | 10.35
15.584 | 3.61 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 18.16 | 4.88 | 20.12 | 4.79 | 14.55 | 4.69 | 17.29 | 9.67 | 9.86 | 10.94 | 9.86 | 21.58 | 0.1 | 17.38 | 0.1 | 17.58 | 0.1 | 14.55 | 4.59 | 7.72 | 4.69 | 4.79 | 5.47 | 4.79 | 11.43 | 9.77 | 8.69 | 9.86 | 6.05
15.616 | 4.39 | 0 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 18.07 | 4.88 | 20.51 | 4.88 | 15.04 | 4.69 | 17.77 | 9.67 | 9.86 | 10.25 | 9.86 | 21.49 | 0.1 | 17.58 | 0.1 | 18.07 | 0.1 | 15.23 | 4.59 | 9.86 | 4.69 | 4.79 | 6.35 | 4.88 | 11.72 | 9.86 | 12.31 | 9.86 | 9.67
15.648 | 4.98 | 0 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 18.07 | 4.88 | 20.9 | 4.88 | 15.53 | 4.69 | 18.16 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 0.1 | 17.58 | 0.1 | 17.97 | 0 | 16.02 | 4.69 | 11.33 | 4.59 | 4.88 | 6.05 | 4.88 | 10.55 | 9.77 | 11.33 | 9.86 | 12.31
15.68 | 5.47 | 0 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 18.07 | 4.88 | 17.68 | 4.88 | 16.02 | 4.69 | 18.65 | 9.67 | 9.86 | 10.06 | 9.86 | 21.49 | 0.1 | 17.58 | 0 | 17.97 | 0.1 | 16.7 | 4.69 | 12.31 | 4.59 | 4.79 | 5.37 | 4.88 | 9.96 | 9.86 | 10.35 | 9.86 | 11.33
15.712 | 6.05 | 0.1 | 0.1 | 0.59 | 0 | 0.78 | 0.1 | 10.94 | 4.88 | 18.46 | 4.88 | 16.5 | 4.79 | 19.14 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 17.29 | 4.69 | 13.28 | 4.59 | 4.79 | 5.18 | 4.79 | 10.25 | 9.77 | 9.96 | 9.77 | 10.16
15.744 | 6.64 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.52 | 4.88 | 18.07 | 4.88 | 16.9 | 4.79 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0 | 17.68 | 0 | 17.87 | 0.1 | 18.07 | 4.69 | 13.97 | 4.59 | 4.79 | 5.08 | 4.79 | 10.55 | 9.77 | 10.35 | 9.86 | 9.86
15.776 | 7.23 | 0 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.13 | 4.88 | 18.07 | 4.88 | 17.38 | 4.79 | 20.02 | 9.67 | 9.86 | 9.86 | 9.86 | 15.82 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 18.65 | 4.69 | 14.84 | 4.59 | 4.79 | 5.18 | 4.88 | 10.45 | 9.86 | 10.45 | 9.86 | 10.45
15.808 | 7.81 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 3.81 | 4.88 | 18.16 | 4.88 | 17.87 | 4.79 | 20.31 | 9.67 | 9.86 | 9.86 | 9.86 | 8.3 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 19.24 | 4.69 | 15.53 | 4.59 | 4.98 | 4.98 | 4.79 | 10.06 | 9.77 | 10.35 | 9.86 | 10.45
15.84 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 4.79 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 7.42 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 19.92 | 4.79 | 16.21 | 4.49 | 6.05 | 5.08 | 4.79 | 10.16 | 9.77 | 10.16 | 9.77 | 10.35
15.872 | 8.98 | 0.1 | 0.1 | 0.2 | 0.1 | 0.29 | 0.1 | 5.86 | 4.88 | 18.16 | 4.88 | 18.75 | 4.79 | 21.19 | 9.67 | 9.86 | 9.86 | 9.86 | 10.45 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 20.51 | 4.79 | 16.9 | 4.49 | 7.62 | 4.98 | 4.79 | 10.16 | 9.77 | 9.96 | 9.77 | 10.06
15.904 | 9.67 | 0.1 | 0.88 | 0.2 | 0.1 | 0.29 | 0.1 | 5.76 | 4.88 | 18.07 | 4.88 | 19.24 | 4.79 | 21.58 | 9.67 | 9.86 | 9.77 | 9.86 | 11.52 | 0 | 17.58 | 0.1 | 17.77 | 0.1 | 21.09 | 4.79 | 17.58 | 4.59 | 10.06 | 4.98 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 10.06
15.936 | 10.25 | 0 | 1.86 | 0.2 | 0.1 | 0.2 | 0.1 | 5.27 | 4.88 | 18.07 | 4.88 | 19.63 | 4.79 | 22.07 | 9.67 | 10.55 | 9.77 | 9.86 | 10.64 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.68 | 4.79 | 18.26 | 4.59 | 11.23 | 4.79 | 4.79 | 10.06 | 9.86 | 10.16 | 9.86 | 10.25
15.968 | 10.74 | 0 | 3.03 | 0.1 | 0.1 | 0.2 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 22.46 | 9.67 | 11.33 | 9.67 | 9.86 | 10.06 | 0.1 | 5.27 | 0.1 | 17.68 | 0 | 22.27 | 4.79 | 18.85 | 4.59 | 12.31 | 4.79 | 4.79 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
16 | 11.43 | 0.1 | 4 | 0.1 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 22.75 | 9.77 | 12.5 | 9.77 | 9.86 | 9.86 | 0.1 | 1.76 | 0 | 8.11 | 0.1 | 22.85 | 4.79 | 19.43 | 4.59 | 13.18 | 4.69 | 4.79 | 9.96 | 9.86 | 9.86 | 9.86 | 9.96
16.032 | 11.91 | 0 | 4.79 | 0 | 0.2 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 23.15 | 9.77 | 13.57 | 9.67 | 9.86 | 10.16 | 0.1 | 2.25 | 0 | 1.86 | 0.1 | 19.24 | 4.79 | 20.12 | 4.59 | 13.97 | 4.69 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
16.064 | 12.5 | 0.1 | 5.27 | 0 | 1.17 | 0.1 | 0.1 | 4.88 | 4.88 | 17.97 | 4.88 | 17.68 | 4.79 | 23.63 | 9.77 | 14.26 | 9.67 | 9.86 | 10.25 | 0.1 | 1.95 | 0.1 | 2.34 | 0.1 | 18.56 | 4.79 | 20.7 | 4.69 | 14.75 | 4.69 | 4.79 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
16.096 | 13.09 | 0 | 5.96 | 0.1 | 2.25 | 0 | 0.29 | 4.79 | 4.88 | 17.19 | 4.88 | 18.07 | 4.88 | 23.93 | 9.67 | 14.94 | 9.67 | 9.86 | 10.06 | 0.1 | 1.76 | 0.1 | 2.15 | 0 | 18.07 | 4.79 | 21.29 | 4.69 | 15.43 | 4.59 | 4.88 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
16.128 | 13.57 | 0 | 6.45 | 0.1 | 3.52 | 0 | 1.17 | 4.79 | 4.88 | 4.69 | 4.88 | 18.07 | 4.88 | 24.32 | 9.77 | 15.53 | 9.67 | 9.86 | 9.86 | 0.1 | 1.56 | 0.1 | 1.66 | 0 | 18.07 | 4.79 | 21.88 | 4.69 | 16.21 | 4.59 | 5.27 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
16.16 | 14.16 | 0.1 | 7.03 | 0.1 | 4.49 | 0.1 | 2.05 | 4.79 | 4.88 | 2.93 | 4.88 | 18.07 | 4.88 | 21.19 | 9.77 | 16.02 | 9.67 | 9.86 | 9.96 | 0.1 | 1.46 | 0 | 1.76 | 0.1 | 17.97 | 4.79 | 22.46 | 4.69 | 16.9 | 4.49 | 6.35 | 9.77 | 9.86 | 9.77 | 9.96 | 9.77
16.192 | 14.65 | 0.1 | 7.72 | 0 | 4.98 | 0.1 | 3.52 | 4.69 | 4.88 | 3.52 | 4.88 | 18.16 | 4.88 | 21.78 | 9.77 | 16.5 | 9.67 | 9.86 | 9.86 | 0.1 | 1.37 | 0 | 1.46 | 0.1 | 17.87 | 4.88 | 22.95 | 4.69 | 17.58 | 4.49 | 8.01 | 9.77 | 9.96 | 9.77 | 10.84 | 9.77
16.224 | 15.14 | 0.1 | 8.3 | 0.1 | 5.47 | 0.1 | 4.3 | 4.69 | 4.88 | 4.49 | 4.98 | 18.07 | 4.79 | 21.49 | 9.77 | 16.99 | 9.67 | 9.86 | 9.86 | 0 | 1.27 | 0.1 | 1.56 | 0.1 | 17.77 | 4.88 | 23.54 | 4.79 | 18.26 | 4.59 | 9.96 | 9.67 | 10.84 | 9.77 | 12.31 | 9.77
16.256 | 15.63 | 0 | 8.98 | 0.1 | 6.05 | 0 | 4.98 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 4.79 | 21.58 | 9.77 | 17.58 | 9.57 | 9.86 | 9.77 | 0.1 | 1.17 | 0.1 | 1.27 | 0 | 17.77 | 4.79 | 24.12 | 4.79 | 18.85 | 4.59 | 11.43 | 9.67 | 12.21 | 9.67 | 14.36 | 9.67
16.288 | 16.11 | 0 | 9.47 | 0.1 | 6.64 | 0 | 5.47 | 4.59 | 4.88 | 5.76 | 4.98 | 18.16 | 4.88 | 21.49 | 9.77 | 17.97 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 1.27 | 0 | 17.77 | 4.79 | 24.61 | 4.79 | 19.53 | 4.59 | 12.4 | 9.67 | 14.16 | 9.67 | 15.82 | 9.67
16.32 | 16.6 | 0.1 | 10.16 | 0.1 | 7.23 | 0.1 | 6.05 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 18.46 | 9.57 | 9.86 | 9.77 | 0.1 | 0.88 | 0 | 1.27 | 0.1 | 17.77 | 4.88 | 25.2 | 4.79 | 20.12 | 4.49 | 13.28 | 9.67 | 15.72 | 9.67 | 16.99 | 9.67
16.352 | 17.09 | 0 | 10.64 | 0 | 7.81 | 0.1 | 6.64 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 18.85 | 9.67 | 9.86 | 9.77 | 0.1 | 0.88 | 0 | 0.88 | 0.1 | 17.68 | 4.88 | 25.68 | 4.79 | 20.7 | 4.49 | 14.06 | 9.67 | 16.9 | 9.67 | 17.68 | 9.67
16.384 | 13.67 | 0.1 | 11.23 | 0 | 8.4 | 0.1 | 7.23 | 4.59 | 4.88 | 4.88 | 4.98 | 18.07 | 4.88 | 21.49 | 9.77 | 19.34 | 9.67 | 9.86 | 9.67 | 0 | 0.88 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 26.08 | 4.79 | 21.19 | 4.49 | 14.84 | 9.67 | 17.68 | 9.67 | 18.46 | 9.67
16.416 | 14.06 | 0 | 11.91 | 0.1 | 8.98 | 0 | 7.81 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.73 | 9.67 | 9.86 | 9.67 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 17.68 | 4.88 | 21.58 | 4.79 | 21.78 | 4.59 | 15.63 | 9.67 | 18.46 | 9.57 | 19.14 | 9.67
16.448 | 14.16 | 0 | 12.4 | 0.1 | 9.67 | 0 | 8.4 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.12 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.78 | 0 | 8.98 | 4.79 | 20.31 | 4.79 | 22.36 | 4.59 | 16.21 | 9.67 | 19.14 | 9.57 | 19.92 | 9.57
16.48 | 14.45 | 0.1 | 12.99 | 0.1 | 10.25 | 0.1 | 9.08 | 4.49 | 4.88 | 4.79 | 4.88 | 17.19 | 4.88 | 21.39 | 9.86 | 20.61 | 9.67 | 10.74 | 9.67 | 0.1 | 0.49 | 0 | 0.68 | 0.1 | 2.25 | 4.88 | 21 | 4.79 | 22.95 | 4.59 | 16.99 | 9.67 | 19.63 | 9.67 | 20.31 | 9.57
16.512 | 14.36 | 0.1 | 13.48 | 0 | 10.84 | 0.1 | 9.67 | 4.49 | 4.88 | 4.79 | 4.88 | 4.39 | 4.88 | 21.39 | 9.86 | 20.9 | 9.67 | 11.72 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.25 | 4.88 | 21.19 | 4.79 | 23.54 | 4.59 | 17.68 | 9.67 | 20.41 | 9.67 | 20.9 | 9.67
16.544 | 14.36 | 0.1 | 13.97 | 0 | 11.43 | 0.1 | 10.25 | 4.59 | 4.88 | 4.69 | 4.88 | 3.03 | 4.79 | 21.39 | 9.86 | 21.39 | 9.67 | 13.09 | 9.67 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.15 | 4.88 | 21.19 | 4.88 | 24.02 | 4.69 | 18.26 | 9.67 | 20.9 | 9.67 | 21.58 | 9.67
16.576 | 14.45 | 0 | 14.55 | 0.1 | 12.01 | 0 | 10.84 | 4.59 | 4.88 | 4.69 | 4.88 | 3.71 | 4.88 | 16.9 | 9.86 | 21.78 | 9.67 | 13.97 | 9.67 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 1.66 | 4.79 | 21.29 | 4.79 | 24.51 | 4.69 | 18.85 | 9.67 | 21.58 | 9.67 | 22.17 | 9.67
16.608 | 14.36 | 0.1 | 15.04 | 0.1 | 12.6 | 0 | 11.43 | 4.59 | 4.88 | 4.59 | 4.88 | 4.69 | 4.88 | 8.5 | 9.86 | 22.17 | 9.67 | 14.65 | 9.57 | 0.1 | 0.29 | 0.1 | 0.49 | 0 | 1.76 | 4.79 | 21.29 | 4.79 | 25 | 4.69 | 19.43 | 9.67 | 22.17 | 9.67 | 22.75 | 9.67
16.64 | 14.45 | 0.1 | 15.53 | 0.1 | 13.09 | 0.1 | 12.01 | 4.69 | 4.88 | 4.59 | 4.88 | 5.76 | 4.88 | 7.23 | 9.86 | 22.56 | 9.67 | 15.23 | 9.67 | 0.1 | 0.29 | 0.2 | 0.39 | 0.1 | 1.46 | 4.79 | 21.29 | 4.79 | 25.59 | 4.69 | 20.12 | 9.67 | 22.66 | 9.67 | 23.34 | 9.67
16.672 | 14.45 | 0.1 | 16.02 | 0 | 13.57 | 0 | 12.5 | 4.69 | 4.88 | 4.49 | 4.88 | 5.86 | 4.88 | 10.35 | 9.86 | 22.95 | 9.67 | 15.72 | 9.67 | 0.1 | 0.2 | 1.46 | 0.29 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 26.08 | 4.69 | 20.9 | 9.67 | 23.24 | 9.67 | 23.73 | 9.67
16.704 | 14.36 | 0.1 | 16.5 | 0.1 | 14.16 | 0.1 | 13.09 | 4.69 | 4.88 | 4.49 | 4.88 | 5.37 | 4.88 | 11.52 | 9.86 | 23.34 | 9.67 | 16.31 | 9.67 | 0.49 | 0.2 | 2.73 | 0.29 | 0.1 | 1.37 | 4.88 | 21.29 | 4.88 | 21.78 | 4.79 | 21.39 | 9.67 | 23.73 | 9.67 | 24.41 | 9.67
16.736 | 14.36 | 0 | 16.99 | 0.1 | 14.65 | 0 | 13.57 | 4.69 | 5.66 | 4.49 | 4.88 | 5.08 | 4.88 | 10.45 | 9.86 | 23.73 | 9.67 | 16.7 | 9.67 | 1.56 | 0.2 | 4.69 | 0.29 | 0 | 1.37 | 4.79 | 21.19 | 4.88 | 20.22 | 4.79 | 21.97 | 9.67 | 24.22 | 9.67 | 24.9 | 9.67
16.768 | 14.36 | 0 | 16.6 | 0.1 | 15.14 | 0 | 14.16 | 4.69 | 6.15 | 4.49 | 4.88 | 4.88 | 4.88 | 9.86 | 9.86 | 24.02 | 9.77 | 17.19 | 9.67 | 2.93 | 0.2 | 6.15 | 0.2 | 0 | 1.07 | 4.79 | 21.19 | 4.88 | 20.9 | 4.79 | 22.46 | 9.67 | 24.71 | 9.67 | 25.39 | 9.67
16.8 | 14.36 | 0.1 | 14.65 | 0.1 | 15.72 | 0.1 | 14.75 | 4.79 | 7.32 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 24.51 | 9.77 | 17.68 | 9.67 | 4.88 | 0.1 | 7.03 | 0.2 | 0.29 | 1.07 | 4.79 | 21.19 | 4.88 | 21.19 | 4.79 | 23.05 | 9.77 | 25.29 | 9.77 | 25.78 | 9.77
16.832 | 4.98 | 0 | 14.65 | 0 | 16.02 | 0 | 15.14 | 4.79 | 8.5 | 4.49 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 22.36 | 9.77 | 18.16 | 9.67 | 6.25 | 0.1 | 7.81 | 0.2 | 1.56 | 1.07 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 23.54 | 9.77 | 25.88 | 9.77 | 26.47 | 9.77
16.864 | 1.86 | 0.1 | 14.75 | 0 | 16.7 | 0 | 15.72 | 4.79 | 9.47 | 4.49 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 22.17 | 9.77 | 18.56 | 9.67 | 7.13 | 0.1 | 8.59 | 0.1 | 2.83 | 0.78 | 4.88 | 5.47 | 4.88 | 21.19 | 4.79 | 24.02 | 9.77 | 26.47 | 9.77 | 26.86 | 9.77
16.896 | 1.86 | 0 | 14.45 | 0.1 | 17.09 | 0 | 16.11 | 4.79 | 10.16 | 4.49 | 4.88 | 4.79 | 4.88 | 9.96 | 9.86 | 21.88 | 9.77 | 19.04 | 9.67 | 7.81 | 0.1 | 9.38 | 0.1 | 4.98 | 0.78 | 4.79 | 2.83 | 4.88 | 21.19 | 4.79 | 24.61 | 9.77 | 26.76 | 9.67 | 27.34 | 9.67
16.928 | 1.66 | 0.1 | 14.55 | 0.1 | 13.67 | 0 | 16.6 | 4.79 | 10.74 | 4.49 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.68 | 9.77 | 19.43 | 9.67 | 8.59 | 0.1 | 10.25 | 0.1 | 6.25 | 0.59 | 4.79 | 3.42 | 4.88 | 21.29 | 4.79 | 25.1 | 9.77 | 27.25 | 9.67 | 27.74 | 9.77
16.96 | 1.46 | 0.1 | 14.55 | 0.1 | 14.36 | 0.1 | 17.19 | 4.79 | 11.33 | 4.59 | 4.88 | 4.69 | 4.88 | 9.96 | 9.86 | 21.68 | 9.77 | 19.92 | 9.67 | 9.38 | 0.1 | 11.04 | 0.1 | 7.23 | 0.59 | 4.88 | 4.2 | 4.88 | 21.19 | 4.79 | 25.68 | 9.77 | 27.74 | 9.77 | 28.32 | 9.77
16.992 | 1.27 | 0 | 14.45 | 0 | 14.16 | 0.1 | 13.97 | 4.88 | 11.91 | 4.49 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.58 | 9.77 | 20.22 | 9.77 | 10.35 | 0.1 | 11.72 | 0.1 | 8.01 | 0.59 | 4.88 | 5.47 | 4.79 | 21.29 | 4.79 | 26.08 | 9.77 | 28.22 | 9.77 | 28.71 | 9.77
17.024 | 1.17 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 13.97 | 4.79 | 12.5 | 4.59 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 20.8 | 9.67 | 10.94 | 0.1 | 12.5 | 0.1 | 8.69 | 0.49 | 4.88 | 6.25 | 4.88 | 21.19 | 4.79 | 21.78 | 9.77 | 28.61 | 9.77 | 29.2 | 9.77
17.056 | 1.07 | 0 | 14.45 | 0.1 | 14.36 | 0 | 14.16 | 4.79 | 13.09 | 4.69 | 4.88 | 4.59 | 4.88 | 9.67 | 9.86 | 21.49 | 9.77 | 21.19 | 9.67 | 11.72 | 0.1 | 13.28 | 0.1 | 9.57 | 0.39 | 4.79 | 5.86 | 4.88 | 21.19 | 4.79 | 20.7 | 9.77 | 29.1 | 9.77 | 25 | 9.77
17.088 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 13.57 | 4.69 | 4.88 | 4.59 | 4.88 | 9.67 | 9.86 | 21.49 | 9.77 | 21.58 | 9.67 | 12.4 | 0.1 | 14.06 | 0.1 | 10.25 | 0.29 | 4.79 | 5.47 | 4.88 | 21.19 | 4.79 | 21.19 | 9.77 | 26.47 | 9.77 | 23.34 | 9.77
17.12 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 14.16 | 4.69 | 5.18 | 4.59 | 4.88 | 9.77 | 9.86 | 21.39 | 9.86 | 21.97 | 9.77 | 13.18 | 0 | 14.75 | 0.1 | 11.04 | 0.29 | 4.88 | 5.18 | 4.79 | 21.19 | 4.79 | 21.19 | 9.77 | 23.63 | 9.77 | 23.93 | 9.86
17.152 | 0.78 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 14.55 | 4.69 | 5.96 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.77 | 22.36 | 9.77 | 13.97 | 0 | 15.43 | 0.1 | 11.82 | 0.29 | 4.88 | 5.08 | 4.79 | 6.15 | 4.79 | 21.39 | 9.77 | 24.02 | 9.77 | 24.12 | 9.86
17.184 | 0.78 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 15.14 | 4.69 | 6.93 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 22.75 | 9.77 | 14.75 | 0.1 | 16.11 | 0.1 | 12.7 | 0.2 | 4.79 | 5.08 | 4.88 | 2.93 | 4.79 | 21.29 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86
17.216 | 0.59 | 0 | 6.15 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 15.63 | 4.79 | 8.4 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.77 | 23.05 | 9.77 | 15.43 | 0.1 | 16.8 | 0.1 | 13.57 | 0.2 | 5.76 | 5.08 | 4.88 | 3.32 | 4.88 | 21.29 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
17.248 | 0.59 | 0 | 2.05 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 16.11 | 4.79 | 9.38 | 4.49 | 4.88 | 9.67 | 9.86 | 14.06 | 9.86 | 23.44 | 9.77 | 16.11 | 0.1 | 17.48 | 0.1 | 14.16 | 0.1 | 6.54 | 4.98 | 4.88 | 4.1 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
17.28 | 0.49 | 0.1 | 2.15 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 16.6 | 4.79 | 9.96 | 4.59 | 4.88 | 9.67 | 9.86 | 7.03 | 9.86 | 23.83 | 9.77 | 16.8 | 0 | 18.26 | 0.1 | 14.94 | 0.2 | 8.3 | 4.98 | 4.88 | 5.47 | 4.79 | 21.19 | 9.86 | 24.32 | 9.77 | 24.22 | 9.86
17.312 | 0.49 | 0.1 | 1.86 | 0 | 14.36 | 0 | 14.45 | 4.88 | 17.09 | 4.79 | 10.55 | 4.49 | 5.57 | 9.67 | 9.86 | 8.01 | 9.86 | 24.22 | 9.77 | 17.58 | 0 | 18.85 | 0.1 | 15.53 | 0.1 | 10.25 | 4.88 | 4.79 | 6.15 | 4.79 | 21.29 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
17.344 | 0.29 | 0 | 1.46 | 0 | 14.36 | 0 | 14.45 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 6.35 | 9.57 | 9.86 | 10.94 | 9.86 | 21 | 9.77 | 18.16 | 0.1 | 19.53 | 0.1 | 16.21 | 0.1 | 11.62 | 4.79 | 4.88 | 5.96 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
17.376 | 0.29 | 0 | 1.56 | 0.1 | 4.98 | 0 | 14.36 | 4.88 | 17.97 | 4.88 | 11.82 | 4.59 | 7.52 | 9.57 | 9.86 | 11.33 | 9.86 | 20.9 | 9.77 | 18.65 | 0.1 | 20.02 | 0.1 | 17.09 | 0.1 | 12.7 | 4.79 | 4.79 | 5.47 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
17.408 | 0.2 | 0.1 | 1.27 | 0.1 | 1.66 | 0 | 14.45 | 4.79 | 18.46 | 4.79 | 12.4 | 4.69 | 8.69 | 9.57 | 9.86 | 10.45 | 9.86 | 21.19 | 9.77 | 19.43 | 0.1 | 20.61 | 0.1 | 17.68 | 0.1 | 13.48 | 4.69 | 4.88 | 5.27 | 4.88 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
17.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.95 | 0.1 | 5.47 | 4.88 | 18.85 | 4.79 | 12.99 | 4.69 | 9.57 | 9.57 | 9.86 | 9.86 | 9.86 | 21.39 | 9.86 | 20.02 | 0.1 | 21.19 | 0.1 | 18.26 | 0.1 | 14.26 | 4.69 | 4.79 | 5.18 | 4.79 | 20.31 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
17.472 | 0.2 | 0.1 | 0.98 | 0 | 1.66 | 0 | 1.76 | 4.88 | 19.24 | 4.79 | 13.38 | 4.69 | 10.25 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 9.86 | 20.7 | 0.1 | 21.78 | 0.1 | 18.95 | 0.1 | 14.94 | 4.69 | 4.88 | 5.18 | 4.79 | 3.91 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
17.504 | 0.1 | 0.1 | 0.98 | 0 | 1.56 | 0 | 1.95 | 4.88 | 19.73 | 4.79 | 13.97 | 4.69 | 10.84 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 9.86 | 21.29 | 0.1 | 22.46 | 0.1 | 19.53 | 0.1 | 15.63 | 4.59 | 4.79 | 5.08 | 4.88 | 2.93 | 9.77 | 24.12 | 9.86 | 11.04 | 9.86
17.536 | 0 | 0 | 0.88 | 0.1 | 1.37 | 0 | 1.66 | 4.88 | 20.12 | 4.88 | 14.55 | 4.79 | 11.43 | 9.57 | 9.86 | 10.16 | 9.86 | 21.49 | 9.86 | 21.78 | 0.1 | 17.87 | 0.1 | 20.12 | 0.1 | 16.41 | 4.59 | 4.79 | 4.98 | 4.88 | 3.42 | 9.86 | 11.43 | 9.77 | 5.66 | 9.86
17.568 | 0 | 0.1 | 0.68 | 0.1 | 1.17 | 0 | 1.56 | 4.79 | 20.61 | 4.88 | 15.04 | 4.79 | 12.11 | 9.67 | 10.55 | 10.06 | 9.86 | 21.49 | 9.86 | 22.36 | 0.1 | 16.7 | 0.1 | 20.8 | 0.1 | 17.09 | 4.59 | 4.79 | 4.98 | 4.88 | 4.3 | 9.77 | 5.66 | 9.77 | 8.98 | 9.86
17.6 | 0.1 | 1.17 | 0.68 | 0.1 | 1.07 | 0.1 | 1.27 | 4.88 | 21 | 4.88 | 15.53 | 4.79 | 12.6 | 9.67 | 11.33 | 9.86 | 9.86 | 21.39 | 9.86 | 17.87 | 0 | 17.38 | 0.1 | 21.29 | 0.1 | 17.77 | 4.59 | 4.79 | 4.88 | 4.79 | 5.66 | 9.77 | 8.69 | 9.86 | 12.4 | 9.86
17.632 | 0.1 | 1.95 | 0.59 | 0 | 0.88 | 0.1 | 1.27 | 4.88 | 18.16 | 4.79 | 16.02 | 4.79 | 13.09 | 9.67 | 12.31 | 9.86 | 9.96 | 21.39 | 9.86 | 16.8 | 0.1 | 17.58 | 0.1 | 21.88 | 0.1 | 18.46 | 4.59 | 4.79 | 4.79 | 4.79 | 6.15 | 9.77 | 12.11 | 9.86 | 11.33 | 9.86
17.664 | 0.1 | 3.32 | 0.59 | 0 | 0.88 | 0.1 | 1.07 | 4.88 | 18.85 | 4.79 | 16.5 | 4.79 | 13.67 | 9.67 | 13.48 | 9.96 | 9.86 | 21.39 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 22.46 | 0.1 | 18.95 | 4.59 | 4.88 | 4.79 | 4.88 | 5.76 | 9.77 | 11.33 | 9.86 | 10.25 | 9.86
17.696 | 0 | 4.3 | 0.49 | 0.1 | 0.78 | 0 | 0.98 | 4.88 | 18.36 | 4.88 | 16.99 | 4.79 | 14.16 | 9.67 | 14.36 | 9.77 | 9.86 | 21.39 | 9.86 | 17.48 | 0.1 | 17.58 | 0.1 | 17.87 | 0 | 19.73 | 4.59 | 5.76 | 4.69 | 4.88 | 5.37 | 9.77 | 10.25 | 9.77 | 10.06 | 9.86
17.728 | 0 | 4.88 | 0.39 | 0.1 | 0.68 | 0 | 0.88 | 4.88 | 18.26 | 4.88 | 17.38 | 4.79 | 14.75 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 16.7 | 0.1 | 20.22 | 4.49 | 7.23 | 4.69 | 4.79 | 5.18 | 9.77 | 9.96 | 9.77 | 10.35 | 9.86
17.76 | 0.1 | 5.37 | 0.39 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 18.46 | 4.88 | 17.87 | 4.79 | 15.23 | 9.67 | 15.43 | 9.77 | 9.96 | 21.39 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.48 | 0.1 | 20.9 | 4.59 | 9.38 | 4.69 | 4.79 | 5.18 | 9.86 | 10.25 | 9.86 | 10.45 | 9.86
17.792 | 0.1 | 5.96 | 0.29 | 0 | 0.59 | 0.1 | 0.78 | 4.88 | 18.26 | 4.79 | 18.26 | 4.79 | 15.72 | 9.67 | 15.92 | 9.67 | 9.96 | 10.64 | 9.86 | 17.58 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 21.39 | 4.59 | 10.74 | 4.59 | 4.79 | 5.08 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86
17.824 | 0.1 | 6.54 | 0.29 | 0 | 0.49 | 0 | 0.68 | 4.88 | 18.26 | 4.79 | 18.75 | 4.88 | 16.21 | 9.67 | 16.5 | 9.67 | 9.86 | 6.54 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 22.07 | 4.59 | 12.01 | 4.59 | 4.88 | 5.18 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86
17.856 | 0 | 7.13 | 0.2 | 0.1 | 0.49 | 0 | 0.49 | 4.88 | 18.26 | 4.88 | 19.24 | 4.88 | 16.7 | 9.67 | 16.99 | 9.67 | 9.86 | 9.28 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.56 | 4.59 | 12.99 | 4.49 | 4.88 | 4.98 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
17.888 | 0.1 | 7.72 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.88 | 18.16 | 4.88 | 19.73 | 4.88 | 17.09 | 9.67 | 17.48 | 9.67 | 9.86 | 11.52 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0 | 23.05 | 4.59 | 13.77 | 4.59 | 4.79 | 4.98 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
17.92 | 0.1 | 8.3 | 0.2 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.16 | 4.88 | 20.12 | 4.88 | 17.58 | 9.77 | 17.87 | 9.67 | 9.96 | 10.84 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 23.73 | 4.69 | 14.55 | 4.59 | 4.79 | 4.98 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
17.952 | 0.1 | 8.98 | 0.1 | 0 | 0.2 | 0 | 0.39 | 4.88 | 18.07 | 4.88 | 20.51 | 4.88 | 18.07 | 9.77 | 18.36 | 9.67 | 9.86 | 10.16 | 9.86 | 17.58 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 24.22 | 4.69 | 15.23 | 4.59 | 4.79 | 4.88 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
17.984 | 0.1 | 9.57 | 0.1 | 0.68 | 0.2 | 0.1 | 0.39 | 4.88 | 18.07 | 4.79 | 20.9 | 4.88 | 18.56 | 9.77 | 18.75 | 9.67 | 9.96 | 9.86 | 9.96 | 17.58 | 0.1 | 5.96 | 0.1 | 17.77 | 0.1 | 24.71 | 4.69 | 16.02 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
18.016 | 0 | 10.16 | 0.1 | 1.37 | 0.2 | 0.68 | 0.29 | 4.88 | 18.07 | 4.88 | 17.58 | 4.88 | 18.95 | 9.77 | 19.24 | 9.57 | 9.86 | 10.06 | 9.86 | 17.29 | 0.1 | 1.66 | 0.1 | 17.68 | 0 | 25.29 | 4.69 | 16.7 | 4.49 | 4.79 | 4.69 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
18.048 | 0 | 10.74 | 0.1 | 2.64 | 0.1 | 1.27 | 0.2 | 4.88 | 14.36 | 4.88 | 17.77 | 4.88 | 19.34 | 9.77 | 19.73 | 9.57 | 9.86 | 10.16 | 9.86 | 4.39 | 0.1 | 2.25 | 0.1 | 17.68 | 0 | 25.68 | 4.69 | 17.29 | 4.49 | 4.88 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
18.08 | 0.1 | 11.43 | 0.1 | 3.81 | 0.1 | 2.34 | 0.2 | 4.88 | 3.71 | 4.88 | 18.07 | 4.88 | 19.82 | 9.77 | 20.12 | 9.57 | 9.86 | 10.06 | 9.96 | 2.05 | 0 | 1.95 | 0.1 | 17.68 | 0.1 | 26.27 | 4.79 | 17.97 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
18.112 | 0.1 | 11.91 | 0.1 | 4.49 | 0 | 3.71 | 0.2 | 4.88 | 3.13 | 4.79 | 18.07 | 4.88 | 20.22 | 9.77 | 20.51 | 9.57 | 9.96 | 9.86 | 9.96 | 2.25 | 0.1 | 1.76 | 0.1 | 17.68 | 0.1 | 22.75 | 4.79 | 18.65 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
18.144 | 0.1 | 12.5 | 0.1 | 5.08 | 0.1 | 4.49 | 0.1 | 4.88 | 3.61 | 4.88 | 18.07 | 4.88 | 20.61 | 9.77 | 20.9 | 9.57 | 10.64 | 9.96 | 9.96 | 1.95 | 0.1 | 1.56 | 0.1 | 5.37 | 0.1 | 22.07 | 4.79 | 19.34 | 4.59 | 5.86 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
18.176 | 0 | 12.99 | 0.1 | 5.66 | 0.1 | 5.08 | 0.1 | 4.88 | 4.59 | 4.88 | 18.07 | 4.88 | 21 | 9.77 | 21.39 | 9.67 | 11.62 | 9.86 | 9.86 | 1.66 | 0.1 | 1.46 | 0.1 | 1.86 | 0 | 21.78 | 4.79 | 19.92 | 4.59 | 7.23 | 4.59 | 10.94 | 9.77 | 9.86 | 9.77 | 9.86
18.208 | 0 | 13.57 | 0.1 | 6.25 | 0.1 | 5.47 | 0.1 | 4.79 | 5.86 | 4.88 | 18.16 | 4.88 | 19.04 | 9.77 | 21.88 | 9.67 | 12.99 | 9.86 | 9.86 | 1.66 | 0.1 | 1.37 | 0.1 | 2.25 | 0.1 | 21.49 | 4.79 | 20.51 | 4.69 | 9.08 | 4.59 | 11.82 | 9.77 | 10.74 | 9.77 | 9.96
18.24 | 0.1 | 14.06 | 0.1 | 6.84 | 0.1 | 6.15 | 0.1 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 18.85 | 9.77 | 22.17 | 9.67 | 13.87 | 9.67 | 9.96 | 1.46 | 0 | 1.27 | 0.1 | 2.05 | 0.1 | 21.58 | 4.79 | 21.19 | 4.69 | 11.04 | 4.59 | 13.48 | 9.67 | 12.11 | 9.67 | 10.94
18.272 | 0.1 | 14.65 | 0.1 | 7.42 | 0.1 | 6.74 | 0.1 | 4.88 | 5.37 | 4.79 | 18.07 | 4.88 | 18.56 | 9.77 | 22.56 | 9.67 | 14.55 | 9.77 | 9.86 | 1.37 | 0.1 | 1.17 | 0.1 | 1.76 | 0.1 | 21.49 | 4.79 | 21.78 | 4.69 | 12.01 | 4.49 | 15.23 | 9.67 | 13.97 | 9.67 | 12.31
18.304 | 0.1 | 15.04 | 0.1 | 8.11 | 0.1 | 7.32 | 0.1 | 4.88 | 5.18 | 4.88 | 18.16 | 4.88 | 18.26 | 9.77 | 22.85 | 9.67 | 15.14 | 9.77 | 9.96 | 1.27 | 0.1 | 1.07 | 0.1 | 1.76 | 0.1 | 21.39 | 4.79 | 22.27 | 4.69 | 12.99 | 4.49 | 16.5 | 9.67 | 15.53 | 9.67 | 14.26
18.336 | 0 | 15.63 | 0.1 | 8.69 | 0.1 | 8.01 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 18.36 | 9.77 | 23.34 | 9.67 | 15.72 | 9.67 | 9.86 | 1.17 | 0.1 | 0.88 | 0.1 | 1.37 | 0 | 21.39 | 4.79 | 22.75 | 4.69 | 13.87 | 4.59 | 17.38 | 9.67 | 16.7 | 9.67 | 15.92
18.368 | 0 | 16.11 | 0.1 | 9.28 | 0.1 | 8.5 | 0.1 | 4.79 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 9.77 | 23.63 | 9.67 | 16.21 | 9.67 | 9.86 | 0.98 | 0.1 | 0.88 | 0.1 | 1.37 | 0 | 21.29 | 4.79 | 23.34 | 4.79 | 14.65 | 4.59 | 18.07 | 9.67 | 17.58 | 9.67 | 16.9
18.4 | 0.1 | 16.5 | 0.1 | 9.86 | 0.1 | 9.18 | 0.1 | 5.57 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 24.12 | 9.67 | 16.7 | 9.67 | 9.96 | 0.98 | 0 | 0.68 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 23.93 | 4.79 | 15.23 | 4.59 | 18.85 | 9.67 | 18.36 | 9.67 | 17.68
18.432 | 0.1 | 16.99 | 0.1 | 10.45 | 0.1 | 9.67 | 0.1 | 6.15 | 4.79 | 4.88 | 17.77 | 4.88 | 18.16 | 9.77 | 24.41 | 9.67 | 17.19 | 9.67 | 9.96 | 0.88 | 0.1 | 0.68 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.41 | 4.79 | 15.92 | 4.59 | 19.43 | 9.57 | 19.04 | 9.57 | 18.46
18.464 | 0.1 | 13.97 | 0.1 | 11.04 | 0.1 | 10.35 | 0.1 | 7.32 | 4.79 | 4.79 | 5.18 | 4.88 | 18.07 | 9.77 | 22.17 | 9.77 | 17.68 | 9.57 | 10.35 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 25 | 4.79 | 16.7 | 4.59 | 20.12 | 9.57 | 19.63 | 9.67 | 19.04
18.496 | 0 | 13.97 | 0.1 | 11.52 | 0.1 | 10.94 | 0.1 | 8.59 | 4.79 | 4.88 | 3.03 | 4.88 | 18.07 | 9.77 | 22.27 | 9.67 | 18.07 | 9.57 | 11.13 | 0.68 | 0.1 | 0.49 | 0.1 | 0.98 | 0 | 21.19 | 4.79 | 25.49 | 4.79 | 17.29 | 4.69 | 20.8 | 9.67 | 20.41 | 9.67 | 19.82
18.528 | 0.1 | 14.16 | 0.1 | 12.21 | 0.1 | 11.52 | 0.1 | 9.47 | 4.69 | 4.88 | 3.52 | 4.88 | 18.07 | 9.77 | 21.78 | 9.77 | 18.56 | 9.57 | 12.01 | 0.59 | 0.1 | 0.59 | 0.1 | 0.78 | 0 | 15.23 | 4.79 | 25.88 | 4.79 | 18.07 | 4.69 | 21.29 | 9.67 | 20.9 | 9.67 | 20.41
18.56 | 0.1 | 14.45 | 0.1 | 12.7 | 0.1 | 12.01 | 0.1 | 10.16 | 4.59 | 4.88 | 4.69 | 4.88 | 18.07 | 9.77 | 21.78 | 9.67 | 19.04 | 9.57 | 13.38 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 0.1 | 3.61 | 4.88 | 23.44 | 4.79 | 18.75 | 4.69 | 21.88 | 9.67 | 21.49 | 9.67 | 21
18.592 | 0.1 | 14.36 | 0.1 | 13.28 | 0.1 | 12.6 | 0.1 | 10.64 | 4.59 | 4.88 | 5.66 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 19.43 | 9.57 | 14.26 | 0.49 | 0.1 | 0.29 | 0.1 | 0.68 | 0.1 | 3.13 | 4.88 | 20.9 | 4.79 | 19.24 | 4.69 | 22.46 | 9.67 | 21.97 | 9.67 | 21.58
18.624 | 0.1 | 14.36 | 0.1 | 13.77 | 0.1 | 13.18 | 0.1 | 11.33 | 4.59 | 4.79 | 5.86 | 4.88 | 9.28 | 9.77 | 21.58 | 9.77 | 19.92 | 9.57 | 14.84 | 0.49 | 0.2 | 0.29 | 0.1 | 0.59 | 0.1 | 3.52 | 4.88 | 21.39 | 4.79 | 20.02 | 4.69 | 23.05 | 9.67 | 22.56 | 9.67 | 22.07
18.656 | 0 | 14.45 | 0.1 | 14.26 | 0.1 | 13.67 | 0.1 | 12.01 | 4.59 | 4.88 | 5.47 | 4.88 | 3.22 | 9.77 | 21.58 | 9.77 | 20.31 | 9.57 | 15.33 | 0.39 | 1.56 | 0.29 | 0.1 | 0.59 | 0 | 4.39 | 4.79 | 21.39 | 4.79 | 20.51 | 4.69 | 23.63 | 9.67 | 23.05 | 9.67 | 22.75
18.688 | 0 | 14.36 | 0.1 | 14.84 | 0.1 | 14.16 | 0.1 | 12.5 | 4.59 | 4.88 | 5.18 | 4.88 | 3.22 | 9.77 | 21.49 | 9.77 | 20.8 | 9.57 | 15.82 | 0.39 | 2.73 | 0.29 | 1.17 | 0.39 | 0 | 5.76 | 4.79 | 21.29 | 4.79 | 21.19 | 4.79 | 24.02 | 9.67 | 23.63 | 9.67 | 23.15
18.72 | 0.1 | 14.36 | 0.1 | 15.33 | 0 | 14.75 | 0.1 | 13.09 | 4.49 | 4.88 | 4.98 | 4.88 | 4 | 9.77 | 21.49 | 9.77 | 21.09 | 9.57 | 16.31 | 0.29 | 4.88 | 0.2 | 2.25 | 0.49 | 0.1 | 6.25 | 4.88 | 21.19 | 4.79 | 21.68 | 4.79 | 24.61 | 9.67 | 24.12 | 9.67 | 23.73
18.752 | 0.1 | 14.45 | 0.1 | 15.82 | 0 | 15.23 | 0.1 | 13.67 | 4.49 | 4.79 | 4.88 | 4.88 | 5.18 | 9.77 | 21.49 | 9.77 | 21.49 | 9.57 | 16.8 | 0.29 | 6.15 | 0.2 | 4.1 | 0.39 | 0.1 | 5.76 | 4.88 | 21.29 | 4.79 | 22.17 | 4.79 | 25.2 | 9.67 | 24.81 | 9.67 | 24.32
18.784 | 0.1 | 14.36 | 0.1 | 16.31 | 0.1 | 15.72 | 0.1 | 14.16 | 4.59 | 5.37 | 4.88 | 4.88 | 5.86 | 9.77 | 21.49 | 9.77 | 21.97 | 9.67 | 17.29 | 0.2 | 7.03 | 0.2 | 5.66 | 0.29 | 0.1 | 5.47 | 4.88 | 21.29 | 4.79 | 22.85 | 4.79 | 25.68 | 9.67 | 25.29 | 9.67 | 24.9
18.816 | 0 | 14.36 | 0.1 | 16.8 | 0.1 | 16.21 | 0.1 | 14.75 | 4.49 | 5.96 | 4.88 | 4.88 | 5.76 | 9.77 | 21.39 | 9.77 | 22.36 | 9.67 | 17.77 | 0.2 | 7.81 | 0.1 | 6.74 | 0.29 | 0 | 5.08 | 4.88 | 21.29 | 4.79 | 23.34 | 4.79 | 26.17 | 9.67 | 25.78 | 9.67 | 25.39
18.848 | 0 | 14.36 | 0.1 | 17.29 | 0.1 | 16.7 | 0.1 | 15.14 | 4.49 | 6.84 | 4.79 | 4.88 | 5.27 | 9.77 | 21.39 | 9.77 | 22.75 | 9.67 | 18.07 | 0.2 | 8.59 | 0.1 | 7.62 | 0.2 | 0.1 | 5.08 | 4.79 | 21.29 | 4.79 | 23.93 | 4.79 | 26.76 | 9.77 | 26.27 | 9.77 | 25.78
18.88 | 0.1 | 14.36 | 0.1 | 15.23 | 0 | 17.19 | 0.1 | 15.63 | 4.49 | 8.3 | 4.79 | 4.88 | 4.98 | 9.77 | 15.92 | 9.77 | 23.15 | 9.67 | 18.65 | 0.2 | 9.38 | 0.1 | 8.4 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 24.41 | 4.79 | 27.05 | 9.77 | 26.76 | 9.77 | 26.37
18.912 | 0.1 | 5.47 | 0.1 | 15.23 | 0.1 | 14.16 | 0.1 | 16.11 | 4.49 | 9.28 | 4.69 | 4.88 | 4.88 | 9.77 | 7.42 | 9.77 | 23.44 | 9.67 | 19.14 | 0.1 | 10.16 | 0.1 | 9.18 | 0.2 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 25 | 4.79 | 27.64 | 9.77 | 27.34 | 9.67 | 26.76
18.944 | 0.1 | 1.66 | 0.1 | 14.84 | 0.1 | 14.84 | 0.1 | 16.6 | 4.59 | 9.86 | 4.69 | 4.88 | 4.88 | 9.77 | 7.72 | 9.86 | 23.83 | 9.67 | 19.53 | 0.1 | 10.94 | 0.1 | 9.96 | 0.1 | 0.1 | 4.98 | 4.88 | 21.19 | 4.79 | 25.49 | 4.79 | 28.13 | 9.77 | 27.74 | 9.77 | 27.25
18.976 | 0 | 1.95 | 0.1 | 14.75 | 0.1 | 14.55 | 0.1 | 17.09 | 4.59 | 10.55 | 4.59 | 4.88 | 4.88 | 9.77 | 10.94 | 9.86 | 24.22 | 9.67 | 20.02 | 0.1 | 11.72 | 0.1 | 10.84 | 0.1 | 0 | 4.98 | 4.79 | 21.19 | 4.79 | 25.98 | 4.88 | 28.61 | 9.77 | 28.22 | 9.77 | 27.83
19.008 | 0 | 1.56 | 0.1 | 14.75 | 0.1 | 14.45 | 0.1 | 17.58 | 4.69 | 11.13 | 4.59 | 4.88 | 4.79 | 9.77 | 11.43 | 9.86 | 22.27 | 9.67 | 20.41 | 0.1 | 12.5 | 0.1 | 11.62 | 0.1 | 0 | 4.79 | 4.79 | 8.01 | 4.88 | 24.22 | 4.88 | 29.01 | 9.77 | 28.61 | 9.77 | 28.22
19.04 | 0.1 | 1.46 | 0.1 | 14.55 | 0 | 14.55 | 0.1 | 17.97 | 4.59 | 11.72 | 4.59 | 4.88 | 4.79 | 9.77 | 10.45 | 9.86 | 21 | 9.67 | 20.8 | 0.1 | 13.38 | 0 | 12.31 | 0.1 | 0.1 | 4.79 | 4.88 | 2.25 | 4.79 | 20.61 | 4.79 | 29.49 | 9.77 | 29.1 | 9.77 | 28.71
19.072 | 0.1 | 1.27 | 0.1 | 14.55 | 0 | 14.45 | 0.1 | 18.46 | 4.69 | 12.31 | 4.59 | 4.88 | 4.79 | 9.77 | 9.86 | 9.86 | 21.29 | 9.67 | 21.19 | 0.1 | 14.06 | 0 | 13.18 | 0.1 | 0.1 | 4.79 | 4.88 | 3.32 | 4.79 | 21.39 | 4.79 | 26.27 | 9.77 | 28.52 | 9.77 | 29.1
19.104 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.85 | 4.69 | 12.89 | 4.59 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.49 | 9.77 | 21.58 | 0.1 | 14.75 | 0.1 | 13.87 | 0.1 | 0.1 | 4.69 | 4.88 | 3.81 | 4.79 | 21.39 | 4.79 | 25.29 | 9.77 | 23.63 | 9.77 | 26.66
19.136 | 0 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.34 | 4.79 | 13.38 | 4.59 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.39 | 9.77 | 22.07 | 0.1 | 15.33 | 0.1 | 14.55 | 0.1 | 0.98 | 4.69 | 4.79 | 4.98 | 4.79 | 21.29 | 4.88 | 24.71 | 9.77 | 24.32 | 9.77 | 23.44
19.168 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 19.73 | 4.79 | 13.97 | 4.49 | 4.98 | 4.59 | 9.77 | 10.25 | 9.86 | 21.49 | 9.77 | 22.36 | 0.1 | 16.11 | 0.1 | 15.33 | 0 | 1.95 | 4.59 | 4.79 | 6.25 | 4.88 | 21.29 | 4.88 | 24.71 | 9.77 | 24.51 | 9.77 | 24.12
19.2 | 0.1 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 20.12 | 4.79 | 14.45 | 4.49 | 5.86 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 22.75 | 0.1 | 16.8 | 0 | 16.02 | 0.1 | 3.71 | 4.59 | 4.88 | 5.96 | 4.88 | 21.29 | 4.79 | 24.61 | 9.77 | 24.41 | 9.77 | 24.32
19.232 | 0.1 | 0.78 | 0.1 | 14.36 | 0 | 14.36 | 0.1 | 20.61 | 4.79 | 15.04 | 4.49 | 6.84 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 23.24 | 0.1 | 17.38 | 0 | 16.8 | 0.1 | 5.37 | 4.59 | 4.88 | 5.66 | 4.79 | 21.19 | 4.79 | 24.51 | 9.77 | 24.32 | 9.77 | 24.22
19.264 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 21 | 4.79 | 15.53 | 4.59 | 8.2 | 4.49 | 9.96 | 9.96 | 9.86 | 21.49 | 9.77 | 23.63 | 0.1 | 18.07 | 0.1 | 17.38 | 0.1 | 6.64 | 4.59 | 4.88 | 5.27 | 4.79 | 21.29 | 4.79 | 24.41 | 9.77 | 24.32 | 9.86 | 24.22
19.296 | 0 | 0.59 | 0.1 | 10.45 | 0.1 | 14.36 | 0.1 | 18.95 | 4.79 | 16.02 | 4.59 | 9.18 | 4.49 | 10.94 | 9.96 | 9.86 | 21.39 | 9.77 | 23.93 | 0.1 | 18.75 | 0.1 | 18.07 | 0.1 | 7.52 | 4.49 | 4.79 | 5.08 | 4.79 | 21.29 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
19.328 | 0 | 0.49 | 0.1 | 3.03 | 0.1 | 11.82 | 0.1 | 18.85 | 4.79 | 16.5 | 4.59 | 9.86 | 4.49 | 11.82 | 9.86 | 9.86 | 21.39 | 9.77 | 24.32 | 0.1 | 19.43 | 0.1 | 18.65 | 0 | 8.2 | 4.49 | 4.79 | 5.18 | 4.79 | 21.19 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
19.36 | 0.1 | 0.49 | 0.1 | 2.15 | 0.1 | 3.91 | 0.1 | 18.46 | 4.79 | 16.9 | 4.59 | 10.55 | 4.49 | 13.18 | 9.86 | 9.86 | 21.39 | 9.77 | 21 | 0.1 | 20.02 | 0 | 19.34 | 0.1 | 8.98 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22
19.392 | 0.1 | 0.49 | 0.1 | 1.76 | 0.1 | 1.76 | 0.1 | 18.46 | 4.79 | 17.38 | 4.59 | 11.13 | 4.49 | 13.97 | 9.77 | 9.86 | 21.39 | 9.77 | 21.39 | 0.1 | 20.7 | 0 | 19.92 | 0.1 | 9.86 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22
19.424 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 18.46 | 4.88 | 17.77 | 4.69 | 11.72 | 4.49 | 14.65 | 9.77 | 9.86 | 21.39 | 9.77 | 21.39 | 0.1 | 21.19 | 0.1 | 20.61 | 0.1 | 10.55 | 4.59 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
19.456 | 0 | 0.29 | 0.1 | 1.46 | 0.1 | 1.46 | 0.1 | 18.26 | 4.88 | 18.26 | 4.69 | 12.31 | 4.59 | 15.23 | 9.67 | 9.86 | 10.94 | 9.77 | 21.49 | 0.1 | 21.78 | 0.1 | 21.19 | 0 | 11.43 | 4.49 | 4.79 | 4.88 | 4.79 | 7.91 | 4.88 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
19.488 | 0.1 | 0.2 | 0.1 | 1.17 | 0.1 | 1.46 | 0.1 | 18.26 | 4.88 | 18.75 | 4.69 | 12.89 | 4.59 | 15.72 | 9.67 | 9.86 | 6.54 | 9.77 | 21.49 | 0.1 | 22.46 | 0.1 | 21.68 | 0 | 12.21 | 4.59 | 4.79 | 4.79 | 4.79 | 2.44 | 4.79 | 15.33 | 9.86 | 24.22 | 9.86 | 24.12
19.52 | 1.07 | 0.2 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 18.16 | 4.88 | 19.14 | 4.69 | 13.38 | 4.59 | 16.31 | 9.67 | 9.86 | 8.89 | 9.77 | 21.49 | 0.1 | 17.77 | 0 | 22.27 | 0.1 | 12.89 | 4.59 | 4.98 | 4.79 | 4.79 | 3.32 | 4.88 | 6.64 | 9.77 | 12.7 | 9.77 | 24.22
19.552 | 1.95 | 0.1 | 0.1 | 1.07 | 0.1 | 1.07 | 0.1 | 18.16 | 4.79 | 19.63 | 4.79 | 13.87 | 4.69 | 16.8 | 9.67 | 9.86 | 11.33 | 9.86 | 21.49 | 0.1 | 16.7 | 0 | 21.78 | 0.1 | 13.67 | 4.59 | 5.96 | 4.69 | 4.79 | 3.91 | 4.79 | 7.32 | 9.77 | 5.76 | 9.77 | 11.72
19.584 | 3.32 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 18.16 | 4.88 | 20.02 | 4.79 | 14.45 | 4.69 | 17.19 | 9.67 | 9.86 | 11.13 | 9.86 | 21.49 | 0.1 | 17.38 | 0.1 | 18.16 | 0.1 | 14.36 | 4.69 | 7.42 | 4.69 | 4.79 | 5.18 | 4.88 | 11.23 | 9.77 | 8.11 | 9.86 | 5.66
19.616 | 4.3 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 18.07 | 4.88 | 20.51 | 4.79 | 14.94 | 4.69 | 17.58 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.58 | 0.1 | 18.36 | 0.1 | 15.14 | 4.69 | 9.57 | 4.69 | 4.79 | 6.15 | 4.88 | 11.72 | 9.86 | 12.01 | 9.86 | 8.69
19.648 | 4.88 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 0.1 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 18.16 | 9.57 | 9.86 | 9.77 | 9.86 | 21.49 | 0.1 | 17.58 | 0.1 | 18.07 | 0 | 15.82 | 4.59 | 11.04 | 4.69 | 4.79 | 5.96 | 4.88 | 10.45 | 9.86 | 11.72 | 9.86 | 12.21
19.68 | 5.37 | 0.1 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 18.07 | 4.88 | 17.48 | 4.79 | 15.92 | 4.69 | 18.65 | 9.57 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 17.58 | 0 | 17.97 | 0.1 | 16.5 | 4.69 | 12.31 | 4.59 | 4.79 | 5.47 | 4.79 | 9.96 | 9.86 | 10.35 | 9.77 | 11.52
19.712 | 5.96 | 0 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 13.28 | 4.88 | 17.97 | 4.79 | 16.41 | 4.79 | 19.14 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 17.58 | 0 | 17.87 | 0.1 | 17.29 | 4.69 | 13.09 | 4.59 | 4.79 | 5.27 | 4.88 | 10.25 | 9.77 | 9.96 | 9.77 | 10.35
19.744 | 6.54 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.91 | 4.88 | 17.87 | 4.88 | 16.9 | 4.79 | 19.43 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 17.97 | 4.79 | 13.97 | 4.59 | 4.79 | 5.08 | 4.88 | 10.35 | 9.77 | 10.25 | 9.86 | 10.06
19.776 | 7.13 | 0 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.22 | 4.88 | 18.07 | 4.88 | 17.38 | 4.79 | 19.92 | 9.67 | 9.86 | 9.96 | 9.86 | 17.77 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 18.56 | 4.69 | 14.75 | 4.59 | 4.79 | 5.27 | 4.88 | 10.45 | 9.86 | 10.45 | 9.86 | 10.35
19.808 | 7.72 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 3.61 | 4.88 | 18.07 | 4.88 | 17.87 | 4.79 | 20.31 | 9.67 | 9.86 | 9.96 | 9.86 | 8.79 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 19.14 | 4.69 | 15.33 | 4.59 | 4.98 | 5.08 | 4.88 | 10.25 | 9.86 | 10.45 | 9.86 | 10.45
19.84 | 8.3 | 0.1 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 18.26 | 4.79 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 7.23 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 19.92 | 4.79 | 16.11 | 4.59 | 5.96 | 5.08 | 4.88 | 10.06 | 9.86 | 10.16 | 9.86 | 10.45
19.872 | 8.89 | 0.1 | 0.1 | 0.2 | 0.1 | 0.29 | 0.1 | 5.76 | 4.88 | 18.16 | 4.88 | 18.75 | 4.79 | 21.19 | 9.67 | 9.86 | 9.86 | 9.86 | 10.25 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 20.51 | 4.79 | 16.8 | 4.59 | 7.52 | 4.98 | 4.79 | 10.16 | 9.86 | 10.25 | 9.77 | 10.16
19.904 | 9.57 | 0.1 | 0.78 | 0.2 | 0.1 | 0.29 | 0.1 | 5.76 | 4.88 | 18.07 | 4.88 | 19.14 | 4.79 | 21.58 | 9.67 | 9.86 | 9.77 | 9.86 | 11.52 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21 | 4.79 | 17.48 | 4.49 | 9.86 | 5.08 | 4.79 | 10.06 | 9.77 | 10.16 | 9.86 | 9.96
19.936 | 10.16 | 0 | 1.66 | 0.2 | 0.1 | 0.2 | 0.1 | 5.37 | 4.88 | 18.07 | 4.88 | 19.63 | 4.79 | 21.97 | 9.67 | 10.55 | 9.67 | 9.86 | 10.74 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.58 | 4.79 | 18.16 | 4.59 | 11.13 | 4.88 | 4.88 | 9.96 | 9.86 | 10.06 | 9.86 | 10.16
19.968 | 10.64 | 0 | 2.83 | 0.2 | 0.1 | 0.2 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.02 | 4.88 | 22.36 | 9.67 | 11.33 | 9.77 | 9.86 | 10.06 | 0.1 | 5.76 | 0.1 | 17.68 | 0 | 22.17 | 4.79 | 18.75 | 4.59 | 12.21 | 4.79 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
20 | 11.33 | 0.1 | 3.91 | 0.1 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 22.75 | 9.67 | 12.5 | 9.77 | 9.86 | 9.77 | 0.1 | 1.86 | 0 | 8.5 | 0.1 | 22.75 | 4.88 | 19.34 | 4.59 | 13.18 | 4.69 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
20.032 | 11.91 | 0.1 | 4.69 | 0 | 0.2 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 23.15 | 9.67 | 13.48 | 9.67 | 9.86 | 9.96 | 0.1 | 2.25 | 0 | 1.86 | 0.1 | 19.73 | 4.88 | 19.92 | 4.59 | 13.97 | 4.69 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
20.064 | 12.4 | 0.1 | 5.27 | 0 | 1.17 | 0.1 | 0.1 | 4.88 | 4.88 | 17.97 | 4.88 | 17.58 | 4.79 | 23.54 | 9.77 | 14.26 | 9.67 | 9.86 | 10.25 | 0.1 | 1.95 | 0.1 | 2.25 | 0.1 | 18.75 | 4.88 | 20.7 | 4.69 | 14.75 | 4.69 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
20.096 | 12.99 | 0 | 5.76 | 0.1 | 2.05 | 0 | 0.2 | 4.79 | 4.88 | 17.87 | 4.88 | 17.58 | 4.79 | 23.93 | 9.77 | 14.94 | 9.67 | 9.86 | 9.96 | 0.1 | 1.76 | 0.1 | 2.05 | 0.1 | 18.26 | 4.79 | 21.19 | 4.69 | 15.33 | 4.69 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
20.128 | 13.57 | 0 | 6.35 | 0.1 | 3.42 | 0 | 1.17 | 4.79 | 4.88 | 5.27 | 4.88 | 17.77 | 4.88 | 24.32 | 9.67 | 15.43 | 9.67 | 9.96 | 9.86 | 0.1 | 1.66 | 0.1 | 1.66 | 0 | 18.07 | 4.79 | 21.88 | 4.69 | 16.11 | 4.59 | 5.08 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
20.16 | 14.06 | 0.1 | 7.03 | 0.1 | 4.49 | 0 | 2.05 | 4.69 | 4.88 | 2.83 | 4.88 | 18.07 | 4.88 | 20.9 | 9.77 | 15.92 | 9.67 | 9.86 | 9.86 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 18.07 | 4.88 | 22.36 | 4.69 | 16.8 | 4.59 | 6.25 | 9.77 | 9.86 | 9.77 | 9.96 | 9.77
20.192 | 14.55 | 0.1 | 7.52 | 0 | 4.98 | 0 | 3.42 | 4.69 | 4.88 | 3.61 | 4.88 | 18.07 | 4.88 | 21 | 9.77 | 16.5 | 9.67 | 9.96 | 9.86 | 0.1 | 1.27 | 0 | 1.46 | 0.1 | 17.97 | 4.88 | 22.95 | 4.69 | 17.38 | 4.49 | 7.81 | 9.77 | 9.86 | 9.77 | 10.84 | 9.77
20.224 | 15.04 | 0.1 | 8.2 | 0 | 5.47 | 0.1 | 4.3 | 4.69 | 4.88 | 4.59 | 4.88 | 18.07 | 4.79 | 21.29 | 9.77 | 16.9 | 9.57 | 9.86 | 9.77 | 0.1 | 1.27 | 0.1 | 1.56 | 0.1 | 17.87 | 4.88 | 23.44 | 4.79 | 18.16 | 4.59 | 9.86 | 9.67 | 10.74 | 9.77 | 12.21 | 9.67
20.256 | 15.53 | 0 | 8.89 | 0.1 | 6.05 | 0 | 4.98 | 4.59 | 4.88 | 5.66 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 17.38 | 9.57 | 9.86 | 9.77 | 0.1 | 1.17 | 0.1 | 1.37 | 0 | 17.77 | 4.88 | 24.02 | 4.79 | 18.75 | 4.59 | 11.43 | 9.67 | 12.11 | 9.67 | 14.06 | 9.67
20.288 | 16.02 | 0 | 9.38 | 0.1 | 6.54 | 0 | 5.47 | 4.49 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 17.87 | 9.57 | 9.96 | 9.77 | 0.1 | 0.98 | 0.1 | 1.17 | 0 | 17.77 | 4.79 | 24.51 | 4.79 | 19.34 | 4.59 | 12.4 | 9.67 | 14.06 | 9.67 | 15.72 | 9.67
20.32 | 16.5 | 0.1 | 9.96 | 0.1 | 7.13 | 0.1 | 6.05 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 18.46 | 9.57 | 9.86 | 9.77 | 0.1 | 1.07 | 0 | 1.17 | 0.1 | 17.77 | 4.88 | 25.1 | 4.79 | 20.02 | 4.49 | 13.38 | 9.67 | 15.63 | 9.67 | 16.8 | 9.67
20.352 | 16.99 | 0.1 | 10.55 | 0 | 7.81 | 0.1 | 6.54 | 4.59 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 18.85 | 9.57 | 9.86 | 9.67 | 0.1 | 0.88 | 0 | 0.88 | 0.1 | 17.77 | 4.88 | 25.49 | 4.79 | 20.51 | 4.49 | 14.06 | 9.67 | 16.8 | 9.67 | 17.58 | 9.67
20.384 | 14.06 | 0.1 | 11.13 | 0 | 8.4 | 0.1 | 7.23 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.34 | 9.57 | 9.96 | 9.67 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 26.08 | 4.79 | 21.29 | 4.59 | 14.84 | 9.67 | 17.58 | 9.67 | 18.36 | 9.67
20.416 | 13.87 | 0 | 11.82 | 0.1 | 8.98 | 0 | 7.81 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.73 | 9.57 | 9.86 | 9.67 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 21.58 | 4.79 | 21.78 | 4.59 | 15.53 | 9.57 | 18.16 | 9.67 | 19.14 | 9.67
20.448 | 14.16 | 0 | 12.31 | 0.1 | 9.67 | 0 | 8.3 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 20.12 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.68 | 0 | 10.45 | 4.88 | 20.31 | 4.79 | 22.27 | 4.59 | 16.31 | 9.57 | 19.04 | 9.67 | 19.73 | 9.57
20.48 | 14.45 | 0.1 | 12.89 | 0.1 | 10.16 | 0 | 8.98 | 4.59 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 21.39 | 9.86 | 20.51 | 9.67 | 10.74 | 9.67 | 0.1 | 0.49 | 0 | 0.68 | 0.1 | 2.54 | 4.88 | 21 | 4.79 | 22.85 | 4.59 | 16.9 | 9.67 | 19.73 | 9.67 | 20.31 | 9.57
20.512 | 14.36 | 0.1 | 13.38 | 0 | 10.74 | 0 | 9.67 | 4.49 | 4.88 | 4.69 | 4.88 | 5.37 | 4.88 | 21.39 | 9.86 | 20.9 | 9.67 | 11.72 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.25 | 4.88 | 21.19 | 4.79 | 23.44 | 4.69 | 17.68 | 9.67 | 20.22 | 9.67 | 20.9 | 9.67
20.544 | 14.36 | 0.1 | 13.87 | 0 | 11.43 | 0.1 | 10.25 | 4.49 | 4.88 | 4.69 | 4.88 | 2.73 | 4.79 | 21.39 | 9.86 | 21.39 | 9.67 | 13.09 | 9.67 | 0 | 0.49 | 0.1 | 0.49 | 0.1 | 2.15 | 4.88 | 21.19 | 4.88 | 23.93 | 4.69 | 18.26 | 9.67 | 20.8 | 9.67 | 21.58 | 9.67
20.576 | 14.45 | 0 | 14.45 | 0.1 | 12.01 | 0 | 10.84 | 4.59 | 4.88 | 4.69 | 4.88 | 3.52 | 4.88 | 19.73 | 9.86 | 21.68 | 9.67 | 13.97 | 9.57 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 1.76 | 4.88 | 21.19 | 4.88 | 24.51 | 4.69 | 18.85 | 9.67 | 21.49 | 9.67 | 22.07 | 9.67
20.608 | 14.36 | 0.1 | 14.94 | 0.1 | 12.5 | 0 | 11.43 | 4.59 | 4.88 | 4.69 | 4.88 | 4.59 | 4.88 | 9.38 | 9.86 | 22.07 | 9.67 | 14.65 | 9.57 | 0.1 | 0.39 | 0.1 | 0.49 | 0 | 1.86 | 4.88 | 21.19 | 4.88 | 25.1 | 4.79 | 19.53 | 9.67 | 22.07 | 9.67 | 22.56 | 9.67
20.64 | 14.45 | 0.1 | 15.43 | 0.1 | 13.09 | 0.1 | 12.01 | 4.59 | 4.88 | 4.59 | 4.88 | 5.66 | 4.88 | 6.93 | 9.86 | 22.56 | 9.67 | 15.23 | 9.67 | 0.1 | 0.2 | 0.2 | 0.29 | 0.1 | 1.46 | 4.88 | 21.29 | 4.88 | 25.39 | 4.69 | 20.22 | 9.67 | 22.56 | 9.67 | 23.15 | 9.67
20.672 | 14.45 | 0.1 | 15.92 | 0 | 13.57 | 0 | 12.5 | 4.59 | 4.88 | 4.49 | 4.88 | 5.86 | 4.88 | 9.77 | 9.86 | 22.85 | 9.67 | 15.72 | 9.57 | 0.1 | 0.29 | 1.46 | 0.29 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 26.08 | 4.69 | 20.8 | 9.67 | 23.15 | 9.67 | 23.73 | 9.67
20.704 | 14.36 | 0.1 | 16.41 | 0 | 14.16 | 0 | 12.99 | 4.69 | 4.88 | 4.49 | 4.88 | 5.47 | 4.88 | 11.52 | 9.86 | 23.34 | 9.67 | 16.21 | 9.67 | 0.49 | 0.2 | 2.73 | 0.29 | 0.1 | 1.27 | 4.88 | 21.29 | 4.88 | 21.49 | 4.79 | 21.29 | 9.67 | 23.63 | 9.67 | 24.22 | 9.67
20.736 | 14.36 | 0 | 16.9 | 0.1 | 14.65 | 0 | 13.57 | 4.69 | 5.27 | 4.49 | 4.88 | 5.08 | 4.88 | 10.55 | 9.86 | 23.63 | 9.67 | 16.7 | 9.67 | 1.46 | 0.2 | 4.59 | 0.29 | 0 | 1.27 | 4.88 | 21.19 | 4.88 | 20.22 | 4.79 | 21.88 | 9.67 | 24.22 | 9.67 | 24.71 | 9.67
20.768 | 14.36 | 0 | 17.29 | 0.1 | 15.14 | 0 | 14.06 | 4.69 | 6.05 | 4.49 | 4.88 | 4.98 | 4.88 | 10.06 | 9.86 | 24.02 | 9.67 | 17.19 | 9.67 | 2.83 | 0.2 | 6.05 | 0.2 | 0 | 1.07 | 4.88 | 21.19 | 4.88 | 20.9 | 4.79 | 22.46 | 9.67 | 24.71 | 9.77 | 25.29 | 9.67
20.8 | 14.36 | 0.1 | 15.33 | 0.1 | 15.63 | 0.1 | 14.65 | 4.69 | 7.03 | 4.49 | 4.88 | 4.98 | 4.88 | 9.86 | 9.86 | 24.41 | 9.77 | 17.68 | 9.67 | 4.79 | 0.1 | 7.03 | 0.2 | 0.2 | 1.07 | 4.88 | 21.19 | 4.88 | 21.09 | 4.79 | 23.05 | 9.77 | 25.2 | 9.77 | 25.78 | 9.67
20.832 | 5.86 | 0 | 15.04 | 0 | 16.11 | 0.1 | 15.14 | 4.79 | 8.4 | 4.49 | 4.88 | 4.88 | 4.88 | 10.06 | 9.86 | 22.46 | 9.77 | 18.16 | 9.67 | 6.25 | 0.1 | 7.81 | 0.1 | 1.46 | 0.88 | 4.88 | 21.19 | 4.79 | 21.09 | 4.79 | 23.54 | 9.77 | 25.78 | 9.77 | 26.27 | 9.67
20.864 | 1.56 | 0.1 | 14.84 | 0 | 16.6 | 0 | 15.63 | 4.79 | 9.47 | 4.49 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 22.27 | 9.77 | 18.56 | 9.67 | 7.03 | 0.1 | 8.5 | 0.1 | 2.73 | 0.88 | 4.88 | 6.93 | 4.88 | 21.19 | 4.79 | 24.12 | 9.77 | 26.08 | 9.77 | 26.76 | 9.77
20.896 | 1.86 | 0 | 14.65 | 0.1 | 17.09 | 0 | 16.11 | 4.79 | 10.06 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 21.88 | 9.77 | 19.04 | 9.67 | 7.81 | 0.1 | 9.28 | 0.1 | 4.98 | 0.78 | 4.88 | 2.44 | 4.88 | 21.19 | 4.88 | 24.61 | 9.77 | 26.66 | 9.77 | 27.25 | 9.77
20.928 | 1.76 | 0 | 14.65 | 0.1 | 13.77 | 0 | 16.6 | 4.79 | 10.64 | 4.49 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.68 | 9.77 | 19.43 | 9.67 | 8.5 | 0.1 | 10.16 | 0.1 | 6.15 | 0.68 | 4.88 | 3.32 | 4.88 | 21.29 | 4.79 | 25.1 | 9.77 | 27.25 | 9.77 | 27.74 | 9.67
20.96 | 1.37 | 0.1 | 14.55 | 0.1 | 14.16 | 0 | 17.09 | 4.79 | 11.23 | 4.49 | 4.88 | 4.69 | 4.88 | 9.86 | 9.86 | 21.68 | 9.77 | 19.92 | 9.67 | 9.28 | 0 | 11.04 | 0.1 | 7.13 | 0.59 | 4.88 | 4.1 | 4.79 | 21.19 | 4.79 | 25.59 | 9.77 | 27.74 | 9.77 | 28.22 | 9.77
20.992 | 1.46 | 0.1 | 14.55 | 0 | 14.06 | 0.1 | 14.06 | 4.79 | 11.82 | 4.49 | 4.88 | 4.59 | 4.88 | 9.96 | 9.86 | 21.58 | 9.77 | 20.31 | 9.67 | 10.06 | 0.1 | 11.72 | 0.1 | 8.01 | 0.59 | 4.88 | 5.27 | 4.79 | 21.19 | 4.79 | 26.08 | 9.77 | 28.13 | 9.77 | 28.61 | 9.77
21.024 | 1.27 | 0.1 | 14.55 | 0 | 14.45 | 0.1 | 13.97 | 4.79 | 12.4 | 4.59 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 20.8 | 9.67 | 11.04 | 0.1 | 12.5 | 0.1 | 8.69 | 0.49 | 4.88 | 6.25 | 4.88 | 21.19 | 4.88 | 21.49 | 9.77 | 28.61 | 9.86 | 29.1 | 9.77
21.056 | 1.07 | 0 | 14.45 | 0.1 | 14.36 | 0 | 14.16 | 4.79 | 12.99 | 4.59 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 9.67 | 11.72 | 0.1 | 13.28 | 0.1 | 9.57 | 0.39 | 4.79 | 5.96 | 4.88 | 21.19 | 4.79 | 20.51 | 9.77 | 29.1 | 9.77 | 27.25 | 9.77
21.088 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 13.48 | 4.69 | 4.88 | 4.59 | 4.88 | 9.67 | 9.86 | 21.49 | 9.77 | 21.58 | 9.67 | 12.6 | 0.1 | 14.06 | 0.1 | 10.35 | 0.29 | 4.88 | 5.57 | 4.88 | 21.19 | 4.88 | 21 | 9.77 | 28.42 | 9.77 | 23.93 | 9.77
21.12 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 14.06 | 4.69 | 5.08 | 4.59 | 4.88 | 9.67 | 9.86 | 21.39 | 9.77 | 21.88 | 9.77 | 13.18 | 0.1 | 14.65 | 0.1 | 11.04 | 0.29 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 21.19 | 9.77 | 24.71 | 9.86 | 24.12 | 9.77
21.152 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 0 | 14.45 | 4.88 | 14.55 | 4.69 | 5.76 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.77 | 22.27 | 9.77 | 13.97 | 0.1 | 15.33 | 0.1 | 11.82 | 0.29 | 4.88 | 5.18 | 4.79 | 6.64 | 4.79 | 21.29 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86
21.184 | 0.78 | 0 | 14.45 | 0 | 14.45 | 0 | 14.45 | 4.88 | 15.14 | 4.69 | 6.74 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.77 | 22.66 | 9.77 | 14.75 | 0.1 | 16.21 | 0.1 | 12.7 | 0.2 | 4.88 | 5.18 | 4.88 | 2.64 | 4.88 | 21.29 | 9.77 | 24.51 | 9.86 | 24.32 | 9.86
21.216 | 0.59 | 0 | 7.23 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 15.53 | 4.69 | 8.2 | 4.59 | 4.88 | 9.67 | 9.86 | 21.39 | 9.77 | 23.15 | 9.77 | 15.33 | 0.1 | 16.8 | 0.1 | 13.48 | 0.2 | 5.76 | 5.08 | 4.88 | 3.61 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.22 | 9.77
21.248 | 0.59 | 0 | 2.44 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 16.11 | 4.69 | 9.18 | 4.49 | 4.88 | 9.67 | 9.86 | 14.75 | 9.86 | 23.54 | 9.77 | 16.11 | 0.1 | 17.48 | 0.1 | 14.06 | 0.1 | 6.45 | 4.98 | 4.88 | 4.1 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.32 | 9.77
21.28 | 0.49 | 0.1 | 2.05 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 16.6 | 4.79 | 9.96 | 4.49 | 4.88 | 9.67 | 9.86 | 7.13 | 9.86 | 23.83 | 9.77 | 16.8 | 0.1 | 18.16 | 0.1 | 14.75 | 0.2 | 8.11 | 4.98 | 4.88 | 5.37 | 4.79 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
21.312 | 0.49 | 0.1 | 1.76 | 0 | 14.36 | 0 | 14.45 | 4.88 | 16.99 | 4.69 | 10.45 | 4.49 | 5.47 | 9.67 | 9.86 | 7.91 | 9.86 | 24.22 | 9.77 | 17.38 | 0.1 | 18.75 | 0.1 | 15.53 | 0.1 | 10.25 | 4.88 | 4.79 | 6.15 | 4.88 | 21.29 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
21.344 | 0.39 | 0.1 | 1.46 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.58 | 4.69 | 11.23 | 4.59 | 6.25 | 9.57 | 9.86 | 10.84 | 9.86 | 21.19 | 9.77 | 18.07 | 0.1 | 19.43 | 0.1 | 16.31 | 0.1 | 11.52 | 4.88 | 4.88 | 5.96 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
21.376 | 0.29 | 0 | 1.37 | 0.1 | 5.27 | 0 | 14.36 | 4.88 | 17.97 | 4.79 | 11.72 | 4.59 | 7.32 | 9.57 | 9.86 | 11.23 | 9.86 | 20.9 | 9.77 | 18.75 | 0.1 | 19.92 | 0.1 | 16.9 | 0.1 | 12.6 | 4.79 | 4.88 | 5.57 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
21.408 | 0.29 | 0.1 | 1.27 | 0.1 | 1.76 | 0 | 14.36 | 4.88 | 18.46 | 4.79 | 12.31 | 4.59 | 8.59 | 9.57 | 9.86 | 10.45 | 9.86 | 21.19 | 9.77 | 19.34 | 0.1 | 20.61 | 0.1 | 17.58 | 0 | 13.38 | 4.69 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
21.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.86 | 0 | 5.76 | 4.88 | 18.85 | 4.79 | 12.89 | 4.59 | 9.47 | 9.67 | 9.86 | 9.77 | 9.86 | 21.39 | 9.77 | 20.02 | 0 | 21.19 | 0.1 | 18.26 | 0.1 | 14.26 | 4.69 | 4.88 | 5.18 | 4.88 | 20.9 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
21.472 | 0.2 | 0 | 1.07 | 0 | 1.66 | 0 | 1.76 | 4.88 | 19.24 | 4.79 | 13.38 | 4.69 | 10.16 | 9.67 | 9.86 | 9.96 | 9.86 | 21.39 | 9.77 | 20.51 | 0.1 | 21.78 | 0.1 | 18.95 | 0.1 | 14.94 | 4.69 | 4.88 | 5.08 | 4.79 | 4.39 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
21.504 | 0.1 | 0.1 | 0.88 | 0 | 1.46 | 0.1 | 1.95 | 4.88 | 19.73 | 4.79 | 13.87 | 4.69 | 10.74 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 9.77 | 21.19 | 0.1 | 22.36 | 0.1 | 19.53 | 0.1 | 15.72 | 4.59 | 4.88 | 5.08 | 4.88 | 2.83 | 9.86 | 24.12 | 9.86 | 11.91 | 9.86
21.536 | 0.1 | 0 | 0.88 | 0.1 | 1.27 | 0 | 1.76 | 4.88 | 20.12 | 4.79 | 14.45 | 4.69 | 11.33 | 9.57 | 9.86 | 10.16 | 9.86 | 21.49 | 9.77 | 21.68 | 0.1 | 18.16 | 0.1 | 20.12 | 0 | 16.41 | 4.49 | 4.88 | 4.98 | 4.88 | 3.42 | 9.86 | 12.6 | 9.77 | 5.76 | 9.86
21.568 | 0.1 | 0.2 | 0.78 | 0.1 | 1.27 | 0 | 1.46 | 4.88 | 20.61 | 4.79 | 14.94 | 4.69 | 11.91 | 9.67 | 10.55 | 9.96 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 16.9 | 0.1 | 20.61 | 0 | 16.9 | 4.49 | 4.88 | 4.98 | 4.88 | 4.3 | 9.77 | 5.86 | 9.86 | 8.2 | 9.86
21.6 | 0.1 | 1.17 | 0.68 | 0.1 | 1.17 | 0 | 1.37 | 4.88 | 21 | 4.79 | 15.43 | 4.79 | 12.5 | 9.67 | 11.23 | 9.86 | 9.86 | 21.39 | 9.86 | 17.58 | 0 | 17.38 | 0.1 | 21.29 | 0.1 | 17.68 | 4.59 | 4.79 | 4.88 | 4.88 | 5.57 | 9.86 | 8.01 | 9.86 | 12.01 | 9.86
21.632 | 0.1 | 1.95 | 0.59 | 0 | 0.88 | 0 | 1.37 | 4.88 | 17.97 | 4.79 | 15.92 | 4.79 | 13.09 | 9.67 | 12.21 | 9.96 | 9.86 | 21.49 | 9.86 | 16.6 | 0.1 | 17.58 | 0.1 | 21.88 | 0.1 | 18.36 | 4.59 | 4.79 | 4.79 | 4.79 | 6.15 | 9.77 | 12.01 | 9.86 | 11.52 | 9.86
21.664 | 0.1 | 3.42 | 0.49 | 0 | 0.88 | 0 | 1.17 | 4.88 | 18.07 | 4.79 | 16.41 | 4.79 | 13.57 | 9.67 | 13.48 | 9.86 | 9.86 | 21.39 | 9.86 | 17.29 | 0.1 | 17.58 | 0.1 | 22.46 | 0.1 | 19.04 | 4.59 | 4.88 | 4.69 | 4.88 | 5.86 | 9.77 | 11.62 | 9.86 | 10.25 | 9.86
21.696 | 0 | 4.3 | 0.49 | 0.1 | 0.78 | 0 | 0.98 | 4.88 | 18.07 | 4.79 | 16.9 | 4.79 | 14.06 | 9.67 | 14.26 | 9.86 | 9.86 | 21.39 | 9.86 | 17.48 | 0.1 | 17.58 | 0.1 | 17.87 | 0.1 | 19.73 | 4.49 | 5.76 | 4.69 | 4.88 | 5.47 | 9.77 | 10.35 | 9.86 | 9.96 | 9.86
21.728 | 0 | 4.88 | 0.39 | 0.1 | 0.68 | 0 | 0.88 | 4.88 | 18.16 | 4.79 | 17.38 | 4.79 | 14.55 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 9.86 | 17.48 | 0.1 | 17.68 | 0.1 | 16.8 | 0.1 | 20.31 | 4.49 | 7.13 | 4.69 | 4.88 | 5.27 | 9.77 | 9.86 | 9.86 | 10.25 | 9.86
21.76 | 0.1 | 5.37 | 0.39 | 0.1 | 0.68 | 0 | 0.88 | 4.88 | 18.26 | 4.79 | 17.77 | 4.88 | 15.14 | 9.67 | 15.33 | 9.77 | 9.86 | 21.39 | 9.86 | 17.48 | 0 | 17.68 | 0.1 | 17.48 | 0.1 | 20.9 | 4.59 | 9.28 | 4.69 | 4.88 | 5.08 | 9.86 | 10.25 | 9.86 | 10.45 | 9.86
21.792 | 0.1 | 5.96 | 0.29 | 0 | 0.59 | 0 | 0.68 | 4.88 | 18.16 | 4.79 | 18.16 | 4.79 | 15.63 | 9.67 | 15.92 | 9.67 | 9.86 | 11.04 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 21.49 | 4.59 | 10.74 | 4.59 | 4.79 | 5.27 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86
21.824 | 0.1 | 6.54 | 0.29 | 0 | 0.49 | 0 | 0.68 | 4.88 | 18.26 | 4.79 | 18.65 | 4.79 | 16.11 | 9.67 | 16.41 | 9.67 | 9.86 | 6.54 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 22.07 | 4.59 | 12.01 | 4.49 | 4.88 | 5.08 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86
21.856 | 0 | 7.13 | 0.2 | 0.1 | 0.49 | 0 | 0.59 | 4.88 | 18.16 | 4.88 | 19.14 | 4.88 | 16.6 | 9.67 | 16.9 | 9.67 | 9.86 | 8.98 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 22.66 | 4.59 | 12.99 | 4.49 | 4.88 | 5.08 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
21.888 | 0 | 7.72 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 4.88 | 18.16 | 4.88 | 19.53 | 4.88 | 16.99 | 9.67 | 17.38 | 9.67 | 9.86 | 11.52 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 23.15 | 4.59 | 13.77 | 4.59 | 4.88 | 4.98 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
21.92 | 0.1 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.16 | 4.79 | 20.02 | 4.88 | 17.58 | 9.77 | 17.87 | 9.67 | 9.96 | 10.94 | 9.86 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 0.1 | 23.63 | 4.59 | 14.55 | 4.59 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 10.16 | 9.86
21.952 | 0.1 | 8.89 | 0.1 | 0 | 0.2 | 0 | 0.39 | 4.88 | 18.07 | 4.79 | 20.51 | 4.88 | 17.97 | 9.77 | 18.36 | 9.67 | 9.96 | 10.16 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 24.22 | 4.69 | 15.23 | 4.49 | 4.79 | 4.88 | 9.86 | 10.06 | 9.86 | 9.96 | 9.86
21.984 | 0.1 | 9.57 | 0.1 | 0.68 | 0.2 | 0 | 0.29 | 4.88 | 18.07 | 4.79 | 20.9 | 4.88 | 18.46 | 9.77 | 18.65 | 9.67 | 9.96 | 9.86 | 9.86 | 17.58 | 0.1 | 6.45 | 0.1 | 17.77 | 0.1 | 24.81 | 4.69 | 15.82 | 4.49 | 4.88 | 4.79 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
22.016 | 0 | 10.16 | 0.1 | 1.37 | 0.2 | 0.29 | 0.29 | 4.88 | 18.07 | 4.88 | 17.48 | 4.88 | 18.95 | 9.77 | 19.24 | 9.57 | 9.86 | 9.86 | 9.86 | 17.48 | 0.1 | 1.66 | 0.1 | 17.68 | 0 | 25.29 | 4.69 | 16.6 | 4.49 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
22.048 | 0 | 10.64 | 0.1 | 2.64 | 0.1 | 1.17 | 0.2 | 4.88 | 17.77 | 4.88 | 17.58 | 4.88 | 19.24 | 9.77 | 19.63 | 9.67 | 9.86 | 10.25 | 9.86 | 5.27 | 0.1 | 2.25 | 0.1 | 17.68 | 0 | 25.68 | 4.69 | 17.38 | 4.59 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
22.08 | 0.1 | 11.33 | 0.1 | 3.81 | 0.1 | 2.15 | 0.2 | 4.88 | 4.69 | 4.88 | 17.77 | 4.88 | 19.73 | 9.77 | 20.12 | 9.57 | 9.96 | 10.06 | 9.86 | 1.76 | 0 | 1.95 | 0.1 | 17.68 | 0.1 | 26.27 | 4.79 | 17.97 | 4.59 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
22.112 | 0.1 | 11.91 | 0.1 | 4.49 | 0 | 3.61 | 0.2 | 4.88 | 2.83 | 4.79 | 18.07 | 4.88 | 20.12 | 9.77 | 20.51 | 9.57 | 9.96 | 9.96 | 9.86 | 2.25 | 0 | 1.66 | 0.1 | 17.68 | 0.1 | 22.17 | 4.79 | 18.56 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
22.144 | 0.1 | 12.4 | 0.1 | 5.08 | 0.1 | 4.39 | 0.1 | 4.88 | 3.61 | 4.79 | 18.07 | 4.88 | 20.61 | 9.77 | 20.9 | 9.57 | 10.55 | 9.96 | 9.86 | 1.95 | 0.1 | 1.56 | 0.1 | 5.86 | 0.1 | 21.49 | 4.79 | 19.24 | 4.59 | 5.86 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
22.176 | 0 | 12.99 | 0.1 | 5.66 | 0.1 | 4.98 | 0.1 | 4.88 | 4.59 | 4.88 | 18.07 | 4.88 | 21 | 9.77 | 21.29 | 9.67 | 11.62 | 9.86 | 9.86 | 1.66 | 0.1 | 1.46 | 0.1 | 1.86 | 0 | 21.58 | 4.79 | 19.92 | 4.59 | 7.03 | 4.59 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86
22.208 | 0 | 13.57 | 0.1 | 6.25 | 0.1 | 5.47 | 0.1 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 18.95 | 9.77 | 21.68 | 9.67 | 12.89 | 9.86 | 9.86 | 1.66 | 0.1 | 1.56 | 0.1 | 2.15 | 0 | 21.39 | 4.79 | 20.51 | 4.59 | 8.89 | 4.59 | 11.62 | 9.67 | 11.04 | 9.77 | 9.96
22.24 | 0.1 | 14.06 | 0.1 | 6.84 | 0 | 6.05 | 0.1 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 18.95 | 9.77 | 22.07 | 9.67 | 13.87 | 9.77 | 9.96 | 1.37 | 0 | 1.27 | 0.1 | 1.95 | 0.1 | 21.49 | 4.79 | 21.19 | 4.69 | 10.74 | 4.59 | 13.28 | 9.67 | 12.01 | 9.67 | 10.84
22.272 | 0.1 | 14.65 | 0.1 | 7.42 | 0.1 | 6.64 | 0.1 | 4.88 | 5.37 | 4.79 | 18.07 | 4.88 | 18.46 | 9.77 | 22.56 | 9.67 | 14.55 | 9.77 | 9.96 | 1.46 | 0 | 1.17 | 0 | 1.66 | 0.1 | 21.39 | 4.79 | 21.78 | 4.69 | 11.91 | 4.49 | 15.14 | 9.67 | 13.77 | 9.67 | 12.11
22.304 | 0.1 | 15.04 | 0.1 | 8.01 | 0.1 | 7.23 | 0.1 | 4.88 | 5.08 | 4.79 | 18.16 | 4.88 | 18.36 | 9.77 | 22.85 | 9.67 | 15.14 | 9.77 | 9.96 | 1.27 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 21.39 | 4.79 | 22.27 | 4.69 | 12.89 | 4.59 | 16.41 | 9.67 | 15.43 | 9.67 | 14.06
22.336 | 0 | 15.63 | 0.1 | 8.59 | 0.1 | 7.81 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 18.36 | 9.77 | 23.34 | 9.67 | 15.63 | 9.67 | 9.86 | 1.17 | 0.1 | 0.88 | 0.1 | 1.46 | 0 | 21.39 | 4.79 | 22.75 | 4.69 | 13.77 | 4.59 | 17.38 | 9.67 | 16.7 | 9.67 | 15.72
22.368 | 0 | 16.02 | 0.1 | 9.28 | 0.1 | 8.4 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 23.63 | 9.67 | 16.11 | 9.67 | 9.86 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 0 | 21.29 | 4.79 | 23.34 | 4.69 | 14.55 | 4.59 | 17.97 | 9.67 | 17.38 | 9.67 | 16.8
22.4 | 0.1 | 16.5 | 0.1 | 9.77 | 0 | 8.98 | 0.1 | 5.66 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 24.02 | 9.67 | 16.6 | 9.67 | 9.96 | 0.88 | 0 | 0.78 | 0.1 | 1.27 | 0.1 | 21.39 | 4.88 | 23.83 | 4.69 | 15.33 | 4.59 | 18.75 | 9.67 | 18.16 | 9.57 | 17.58
22.432 | 0.1 | 16.99 | 0.1 | 10.45 | 0.1 | 9.67 | 0.1 | 6.05 | 4.88 | 4.79 | 18.07 | 4.88 | 18.16 | 9.77 | 24.32 | 9.67 | 17.19 | 9.67 | 9.96 | 0.88 | 0.1 | 0.68 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.32 | 4.79 | 15.92 | 4.49 | 19.43 | 9.57 | 18.95 | 9.67 | 18.36
22.464 | 0.1 | 14.06 | 0.1 | 10.94 | 0.1 | 10.25 | 0.1 | 7.13 | 4.79 | 4.79 | 6.35 | 4.88 | 18.07 | 9.77 | 22.07 | 9.67 | 17.58 | 9.67 | 10.06 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 24.81 | 4.69 | 16.7 | 4.59 | 20.12 | 9.57 | 19.43 | 9.67 | 19.14
22.496 | 0 | 13.97 | 0.1 | 11.62 | 0.1 | 10.84 | 0.1 | 8.59 | 4.69 | 4.79 | 2.73 | 4.88 | 18.07 | 9.77 | 22.27 | 9.67 | 18.07 | 9.57 | 11.04 | 0.78 | 0.1 | 0.59 | 0.1 | 0.88 | 0 | 21.19 | 4.79 | 25.39 | 4.79 | 17.38 | 4.69 | 20.7 | 9.57 | 20.22 | 9.67 | 19.73
22.528 | 0 | 14.16 | 0.1 | 12.11 | 0.1 | 11.43 | 0.1 | 9.38 | 4.69 | 4.88 | 3.52 | 4.88 | 18.07 | 9.77 | 21.78 | 9.67 | 18.46 | 9.57 | 11.91 | 0.68 | 0.1 | 0.49 | 0.1 | 0.78 | 0 | 17.68 | 4.79 | 25.98 | 4.79 | 17.97 | 4.59 | 21.19 | 9.67 | 20.9 | 9.67 | 20.22
22.56 | 0.1 | 14.45 | 0.1 | 12.7 | 0.1 | 12.01 | 0.1 | 10.06 | 4.69 | 4.88 | 4.49 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 18.95 | 9.57 | 13.18 | 0.59 | 0 | 0.39 | 0.1 | 0.78 | 0.1 | 4 | 4.88 | 24.41 | 4.79 | 18.65 | 4.69 | 21.88 | 9.67 | 21.49 | 9.67 | 20.9
22.592 | 0.1 | 14.36 | 0.1 | 13.28 | 0 | 12.6 | 0.1 | 10.64 | 4.59 | 4.79 | 5.57 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 19.34 | 9.57 | 14.16 | 0.49 | 0 | 0.29 | 0 | 0.68 | 0.1 | 3.13 | 4.88 | 21.78 | 4.79 | 19.24 | 4.69 | 22.36 | 9.67 | 21.97 | 9.67 | 21.49
22.624 | 0.1 | 14.36 | 0.1 | 13.77 | 0.1 | 13.09 | 0.1 | 11.33 | 4.59 | 4.79 | 5.96 | 4.88 | 10.74 | 9.77 | 21.58 | 9.77 | 19.92 | 9.57 | 14.75 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 0.1 | 3.61 | 4.88 | 21.68 | 4.79 | 19.92 | 4.69 | 22.95 | 9.57 | 22.56 | 9.67 | 22.07
22.656 | 0 | 14.45 | 0.1 | 14.26 | 0.1 | 13.67 | 0.1 | 11.91 | 4.49 | 4.88 | 5.47 | 4.88 | 3.42 | 9.77 | 21.49 | 9.77 | 20.22 | 9.57 | 15.33 | 0.39 | 1.27 | 0.29 | 0.1 | 0.59 | 0 | 4.3 | 4.79 | 21.58 | 4.79 | 20.51 | 4.69 | 23.54 | 9.67 | 23.15 | 9.67 | 22.75
22.688 | 0 | 14.36 | 0.1 | 14.75 | 0.1 | 14.06 | 0.1 | 12.4 | 4.59 | 4.88 | 5.27 | 4.88 | 3.42 | 9.77 | 21.49 | 9.77 | 20.7 | 9.67 | 15.82 | 0.39 | 2.64 | 0.29 | 1.07 | 0.49 | 0 | 5.66 | 4.79 | 21.49 | 4.79 | 21 | 4.69 | 24.02 | 9.67 | 23.63 | 9.67 | 23.24
22.72 | 0.1 | 14.36 | 0.1 | 15.23 | 0 | 14.65 | 0.1 | 12.99 | 4.49 | 4.88 | 4.88 | 4.88 | 3.71 | 9.77 | 21.49 | 9.77 | 21 | 9.57 | 16.31 | 0.29 | 4.49 | 0.2 | 2.05 | 0.49 | 0.1 | 6.15 | 4.88 | 21.29 | 4.79 | 21.68 | 4.79 | 24.61 | 9.67 | 24.22 | 9.67 | 23.73
22.752 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 15.14 | 0.1 | 13.57 | 4.49 | 4.79 | 4.88 | 4.88 | 4.98 | 9.77 | 21.49 | 9.77 | 21.49 | 9.57 | 16.8 | 0.29 | 5.86 | 0.2 | 3.91 | 0.39 | 0.1 | 5.96 | 4.88 | 21.29 | 4.79 | 22.17 | 4.69 | 25 | 9.67 | 24.81 | 9.67 | 24.32
22.784 | 0.1 | 14.36 | 0.1 | 16.31 | 0.1 | 15.72 | 0.1 | 14.06 | 4.49 | 4.98 | 4.98 | 4.88 | 5.96 | 9.77 | 21.49 | 9.77 | 21.88 | 9.67 | 17.38 | 0.2 | 6.93 | 0.2 | 5.57 | 0.29 | 0.1 | 5.37 | 4.88 | 21.29 | 4.79 | 22.75 | 4.79 | 25.49 | 9.67 | 25.2 | 9.67 | 24.9
22.816 | 0 | 14.36 | 0.1 | 16.8 | 0.1 | 16.21 | 0.1 | 14.65 | 4.49 | 5.76 | 4.88 | 4.88 | 5.76 | 9.77 | 21.39 | 9.77 | 22.27 | 9.67 | 17.77 | 0.2 | 7.72 | 0.2 | 6.64 | 0.29 | 0 | 5.18 | 4.88 | 21.29 | 4.79 | 23.44 | 4.79 | 26.17 | 9.67 | 25.68 | 9.77 | 25.29
22.848 | 0 | 14.36 | 0.1 | 17.29 | 0.1 | 16.7 | 0.1 | 15.04 | 4.49 | 6.84 | 4.79 | 4.88 | 5.47 | 9.77 | 21.39 | 9.77 | 22.66 | 9.67 | 18.16 | 0.2 | 8.5 | 0.1 | 7.42 | 0.2 | 0 | 5.08 | 4.79 | 21.29 | 4.88 | 23.83 | 4.79 | 26.56 | 9.67 | 26.27 | 9.77 | 25.68
22.88 | 0.1 | 14.36 | 0.1 | 14.84 | 0 | 17.19 | 0.1 | 15.63 | 4.49 | 8.2 | 4.69 | 4.88 | 4.98 | 9.77 | 16.9 | 9.77 | 23.05 | 9.67 | 18.56 | 0.1 | 9.28 | 0.1 | 8.3 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 24.41 | 4.79 | 27.15 | 9.77 | 26.66 | 9.77 | 26.27
22.912 | 0.1 | 5.76 | 0.1 | 15.23 | 0 | 13.77 | 0.1 | 16.11 | 4.59 | 9.18 | 4.69 | 4.88 | 4.88 | 9.77 | 7.81 | 9.77 | 23.34 | 9.67 | 19.04 | 0.1 | 10.06 | 0 | 9.08 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 25 | 4.79 | 27.64 | 9.67 | 27.15 | 9.77 | 26.66
22.944 | 0.1 | 1.66 | 0.1 | 14.75 | 0.1 | 14.16 | 0.1 | 16.5 | 4.59 | 9.86 | 4.69 | 4.88 | 4.98 | 9.77 | 7.62 | 9.86 | 23.83 | 9.67 | 19.53 | 0.1 | 11.04 | 0.1 | 9.86 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 25.49 | 4.79 | 28.03 | 9.67 | 27.64 | 9.77 | 27.15
22.976 | 0 | 1.95 | 0.1 | 14.65 | 0.1 | 14.36 | 0.1 | 17.09 | 4.59 | 10.55 | 4.59 | 4.88 | 4.88 | 9.77 | 10.74 | 9.86 | 24.12 | 9.67 | 19.92 | 0.1 | 11.72 | 0.1 | 10.64 | 0.1 | 0 | 4.98 | 4.79 | 21.19 | 4.79 | 25.88 | 4.79 | 28.61 | 9.77 | 28.22 | 9.77 | 27.74
23.008 | 0 | 1.76 | 0.1 | 14.65 | 0.1 | 14.36 | 0.1 | 17.58 | 4.69 | 11.13 | 4.59 | 4.88 | 4.79 | 9.77 | 11.43 | 9.86 | 23.15 | 9.67 | 20.31 | 0.1 | 12.5 | 0.1 | 11.62 | 0.1 | 0 | 4.79 | 4.79 | 8.69 | 4.79 | 25.78 | 4.79 | 28.91 | 9.77 | 28.61 | 9.77 | 28.22
23.04 | 0.1 | 1.46 | 0.1 | 14.55 | 0 | 14.45 | 0.1 | 17.97 | 4.69 | 11.72 | 4.59 | 4.88 | 4.79 | 9.77 | 10.45 | 9.77 | 21.29 | 9.67 | 20.8 | 0.1 | 13.18 | 0 | 12.31 | 0.1 | 0.1 | 4.79 | 4.88 | 2.44 | 4.79 | 21.68 | 4.79 | 29.49 | 9.77 | 29.1 | 9.77 | 28.71
23.072 | 0.1 | 1.27 | 0.1 | 14.55 | 0 | 14.45 | 0.1 | 18.46 | 4.69 | 12.31 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.58 | 9.67 | 21.19 | 0.1 | 13.97 | 0 | 13.09 | 0.1 | 0.1 | 4.79 | 4.88 | 3.22 | 4.79 | 21.78 | 4.79 | 25.98 | 9.77 | 29.49 | 9.77 | 29.1
23.104 | 0 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.95 | 4.69 | 12.79 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.68 | 9.67 | 21.58 | 0.1 | 14.75 | 0.1 | 13.77 | 0.1 | 0.1 | 4.69 | 4.88 | 3.81 | 4.79 | 21.58 | 4.79 | 25.1 | 9.77 | 25.78 | 9.77 | 27.64
23.136 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.34 | 4.79 | 13.38 | 4.49 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.39 | 9.77 | 21.97 | 0.1 | 15.33 | 0.1 | 14.45 | 0 | 0.88 | 4.69 | 4.79 | 4.88 | 4.79 | 21.49 | 4.79 | 24.9 | 9.77 | 24.81 | 9.86 | 23.73
23.168 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 19.73 | 4.79 | 13.87 | 4.49 | 4.98 | 4.59 | 9.77 | 10.16 | 9.86 | 21.49 | 9.77 | 22.36 | 0.1 | 16.11 | 0.1 | 15.14 | 0 | 1.76 | 4.59 | 4.79 | 6.25 | 4.79 | 21.49 | 4.79 | 24.61 | 9.77 | 24.71 | 9.77 | 24.41
23.2 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 20.12 | 4.79 | 14.45 | 4.49 | 5.66 | 4.59 | 9.77 | 10.06 | 9.86 | 21.49 | 9.67 | 22.75 | 0.1 | 16.8 | 0 | 16.02 | 0.1 | 3.42 | 4.59 | 4.88 | 6.15 | 4.79 | 21.39 | 4.79 | 24.61 | 9.77 | 24.61 | 9.86 | 24.41
23.232 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 20.7 | 4.79 | 14.94 | 4.49 | 6.74 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 23.15 | 0.1 | 17.38 | 0 | 16.7 | 0.1 | 5.27 | 4.59 | 4.88 | 5.57 | 4.79 | 21.29 | 4.79 | 24.51 | 9.77 | 24.51 | 9.77 | 24.32
23.264 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 21 | 4.79 | 15.43 | 4.59 | 8.11 | 4.49 | 9.96 | 9.96 | 9.86 | 21.49 | 9.77 | 23.54 | 0.1 | 18.07 | 0.1 | 17.29 | 0.1 | 6.54 | 4.59 | 4.88 | 5.27 | 4.79 | 21.29 | 4.79 | 24.41 | 9.77 | 24.41 | 9.86 | 24.22
23.296 | 0 | 0.59 | 0.1 | 11.23 | 0.1 | 14.36 | 0.1 | 18.26 | 4.79 | 15.92 | 4.59 | 9.18 | 4.49 | 10.84 | 9.86 | 9.86 | 21.49 | 9.77 | 23.93 | 0.1 | 18.75 | 0.1 | 17.97 | 0.1 | 7.32 | 4.59 | 4.79 | 5.08 | 4.79 | 21.29 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
23.328 | 0 | 0.59 | 0.1 | 3.13 | 0.1 | 14.16 | 0.1 | 18.46 | 4.79 | 16.41 | 4.59 | 9.86 | 4.49 | 11.72 | 9.86 | 9.86 | 21.39 | 9.77 | 24.32 | 0.1 | 19.34 | 0.1 | 18.46 | 0 | 8.11 | 4.49 | 4.79 | 5.18 | 4.88 | 21.29 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
23.36 | 0.1 | 0.49 | 0.1 | 1.95 | 0 | 4.3 | 0.1 | 18.26 | 4.79 | 16.9 | 4.59 | 10.45 | 4.49 | 13.18 | 9.77 | 9.86 | 21.49 | 9.77 | 20.8 | 0.1 | 19.92 | 0 | 19.24 | 0.1 | 8.89 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
23.392 | 0.1 | 0.49 | 0.1 | 1.76 | 0 | 1.86 | 0.1 | 18.26 | 4.79 | 17.38 | 4.59 | 11.13 | 4.49 | 13.97 | 9.77 | 9.86 | 21.39 | 9.77 | 20.9 | 0.1 | 20.61 | 0.1 | 19.82 | 0.1 | 9.67 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
23.424 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.95 | 0.1 | 18.26 | 4.88 | 17.87 | 4.69 | 11.72 | 4.49 | 14.65 | 9.77 | 9.86 | 21.39 | 9.77 | 21.19 | 0.1 | 21.09 | 0.1 | 20.41 | 0.1 | 10.45 | 4.59 | 4.79 | 4.98 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
23.456 | 0 | 0.29 | 0.1 | 1.56 | 0.1 | 1.56 | 0.1 | 18.26 | 4.88 | 18.26 | 4.69 | 12.31 | 4.59 | 15.14 | 9.67 | 9.86 | 11.52 | 9.77 | 21.39 | 0.1 | 21.68 | 0.1 | 21 | 0 | 11.33 | 4.49 | 4.79 | 4.98 | 4.79 | 9.18 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
23.488 | 0.1 | 0.2 | 0 | 1.27 | 0.1 | 1.37 | 0.1 | 18.26 | 4.88 | 18.75 | 4.69 | 12.89 | 4.59 | 15.72 | 9.67 | 9.86 | 6.74 | 9.86 | 21.39 | 0.1 | 22.27 | 0.1 | 21.68 | 0 | 12.01 | 4.59 | 4.79 | 4.88 | 4.79 | 2.54 | 4.88 | 17.58 | 9.77 | 24.22 | 9.86 | 24.12
23.52 | 1.07 | 0.2 | 0.1 | 1.17 | 0 | 1.37 | 0.1 | 18.16 | 4.88 | 19.24 | 4.69 | 13.38 | 4.59 | 16.21 | 9.67 | 9.86 | 8.59 | 9.77 | 21.39 | 0.1 | 18.16 | 0 | 22.27 | 0.1 | 12.79 | 4.59 | 4.88 | 4.79 | 4.79 | 3.32 | 4.79 | 6.93 | 9.77 | 13.87 | 9.86 | 24.12
23.552 | 1.76 | 0.2 | 0.1 | 0.88 | 0 | 0.98 | 0.1 | 18.16 | 4.88 | 19.53 | 4.69 | 13.97 | 4.59 | 16.7 | 9.67 | 9.86 | 11.33 | 9.77 | 21.49 | 0.1 | 16.99 | 0 | 22.66 | 0.1 | 13.57 | 4.59 | 5.86 | 4.69 | 4.79 | 3.71 | 4.79 | 6.93 | 9.77 | 6.25 | 9.86 | 12.21
23.584 | 3.13 | 0.1 | 0.1 | 0.98 | 0.1 | 1.17 | 0.1 | 18.16 | 4.88 | 20.02 | 4.79 | 14.55 | 4.59 | 17.19 | 9.67 | 9.86 | 11.23 | 9.86 | 21.49 | 0.1 | 17.58 | 0.1 | 19.24 | 0.1 | 14.26 | 4.69 | 7.23 | 4.69 | 4.79 | 4.98 | 4.79 | 11.13 | 9.77 | 7.62 | 9.86 | 5.66
23.616 | 4.2 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 18.16 | 4.88 | 20.41 | 4.79 | 14.94 | 4.69 | 17.68 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 18.65 | 0.1 | 15.04 | 4.69 | 9.38 | 4.69 | 4.79 | 6.15 | 4.88 | 11.91 | 9.77 | 11.52 | 9.86 | 8.4
23.648 | 4.79 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 0.1 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 18.07 | 9.57 | 9.86 | 9.77 | 9.86 | 21.49 | 0.1 | 17.58 | 0.1 | 18.26 | 0 | 15.82 | 4.59 | 10.84 | 4.59 | 4.79 | 6.05 | 4.79 | 10.64 | 9.86 | 11.82 | 9.86 | 11.82
23.68 | 5.37 | 0.1 | 0.1 | 0.68 | 0 | 0.88 | 0.1 | 18.07 | 4.88 | 17.38 | 4.79 | 16.02 | 4.69 | 18.56 | 9.67 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 17.58 | 0 | 18.07 | 0.1 | 16.5 | 4.69 | 12.01 | 4.59 | 4.79 | 5.57 | 4.79 | 9.96 | 9.86 | 10.55 | 9.86 | 11.62
23.712 | 5.86 | 0.1 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 15.33 | 4.88 | 17.77 | 4.79 | 16.41 | 4.69 | 19.04 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 17.58 | 0 | 17.97 | 0.1 | 17.19 | 4.69 | 12.99 | 4.59 | 4.79 | 5.27 | 4.79 | 10.25 | 9.77 | 9.96 | 9.86 | 10.35
23.744 | 6.45 | 0.1 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 4.49 | 4.88 | 17.77 | 4.88 | 16.9 | 4.69 | 19.43 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 17.87 | 4.79 | 13.87 | 4.59 | 4.79 | 5.18 | 4.79 | 10.45 | 9.77 | 10.25 | 9.86 | 9.96
23.776 | 7.03 | 0 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.13 | 4.88 | 18.07 | 4.79 | 17.38 | 4.79 | 19.92 | 9.67 | 9.86 | 9.96 | 9.86 | 21.19 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 18.36 | 4.69 | 14.55 | 4.59 | 4.79 | 5.08 | 4.79 | 10.45 | 9.77 | 10.45 | 9.86 | 10.35
23.808 | 7.62 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.61 | 4.88 | 18.07 | 4.88 | 17.77 | 4.79 | 20.22 | 9.67 | 9.86 | 9.86 | 9.86 | 10.25 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 19.04 | 4.69 | 15.33 | 4.59 | 4.88 | 5.08 | 4.79 | 10.25 | 9.77 | 10.45 | 9.86 | 10.45
23.84 | 8.2 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 18.26 | 4.79 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 6.84 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 19.73 | 4.79 | 16.11 | 4.49 | 5.86 | 5.08 | 4.79 | 10.25 | 9.86 | 10.25 | 9.86 | 10.35
23.872 | 8.89 | 0.1 | 0.1 | 0.2 | 0 | 0.39 | 0.1 | 5.76 | 4.88 | 18.07 | 4.88 | 18.75 | 4.79 | 21.09 | 9.67 | 9.86 | 9.86 | 9.86 | 9.47 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 20.31 | 4.79 | 16.8 | 4.49 | 7.23 | 4.98 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16
23.904 | 9.47 | 0.1 | 0.68 | 0.2 | 0.1 | 0.29 | 0.1 | 5.86 | 4.88 | 18.07 | 4.88 | 19.24 | 4.79 | 21.49 | 9.67 | 9.86 | 9.77 | 9.86 | 11.62 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 20.9 | 4.79 | 17.38 | 4.59 | 9.67 | 5.08 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 10.16
23.936 | 10.16 | 0 | 1.56 | 0.2 | 0.1 | 0.2 | 0.1 | 5.37 | 4.88 | 18.07 | 4.88 | 19.63 | 4.79 | 21.97 | 9.67 | 10.45 | 9.67 | 9.86 | 11.04 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.58 | 4.79 | 18.07 | 4.59 | 10.94 | 4.88 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 10.16
23.968 | 10.64 | 0 | 2.64 | 0.2 | 0.1 | 0.2 | 0.1 | 5.08 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 22.36 | 9.67 | 11.13 | 9.77 | 9.86 | 10.25 | 0.1 | 6.74 | 0.1 | 17.68 | 0 | 22.07 | 4.79 | 18.56 | 4.59 | 12.11 | 4.79 | 4.79 | 9.96 | 9.86 | 10.06 | 9.86 | 10.16
24 | 11.23 | 0.1 | 3.81 | 0.1 | 0 | 0.2 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 22.75 | 9.67 | 12.31 | 9.67 | 9.86 | 9.77 | 0.1 | 1.56 | 0.1 | 9.67 | 0.1 | 22.75 | 4.88 | 19.34 | 4.59 | 12.99 | 4.69 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
24.032 | 11.72 | 0.1 | 4.49 | 0 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 23.05 | 9.67 | 13.48 | 9.67 | 9.86 | 10.16 | 0.1 | 2.25 | 0 | 2.05 | 0.1 | 19.24 | 4.88 | 20.02 | 4.59 | 13.77 | 4.69 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
24.064 | 12.4 | 0.1 | 5.18 | 0 | 1.07 | 0.1 | 0.1 | 4.98 | 4.88 | 17.97 | 4.88 | 17.48 | 4.79 | 23.54 | 9.67 | 14.26 | 9.67 | 9.96 | 10.25 | 0.1 | 1.95 | 0.1 | 2.34 | 0.1 | 18.56 | 4.88 | 20.61 | 4.69 | 14.65 | 4.69 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
24.096 | 12.89 | 0 | 5.66 | 0.1 | 1.95 | 0 | 0.2 | 4.79 | 4.88 | 17.97 | 4.88 | 17.58 | 4.88 | 23.93 | 9.67 | 14.84 | 9.67 | 9.86 | 9.96 | 0.1 | 1.66 | 0.1 | 2.15 | 0.1 | 18.26 | 4.88 | 21.19 | 4.69 | 15.33 | 4.69 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
24.128 | 13.57 | 0.1 | 6.25 | 0.1 | 3.22 | 0 | 1.17 | 4.79 | 4.88 | 5.76 | 4.88 | 17.77 | 4.88 | 24.22 | 9.67 | 15.33 | 9.67 | 9.86 | 9.86 | 0.1 | 1.56 | 0.1 | 1.66 | 0 | 17.97 | 4.79 | 21.78 | 4.69 | 16.02 | 4.69 | 4.98 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
24.16 | 13.97 | 0.1 | 6.84 | 0.1 | 4.3 | 0 | 1.86 | 4.69 | 4.88 | 2.83 | 4.88 | 18.07 | 4.88 | 20.7 | 9.77 | 15.92 | 9.67 | 9.86 | 9.86 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 18.07 | 4.88 | 22.27 | 4.69 | 16.7 | 4.59 | 5.96 | 9.77 | 9.77 | 9.86 | 9.96 | 9.86
24.192 | 14.55 | 0.1 | 7.42 | 0 | 4.88 | 0.1 | 3.22 | 4.69 | 4.88 | 3.52 | 4.88 | 18.07 | 4.88 | 20.9 | 9.77 | 16.41 | 9.67 | 9.86 | 9.86 | 0.1 | 1.46 | 0 | 1.46 | 0.1 | 17.97 | 4.88 | 22.95 | 4.69 | 17.38 | 4.59 | 7.62 | 9.77 | 9.86 | 9.77 | 10.74 | 9.77
24.224 | 15.04 | 0.1 | 8.11 | 0.1 | 5.37 | 0 | 4.1 | 4.59 | 4.88 | 4.49 | 4.88 | 18.07 | 4.88 | 21.09 | 9.77 | 16.9 | 9.57 | 9.86 | 9.86 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 17.87 | 4.88 | 23.34 | 4.79 | 18.07 | 4.59 | 9.67 | 9.67 | 10.84 | 9.77 | 12.11 | 9.77
24.256 | 15.53 | 0 | 8.69 | 0.1 | 5.96 | 0 | 4.88 | 4.59 | 4.88 | 5.57 | 4.88 | 18.16 | 4.88 | 21.39 | 9.77 | 17.38 | 9.57 | 9.86 | 9.77 | 0.1 | 1.17 | 0.1 | 1.27 | 0.1 | 17.77 | 4.88 | 23.93 | 4.79 | 18.65 | 4.59 | 11.23 | 9.67 | 11.91 | 9.67 | 13.97 | 9.67
24.288 | 16.02 | 0.1 | 9.28 | 0.1 | 6.54 | 0 | 5.37 | 4.49 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 21.29 | 9.77 | 17.87 | 9.67 | 9.96 | 9.77 | 0.1 | 1.07 | 0.1 | 1.37 | 0 | 17.77 | 4.88 | 24.32 | 4.79 | 19.24 | 4.59 | 12.31 | 9.67 | 13.87 | 9.67 | 15.53 | 9.67
24.32 | 16.5 | 0.1 | 9.86 | 0.1 | 7.13 | 0.1 | 5.96 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 18.26 | 9.67 | 9.86 | 9.77 | 0.1 | 0.88 | 0 | 1.07 | 0.1 | 17.77 | 4.88 | 25 | 4.79 | 20.02 | 4.49 | 13.18 | 9.67 | 15.43 | 9.67 | 16.7 | 9.67
24.352 | 16.99 | 0.1 | 10.45 | 0 | 7.72 | 0.1 | 6.45 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 18.75 | 9.57 | 9.86 | 9.67 | 0.1 | 0.88 | 0 | 0.88 | 0.1 | 17.68 | 4.88 | 25.49 | 4.79 | 20.61 | 4.59 | 13.97 | 9.67 | 16.7 | 9.67 | 17.58 | 9.67
24.384 | 15.53 | 0.1 | 11.13 | 0 | 8.3 | 0.1 | 7.03 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.24 | 9.57 | 9.96 | 9.67 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 25.98 | 4.79 | 21.09 | 4.59 | 14.75 | 9.67 | 17.58 | 9.67 | 18.26 | 9.67
24.416 | 13.97 | 0.1 | 11.72 | 0.1 | 8.89 | 0 | 7.72 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.63 | 9.67 | 9.86 | 9.67 | 0 | 0.68 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 23.34 | 4.79 | 21.68 | 4.59 | 15.43 | 9.57 | 18.26 | 9.67 | 18.95 | 9.67
24.448 | 14.36 | 0 | 12.21 | 0.1 | 9.57 | 0 | 8.3 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.12 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.78 | 0 | 13.48 | 4.88 | 20.9 | 4.79 | 22.36 | 4.59 | 16.11 | 9.57 | 18.95 | 9.67 | 19.73 | 9.67
24.48 | 14.55 | 0.1 | 12.79 | 0.1 | 10.16 | 0.1 | 8.89 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.51 | 9.67 | 10.55 | 9.67 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 3.13 | 4.88 | 21.39 | 4.79 | 22.85 | 4.59 | 16.8 | 9.67 | 19.63 | 9.67 | 20.22 | 9.67
24.512 | 14.36 | 0.1 | 13.28 | 0 | 10.64 | 0.1 | 9.57 | 4.59 | 4.88 | 4.79 | 4.88 | 5.86 | 4.88 | 21.39 | 9.77 | 20.9 | 9.67 | 11.52 | 9.67 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 2.34 | 4.88 | 21.39 | 4.79 | 23.44 | 4.69 | 17.58 | 9.67 | 20.22 | 9.67 | 20.9 | 9.67
24.544 | 14.45 | 0.1 | 13.87 | 0 | 11.33 | 0.1 | 10.16 | 4.59 | 4.88 | 4.69 | 4.88 | 2.64 | 4.88 | 21.39 | 9.77 | 21.29 | 9.67 | 12.89 | 9.57 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.05 | 4.88 | 21.29 | 4.79 | 23.93 | 4.69 | 18.16 | 9.67 | 20.7 | 9.67 | 21.49 | 9.67
24.576 | 14.36 | 0 | 14.45 | 0.1 | 11.82 | 0 | 10.74 | 4.59 | 4.88 | 4.69 | 4.88 | 3.42 | 4.88 | 21 | 9.77 | 21.68 | 9.67 | 13.87 | 9.57 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 1.76 | 4.88 | 21.19 | 4.79 | 24.51 | 4.69 | 18.85 | 9.67 | 21.39 | 9.67 | 21.97 | 9.67
24.608 | 14.36 | 0.1 | 14.84 | 0.1 | 12.4 | 0 | 11.33 | 4.59 | 4.88 | 4.59 | 4.88 | 4.49 | 4.88 | 9.96 | 9.77 | 22.07 | 9.67 | 14.45 | 9.57 | 0.1 | 0.29 | 0.1 | 0.49 | 0 | 1.76 | 4.88 | 21.29 | 4.79 | 25 | 4.69 | 19.43 | 9.67 | 21.88 | 9.67 | 22.56 | 9.67
24.64 | 14.45 | 0.1 | 15.33 | 0.1 | 12.99 | 0.1 | 11.82 | 4.69 | 4.88 | 4.59 | 4.79 | 5.57 | 4.88 | 6.74 | 9.77 | 22.46 | 9.67 | 15.14 | 9.57 | 0.1 | 0.2 | 0.1 | 0.29 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 25.49 | 4.69 | 20.02 | 9.67 | 22.46 | 9.67 | 23.24 | 9.67
24.672 | 14.36 | 0.1 | 15.92 | 0 | 13.48 | 0.1 | 12.4 | 4.69 | 4.88 | 4.49 | 4.88 | 5.86 | 4.88 | 9.67 | 9.77 | 22.85 | 9.67 | 15.63 | 9.57 | 0.1 | 0.2 | 1.17 | 0.29 | 0.1 | 1.56 | 4.88 | 21.29 | 4.79 | 25.98 | 4.69 | 20.61 | 9.67 | 23.15 | 9.67 | 23.63 | 9.67
24.704 | 14.36 | 0.1 | 16.31 | 0 | 14.06 | 0 | 12.99 | 4.69 | 4.88 | 4.49 | 4.88 | 5.66 | 4.88 | 11.52 | 9.86 | 23.24 | 9.67 | 16.11 | 9.57 | 0.29 | 0.2 | 2.64 | 0.2 | 0.1 | 1.46 | 4.88 | 21.29 | 4.88 | 21.49 | 4.79 | 21.19 | 9.67 | 23.63 | 9.77 | 24.22 | 9.67
24.736 | 14.36 | 0 | 16.9 | 0.1 | 14.55 | 0 | 13.48 | 4.69 | 5.08 | 4.49 | 4.88 | 5.37 | 4.88 | 10.74 | 9.77 | 23.63 | 9.67 | 16.6 | 9.57 | 1.46 | 0.2 | 4.39 | 0.2 | 0.1 | 1.27 | 4.88 | 21.19 | 4.88 | 20.61 | 4.79 | 21.78 | 9.67 | 24.12 | 9.77 | 24.71 | 9.67
24.768 | 14.36 | 0 | 17.29 | 0.1 | 15.04 | 0 | 13.97 | 4.69 | 5.96 | 4.49 | 4.88 | 4.98 | 4.88 | 10.06 | 9.77 | 24.02 | 9.67 | 17.09 | 9.57 | 2.73 | 0.2 | 5.96 | 0.2 | 0 | 1.07 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 22.36 | 9.67 | 24.71 | 9.77 | 25.2 | 9.67
24.8 | 14.36 | 0.1 | 15.43 | 0.1 | 15.53 | 0.1 | 14.55 | 4.79 | 7.03 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 24.41 | 9.77 | 17.58 | 9.67 | 4.69 | 0.1 | 6.93 | 0.2 | 0.1 | 1.17 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 22.85 | 9.77 | 25.2 | 9.77 | 25.88 | 9.67
24.832 | 6.25 | 0.1 | 15.33 | 0 | 16.02 | 0.1 | 15.04 | 4.79 | 8.4 | 4.49 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 21.97 | 9.77 | 17.97 | 9.67 | 6.15 | 0.1 | 7.72 | 0.1 | 1.37 | 0.98 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 23.44 | 9.77 | 25.68 | 9.77 | 26.27 | 9.67
24.864 | 1.76 | 0.1 | 14.84 | 0 | 16.5 | 0 | 15.53 | 4.79 | 9.38 | 4.49 | 4.88 | 4.88 | 4.79 | 10.16 | 9.86 | 22.27 | 9.77 | 18.46 | 9.67 | 7.03 | 0.1 | 8.5 | 0.1 | 2.54 | 0.78 | 4.88 | 7.72 | 4.88 | 21.09 | 4.79 | 23.83 | 9.77 | 26.17 | 9.77 | 26.66 | 9.67
24.896 | 1.86 | 0 | 14.75 | 0.1 | 16.99 | 0 | 16.02 | 4.79 | 9.96 | 4.49 | 4.88 | 4.88 | 4.88 | 9.96 | 9.77 | 21.68 | 9.77 | 18.95 | 9.67 | 7.72 | 0.1 | 9.38 | 0.1 | 4.59 | 0.78 | 4.88 | 2.25 | 4.88 | 21.19 | 4.79 | 24.51 | 9.77 | 26.76 | 9.77 | 27.25 | 9.67
24.928 | 1.66 | 0.1 | 14.75 | 0.1 | 15.23 | 0 | 16.5 | 4.79 | 10.55 | 4.49 | 4.88 | 4.69 | 4.88 | 9.96 | 9.86 | 21.68 | 9.77 | 19.34 | 9.67 | 8.59 | 0.1 | 10.06 | 0.1 | 5.86 | 0.68 | 4.88 | 3.32 | 4.88 | 21.29 | 4.79 | 25.1 | 9.77 | 27.25 | 9.77 | 27.64 | 9.67
24.96 | 1.46 | 0.1 | 14.55 | 0.1 | 13.97 | 0.1 | 16.99 | 4.79 | 11.23 | 4.59 | 4.88 | 4.69 | 4.88 | 9.96 | 9.86 | 21.68 | 9.77 | 19.92 | 9.67 | 9.38 | 0 | 10.94 | 0.1 | 6.93 | 0.59 | 4.88 | 4.1 | 4.79 | 21.19 | 4.79 | 25.59 | 9.77 | 27.74 | 9.77 | 28.13 | 9.77
24.992 | 1.27 | 0.1 | 14.55 | 0 | 14.26 | 0.1 | 15.82 | 4.88 | 11.82 | 4.59 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.58 | 9.77 | 20.22 | 9.67 | 10.25 | 0 | 11.62 | 0.1 | 7.72 | 0.49 | 4.88 | 5.18 | 4.79 | 21.19 | 4.79 | 25.98 | 9.77 | 28.22 | 9.86 | 28.52 | 9.77
25.024 | 1.17 | 0.1 | 14.55 | 0 | 14.55 | 0.1 | 14.16 | 4.79 | 12.4 | 4.59 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.49 | 9.77 | 20.7 | 9.67 | 11.04 | 0.1 | 12.4 | 0.1 | 8.5 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 21.49 | 9.77 | 28.61 | 9.86 | 29.1 | 9.77
25.056 | 1.17 | 0 | 14.45 | 0.1 | 14.45 | 0 | 14.26 | 4.79 | 12.99 | 4.59 | 4.88 | 4.59 | 4.88 | 9.77 | 9.77 | 21.49 | 9.77 | 21 | 9.67 | 11.72 | 0 | 13.18 | 0.1 | 9.38 | 0.39 | 4.88 | 6.05 | 4.79 | 21.19 | 4.88 | 20.22 | 9.77 | 29.1 | 9.86 | 28.61 | 9.77
25.088 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 0 | 14.65 | 4.79 | 13.48 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 21.49 | 9.67 | 12.5 | 0.1 | 13.97 | 0.1 | 10.16 | 0.29 | 4.88 | 5.47 | 4.88 | 21.19 | 4.88 | 20.9 | 9.77 | 29.1 | 9.86 | 24.81 | 9.77
25.12 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 13.97 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.39 | 9.86 | 21.88 | 9.77 | 13.28 | 0 | 14.65 | 0.1 | 10.84 | 0.29 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 21.19 | 9.77 | 25.59 | 9.86 | 24.41 | 9.77
25.152 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 14.45 | 4.69 | 5.57 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.27 | 9.77 | 13.97 | 0 | 15.33 | 0 | 11.82 | 0.29 | 4.88 | 5.08 | 4.79 | 7.42 | 4.79 | 21.09 | 9.77 | 24.61 | 9.86 | 24.51 | 9.77
25.184 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 15.04 | 4.69 | 6.64 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.66 | 9.77 | 14.75 | 0.1 | 16.02 | 0.1 | 12.5 | 0.2 | 4.88 | 5.08 | 4.88 | 2.25 | 4.88 | 21.19 | 9.77 | 24.61 | 9.86 | 24.41 | 9.77
25.216 | 0.59 | 0 | 8.2 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 15.53 | 4.79 | 8.01 | 4.59 | 4.79 | 9.67 | 9.77 | 21.39 | 9.77 | 23.05 | 9.77 | 15.43 | 0.1 | 16.7 | 0.1 | 13.38 | 0.2 | 5.86 | 4.98 | 4.88 | 3.32 | 4.88 | 21.19 | 9.77 | 24.51 | 9.86 | 24.32 | 9.77
25.248 | 0.49 | 0 | 2.73 | 0.1 | 14.36 | 0 | 14.45 | 4.88 | 16.02 | 4.79 | 8.98 | 4.59 | 4.88 | 9.67 | 9.77 | 17.38 | 9.86 | 23.44 | 9.77 | 16.11 | 0.1 | 17.48 | 0.1 | 13.97 | 0.2 | 6.25 | 5.08 | 4.88 | 4.1 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.32 | 9.77
25.28 | 0.49 | 0.1 | 1.86 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 16.5 | 4.79 | 9.77 | 4.49 | 4.88 | 9.67 | 9.86 | 7.62 | 9.86 | 23.83 | 9.77 | 16.8 | 0 | 18.07 | 0.1 | 14.75 | 0.2 | 7.81 | 4.98 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.41 | 9.86 | 24.22 | 9.86
25.312 | 0.49 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 16.9 | 4.69 | 10.45 | 4.59 | 5.66 | 9.67 | 9.86 | 7.42 | 9.86 | 24.12 | 9.77 | 17.38 | 0 | 18.65 | 0 | 15.43 | 0.1 | 9.96 | 4.88 | 4.79 | 6.15 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77
25.344 | 0.29 | 0.1 | 1.46 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 17.48 | 4.79 | 11.04 | 4.59 | 6.15 | 9.67 | 9.86 | 10.55 | 9.86 | 22.95 | 9.77 | 18.07 | 0.1 | 19.34 | 0.1 | 16.21 | 0.1 | 11.43 | 4.88 | 4.88 | 5.86 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
25.376 | 0.29 | 0 | 1.46 | 0.1 | 6.15 | 0 | 14.36 | 4.88 | 17.87 | 4.79 | 11.62 | 4.59 | 7.13 | 9.57 | 9.77 | 11.43 | 9.86 | 21.09 | 9.77 | 18.75 | 0.1 | 19.92 | 0.1 | 16.8 | 0.1 | 12.6 | 4.69 | 4.79 | 5.47 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
25.408 | 0.2 | 0.1 | 1.27 | 0.1 | 1.76 | 0 | 14.36 | 4.88 | 18.46 | 4.79 | 12.21 | 4.59 | 8.5 | 9.57 | 9.77 | 10.64 | 9.86 | 21.49 | 9.77 | 19.24 | 0 | 20.51 | 0.1 | 17.38 | 0 | 13.28 | 4.69 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77
25.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 6.54 | 4.88 | 18.85 | 4.79 | 12.79 | 4.69 | 9.38 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 19.92 | 0 | 21.19 | 0.1 | 18.07 | 0.1 | 14.06 | 4.69 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
25.472 | 0.2 | 0.1 | 1.07 | 0 | 1.76 | 0.1 | 1.86 | 4.88 | 19.24 | 4.79 | 13.28 | 4.69 | 9.96 | 9.67 | 9.86 | 9.86 | 9.86 | 21.39 | 9.77 | 20.61 | 0 | 21.68 | 0.1 | 18.75 | 0.1 | 14.94 | 4.69 | 4.88 | 5.18 | 4.79 | 5.57 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
25.504 | 0.1 | 0.1 | 0.88 | 0 | 1.46 | 0 | 1.86 | 4.88 | 19.73 | 4.79 | 13.87 | 4.69 | 10.64 | 9.67 | 9.86 | 10.25 | 9.86 | 21.49 | 9.86 | 21.09 | 0.1 | 22.36 | 0.1 | 19.43 | 0.1 | 15.63 | 4.69 | 4.88 | 5.08 | 4.88 | 2.44 | 9.86 | 24.12 | 9.86 | 12.79 | 9.86
25.536 | 0.1 | 0 | 0.78 | 0.1 | 1.37 | 0 | 1.76 | 4.88 | 20.12 | 4.79 | 14.36 | 4.69 | 11.23 | 9.67 | 9.77 | 10.16 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 20.02 | 0.1 | 20.12 | 0 | 16.41 | 4.59 | 4.88 | 4.98 | 4.88 | 3.52 | 9.86 | 13.18 | 9.86 | 5.86 | 9.86
25.568 | 0 | 0.1 | 0.78 | 0.1 | 1.17 | 0 | 1.46 | 4.88 | 20.61 | 4.79 | 14.84 | 4.69 | 11.82 | 9.67 | 10.64 | 10.16 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 17.09 | 0.1 | 20.7 | 0 | 16.99 | 4.59 | 4.88 | 4.88 | 4.88 | 4.2 | 9.77 | 5.96 | 9.86 | 8.01 | 9.86
25.6 | 0.1 | 0.88 | 0.68 | 0.1 | 1.17 | 0.1 | 1.37 | 4.88 | 20.9 | 4.88 | 15.43 | 4.79 | 12.4 | 9.67 | 11.04 | 9.86 | 9.86 | 21.39 | 9.86 | 17.58 | 0 | 17.48 | 0.1 | 21.19 | 0.1 | 17.58 | 4.59 | 4.88 | 4.88 | 4.79 | 5.47 | 9.86 | 8.01 | 9.86 | 11.91 | 9.86
25.632 | 0.1 | 1.86 | 0.59 | 0 | 0.98 | 0.1 | 1.27 | 4.88 | 17.58 | 4.79 | 15.82 | 4.79 | 12.99 | 9.67 | 12.01 | 9.96 | 9.96 | 21.39 | 9.86 | 16.6 | 0.1 | 17.68 | 0.1 | 21.78 | 0.1 | 18.36 | 4.59 | 4.79 | 4.88 | 4.79 | 6.15 | 9.77 | 11.72 | 9.86 | 11.62 | 9.86
25.664 | 0.1 | 3.13 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 4.88 | 17.77 | 4.79 | 16.31 | 4.79 | 13.48 | 9.67 | 13.38 | 9.96 | 9.86 | 21.39 | 9.86 | 17.29 | 0.1 | 17.68 | 0.1 | 22.36 | 0.1 | 18.85 | 4.59 | 4.88 | 4.79 | 4.88 | 6.05 | 9.86 | 11.62 | 9.86 | 10.55 | 9.86
25.696 | 0 | 4.1 | 0.49 | 0.1 | 0.78 | 0 | 0.88 | 4.88 | 17.97 | 4.88 | 16.8 | 4.79 | 13.97 | 9.67 | 14.16 | 9.86 | 9.86 | 21.39 | 9.86 | 17.48 | 0 | 17.68 | 0.1 | 18.26 | 0 | 19.53 | 4.49 | 5.86 | 4.69 | 4.88 | 5.47 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86
25.728 | 0 | 4.79 | 0.29 | 0.1 | 0.68 | 0 | 0.88 | 4.88 | 18.16 | 4.88 | 17.29 | 4.79 | 14.55 | 9.67 | 14.75 | 9.77 | 9.86 | 21.39 | 9.86 | 17.48 | 0 | 17.68 | 0.1 | 17.09 | 0.1 | 20.22 | 4.49 | 7.03 | 4.69 | 4.88 | 5.18 | 9.77 | 9.96 | 9.86 | 10.25 | 9.86
25.76 | 0.1 | 5.27 | 0.39 | 0.1 | 0.68 | 0 | 0.78 | 4.88 | 18.16 | 4.88 | 17.77 | 4.79 | 15.04 | 9.67 | 15.33 | 9.77 | 9.86 | 21.39 | 9.86 | 17.48 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 20.8 | 4.49 | 9.08 | 4.69 | 4.79 | 5.18 | 9.77 | 10.25 | 9.86 | 10.45 | 9.86
25.792 | 0.1 | 5.86 | 0.29 | 0 | 0.59 | 0.1 | 0.68 | 4.88 | 18.07 | 4.79 | 18.16 | 4.79 | 15.63 | 9.77 | 15.82 | 9.77 | 9.86 | 11.91 | 9.86 | 17.58 | 0 | 17.68 | 0 | 17.77 | 0.1 | 21.39 | 4.49 | 10.64 | 4.59 | 4.79 | 5.08 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86
25.824 | 0.1 | 6.45 | 0.29 | 0 | 0.49 | 0.1 | 0.59 | 4.88 | 18.16 | 4.79 | 18.65 | 4.79 | 16.02 | 9.77 | 16.41 | 9.67 | 9.86 | 6.84 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 21.97 | 4.59 | 11.91 | 4.59 | 4.88 | 5.18 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86
25.856 | 0 | 7.03 | 0.2 | 0.1 | 0.49 | 0 | 0.59 | 4.88 | 18.16 | 4.88 | 19.14 | 4.88 | 16.6 | 9.67 | 16.8 | 9.67 | 9.86 | 8.59 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0 | 22.56 | 4.59 | 12.89 | 4.59 | 4.88 | 5.08 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
25.888 | 0 | 7.62 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.88 | 18.16 | 4.88 | 19.53 | 4.88 | 17.09 | 9.77 | 17.29 | 9.67 | 9.86 | 11.23 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 0 | 23.15 | 4.59 | 13.77 | 4.59 | 4.88 | 4.98 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
25.92 | 0.1 | 8.3 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.16 | 4.88 | 19.92 | 4.88 | 17.48 | 9.77 | 17.77 | 9.67 | 9.96 | 11.13 | 9.86 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 0.1 | 23.63 | 4.59 | 14.55 | 4.59 | 4.88 | 4.98 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
25.952 | 0.1 | 8.89 | 0.1 | 0 | 0.29 | 0 | 0.39 | 4.88 | 18.07 | 4.79 | 20.31 | 4.79 | 17.87 | 9.77 | 18.26 | 9.67 | 9.86 | 10.35 | 9.86 | 17.58 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 24.12 | 4.69 | 15.14 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
25.984 | 0.1 | 9.47 | 0.1 | 0.68 | 0.2 | 0.1 | 0.29 | 4.88 | 18.16 | 4.79 | 20.8 | 4.88 | 18.46 | 9.77 | 18.65 | 9.67 | 9.86 | 9.77 | 9.86 | 17.58 | 0.1 | 7.03 | 0.1 | 17.68 | 0.1 | 24.71 | 4.69 | 15.92 | 4.49 | 4.88 | 4.79 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86
26.016 | 0 | 10.06 | 0.1 | 1.37 | 0.2 | 0.2 | 0.2 | 4.88 | 18.07 | 4.88 | 18.95 | 4.88 | 18.85 | 9.77 | 19.14 | 9.57 | 9.86 | 9.96 | 9.86 | 17.58 | 0 | 1.56 | 0.1 | 17.68 | 0 | 25.2 | 4.69 | 16.6 | 4.49 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
26.048 | 0 | 10.55 | 0 | 2.44 | 0.2 | 1.07 | 0.2 | 4.88 | 17.87 | 4.88 | 17.68 | 4.88 | 19.24 | 9.77 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 5.76 | 0 | 2.25 | 0.1 | 17.68 | 0 | 25.59 | 4.69 | 17.38 | 4.59 | 4.88 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
26.08 | 0.1 | 11.23 | 0.1 | 3.81 | 0.1 | 2.15 | 0.2 | 4.88 | 4.59 | 4.88 | 17.77 | 4.88 | 19.73 | 9.77 | 20.02 | 9.57 | 9.86 | 9.96 | 9.96 | 1.86 | 0 | 2.25 | 0.1 | 17.68 | 0.1 | 26.17 | 4.79 | 17.97 | 4.59 | 4.88 | 4.69 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
26.112 | 0.1 | 11.82 | 0.1 | 4.49 | 0 | 3.52 | 0.1 | 4.88 | 2.83 | 4.79 | 18.16 | 4.88 | 20.12 | 9.77 | 20.41 | 9.67 | 9.86 | 9.86 | 9.96 | 2.25 | 0 | 1.66 | 0.1 | 17.68 | 0.1 | 21.58 | 4.79 | 18.65 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
26.144 | 0.1 | 12.4 | 0.1 | 5.08 | 0.1 | 4.39 | 0.1 | 4.88 | 3.71 | 4.79 | 18.07 | 4.88 | 20.51 | 9.77 | 20.9 | 9.67 | 10.45 | 9.96 | 9.96 | 1.95 | 0.1 | 1.66 | 0.1 | 6.93 | 0.1 | 20.61 | 4.79 | 19.24 | 4.59 | 5.96 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
26.176 | 0 | 12.99 | 0.1 | 5.57 | 0.1 | 4.98 | 0.1 | 4.88 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 9.77 | 21.19 | 9.67 | 11.52 | 9.86 | 9.86 | 1.76 | 0 | 1.56 | 0.1 | 1.76 | 0 | 21.29 | 4.79 | 20.02 | 4.59 | 6.93 | 4.59 | 10.64 | 9.77 | 9.86 | 9.77 | 9.86
26.208 | 0 | 13.48 | 0.1 | 6.15 | 0.1 | 5.37 | 0.1 | 4.88 | 5.57 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 21.58 | 9.67 | 12.79 | 9.86 | 9.86 | 1.76 | 0.1 | 1.46 | 0.1 | 2.25 | 0.1 | 21.19 | 4.79 | 20.61 | 4.59 | 8.79 | 4.59 | 11.43 | 9.67 | 10.84 | 9.77 | 9.86
26.24 | 0.1 | 13.97 | 0.1 | 6.74 | 0 | 5.96 | 0.1 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 18.75 | 9.77 | 22.07 | 9.67 | 13.77 | 9.77 | 9.96 | 1.46 | 0 | 1.37 | 0.1 | 2.15 | 0.1 | 21.29 | 4.79 | 21.19 | 4.59 | 10.55 | 4.59 | 13.09 | 9.77 | 11.91 | 9.67 | 10.74
26.272 | 0.1 | 14.55 | 0.1 | 7.42 | 0.1 | 6.54 | 0.1 | 4.88 | 5.47 | 4.79 | 18.07 | 4.88 | 18.26 | 9.77 | 22.36 | 9.67 | 14.45 | 9.77 | 9.96 | 1.37 | 0 | 1.07 | 0 | 1.86 | 0.1 | 21.29 | 4.79 | 21.78 | 4.69 | 11.82 | 4.49 | 15.14 | 9.67 | 13.57 | 9.67 | 12.11
26.304 | 0.1 | 15.04 | 0.1 | 8.01 | 0.1 | 7.13 | 0.1 | 4.88 | 5.18 | 4.79 | 18.07 | 4.88 | 18.26 | 9.86 | 22.85 | 9.67 | 15.04 | 9.77 | 9.96 | 1.17 | 0.1 | 1.07 | 0.1 | 1.76 | 0.1 | 21.39 | 4.79 | 22.27 | 4.69 | 12.79 | 4.49 | 16.41 | 9.67 | 15.33 | 9.67 | 13.97
26.336 | 0 | 15.63 | 0.1 | 8.5 | 0.1 | 7.81 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 9.77 | 23.24 | 9.67 | 15.63 | 9.67 | 9.86 | 1.07 | 0.1 | 0.98 | 0.1 | 1.56 | 0 | 21.39 | 4.79 | 22.75 | 4.69 | 13.67 | 4.59 | 17.38 | 9.67 | 16.5 | 9.67 | 15.63
26.368 | 0 | 16.02 | 0 | 9.18 | 0.1 | 8.3 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 23.54 | 9.67 | 16.11 | 9.67 | 9.86 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 0 | 21.29 | 4.79 | 23.34 | 4.79 | 14.45 | 4.59 | 17.97 | 9.67 | 17.38 | 9.67 | 16.7
26.4 | 0.1 | 16.5 | 0.1 | 9.77 | 0 | 8.98 | 0.1 | 5.08 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.86 | 23.93 | 9.67 | 16.6 | 9.67 | 9.96 | 1.07 | 0 | 0.78 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 23.93 | 4.79 | 15.14 | 4.59 | 18.56 | 9.67 | 18.16 | 9.67 | 17.58
26.432 | 0.1 | 16.99 | 0.1 | 10.45 | 0 | 9.67 | 0.1 | 5.96 | 4.88 | 4.79 | 17.97 | 4.88 | 18.07 | 9.86 | 24.32 | 9.67 | 17.09 | 9.67 | 9.96 | 0.88 | 0 | 0.68 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.32 | 4.79 | 15.92 | 4.59 | 19.24 | 9.57 | 18.95 | 9.67 | 18.16
26.464 | 0.1 | 15.92 | 0.1 | 10.94 | 0.1 | 10.25 | 0.1 | 6.93 | 4.79 | 4.79 | 7.23 | 4.88 | 18.07 | 9.86 | 21.29 | 9.67 | 17.58 | 9.67 | 10.06 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 24.9 | 4.69 | 16.6 | 4.59 | 20.02 | 9.57 | 19.53 | 9.67 | 18.95
26.496 | 0 | 14.06 | 0.1 | 11.52 | 0.1 | 10.84 | 0.1 | 8.5 | 4.79 | 4.88 | 2.64 | 4.88 | 18.07 | 9.77 | 22.07 | 9.67 | 18.07 | 9.67 | 10.84 | 0.68 | 0 | 0.59 | 0.1 | 0.88 | 0 | 21.29 | 4.79 | 25.49 | 4.79 | 17.19 | 4.69 | 20.61 | 9.57 | 20.22 | 9.67 | 19.73
26.528 | 0 | 14.45 | 0 | 12.11 | 0.1 | 11.33 | 0.1 | 9.28 | 4.69 | 4.88 | 3.32 | 4.88 | 18.07 | 9.77 | 21.49 | 9.67 | 18.46 | 9.67 | 11.82 | 0.68 | 0 | 0.59 | 0.1 | 0.88 | 0 | 20.9 | 4.88 | 25.88 | 4.79 | 17.87 | 4.69 | 21.19 | 9.67 | 20.8 | 9.67 | 20.31
26.56 | 0.1 | 14.55 | 0.1 | 12.7 | 0.1 | 12.01 | 0.1 | 9.96 | 4.59 | 4.88 | 4.39 | 4.88 | 17.97 | 9.86 | 21.49 | 9.67 | 18.95 | 9.57 | 13.18 | 0.59 | 0 | 0.49 | 0.1 | 0.78 | 0.1 | 4.79 | 4.88 | 25.1 | 4.79 | 18.65 | 4.69 | 21.88 | 9.67 | 21.39 | 9.67 | 20.9
26.592 | 0.1 | 14.45 | 0.1 | 13.28 | 0 | 12.5 | 0.1 | 10.64 | 4.59 | 4.79 | 5.57 | 4.88 | 17.97 | 9.86 | 21.58 | 9.67 | 19.34 | 9.57 | 14.06 | 0.49 | 0 | 0.39 | 0 | 0.68 | 0.1 | 2.93 | 4.88 | 22.07 | 4.79 | 19.24 | 4.69 | 22.36 | 9.57 | 21.88 | 9.67 | 21.49
26.624 | 0.1 | 14.45 | 0.1 | 13.77 | 0.1 | 13.09 | 0.1 | 11.23 | 4.59 | 4.79 | 5.86 | 4.88 | 12.89 | 9.86 | 21.49 | 9.77 | 19.82 | 9.67 | 14.75 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 0.1 | 3.52 | 4.88 | 21.88 | 4.79 | 19.82 | 4.69 | 22.85 | 9.67 | 22.56 | 9.67 | 21.97
26.656 | 0 | 14.45 | 0.1 | 14.26 | 0.1 | 13.67 | 0.1 | 11.82 | 4.59 | 4.88 | 5.57 | 4.88 | 4 | 9.77 | 21.49 | 9.77 | 20.22 | 9.67 | 15.33 | 0.39 | 1.07 | 0.29 | 0.1 | 0.59 | 0 | 4.2 | 4.79 | 21.58 | 4.79 | 20.51 | 4.79 | 23.44 | 9.67 | 23.05 | 9.67 | 22.56
26.688 | 0 | 14.45 | 0.1 | 14.75 | 0.1 | 14.06 | 0 | 12.4 | 4.59 | 4.88 | 5.18 | 4.88 | 3.13 | 9.77 | 21.49 | 9.77 | 20.61 | 9.67 | 15.82 | 0.29 | 2.34 | 0.29 | 0.88 | 0.49 | 0 | 5.47 | 4.79 | 21.58 | 4.79 | 21.09 | 4.79 | 24.02 | 9.67 | 23.44 | 9.67 | 23.15
26.72 | 0.1 | 14.45 | 0.1 | 15.23 | 0 | 14.55 | 0.1 | 12.99 | 4.49 | 4.88 | 5.08 | 4.88 | 3.81 | 9.77 | 21.39 | 9.77 | 21 | 9.57 | 16.21 | 0.29 | 4.3 | 0.2 | 1.86 | 0.49 | 0.1 | 6.25 | 4.88 | 21.39 | 4.79 | 21.68 | 4.79 | 24.51 | 9.67 | 24.02 | 9.67 | 23.54
26.752 | 0.1 | 14.45 | 0.1 | 15.82 | 0 | 15.14 | 0.1 | 13.48 | 4.49 | 4.79 | 4.88 | 4.88 | 4.88 | 9.86 | 21.49 | 9.77 | 21.39 | 9.67 | 16.7 | 0.29 | 5.57 | 0.2 | 3.61 | 0.39 | 0.1 | 5.96 | 4.88 | 21.39 | 4.79 | 22.17 | 4.79 | 25 | 9.67 | 24.71 | 9.67 | 24.12
26.784 | 0 | 14.36 | 0.1 | 16.21 | 0.1 | 15.63 | 0.1 | 13.97 | 4.49 | 4.88 | 4.98 | 4.88 | 5.86 | 9.77 | 21.39 | 9.77 | 21.88 | 9.67 | 17.29 | 0.2 | 6.74 | 0.2 | 5.37 | 0.29 | 0.1 | 5.47 | 4.88 | 21.29 | 4.79 | 22.85 | 4.79 | 25.49 | 9.67 | 25.2 | 9.67 | 24.81
26.816 | 0 | 14.36 | 0.1 | 16.7 | 0.1 | 16.11 | 0.1 | 14.55 | 4.49 | 5.66 | 4.79 | 4.88 | 5.76 | 9.77 | 21.39 | 9.77 | 22.27 | 9.67 | 17.68 | 0.2 | 7.62 | 0.2 | 6.54 | 0.29 | 0 | 5.18 | 4.79 | 21.29 | 4.79 | 23.34 | 4.79 | 25.98 | 9.67 | 25.68 | 9.67 | 25.2
26.848 | 0 | 14.36 | 0.1 | 17.19 | 0.1 | 16.6 | 0 | 14.94 | 4.49 | 6.64 | 4.79 | 4.88 | 5.37 | 9.77 | 21.39 | 9.77 | 22.66 | 9.67 | 18.16 | 0.2 | 8.4 | 0.2 | 7.42 | 0.2 | 0 | 5.18 | 4.79 | 21.29 | 4.79 | 23.83 | 4.79 | 26.47 | 9.67 | 25.98 | 9.77 | 25.59
26.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0 | 17.09 | 0.1 | 15.53 | 4.49 | 8.01 | 4.69 | 4.88 | 4.98 | 9.86 | 19.14 | 9.77 | 23.05 | 9.67 | 18.56 | 0.2 | 9.18 | 0.1 | 8.2 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.79 | 24.32 | 4.79 | 27.05 | 9.67 | 26.56 | 9.67 | 26.17
26.912 | 0.1 | 6.64 | 0.1 | 15.33 | 0.1 | 13.67 | 0.1 | 16.02 | 4.49 | 9.08 | 4.69 | 4.79 | 4.79 | 9.86 | 8.59 | 9.77 | 23.44 | 9.67 | 19.04 | 0.1 | 9.86 | 0.1 | 8.98 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.79 | 24.9 | 4.79 | 27.44 | 9.67 | 27.05 | 9.67 | 26.66
26.944 | 0.1 | 1.86 | 0.1 | 14.65 | 0.1 | 13.97 | 0.1 | 16.5 | 4.59 | 9.77 | 4.69 | 4.88 | 4.88 | 9.77 | 7.32 | 9.77 | 23.83 | 9.67 | 19.43 | 0.1 | 10.74 | 0.1 | 9.77 | 0.2 | 0 | 5.08 | 4.88 | 21.19 | 4.79 | 25.29 | 4.79 | 28.03 | 9.67 | 27.64 | 9.77 | 27.15
26.976 | 0 | 1.86 | 0.1 | 14.65 | 0.1 | 14.26 | 0.1 | 16.99 | 4.59 | 10.45 | 4.59 | 4.88 | 4.88 | 9.77 | 10.45 | 9.77 | 24.22 | 9.67 | 19.92 | 0.1 | 11.52 | 0.1 | 10.55 | 0.1 | 0 | 4.98 | 4.79 | 21.19 | 4.79 | 25.88 | 4.79 | 28.32 | 9.77 | 28.13 | 9.77 | 27.74
27.008 | 0 | 1.76 | 0 | 14.65 | 0.1 | 14.36 | 0.1 | 17.48 | 4.59 | 11.04 | 4.59 | 4.88 | 4.79 | 9.77 | 11.43 | 9.77 | 23.63 | 9.67 | 20.31 | 0.1 | 12.21 | 0.1 | 11.33 | 0.1 | 0 | 4.88 | 4.79 | 9.38 | 4.88 | 26.37 | 4.88 | 28.91 | 9.77 | 28.52 | 9.77 | 28.22
27.04 | 0.1 | 1.46 | 0.1 | 14.55 | 0 | 14.36 | 0.1 | 17.87 | 4.69 | 11.62 | 4.59 | 4.88 | 4.79 | 9.86 | 10.55 | 9.77 | 21.58 | 9.67 | 20.8 | 0.1 | 13.09 | 0.1 | 12.11 | 0.1 | 0.1 | 4.88 | 4.88 | 2.25 | 4.88 | 22.36 | 4.79 | 29.4 | 9.77 | 29.01 | 9.77 | 28.52
27.072 | 0.1 | 1.37 | 0.1 | 14.55 | 0.1 | 14.45 | 0.1 | 18.46 | 4.69 | 12.21 | 4.49 | 4.79 | 4.69 | 9.77 | 9.96 | 9.86 | 21.68 | 9.67 | 21.09 | 0.1 | 13.77 | 0.1 | 12.79 | 0.1 | 0.1 | 4.88 | 4.88 | 3.32 | 4.79 | 22.07 | 4.88 | 25.29 | 9.77 | 29.59 | 9.77 | 29.01
27.104 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.85 | 4.69 | 12.7 | 4.49 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 21.58 | 0.1 | 14.55 | 0.1 | 13.57 | 0.1 | 0.1 | 4.69 | 4.88 | 3.81 | 4.88 | 21.78 | 4.88 | 24.71 | 9.77 | 26.37 | 9.77 | 28.22
27.136 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0 | 19.24 | 4.69 | 13.28 | 4.49 | 4.88 | 4.59 | 9.77 | 10.16 | 9.86 | 21.49 | 9.77 | 21.88 | 0.1 | 15.23 | 0.1 | 14.36 | 0 | 0.78 | 4.69 | 4.79 | 4.88 | 4.79 | 21.58 | 4.88 | 24.81 | 9.77 | 25.2 | 9.77 | 23.83
27.168 | 0 | 0.88 | 0 | 14.45 | 0.1 | 14.36 | 0 | 19.73 | 4.79 | 13.87 | 4.49 | 4.88 | 4.59 | 9.77 | 10.35 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 15.82 | 0.1 | 15.14 | 0 | 1.66 | 4.59 | 4.79 | 6.15 | 4.88 | 21.58 | 4.88 | 24.41 | 9.77 | 24.9 | 9.77 | 24.51
27.2 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 20.12 | 4.79 | 14.36 | 4.49 | 5.66 | 4.59 | 9.86 | 10.16 | 9.77 | 21.49 | 9.67 | 22.66 | 0.1 | 16.5 | 0.1 | 15.63 | 0.1 | 3.22 | 4.59 | 4.88 | 6.15 | 4.88 | 21.39 | 4.88 | 24.51 | 9.77 | 24.61 | 9.77 | 24.41
27.232 | 0.1 | 0.88 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 20.51 | 4.79 | 14.84 | 4.49 | 6.74 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 23.15 | 0.1 | 17.29 | 0.1 | 16.41 | 0.1 | 5.18 | 4.59 | 4.88 | 5.57 | 4.79 | 21.29 | 4.79 | 24.41 | 9.77 | 24.61 | 9.86 | 24.41
27.264 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 21 | 4.79 | 15.33 | 4.49 | 8.11 | 4.49 | 9.96 | 9.96 | 9.86 | 21.49 | 9.77 | 23.54 | 0.1 | 17.97 | 0.1 | 17.19 | 0.1 | 6.45 | 4.59 | 4.88 | 5.27 | 4.79 | 21.39 | 4.88 | 24.32 | 9.77 | 24.51 | 9.86 | 24.22
27.296 | 0 | 0.59 | 0.1 | 11.62 | 0.1 | 14.36 | 0.1 | 17.77 | 4.79 | 15.92 | 4.59 | 8.98 | 4.49 | 10.74 | 9.77 | 9.86 | 21.39 | 9.77 | 23.83 | 0.1 | 18.65 | 0.1 | 17.87 | 0.1 | 7.32 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 4.88 | 24.41 | 9.77 | 24.32 | 9.86 | 24.22
27.328 | 0 | 0.59 | 0.1 | 3.32 | 0.1 | 14.36 | 0 | 17.77 | 4.79 | 16.31 | 4.59 | 9.77 | 4.49 | 11.72 | 9.77 | 9.86 | 21.39 | 9.77 | 24.22 | 0.1 | 19.24 | 0.1 | 18.46 | 0 | 8.01 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.41 | 9.86 | 24.22
27.36 | 0 | 0.49 | 0.1 | 1.95 | 0 | 4.98 | 0.1 | 17.97 | 4.79 | 16.8 | 4.59 | 10.45 | 4.49 | 13.09 | 9.77 | 9.86 | 21.39 | 9.77 | 21 | 0.1 | 19.92 | 0 | 19.14 | 0.1 | 8.79 | 4.59 | 4.88 | 5.18 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22
27.392 | 0.1 | 0.49 | 0.1 | 1.76 | 0 | 1.76 | 0.1 | 18.16 | 4.79 | 17.29 | 4.59 | 11.13 | 4.49 | 13.97 | 9.77 | 9.86 | 21.39 | 9.77 | 20.9 | 0.1 | 20.51 | 0.1 | 19.82 | 0.1 | 9.67 | 4.59 | 4.79 | 4.98 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22
27.424 | 0 | 0.39 | 0.1 | 1.46 | 0.1 | 2.15 | 0.1 | 18.16 | 4.79 | 17.77 | 4.69 | 11.72 | 4.49 | 14.55 | 9.77 | 9.86 | 21.39 | 9.86 | 21.19 | 0.1 | 21 | 0.1 | 20.41 | 0.1 | 10.45 | 4.59 | 4.79 | 5.08 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.12
27.456 | 0 | 0.29 | 0.1 | 1.56 | 0.1 | 1.76 | 0.1 | 18.16 | 4.88 | 18.16 | 4.69 | 12.21 | 4.49 | 15.23 | 9.67 | 9.86 | 11.91 | 9.77 | 21.39 | 0.1 | 21.58 | 0.1 | 21 | 0.1 | 11.33 | 4.59 | 4.79 | 4.98 | 4.79 | 9.96 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
27.488 | 0.1 | 0.29 | 0 | 1.27 | 0.1 | 1.27 | 0 | 18.16 | 4.88 | 18.65 | 4.69 | 12.79 | 4.59 | 15.63 | 9.67 | 9.86 | 7.03 | 9.86 | 21.39 | 0.1 | 22.27 | 0.1 | 21.58 | 0 | 12.01 | 4.59 | 4.79 | 4.88 | 4.79 | 2.73 | 4.88 | 19.53 | 9.77 | 24.22 | 9.86 | 24.12
27.52 | 1.07 | 0.29 | 0.1 | 1.17 | 0 | 1.27 | 0.1 | 18.16 | 4.79 | 19.14 | 4.69 | 13.38 | 4.59 | 16.21 | 9.67 | 9.86 | 8.3 | 9.77 | 21.39 | 0.1 | 20.41 | 0.1 | 22.17 | 0.1 | 12.79 | 4.59 | 4.79 | 4.79 | 4.79 | 3.22 | 4.88 | 7.62 | 9.77 | 15.43 | 9.86 | 24.12
27.552 | 1.86 | 0.2 | 0.1 | 0.98 | 0 | 1.17 | 0.1 | 18.16 | 4.79 | 19.53 | 4.69 | 13.87 | 4.59 | 16.7 | 9.67 | 9.86 | 11.23 | 9.86 | 21.49 | 0.1 | 17.97 | 0.1 | 22.66 | 0.1 | 13.57 | 4.59 | 5.76 | 4.69 | 4.79 | 3.91 | 4.79 | 6.74 | 9.77 | 6.64 | 9.77 | 12.7
27.584 | 3.13 | 0.2 | 0.1 | 0.98 | 0.1 | 1.17 | 0.1 | 18.16 | 4.88 | 20.02 | 4.79 | 14.36 | 4.59 | 17.19 | 9.67 | 9.86 | 11.23 | 9.86 | 21.49 | 0.1 | 17.87 | 0.1 | 19.73 | 0.1 | 14.26 | 4.69 | 7.03 | 4.79 | 4.79 | 4.79 | 4.88 | 10.64 | 9.77 | 7.32 | 9.86 | 5.86
27.616 | 4.1 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 18.16 | 4.88 | 20.51 | 4.79 | 14.84 | 4.69 | 17.68 | 9.57 | 9.86 | 10.35 | 9.86 | 21.39 | 0.1 | 17.87 | 0.1 | 18.75 | 0.1 | 15.04 | 4.69 | 9.08 | 4.69 | 4.79 | 6.15 | 4.88 | 12.11 | 9.77 | 11.33 | 9.86 | 8.4
27.648 | 4.69 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 0 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 18.07 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 18.26 | 0 | 15.72 | 4.69 | 10.74 | 4.69 | 4.79 | 6.15 | 4.88 | 10.64 | 9.77 | 11.91 | 9.86 | 11.72
27.68 | 5.27 | 0.1 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 18.07 | 4.88 | 17.77 | 4.79 | 15.92 | 4.69 | 18.56 | 9.67 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 17.58 | 0 | 18.07 | 0.1 | 16.41 | 4.69 | 12.01 | 4.59 | 4.79 | 5.66 | 4.79 | 9.96 | 9.77 | 10.55 | 9.86 | 11.62
27.712 | 5.86 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 17.97 | 4.88 | 17.58 | 4.79 | 16.31 | 4.69 | 19.04 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 17.09 | 4.79 | 12.89 | 4.59 | 4.79 | 5.27 | 4.79 | 10.16 | 9.77 | 10.16 | 9.77 | 10.35
27.744 | 6.45 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 5.37 | 4.88 | 17.87 | 4.88 | 16.8 | 4.69 | 19.43 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 17.68 | 4.79 | 13.77 | 4.59 | 4.79 | 5.18 | 4.79 | 10.45 | 9.77 | 10.25 | 9.86 | 9.86
27.776 | 7.03 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 0 | 2.93 | 4.88 | 18.16 | 4.88 | 17.29 | 4.79 | 19.92 | 9.57 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.87 | 0 | 18.36 | 4.79 | 14.55 | 4.59 | 4.79 | 5.27 | 4.88 | 10.45 | 9.77 | 10.45 | 9.86 | 10.35
27.808 | 7.62 | 0.1 | 0.1 | 0.39 | 0.1 | 0.39 | 0 | 3.42 | 4.88 | 18.07 | 4.88 | 17.77 | 4.79 | 20.22 | 9.67 | 9.86 | 9.86 | 9.86 | 10.64 | 0.1 | 17.58 | 0.1 | 17.87 | 0 | 19.14 | 4.79 | 15.23 | 4.59 | 4.79 | 5.18 | 4.88 | 10.25 | 9.77 | 10.45 | 9.86 | 10.35
27.84 | 8.3 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 0.1 | 4.59 | 4.88 | 18.07 | 4.79 | 18.16 | 4.79 | 20.7 | 9.67 | 9.86 | 9.86 | 9.86 | 6.54 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 19.73 | 4.79 | 16.02 | 4.49 | 5.76 | 5.08 | 4.88 | 10.16 | 9.77 | 10.25 | 9.86 | 10.35
27.872 | 8.79 | 0.1 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 5.57 | 4.88 | 18.07 | 4.79 | 18.65 | 4.79 | 21.09 | 9.67 | 9.86 | 9.86 | 9.86 | 9.38 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 20.22 | 4.79 | 16.7 | 4.59 | 7.13 | 4.98 | 4.79 | 10.16 | 9.77 | 10.16 | 9.77 | 10.25
27.904 | 9.47 | 0.1 | 0.68 | 0.2 | 0.1 | 0.29 | 0.1 | 5.86 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 21.58 | 9.67 | 9.86 | 9.77 | 9.96 | 11.62 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 20.9 | 4.79 | 17.38 | 4.59 | 9.38 | 4.98 | 4.88 | 10.16 | 9.77 | 10.25 | 9.86 | 10.06
27.936 | 10.06 | 0 | 1.37 | 0.2 | 0.1 | 0.2 | 0.1 | 5.47 | 4.88 | 18.07 | 4.88 | 19.43 | 4.79 | 21.97 | 9.67 | 10.64 | 9.77 | 9.86 | 10.94 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.58 | 4.79 | 18.07 | 4.59 | 10.74 | 4.88 | 4.88 | 10.16 | 9.77 | 10.16 | 9.86 | 10.16
27.968 | 10.64 | 0.1 | 2.54 | 0.2 | 0.1 | 0.2 | 0 | 5.18 | 4.88 | 18.07 | 4.88 | 19.92 | 4.79 | 22.36 | 9.67 | 11.13 | 9.67 | 9.86 | 10.16 | 0.1 | 7.81 | 0.1 | 17.68 | 0 | 22.07 | 4.79 | 18.56 | 4.59 | 12.01 | 4.88 | 4.88 | 9.96 | 9.77 | 10.06 | 9.86 | 10.06
28 | 11.23 | 0.1 | 3.81 | 0.1 | 0 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.79 | 20.31 | 4.79 | 22.75 | 9.67 | 12.21 | 9.67 | 9.86 | 9.77 | 0.1 | 1.66 | 0 | 11.82 | 0.1 | 22.66 | 4.88 | 19.34 | 4.59 | 12.89 | 4.79 | 4.88 | 9.96 | 9.77 | 9.96 | 9.86 | 9.96
28.032 | 11.82 | 0.1 | 4.59 | 0 | 0.1 | 0.1 | 0.1 | 4.88 | 4.79 | 18.07 | 4.79 | 20.8 | 4.79 | 23.15 | 9.67 | 13.48 | 9.67 | 9.86 | 10.06 | 0.1 | 2.34 | 0.1 | 2.64 | 0.1 | 18.56 | 4.88 | 19.92 | 4.59 | 13.67 | 4.69 | 4.79 | 9.96 | 9.77 | 10.06 | 9.86 | 9.96
28.064 | 12.31 | 0.1 | 5.08 | 0.1 | 0.88 | 0.1 | 0.1 | 4.88 | 4.88 | 17.97 | 4.88 | 17.77 | 4.79 | 23.44 | 9.67 | 14.16 | 9.67 | 9.86 | 10.25 | 0.1 | 1.95 | 0.1 | 2.34 | 0.1 | 17.97 | 4.88 | 20.51 | 4.69 | 14.45 | 4.79 | 4.88 | 9.96 | 9.77 | 9.86 | 9.86 | 9.96
28.096 | 12.89 | 0 | 5.66 | 0.1 | 1.86 | 0 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 17.58 | 4.79 | 23.83 | 9.67 | 14.75 | 9.67 | 9.86 | 9.96 | 0.1 | 1.66 | 0.1 | 2.05 | 0 | 18.07 | 4.88 | 21.09 | 4.69 | 15.14 | 4.69 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
28.128 | 13.38 | 0.1 | 6.15 | 0.1 | 3.13 | 0 | 0.88 | 4.79 | 4.88 | 6.93 | 4.88 | 17.87 | 4.88 | 24.22 | 9.67 | 15.33 | 9.67 | 9.96 | 9.96 | 0.1 | 1.66 | 0.1 | 1.66 | 0 | 17.87 | 4.88 | 21.68 | 4.69 | 15.92 | 4.59 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
28.16 | 13.97 | 0.1 | 6.84 | 0.1 | 4.2 | 0.1 | 1.66 | 4.79 | 4.79 | 2.64 | 4.79 | 18.07 | 4.88 | 20.8 | 9.77 | 15.82 | 9.67 | 9.86 | 9.96 | 0.1 | 1.46 | 0.1 | 1.86 | 0.1 | 17.97 | 4.88 | 22.36 | 4.69 | 16.5 | 4.59 | 5.86 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77
28.192 | 14.45 | 0.1 | 7.42 | 0 | 4.69 | 0.1 | 2.93 | 4.69 | 4.88 | 3.42 | 4.79 | 17.97 | 4.88 | 20.9 | 9.77 | 16.31 | 9.57 | 9.86 | 9.86 | 0.1 | 1.37 | 0 | 1.56 | 0.1 | 17.87 | 4.88 | 22.85 | 4.69 | 17.29 | 4.59 | 7.32 | 9.77 | 9.77 | 9.77 | 10.84 | 9.77
28.224 | 15.04 | 0.1 | 8.11 | 0.1 | 5.37 | 0.1 | 4 | 4.59 | 4.88 | 4.39 | 4.88 | 18.07 | 4.79 | 21.09 | 9.77 | 16.8 | 9.57 | 9.96 | 9.86 | 0 | 1.27 | 0.1 | 1.37 | 0.1 | 17.87 | 4.88 | 23.34 | 4.79 | 17.97 | 4.59 | 9.38 | 9.67 | 10.94 | 9.77 | 12.01 | 9.77
28.256 | 15.53 | 0.1 | 8.69 | 0.1 | 5.86 | 0 | 4.69 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 4.79 | 21.39 | 9.77 | 17.38 | 9.57 | 9.96 | 9.77 | 0.1 | 1.17 | 0.1 | 1.46 | 0.1 | 17.77 | 4.88 | 23.83 | 4.79 | 18.56 | 4.59 | 10.94 | 9.67 | 11.91 | 9.77 | 13.87 | 9.67
28.288 | 16.02 | 0.1 | 9.28 | 0.1 | 6.35 | 0 | 5.18 | 4.49 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 21.29 | 9.77 | 17.77 | 9.57 | 9.96 | 9.67 | 0.1 | 0.88 | 0.1 | 1.17 | 0.1 | 17.77 | 4.88 | 24.41 | 4.79 | 19.24 | 4.59 | 12.21 | 9.67 | 13.67 | 9.67 | 15.53 | 9.67
28.32 | 16.5 | 0.1 | 9.86 | 0.1 | 7.13 | 0.1 | 5.86 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 18.26 | 9.67 | 9.86 | 9.77 | 0.1 | 0.98 | 0 | 1.17 | 0.1 | 17.77 | 4.88 | 24.9 | 4.79 | 19.92 | 4.59 | 13.09 | 9.67 | 15.33 | 9.67 | 16.7 | 9.67
28.352 | 16.9 | 0.1 | 10.45 | 0 | 7.62 | 0.1 | 6.35 | 4.59 | 4.88 | 5.27 | 4.79 | 18.07 | 4.88 | 21.49 | 9.77 | 18.65 | 9.67 | 9.86 | 9.67 | 0.1 | 0.78 | 0.1 | 1.07 | 0.1 | 17.68 | 4.88 | 25.59 | 4.79 | 20.51 | 4.59 | 13.87 | 9.67 | 16.6 | 9.67 | 17.58 | 9.67
28.384 | 15.72 | 0.1 | 11.04 | 0 | 8.3 | 0.1 | 7.03 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 21.39 | 9.77 | 19.14 | 9.67 | 9.86 | 9.67 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 17.77 | 4.88 | 25.98 | 4.79 | 21.19 | 4.59 | 14.65 | 9.67 | 17.38 | 9.67 | 18.26 | 9.67
28.416 | 13.97 | 0.1 | 11.72 | 0.1 | 8.89 | 0 | 7.52 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.53 | 9.57 | 9.96 | 9.67 | 0 | 0.68 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 24.51 | 4.79 | 21.58 | 4.59 | 15.43 | 9.57 | 18.16 | 9.67 | 18.95 | 9.67
28.448 | 14.45 | 0.1 | 12.11 | 0.1 | 9.47 | 0 | 8.2 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.02 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.78 | 0 | 15.04 | 4.88 | 21.88 | 4.79 | 22.27 | 4.69 | 16.11 | 9.57 | 18.85 | 9.67 | 19.53 | 9.57
28.48 | 14.55 | 0.1 | 12.79 | 0.1 | 10.16 | 0.1 | 8.79 | 4.59 | 4.88 | 4.88 | 4.79 | 17.97 | 4.88 | 21.39 | 9.86 | 20.31 | 9.67 | 10.45 | 9.67 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 3.71 | 4.88 | 21.68 | 4.79 | 22.85 | 4.69 | 16.7 | 9.67 | 19.43 | 9.67 | 20.22 | 9.67
28.512 | 14.45 | 0.1 | 13.28 | 0 | 10.64 | 0.1 | 9.47 | 4.59 | 4.88 | 4.79 | 4.79 | 6.45 | 4.88 | 21.39 | 9.86 | 20.9 | 9.67 | 11.43 | 9.67 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.25 | 4.88 | 21.58 | 4.79 | 23.34 | 4.69 | 17.38 | 9.67 | 20.12 | 9.67 | 20.8 | 9.67
28.544 | 14.45 | 0.1 | 13.77 | 0 | 11.23 | 0.1 | 9.96 | 4.59 | 4.88 | 4.69 | 4.88 | 2.64 | 4.88 | 21.39 | 9.86 | 21.19 | 9.67 | 12.79 | 9.67 | 0 | 0.49 | 0.1 | 0.59 | 0.1 | 2.15 | 4.88 | 21.49 | 4.88 | 23.93 | 4.69 | 18.07 | 9.67 | 20.7 | 9.67 | 21.39 | 9.67
28.576 | 14.36 | 0 | 14.36 | 0.1 | 11.82 | 0 | 10.55 | 4.59 | 4.88 | 4.69 | 4.88 | 3.32 | 4.79 | 21.19 | 9.77 | 21.58 | 9.67 | 13.77 | 9.57 | 0.1 | 0.39 | 0.1 | 0.49 | 0 | 1.76 | 4.88 | 21.29 | 4.88 | 24.22 | 4.69 | 18.75 | 9.67 | 21.39 | 9.67 | 21.97 | 9.67
28.608 | 14.36 | 0.1 | 14.75 | 0.1 | 12.31 | 0 | 11.13 | 4.59 | 4.88 | 4.69 | 4.88 | 4.49 | 4.88 | 10.45 | 9.86 | 22.07 | 9.67 | 14.45 | 9.57 | 0.1 | 0.29 | 0.1 | 0.39 | 0 | 1.66 | 4.88 | 21.39 | 4.88 | 24.9 | 4.79 | 19.34 | 9.67 | 21.97 | 9.67 | 22.66 | 9.67
28.64 | 14.45 | 0.1 | 15.33 | 0.1 | 12.89 | 0.1 | 11.72 | 4.59 | 4.79 | 4.59 | 4.88 | 5.66 | 4.88 | 6.64 | 9.86 | 22.36 | 9.67 | 15.04 | 9.57 | 0.1 | 0.29 | 0.2 | 0.39 | 0.1 | 1.56 | 4.88 | 21.29 | 4.79 | 25.49 | 4.69 | 20.02 | 9.67 | 22.56 | 9.67 | 23.24 | 9.67
28.672 | 14.36 | 0.1 | 15.82 | 0 | 13.48 | 0.1 | 12.31 | 4.69 | 4.79 | 4.59 | 4.79 | 5.86 | 4.88 | 9.47 | 9.86 | 22.75 | 9.67 | 15.63 | 9.57 | 0.1 | 0.2 | 1.07 | 0.29 | 0.1 | 1.46 | 4.88 | 21.29 | 4.88 | 25.98 | 4.69 | 20.61 | 9.67 | 23.15 | 9.67 | 23.73 | 9.67
28.704 | 14.36 | 0.1 | 16.31 | 0.1 | 13.97 | 0.1 | 12.89 | 4.69 | 4.88 | 4.49 | 4.88 | 5.47 | 4.79 | 11.62 | 9.86 | 23.15 | 9.67 | 16.11 | 9.67 | 0.2 | 0.2 | 2.34 | 0.29 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 23.44 | 4.79 | 21.19 | 9.67 | 23.54 | 9.77 | 24.22 | 9.67
28.736 | 14.36 | 0.1 | 16.8 | 0.1 | 14.55 | 0 | 13.48 | 4.69 | 4.98 | 4.59 | 4.88 | 5.27 | 4.79 | 10.74 | 9.77 | 23.54 | 9.67 | 16.6 | 9.57 | 1.17 | 0.2 | 4.2 | 0.29 | 0 | 1.07 | 4.88 | 21.19 | 4.88 | 21.68 | 4.79 | 21.88 | 9.67 | 24.02 | 9.67 | 24.71 | 9.67
28.768 | 14.36 | 0.1 | 17.19 | 0.1 | 15.04 | 0 | 13.97 | 4.69 | 5.86 | 4.59 | 4.88 | 4.98 | 4.88 | 10.16 | 9.86 | 23.93 | 9.77 | 17.09 | 9.57 | 2.64 | 0.1 | 5.66 | 0.2 | 0 | 1.17 | 4.88 | 21.19 | 4.88 | 21.58 | 4.79 | 22.36 | 9.67 | 24.61 | 9.77 | 25.2 | 9.67
28.8 | 14.36 | 0.1 | 15.23 | 0.1 | 15.53 | 0.1 | 14.45 | 4.69 | 6.84 | 4.49 | 4.88 | 5.08 | 4.88 | 9.96 | 9.86 | 24.32 | 9.77 | 17.58 | 9.67 | 4.39 | 0.1 | 6.93 | 0.2 | 0.1 | 1.07 | 4.88 | 21.19 | 4.79 | 21.49 | 4.79 | 22.95 | 9.77 | 25.1 | 9.77 | 25.78 | 9.67
28.832 | 6.64 | 0.1 | 15.33 | 0 | 16.02 | 0.1 | 14.94 | 4.79 | 8.2 | 4.49 | 4.79 | 4.88 | 4.79 | 10.16 | 9.86 | 21.39 | 9.77 | 17.97 | 9.67 | 5.86 | 0.1 | 7.62 | 0.1 | 1.17 | 0.88 | 4.88 | 21.19 | 4.79 | 21.39 | 4.79 | 23.54 | 9.77 | 25.68 | 9.77 | 26.17 | 9.67
28.864 | 1.86 | 0.1 | 14.84 | 0 | 16.5 | 0.1 | 15.53 | 4.79 | 9.18 | 4.49 | 4.88 | 4.79 | 4.79 | 10.25 | 9.86 | 21.97 | 9.77 | 18.46 | 9.67 | 6.93 | 0.1 | 8.5 | 0.1 | 2.54 | 0.78 | 4.88 | 8.69 | 4.88 | 21.29 | 4.88 | 23.93 | 9.77 | 26.08 | 9.77 | 26.76 | 9.77
28.896 | 1.95 | 0.1 | 14.75 | 0.1 | 16.99 | 0 | 16.02 | 4.79 | 9.86 | 4.49 | 4.88 | 4.79 | 4.79 | 10.06 | 9.77 | 21.49 | 9.77 | 18.95 | 9.67 | 7.52 | 0.1 | 9.28 | 0.1 | 4.39 | 0.78 | 4.88 | 2.25 | 4.88 | 21.29 | 4.88 | 24.51 | 9.77 | 26.56 | 9.77 | 27.15 | 9.77
28.928 | 1.76 | 0.1 | 14.65 | 0.1 | 16.02 | 0 | 16.5 | 4.79 | 10.55 | 4.59 | 4.88 | 4.79 | 4.88 | 9.86 | 9.86 | 21.58 | 9.77 | 19.34 | 9.67 | 8.5 | 0.1 | 10.06 | 0.1 | 5.76 | 0.68 | 4.88 | 3.13 | 4.88 | 21.29 | 4.88 | 25.1 | 9.77 | 27.15 | 9.77 | 27.83 | 9.77
28.96 | 1.37 | 0.1 | 14.55 | 0.1 | 13.97 | 0.1 | 16.9 | 4.79 | 11.23 | 4.59 | 4.79 | 4.69 | 4.88 | 9.96 | 9.86 | 21.58 | 9.77 | 19.82 | 9.67 | 9.28 | 0 | 10.94 | 0.1 | 6.93 | 0.68 | 4.88 | 3.91 | 4.88 | 21.29 | 4.79 | 25.39 | 9.77 | 27.64 | 9.77 | 28.22 | 9.77
28.992 | 1.37 | 0.1 | 14.55 | 0 | 14.55 | 0.1 | 16.9 | 4.79 | 11.72 | 4.59 | 4.79 | 4.59 | 4.88 | 9.86 | 9.86 | 21.49 | 9.77 | 20.22 | 9.67 | 9.96 | 0 | 11.62 | 0.1 | 7.72 | 0.59 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.98 | 9.77 | 28.03 | 9.77 | 28.61 | 9.77
29.024 | 1.27 | 0.1 | 14.55 | 0.1 | 14.65 | 0.1 | 14.94 | 4.88 | 12.31 | 4.59 | 4.88 | 4.59 | 4.79 | 9.86 | 9.86 | 21.49 | 9.77 | 20.7 | 9.67 | 10.84 | 0.1 | 12.4 | 0.1 | 8.59 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 23.24 | 9.77 | 28.52 | 9.86 | 29.1 | 9.77
29.056 | 1.07 | 0 | 14.45 | 0.1 | 14.55 | 0 | 14.65 | 4.79 | 12.89 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 21 | 9.67 | 11.62 | 0.1 | 13.09 | 0.1 | 9.47 | 0.39 | 4.88 | 6.05 | 4.88 | 21.19 | 4.88 | 20.61 | 9.77 | 29.01 | 9.86 | 28.91 | 9.77
29.088 | 0.98 | 0.1 | 14.45 | 0.1 | 14.55 | 0 | 14.75 | 4.79 | 13.38 | 4.69 | 4.88 | 4.59 | 4.79 | 9.77 | 9.86 | 21.39 | 9.77 | 21.49 | 9.67 | 12.4 | 0.1 | 13.77 | 0.1 | 10.06 | 0.39 | 4.79 | 5.57 | 4.88 | 21.19 | 4.88 | 21.19 | 9.77 | 29.4 | 9.86 | 25.29 | 9.77
29.12 | 0.78 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 14.65 | 4.88 | 13.97 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.39 | 9.86 | 21.88 | 9.77 | 13.18 | 0 | 14.55 | 0.1 | 10.94 | 0.29 | 4.88 | 5.27 | 4.79 | 21.19 | 4.88 | 21.29 | 9.77 | 25.78 | 9.86 | 24.51 | 9.77
29.152 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 14.65 | 4.88 | 14.45 | 4.69 | 5.47 | 4.49 | 4.88 | 9.77 | 9.86 | 21.39 | 9.86 | 22.17 | 9.77 | 13.77 | 0 | 15.33 | 0.1 | 11.82 | 0.29 | 4.88 | 5.18 | 4.79 | 8.5 | 4.79 | 21.19 | 9.77 | 24.71 | 9.86 | 24.61 | 9.86
29.184 | 0.68 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 14.94 | 4.69 | 6.64 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.86 | 22.56 | 9.77 | 14.45 | 0.1 | 16.02 | 0.1 | 12.5 | 0.2 | 4.88 | 5.18 | 4.88 | 2.34 | 4.88 | 21.19 | 9.77 | 24.61 | 9.86 | 24.51 | 9.86
29.216 | 0.59 | 0 | 8.79 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 15.43 | 4.79 | 7.81 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 22.95 | 9.77 | 15.14 | 0 | 16.7 | 0.1 | 13.28 | 0.2 | 5.18 | 5.08 | 4.88 | 3.22 | 4.88 | 21.29 | 9.77 | 24.51 | 9.86 | 24.41 | 9.86
29.248 | 0.49 | 0.1 | 2.93 | 0.1 | 14.45 | 0 | 14.45 | 4.79 | 15.92 | 4.79 | 8.79 | 4.49 | 4.88 | 9.67 | 9.77 | 17.87 | 9.86 | 23.34 | 9.77 | 15.82 | 0.1 | 17.38 | 0.1 | 13.97 | 0.2 | 5.96 | 5.08 | 4.88 | 3.91 | 4.88 | 21.29 | 9.77 | 24.51 | 9.86 | 24.32 | 9.77
29.28 | 0.49 | 0.1 | 2.05 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 16.41 | 4.79 | 9.67 | 4.49 | 4.88 | 9.67 | 9.86 | 7.91 | 9.86 | 23.83 | 9.77 | 16.6 | 0 | 18.07 | 0.1 | 14.65 | 0.1 | 7.52 | 4.98 | 4.79 | 5.08 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.22 | 9.86
29.312 | 0.39 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 16.9 | 4.79 | 10.35 | 4.49 | 5.57 | 9.67 | 9.86 | 7.52 | 9.86 | 24.12 | 9.77 | 17.19 | 0 | 18.65 | 0.1 | 15.53 | 0.1 | 9.67 | 4.88 | 4.79 | 6.05 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
29.344 | 0.29 | 0.1 | 1.46 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.38 | 4.79 | 11.04 | 4.59 | 6.05 | 9.67 | 9.86 | 10.45 | 9.86 | 24.02 | 9.77 | 17.97 | 0.1 | 19.34 | 0.1 | 16.02 | 0.1 | 11.23 | 4.79 | 4.88 | 6.05 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
29.376 | 0.29 | 0 | 1.46 | 0.1 | 6.74 | 0 | 14.36 | 4.88 | 17.87 | 4.79 | 11.52 | 4.59 | 7.03 | 9.57 | 9.86 | 11.43 | 9.86 | 21.97 | 9.77 | 18.56 | 0 | 20.02 | 0.1 | 16.9 | 0.1 | 12.4 | 4.79 | 4.88 | 5.57 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
29.408 | 0.29 | 0.1 | 1.27 | 0.1 | 1.95 | 0 | 14.45 | 4.88 | 18.26 | 4.79 | 12.11 | 4.59 | 8.5 | 9.57 | 9.86 | 10.55 | 9.86 | 21.88 | 9.77 | 19.24 | 0 | 20.61 | 0.1 | 17.48 | 0 | 13.18 | 4.79 | 4.88 | 5.37 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
29.44 | 0.2 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 7.52 | 4.88 | 18.75 | 4.79 | 12.79 | 4.59 | 9.38 | 9.67 | 9.86 | 9.86 | 9.86 | 21.68 | 9.86 | 19.92 | 0 | 21.19 | 0.1 | 18.07 | 0.1 | 14.06 | 4.69 | 4.79 | 5.18 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
29.472 | 0.2 | 0.1 | 1.07 | 0 | 1.76 | 0.1 | 2.15 | 4.88 | 19.24 | 4.79 | 13.18 | 4.69 | 9.96 | 9.67 | 9.86 | 9.86 | 9.86 | 21.58 | 9.86 | 20.41 | 0 | 21.78 | 0.1 | 18.65 | 0.1 | 14.75 | 4.69 | 4.88 | 5.08 | 4.79 | 7.03 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
29.504 | 0.1 | 0.1 | 0.88 | 0 | 1.46 | 0.1 | 1.95 | 4.88 | 19.53 | 4.79 | 13.77 | 4.69 | 10.64 | 9.67 | 9.86 | 10.16 | 9.86 | 21.68 | 9.86 | 21 | 0.1 | 22.36 | 0.1 | 19.34 | 0.1 | 15.53 | 4.59 | 4.88 | 5.08 | 4.88 | 2.34 | 9.86 | 24.22 | 9.86 | 13.18 | 9.86
29.536 | 0 | 0 | 0.78 | 0.1 | 1.46 | 0 | 1.66 | 4.88 | 20.02 | 4.88 | 14.36 | 4.79 | 11.23 | 9.67 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 21.58 | 0.1 | 21.49 | 0.1 | 20.02 | 0 | 16.31 | 4.59 | 4.79 | 4.98 | 4.88 | 3.32 | 9.86 | 13.57 | 9.86 | 5.96 | 9.86
29.568 | 0 | 0.1 | 0.68 | 0.1 | 1.17 | 0 | 1.56 | 4.88 | 20.51 | 4.88 | 14.84 | 4.69 | 11.82 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 9.77 | 22.17 | 0.1 | 17.97 | 0.1 | 20.61 | 0 | 16.9 | 4.59 | 4.79 | 4.98 | 4.88 | 4 | 9.77 | 6.15 | 9.86 | 7.81 | 9.86
29.6 | 0.1 | 0.78 | 0.68 | 0.1 | 1.07 | 0.1 | 1.37 | 4.88 | 20.9 | 4.88 | 15.33 | 4.79 | 12.4 | 9.67 | 11.04 | 9.86 | 9.96 | 21.49 | 9.86 | 19.53 | 0 | 17.77 | 0.1 | 21.19 | 0.1 | 17.58 | 4.59 | 4.79 | 4.88 | 4.88 | 5.37 | 9.86 | 7.52 | 9.86 | 11.72 | 9.86
29.632 | 0.1 | 1.66 | 0.59 | 0 | 0.88 | 0.1 | 1.27 | 4.88 | 17.38 | 4.88 | 15.82 | 4.79 | 12.99 | 9.67 | 11.91 | 9.96 | 9.96 | 21.49 | 9.86 | 16.9 | 0 | 17.87 | 0.1 | 21.88 | 0.1 | 18.26 | 4.59 | 4.79 | 4.88 | 4.79 | 6.35 | 9.77 | 11.72 | 9.86 | 11.62 | 9.86
29.664 | 0 | 2.93 | 0.59 | 0 | 0.88 | 0.1 | 1.07 | 4.88 | 17.77 | 4.79 | 16.31 | 4.79 | 13.48 | 9.67 | 13.28 | 9.86 | 9.96 | 21.49 | 9.86 | 17.29 | 0.1 | 17.87 | 0.1 | 22.36 | 0.1 | 18.95 | 4.59 | 4.88 | 4.79 | 4.88 | 5.96 | 9.77 | 11.72 | 9.86 | 10.45 | 9.86
29.696 | 0 | 4 | 0.49 | 0.1 | 0.78 | 0 | 0.88 | 4.88 | 17.87 | 4.88 | 16.8 | 4.79 | 13.97 | 9.67 | 14.16 | 9.77 | 9.86 | 21.49 | 9.86 | 17.58 | 0.1 | 17.77 | 0.1 | 20.31 | 0 | 19.43 | 4.49 | 5.96 | 4.69 | 4.88 | 5.57 | 9.77 | 10.45 | 9.86 | 9.96 | 9.86
29.728 | 0 | 4.69 | 0.29 | 0.1 | 0.78 | 0 | 0.88 | 4.88 | 18.07 | 4.88 | 17.29 | 4.79 | 14.45 | 9.67 | 14.75 | 9.77 | 9.86 | 21.39 | 9.86 | 17.58 | 0.1 | 17.77 | 0.1 | 17.58 | 0 | 20.02 | 4.49 | 6.93 | 4.69 | 4.88 | 5.18 | 9.77 | 9.96 | 9.86 | 10.25 | 9.86
29.76 | 0.1 | 5.27 | 0.39 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 17.68 | 4.79 | 15.04 | 9.67 | 15.23 | 9.77 | 9.86 | 21.39 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.97 | 0.1 | 20.7 | 4.59 | 8.79 | 4.69 | 4.88 | 5.18 | 9.77 | 10.25 | 9.86 | 10.45 | 9.86
29.792 | 0.1 | 5.86 | 0.29 | 0 | 0.49 | 0.1 | 0.78 | 4.88 | 18.07 | 4.79 | 18.07 | 4.79 | 15.53 | 9.77 | 15.72 | 9.67 | 9.86 | 12.79 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.97 | 0.1 | 21.29 | 4.59 | 10.45 | 4.59 | 4.79 | 5.08 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86
29.824 | 0 | 6.35 | 0.2 | 0 | 0.49 | 0.1 | 0.59 | 4.88 | 18.16 | 4.79 | 18.56 | 4.79 | 16.02 | 9.77 | 16.31 | 9.67 | 9.86 | 6.84 | 9.86 | 17.58 | 0.1 | 17.77 | 0.1 | 17.87 | 0.1 | 21.97 | 4.59 | 11.82 | 4.49 | 4.88 | 5.08 | 9.77 | 10.45 | 9.86 | 10.16 | 9.86
29.856 | 0 | 7.03 | 0.2 | 0.1 | 0.39 | 0 | 0.59 | 4.88 | 18.07 | 4.88 | 19.14 | 4.88 | 16.5 | 9.67 | 16.8 | 9.67 | 9.86 | 8.2 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.68 | 0 | 22.46 | 4.59 | 12.7 | 4.49 | 4.88 | 5.08 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
29.888 | 0 | 7.62 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.88 | 18.07 | 4.88 | 19.43 | 4.88 | 16.99 | 9.67 | 17.19 | 9.67 | 9.86 | 11.04 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 0 | 22.95 | 4.59 | 13.67 | 4.59 | 4.88 | 4.98 | 9.77 | 10.06 | 9.86 | 10.25 | 9.86
29.92 | 0.1 | 8.2 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 19.92 | 4.88 | 17.48 | 9.77 | 17.77 | 9.67 | 9.86 | 11.13 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 23.54 | 4.69 | 14.45 | 4.59 | 4.88 | 4.88 | 9.86 | 10.06 | 9.86 | 10.16 | 9.86
29.952 | 0 | 8.79 | 0.1 | 0 | 0.29 | 0.1 | 0.39 | 4.88 | 18.07 | 4.79 | 20.31 | 4.79 | 17.87 | 9.77 | 18.26 | 9.67 | 9.86 | 10.35 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 24.22 | 4.69 | 15.14 | 4.59 | 4.79 | 4.88 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86
29.984 | 0.1 | 9.47 | 0.1 | 0.59 | 0.2 | 0.1 | 0.29 | 4.88 | 18.07 | 4.88 | 20.8 | 4.88 | 18.26 | 9.77 | 18.65 | 9.67 | 9.86 | 9.96 | 9.86 | 17.58 | 0.1 | 8.01 | 0.1 | 17.77 | 0.1 | 24.61 | 4.69 | 15.92 | 4.49 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
30.016 | 0 | 10.06 | 0.1 | 1.37 | 0.2 | 0.2 | 0.2 | 4.88 | 18.07 | 4.88 | 19.73 | 4.88 | 18.75 | 9.77 | 19.14 | 9.67 | 9.86 | 9.96 | 9.86 | 17.48 | 0 | 1.76 | 0.1 | 17.68 | 0 | 25.1 | 4.69 | 16.6 | 4.49 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
30.048 | 0 | 10.55 | 0 | 2.25 | 0.1 | 1.07 | 0.2 | 4.88 | 18.07 | 4.88 | 17.58 | 4.88 | 19.24 | 9.77 | 19.43 | 9.57 | 9.86 | 10.25 | 9.86 | 6.74 | 0.1 | 2.34 | 0.1 | 17.68 | 0 | 25.59 | 4.69 | 17.19 | 4.49 | 4.88 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
30.08 | 0.1 | 11.23 | 0.1 | 3.71 | 0.1 | 2.05 | 0.2 | 4.88 | 5.18 | 4.88 | 17.97 | 4.88 | 19.63 | 9.77 | 20.02 | 9.57 | 9.86 | 10.16 | 9.96 | 1.66 | 0 | 2.15 | 0.1 | 17.68 | 0.1 | 26.17 | 4.79 | 17.97 | 4.49 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
30.112 | 0.1 | 11.82 | 0.1 | 4.49 | 0.1 | 3.52 | 0.1 | 4.88 | 2.93 | 4.79 | 18.26 | 4.79 | 20.02 | 9.77 | 20.41 | 9.57 | 9.86 | 9.86 | 9.96 | 2.15 | 0 | 1.76 | 0.1 | 17.68 | 0 | 21.58 | 4.79 | 18.56 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
30.144 | 0 | 12.31 | 0.1 | 5.08 | 0.1 | 4.3 | 0.1 | 4.88 | 3.52 | 4.79 | 18.16 | 4.88 | 20.61 | 9.77 | 20.8 | 9.57 | 10.64 | 9.86 | 9.96 | 2.05 | 0 | 1.86 | 0.1 | 7.42 | 0.1 | 20.31 | 4.79 | 19.14 | 4.59 | 5.96 | 4.59 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86
30.176 | 0 | 12.99 | 0.1 | 5.57 | 0.1 | 4.98 | 0.1 | 4.88 | 4.49 | 4.88 | 18.16 | 4.88 | 20.9 | 9.77 | 21.19 | 9.67 | 11.43 | 9.96 | 9.96 | 1.76 | 0 | 1.56 | 0.1 | 1.66 | 0 | 21 | 4.79 | 19.82 | 4.59 | 6.93 | 4.59 | 10.64 | 9.77 | 9.86 | 9.77 | 9.86
30.208 | 0 | 13.48 | 0.1 | 6.05 | 0.1 | 5.47 | 0 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 17.77 | 9.77 | 21.58 | 9.67 | 12.6 | 9.86 | 9.86 | 1.66 | 0 | 1.46 | 0.1 | 2.44 | 0 | 21.19 | 4.79 | 20.31 | 4.59 | 8.69 | 4.59 | 11.43 | 9.67 | 10.64 | 9.77 | 9.86
30.24 | 0.1 | 13.97 | 0.1 | 6.64 | 0 | 5.96 | 0.1 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 18.26 | 9.86 | 22.07 | 9.67 | 13.67 | 9.77 | 9.96 | 1.46 | 0 | 1.27 | 0.1 | 2.05 | 0.1 | 21.19 | 4.79 | 21.09 | 4.59 | 10.55 | 4.59 | 12.99 | 9.67 | 11.72 | 9.67 | 11.04
30.272 | 0.1 | 14.45 | 0.1 | 7.32 | 0 | 6.64 | 0.1 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 18.07 | 9.86 | 22.46 | 9.67 | 14.36 | 9.77 | 9.96 | 1.46 | 0 | 1.17 | 0.1 | 1.56 | 0 | 21.29 | 4.79 | 21.58 | 4.69 | 11.72 | 4.59 | 14.94 | 9.67 | 13.38 | 9.67 | 11.91
30.304 | 0.1 | 14.94 | 0.1 | 8.01 | 0.1 | 7.23 | 0.1 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 18.16 | 9.86 | 22.75 | 9.67 | 14.94 | 9.77 | 9.96 | 1.27 | 0.1 | 1.17 | 0.1 | 1.76 | 0.1 | 21.29 | 4.88 | 22.17 | 4.69 | 12.79 | 4.49 | 16.31 | 9.67 | 15.14 | 9.67 | 13.77
30.336 | 0 | 15.53 | 0.1 | 8.5 | 0.1 | 7.72 | 0 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 23.15 | 9.67 | 15.53 | 9.77 | 9.96 | 1.17 | 0 | 1.17 | 0.1 | 1.37 | 0 | 21.29 | 4.79 | 22.75 | 4.69 | 13.67 | 4.49 | 17.29 | 9.67 | 16.41 | 9.67 | 15.53
30.368 | 0 | 15.92 | 0 | 9.08 | 0.1 | 8.4 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 23.54 | 9.67 | 16.02 | 9.67 | 9.96 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 0 | 21.29 | 4.79 | 23.24 | 4.69 | 14.45 | 4.59 | 17.87 | 9.67 | 17.29 | 9.67 | 16.6
30.4 | 0.1 | 16.41 | 0.1 | 9.67 | 0 | 9.08 | 0.1 | 5.08 | 4.88 | 4.88 | 17.97 | 4.88 | 18.16 | 9.86 | 23.93 | 9.67 | 16.5 | 9.67 | 9.96 | 0.88 | 0 | 0.78 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 23.83 | 4.69 | 15.23 | 4.59 | 18.65 | 9.67 | 18.07 | 9.57 | 17.38
30.432 | 0 | 16.9 | 0.1 | 10.35 | 0 | 9.67 | 0.1 | 6.05 | 4.88 | 4.79 | 17.97 | 4.88 | 18.07 | 9.86 | 24.32 | 9.67 | 16.99 | 9.67 | 9.96 | 0.88 | 0 | 0.78 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 24.32 | 4.69 | 15.82 | 4.59 | 19.34 | 9.57 | 18.85 | 9.57 | 18.16
30.464 | 0.1 | 16.9 | 0.1 | 10.84 | 0.1 | 10.25 | 0.1 | 6.93 | 4.79 | 4.79 | 7.81 | 4.88 | 18.07 | 9.86 | 21 | 9.67 | 17.58 | 9.67 | 10.06 | 0.78 | 0 | 0.68 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 24.81 | 4.69 | 16.5 | 4.59 | 20.02 | 9.57 | 19.43 | 9.57 | 18.85
30.496 | 0 | 14.65 | 0.1 | 11.52 | 0.1 | 10.74 | 0.1 | 8.4 | 4.69 | 4.88 | 2.73 | 4.88 | 18.07 | 9.77 | 21.68 | 9.67 | 17.97 | 9.67 | 10.74 | 0.78 | 0 | 0.59 | 0.1 | 0.88 | 0 | 21.19 | 4.79 | 25.39 | 4.69 | 17.29 | 4.69 | 20.61 | 9.67 | 20.12 | 9.67 | 19.63
30.528 | 0 | 14.84 | 0 | 12.01 | 0.1 | 11.33 | 0 | 9.28 | 4.69 | 4.88 | 3.22 | 4.88 | 18.07 | 9.86 | 21.39 | 9.77 | 18.46 | 9.67 | 11.72 | 0.59 | 0 | 0.59 | 0.1 | 0.88 | 0 | 21.19 | 4.79 | 25.78 | 4.79 | 17.87 | 4.69 | 21.19 | 9.57 | 20.7 | 9.67 | 20.12
30.56 | 0.1 | 14.75 | 0.1 | 12.6 | 0.1 | 11.91 | 0.1 | 9.96 | 4.59 | 4.88 | 4.39 | 4.88 | 18.07 | 9.86 | 21.39 | 9.67 | 18.85 | 9.57 | 13.09 | 0.59 | 0 | 0.49 | 0.1 | 0.78 | 0.1 | 6.45 | 4.88 | 26.27 | 4.79 | 18.56 | 4.69 | 21.88 | 9.67 | 21.29 | 9.67 | 20.7
30.592 | 0.1 | 14.65 | 0.1 | 13.18 | 0 | 12.5 | 0.1 | 10.55 | 4.59 | 4.79 | 5.37 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 19.34 | 9.57 | 14.06 | 0.49 | 0 | 0.39 | 0.1 | 0.78 | 0 | 2.64 | 4.88 | 23.05 | 4.79 | 19.24 | 4.69 | 22.27 | 9.57 | 21.97 | 9.67 | 21.39
30.624 | 0.1 | 14.65 | 0.1 | 13.67 | 0.1 | 12.99 | 0.1 | 11.13 | 4.59 | 4.88 | 6.05 | 4.88 | 13.38 | 9.86 | 21.49 | 9.77 | 19.73 | 9.67 | 14.65 | 0.49 | 0 | 0.39 | 0.1 | 0.68 | 0.1 | 3.52 | 4.88 | 22.27 | 4.79 | 19.82 | 4.69 | 22.95 | 9.67 | 22.46 | 9.67 | 21.97
30.656 | 0 | 14.45 | 0.1 | 14.26 | 0.1 | 13.57 | 0 | 11.82 | 4.49 | 4.88 | 5.57 | 4.88 | 4.3 | 9.86 | 21.49 | 9.77 | 20.22 | 9.67 | 15.14 | 0.39 | 1.07 | 0.29 | 0.1 | 0.59 | 0 | 4.2 | 4.79 | 21.78 | 4.79 | 20.51 | 4.69 | 23.44 | 9.67 | 22.85 | 9.67 | 22.56
30.688 | 0 | 14.45 | 0 | 14.75 | 0.1 | 14.06 | 0.1 | 12.31 | 4.49 | 4.88 | 5.18 | 4.88 | 3.13 | 9.86 | 21.49 | 9.77 | 20.61 | 9.67 | 15.72 | 0.29 | 2.15 | 0.29 | 0.78 | 0.49 | 0 | 5.27 | 4.79 | 21.58 | 4.79 | 21 | 4.79 | 23.93 | 9.67 | 23.54 | 9.67 | 23.05
30.72 | 0.1 | 14.55 | 0.1 | 15.23 | 0 | 14.65 | 0.1 | 12.89 | 4.49 | 4.88 | 4.88 | 4.88 | 3.61 | 9.86 | 21.39 | 9.77 | 21 | 9.67 | 16.31 | 0.29 | 4 | 0.2 | 1.76 | 0.49 | 0.1 | 6.15 | 4.88 | 21.49 | 4.79 | 21.58 | 4.79 | 24.51 | 9.67 | 24.02 | 9.67 | 23.54
30.752 | 0 | 14.45 | 0.1 | 15.72 | 0 | 15.14 | 0.1 | 13.38 | 4.49 | 4.79 | 4.88 | 4.88 | 4.79 | 9.86 | 21.49 | 9.77 | 21.39 | 9.67 | 16.7 | 0.29 | 5.47 | 0.2 | 3.42 | 0.49 | 0.1 | 6.05 | 4.88 | 21.39 | 4.79 | 22.17 | 4.69 | 25 | 9.67 | 24.61 | 9.67 | 24.12
30.784 | 0 | 14.45 | 0.1 | 16.11 | 0.1 | 15.63 | 0.1 | 13.97 | 4.49 | 4.98 | 4.88 | 4.88 | 5.76 | 9.86 | 21.39 | 9.77 | 21.78 | 9.67 | 17.29 | 0.2 | 6.64 | 0.2 | 5.37 | 0.39 | 0 | 5.57 | 4.88 | 21.29 | 4.79 | 22.66 | 4.79 | 25.49 | 9.67 | 25.1 | 9.67 | 24.51
30.816 | 0 | 14.36 | 0.1 | 16.7 | 0.1 | 16.11 | 0 | 14.55 | 4.49 | 5.57 | 4.98 | 4.88 | 5.76 | 9.86 | 21.39 | 9.77 | 22.17 | 9.67 | 17.68 | 0.2 | 7.52 | 0.2 | 6.45 | 0.29 | 0 | 5.18 | 4.79 | 21.39 | 4.79 | 23.24 | 4.79 | 25.98 | 9.67 | 25.68 | 9.77 | 25.2
30.848 | 0 | 14.36 | 0.1 | 17.19 | 0.1 | 16.6 | 0.1 | 14.94 | 4.49 | 6.54 | 4.88 | 4.88 | 5.37 | 9.86 | 21.39 | 9.77 | 22.56 | 9.67 | 18.07 | 0.2 | 8.2 | 0.2 | 7.32 | 0.2 | 0 | 5.08 | 4.79 | 21.29 | 4.79 | 23.83 | 4.79 | 26.47 | 9.67 | 26.17 | 9.77 | 25.59
30.88 | 0.1 | 14.45 | 0.1 | 13.97 | 0 | 17.09 | 0.1 | 15.53 | 4.49 | 8.01 | 4.69 | 4.88 | 4.98 | 9.86 | 20.9 | 9.77 | 23.05 | 9.67 | 18.65 | 0.1 | 8.98 | 0.1 | 8.01 | 0.2 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 24.41 | 4.79 | 27.05 | 9.67 | 26.56 | 9.77 | 26.17
30.912 | 0.1 | 7.72 | 0.1 | 14.75 | 0 | 13.87 | 0.1 | 16.02 | 4.59 | 8.98 | 4.69 | 4.88 | 4.98 | 9.86 | 8.79 | 9.86 | 23.44 | 9.67 | 19.04 | 0.1 | 9.77 | 0.1 | 8.79 | 0.2 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 24.9 | 4.79 | 27.64 | 9.67 | 27.05 | 9.77 | 26.66
30.944 | 0 | 2.15 | 0.1 | 14.36 | 0.1 | 13.97 | 0.1 | 16.5 | 4.59 | 9.67 | 4.69 | 4.88 | 4.88 | 9.86 | 7.13 | 9.86 | 23.73 | 9.67 | 19.43 | 0.1 | 10.55 | 0.1 | 9.67 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 25.29 | 4.79 | 28.03 | 9.67 | 27.54 | 9.77 | 27.15
30.976 | 0 | 1.95 | 0.1 | 14.45 | 0.1 | 14.16 | 0 | 16.99 | 4.69 | 10.45 | 4.59 | 4.88 | 4.88 | 9.86 | 10.25 | 9.86 | 24.02 | 9.67 | 19.92 | 0.1 | 11.43 | 0.1 | 10.45 | 0.1 | 0 | 4.98 | 4.79 | 21.19 | 4.79 | 25.78 | 4.79 | 28.52 | 9.77 | 28.03 | 9.77 | 27.64
31.008 | 0 | 1.76 | 0 | 14.45 | 0.1 | 14.45 | 0 | 17.38 | 4.69 | 11.04 | 4.59 | 4.88 | 4.79 | 9.86 | 11.43 | 9.86 | 24.41 | 9.77 | 20.31 | 0.1 | 12.21 | 0.1 | 11.23 | 0.1 | 0 | 4.98 | 4.79 | 10.45 | 4.79 | 26.27 | 4.79 | 28.91 | 9.77 | 28.52 | 9.77 | 28.13
31.04 | 0.1 | 1.56 | 0.1 | 14.45 | 0 | 14.36 | 0.1 | 17.87 | 4.69 | 11.62 | 4.59 | 4.88 | 4.79 | 9.86 | 10.64 | 9.86 | 22.27 | 9.67 | 20.8 | 0.1 | 12.99 | 0.1 | 12.11 | 0.1 | 0.1 | 4.88 | 4.88 | 2.64 | 4.88 | 22.75 | 4.88 | 29.4 | 9.77 | 29.01 | 9.77 | 28.52
31.072 | 0 | 1.46 | 0.1 | 14.55 | 0 | 14.45 | 0.1 | 18.26 | 4.69 | 12.21 | 4.49 | 4.88 | 4.69 | 9.86 | 10.06 | 9.86 | 22.07 | 9.67 | 21.19 | 0.1 | 13.67 | 0.1 | 12.79 | 0.1 | 0.1 | 4.79 | 4.88 | 3.32 | 4.79 | 22.17 | 4.79 | 25.2 | 9.77 | 29.49 | 9.77 | 29.01
31.104 | 0 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.75 | 4.69 | 12.79 | 4.49 | 4.88 | 4.69 | 9.86 | 9.96 | 9.86 | 21.78 | 9.77 | 21.58 | 0.1 | 14.36 | 0.1 | 13.48 | 0.1 | 0.1 | 4.79 | 4.79 | 3.61 | 4.79 | 21.78 | 4.79 | 24.32 | 9.77 | 25.88 | 9.77 | 29.4
31.136 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0 | 19.24 | 4.79 | 13.38 | 4.59 | 4.88 | 4.59 | 9.86 | 10.06 | 9.86 | 21.58 | 9.77 | 21.97 | 0.1 | 15.04 | 0.1 | 14.36 | 0.1 | 0.49 | 4.69 | 4.79 | 4.69 | 4.79 | 21.58 | 4.79 | 24.71 | 9.77 | 25.2 | 9.86 | 25.2
31.168 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.63 | 4.79 | 13.87 | 4.49 | 4.88 | 4.59 | 9.86 | 10.16 | 9.86 | 21.68 | 9.77 | 22.27 | 0.1 | 15.82 | 0.1 | 15.04 | 0 | 1.46 | 4.69 | 4.79 | 6.15 | 4.79 | 21.58 | 4.79 | 24.41 | 9.77 | 24.71 | 9.86 | 25
31.2 | 0.1 | 1.07 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 20.12 | 4.79 | 14.36 | 4.49 | 5.57 | 4.59 | 9.86 | 9.96 | 9.86 | 21.58 | 9.67 | 22.75 | 0.1 | 16.6 | 0.1 | 15.72 | 0.1 | 3.03 | 4.59 | 4.88 | 6.05 | 4.79 | 21.39 | 4.79 | 24.51 | 9.77 | 24.51 | 9.86 | 24.61
31.232 | 0 | 0.88 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 20.61 | 4.79 | 14.84 | 4.49 | 6.54 | 4.59 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 23.15 | 0.1 | 17.19 | 0.1 | 16.41 | 0.1 | 5.08 | 4.59 | 4.88 | 5.57 | 4.88 | 21.29 | 4.79 | 24.41 | 9.77 | 24.51 | 9.77 | 24.51
31.264 | 0 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 20.9 | 4.79 | 15.33 | 4.59 | 7.91 | 4.49 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 23.44 | 0.1 | 17.87 | 0.1 | 17.09 | 0.1 | 6.35 | 4.59 | 4.88 | 5.18 | 4.79 | 21.39 | 4.79 | 24.32 | 9.77 | 24.41 | 9.86 | 24.41
31.296 | 0 | 0.59 | 0.1 | 14.16 | 0.1 | 14.36 | 0 | 17.58 | 4.79 | 15.92 | 4.59 | 8.89 | 4.49 | 10.64 | 9.86 | 9.86 | 21.49 | 9.77 | 23.83 | 0.1 | 18.46 | 0.1 | 17.68 | 0.1 | 7.23 | 4.59 | 4.79 | 5.27 | 4.79 | 21.29 | 4.88 | 24.41 | 9.77 | 24.32 | 9.86 | 24.32
31.328 | 0 | 0.68 | 0 | 4 | 0.1 | 14.36 | 0.1 | 17.68 | 4.79 | 16.41 | 4.59 | 9.67 | 4.49 | 11.52 | 9.86 | 9.86 | 21.39 | 9.77 | 24.12 | 0.1 | 19.14 | 0.1 | 18.36 | 0 | 8.01 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 4.79 | 24.32 | 9.77 | 24.41 | 9.86 | 24.22
31.36 | 0.1 | 0.49 | 0.1 | 1.86 | 0.1 | 5.37 | 0.1 | 17.87 | 4.79 | 16.9 | 4.59 | 10.45 | 4.49 | 12.99 | 9.77 | 9.86 | 21.39 | 9.77 | 21.19 | 0.1 | 19.82 | 0.1 | 18.95 | 0.1 | 8.79 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.32
31.392 | 0 | 0.49 | 0.1 | 1.76 | 0 | 1.76 | 0.1 | 18.16 | 4.79 | 17.29 | 4.59 | 11.04 | 4.49 | 13.87 | 9.77 | 9.86 | 21.39 | 9.77 | 20.9 | 0.1 | 20.31 | 0.1 | 19.63 | 0.1 | 9.57 | 4.59 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 24.32 | 9.77 | 24.22
31.424 | 0.1 | 0.49 | 0.1 | 1.56 | 0.1 | 1.95 | 0.1 | 18.07 | 4.88 | 17.68 | 4.69 | 11.62 | 4.49 | 14.55 | 9.77 | 9.86 | 21.39 | 9.77 | 21.19 | 0.1 | 20.9 | 0.1 | 20.22 | 0.1 | 10.35 | 4.59 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
31.456 | 0 | 0.29 | 0.1 | 1.46 | 0.1 | 1.66 | 0.1 | 18.16 | 4.88 | 18.26 | 4.69 | 12.21 | 4.49 | 15.14 | 9.67 | 9.86 | 12.79 | 9.86 | 21.39 | 0.1 | 21.58 | 0.1 | 20.8 | 0 | 11.23 | 4.59 | 4.79 | 4.98 | 4.79 | 10.45 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
31.488 | 0 | 0.29 | 0 | 1.27 | 0.1 | 1.37 | 0.1 | 18.16 | 4.88 | 18.65 | 4.69 | 12.79 | 4.59 | 15.63 | 9.67 | 9.86 | 6.93 | 9.86 | 21.39 | 0.1 | 22.07 | 0.1 | 21.49 | 0 | 12.01 | 4.59 | 4.79 | 4.88 | 4.79 | 3.03 | 4.88 | 20.8 | 9.77 | 24.22 | 9.86 | 24.12
31.52 | 0.78 | 0.2 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 18.16 | 4.88 | 19.14 | 4.69 | 13.28 | 4.59 | 16.11 | 9.67 | 9.86 | 8.11 | 9.77 | 21.49 | 0.1 | 21.78 | 0.1 | 21.97 | 0.1 | 12.7 | 4.59 | 4.79 | 4.79 | 4.79 | 3.13 | 4.88 | 7.72 | 9.77 | 18.75 | 9.86 | 24.12
31.552 | 1.66 | 0.2 | 0.1 | 0.98 | 0 | 1.27 | 0.1 | 18.16 | 4.79 | 19.43 | 4.69 | 13.87 | 4.59 | 16.7 | 9.67 | 9.86 | 11.04 | 9.77 | 21.49 | 0.1 | 18.75 | 0.1 | 22.66 | 0.1 | 13.48 | 4.59 | 5.86 | 4.69 | 4.79 | 3.61 | 4.79 | 6.54 | 9.77 | 7.23 | 9.86 | 13.77
31.584 | 2.93 | 0.2 | 0.1 | 0.98 | 0.1 | 1.07 | 0.1 | 18.16 | 4.88 | 19.92 | 4.79 | 14.36 | 4.59 | 17.19 | 9.67 | 9.86 | 11.23 | 9.86 | 21.49 | 0.1 | 18.26 | 0.1 | 19.53 | 0.1 | 14.26 | 4.69 | 7.03 | 4.69 | 4.79 | 4.79 | 4.88 | 10.45 | 9.77 | 6.93 | 9.86 | 6.35
31.616 | 4 | 0.1 | 0.1 | 0.88 | 0.1 | 0.98 | 0 | 18.07 | 4.88 | 20.31 | 4.79 | 14.94 | 4.69 | 17.58 | 9.67 | 9.86 | 10.35 | 9.86 | 21.39 | 0.1 | 17.97 | 0.1 | 18.26 | 0.1 | 14.94 | 4.69 | 8.98 | 4.69 | 4.79 | 6.15 | 4.88 | 12.11 | 9.77 | 11.04 | 9.86 | 7.72
31.648 | 4.69 | 0.1 | 0 | 0.88 | 0.1 | 0.88 | 0 | 18.07 | 4.88 | 20.9 | 4.79 | 15.33 | 4.69 | 18.07 | 9.57 | 9.86 | 9.77 | 9.86 | 21.49 | 0.1 | 17.87 | 0.1 | 18.07 | 0 | 15.72 | 4.69 | 10.55 | 4.69 | 4.79 | 6.15 | 4.88 | 10.74 | 9.77 | 12.21 | 9.86 | 11.43
31.68 | 5.18 | 0.1 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 18.07 | 4.88 | 19.34 | 4.79 | 15.92 | 4.69 | 18.56 | 9.57 | 9.86 | 9.96 | 9.77 | 21.39 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 16.31 | 4.69 | 11.91 | 4.59 | 4.79 | 5.66 | 4.88 | 10.16 | 9.77 | 10.74 | 9.86 | 11.82
31.712 | 5.76 | 0.1 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 18.07 | 4.88 | 17.68 | 4.79 | 16.41 | 4.69 | 18.95 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 16.99 | 4.69 | 12.89 | 4.59 | 4.79 | 5.37 | 4.79 | 10.16 | 9.77 | 9.86 | 9.86 | 10.45
31.744 | 6.35 | 0.1 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 5.96 | 4.88 | 17.97 | 4.88 | 16.9 | 4.69 | 19.43 | 9.57 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 17.58 | 4.79 | 13.67 | 4.59 | 4.79 | 5.27 | 4.88 | 10.45 | 9.77 | 10.06 | 9.86 | 9.96
31.776 | 7.03 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 2.93 | 4.88 | 18.26 | 4.88 | 17.29 | 4.69 | 19.92 | 9.57 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 18.26 | 4.79 | 14.55 | 4.59 | 4.79 | 5.08 | 4.88 | 10.45 | 9.77 | 10.55 | 9.86 | 10.25
31.808 | 7.62 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 0 | 3.52 | 4.88 | 18.16 | 4.88 | 17.77 | 4.79 | 20.31 | 9.57 | 9.86 | 9.86 | 9.86 | 11.13 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 18.95 | 4.69 | 15.23 | 4.59 | 4.79 | 5.08 | 4.88 | 10.16 | 9.77 | 10.55 | 9.86 | 10.45
31.84 | 8.2 | 0.1 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 18.16 | 4.79 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 6.45 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 19.53 | 4.79 | 16.02 | 4.49 | 5.86 | 5.08 | 4.88 | 10.16 | 9.77 | 10.35 | 9.86 | 10.45
31.872 | 8.79 | 0.1 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 5.66 | 4.88 | 18.07 | 4.79 | 18.65 | 4.79 | 21.09 | 9.67 | 9.86 | 9.77 | 9.86 | 8.98 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 20.22 | 4.79 | 16.6 | 4.49 | 6.93 | 4.98 | 4.79 | 10.16 | 9.77 | 9.96 | 9.86 | 10.16
31.904 | 9.38 | 0.1 | 0.68 | 0.2 | 0.1 | 0.29 | 0.1 | 5.86 | 4.88 | 18.07 | 4.88 | 19.14 | 4.79 | 21.58 | 9.67 | 9.86 | 9.77 | 9.86 | 11.43 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 20.8 | 4.79 | 17.38 | 4.59 | 9.08 | 4.98 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 9.96
31.936 | 10.06 | 0.1 | 1.37 | 0.2 | 0.1 | 0.2 | 0.1 | 5.47 | 4.88 | 18.16 | 4.88 | 19.53 | 4.79 | 21.97 | 9.67 | 10.64 | 9.77 | 9.86 | 10.94 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.39 | 4.79 | 18.07 | 4.59 | 10.55 | 4.88 | 4.79 | 9.96 | 9.77 | 10.06 | 9.86 | 10.16
31.968 | 10.55 | 0.1 | 2.54 | 0.2 | 0 | 0.2 | 0 | 5.08 | 4.88 | 18.16 | 4.88 | 20.02 | 4.79 | 22.27 | 9.67 | 11.13 | 9.67 | 9.86 | 10.16 | 0.1 | 8.5 | 0.1 | 17.68 | 0 | 21.97 | 4.79 | 18.56 | 4.59 | 11.82 | 4.88 | 4.88 | 10.06 | 9.77 | 10.16 | 9.86 | 9.96
32 | 11.13 | 0.1 | 3.81 | 0.1 | 0 | 0.2 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 20.41 | 4.79 | 22.66 | 9.67 | 12.21 | 9.77 | 9.86 | 9.86 | 0.1 | 1.86 | 0.1 | 14.26 | 0.1 | 22.46 | 4.88 | 19.24 | 4.59 | 12.79 | 4.79 | 4.88 | 9.86 | 9.77 | 10.06 | 9.86 | 9.96
32.032 | 11.72 | 0.1 | 4.49 | 0 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.8 | 4.79 | 23.15 | 9.67 | 13.38 | 9.67 | 9.86 | 9.96 | 0.1 | 2.34 | 0.1 | 3.22 | 0.1 | 17.97 | 4.88 | 19.92 | 4.59 | 13.67 | 4.69 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86
32.064 | 12.31 | 0.1 | 5.08 | 0 | 0.78 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 19.63 | 4.79 | 23.44 | 9.67 | 14.26 | 9.67 | 9.86 | 10.25 | 0.1 | 2.05 | 0.1 | 2.25 | 0.1 | 16.99 | 4.88 | 20.51 | 4.69 | 14.55 | 4.79 | 4.79 | 9.96 | 9.77 | 9.86 | 9.86 | 9.86
32.096 | 12.79 | 0.1 | 5.66 | 0.1 | 1.76 | 0 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 17.77 | 4.79 | 23.93 | 9.67 | 14.84 | 9.67 | 9.96 | 10.16 | 0.1 | 1.66 | 0.1 | 2.15 | 0.1 | 17.77 | 4.88 | 21.19 | 4.69 | 15.14 | 4.69 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
32.128 | 13.48 | 0.1 | 6.25 | 0.1 | 3.03 | 0 | 0.78 | 4.79 | 4.88 | 7.42 | 4.88 | 18.16 | 4.88 | 24.32 | 9.67 | 15.33 | 9.57 | 9.86 | 9.86 | 0.1 | 1.76 | 0.1 | 1.76 | 0 | 17.68 | 4.79 | 21.58 | 4.69 | 15.92 | 4.69 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77
32.16 | 13.87 | 0.1 | 6.84 | 0.1 | 4.2 | 0.1 | 1.56 | 4.79 | 4.88 | 2.73 | 4.88 | 18.26 | 4.79 | 22.27 | 9.77 | 15.92 | 9.57 | 9.86 | 9.86 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 17.68 | 4.88 | 22.17 | 4.69 | 16.6 | 4.59 | 5.76 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
32.192 | 14.45 | 0.1 | 7.42 | 0 | 4.69 | 0.1 | 2.73 | 4.69 | 4.88 | 3.32 | 4.88 | 18.07 | 4.79 | 21 | 9.77 | 16.31 | 9.57 | 9.86 | 9.96 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 17.68 | 4.88 | 22.85 | 4.69 | 17.38 | 4.59 | 7.13 | 9.67 | 9.77 | 9.77 | 11.04 | 9.77
32.224 | 14.84 | 0.1 | 8.01 | 0.1 | 5.27 | 0.1 | 3.91 | 4.69 | 4.88 | 4.39 | 4.88 | 18.16 | 4.79 | 21.39 | 9.77 | 16.9 | 9.57 | 9.96 | 9.86 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 17.77 | 4.88 | 23.44 | 4.79 | 18.07 | 4.59 | 8.98 | 9.67 | 10.74 | 9.77 | 11.91 | 9.77
32.256 | 15.43 | 0.1 | 8.69 | 0.1 | 5.86 | 0.1 | 4.59 | 4.59 | 4.88 | 5.57 | 4.88 | 18.16 | 4.79 | 21.49 | 9.77 | 17.38 | 9.57 | 9.96 | 9.67 | 0.1 | 1.17 | 0.1 | 1.27 | 0 | 17.77 | 4.88 | 23.93 | 4.79 | 18.65 | 4.59 | 10.74 | 9.67 | 11.82 | 9.67 | 13.67 | 9.77
32.288 | 15.92 | 0.1 | 9.18 | 0.1 | 6.45 | 0 | 5.18 | 4.59 | 4.88 | 5.96 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 17.77 | 9.57 | 9.96 | 9.77 | 0.1 | 1.07 | 0.1 | 1.27 | 0 | 17.68 | 4.88 | 24.41 | 4.79 | 19.24 | 4.59 | 11.91 | 9.67 | 13.48 | 9.67 | 15.43 | 9.67
32.32 | 16.41 | 0.1 | 9.86 | 0.1 | 7.03 | 0.1 | 5.76 | 4.59 | 4.79 | 5.57 | 4.88 | 18.16 | 4.88 | 21.49 | 9.77 | 18.26 | 9.57 | 9.86 | 9.77 | 0.1 | 0.98 | 0.1 | 1.07 | 0.1 | 17.77 | 4.88 | 25 | 4.79 | 19.92 | 4.59 | 12.89 | 9.67 | 15.23 | 9.67 | 16.6 | 9.67
32.352 | 16.9 | 0.1 | 10.45 | 0 | 7.62 | 0.1 | 6.35 | 4.59 | 4.79 | 5.18 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 18.75 | 9.57 | 9.86 | 9.67 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 17.68 | 4.88 | 25.39 | 4.79 | 20.41 | 4.59 | 13.77 | 9.67 | 16.5 | 9.67 | 17.38 | 9.67
32.384 | 16.9 | 0.1 | 10.94 | 0.1 | 8.3 | 0.1 | 6.93 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 19.24 | 9.57 | 9.96 | 9.67 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 25.98 | 4.79 | 21 | 4.59 | 14.55 | 9.67 | 17.29 | 9.67 | 18.26 | 9.67
32.416 | 14.55 | 0.1 | 11.62 | 0.1 | 8.89 | 0.1 | 7.52 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 19.53 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 25.2 | 4.79 | 21.49 | 4.59 | 15.23 | 9.67 | 18.07 | 9.67 | 18.95 | 9.67
32.448 | 14.75 | 0.1 | 12.11 | 0.1 | 9.47 | 0.1 | 8.11 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 20.02 | 9.67 | 9.96 | 9.67 | 0.1 | 0.59 | 0.1 | 0.78 | 0 | 17.48 | 4.88 | 22.17 | 4.79 | 22.17 | 4.59 | 15.92 | 9.57 | 18.65 | 9.67 | 19.53 | 9.57
32.48 | 14.75 | 0.1 | 12.7 | 0.1 | 10.06 | 0.1 | 8.69 | 4.59 | 4.79 | 4.98 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 20.51 | 9.67 | 10.55 | 9.67 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 4.59 | 4.88 | 21.88 | 4.79 | 22.66 | 4.59 | 16.7 | 9.67 | 19.43 | 9.67 | 20.22 | 9.67
32.512 | 14.55 | 0.1 | 13.18 | 0 | 10.55 | 0.1 | 9.38 | 4.49 | 4.79 | 4.88 | 4.88 | 7.23 | 4.88 | 21.49 | 9.77 | 20.9 | 9.67 | 11.33 | 9.57 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 1.86 | 4.88 | 21.58 | 4.79 | 23.24 | 4.69 | 17.38 | 9.67 | 20.12 | 9.67 | 20.7 | 9.57
32.544 | 14.55 | 0.1 | 13.67 | 0 | 11.23 | 0.1 | 9.96 | 4.59 | 4.88 | 4.69 | 4.88 | 2.73 | 4.79 | 21.49 | 9.86 | 21.19 | 9.67 | 12.6 | 9.67 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 2.34 | 4.88 | 21.49 | 4.88 | 23.83 | 4.69 | 17.97 | 9.67 | 20.61 | 9.67 | 21.29 | 9.67
32.576 | 14.45 | 0.1 | 14.36 | 0.1 | 11.82 | 0 | 10.55 | 4.59 | 4.88 | 4.79 | 4.88 | 3.32 | 4.79 | 21.39 | 9.77 | 21.68 | 9.67 | 13.67 | 9.57 | 0 | 0.39 | 0.1 | 0.59 | 0 | 1.66 | 4.88 | 21.29 | 4.88 | 24.32 | 4.69 | 18.65 | 9.67 | 21.29 | 9.67 | 21.97 | 9.67
32.608 | 14.45 | 0.1 | 14.84 | 0.1 | 12.31 | 0 | 11.13 | 4.59 | 4.79 | 4.69 | 4.88 | 4.2 | 4.79 | 11.52 | 9.77 | 22.07 | 9.67 | 14.36 | 9.57 | 0.1 | 0.29 | 0.1 | 0.49 | 0 | 1.66 | 4.88 | 21.39 | 4.88 | 24.81 | 4.69 | 19.34 | 9.67 | 21.97 | 9.67 | 22.56 | 9.67
32.64 | 14.45 | 0.1 | 15.33 | 0.1 | 12.89 | 0.1 | 11.72 | 4.59 | 4.79 | 4.59 | 4.88 | 5.37 | 4.79 | 6.54 | 9.86 | 22.46 | 9.67 | 15.04 | 9.57 | 0.1 | 0.2 | 0.1 | 0.39 | 0.1 | 1.56 | 4.88 | 21.29 | 4.79 | 25.39 | 4.69 | 19.82 | 9.67 | 22.46 | 9.67 | 23.05 | 9.67
32.672 | 14.36 | 0.1 | 15.82 | 0 | 13.38 | 0.1 | 12.31 | 4.59 | 4.79 | 4.59 | 4.88 | 5.86 | 4.79 | 8.98 | 9.86 | 22.85 | 9.67 | 15.63 | 9.57 | 0.1 | 0.2 | 1.17 | 0.29 | 0.1 | 1.46 | 4.88 | 21.29 | 4.88 | 25.78 | 4.69 | 20.51 | 9.67 | 23.05 | 9.67 | 23.54 | 9.67
32.704 | 14.36 | 0.1 | 16.31 | 0.1 | 13.97 | 0.1 | 12.79 | 4.69 | 4.88 | 4.59 | 4.88 | 5.66 | 4.79 | 11.52 | 9.86 | 23.24 | 9.67 | 16.11 | 9.57 | 0.2 | 0.2 | 2.15 | 0.29 | 0.1 | 1.27 | 4.88 | 21.29 | 4.88 | 24.32 | 4.79 | 21 | 9.67 | 23.54 | 9.67 | 24.22 | 9.67
32.736 | 14.36 | 0.1 | 16.8 | 0.1 | 14.45 | 0.1 | 13.38 | 4.69 | 4.88 | 4.59 | 4.88 | 5.18 | 4.79 | 10.94 | 9.77 | 23.54 | 9.67 | 16.6 | 9.57 | 1.17 | 0.2 | 3.91 | 0.2 | 0.1 | 1.27 | 4.88 | 21.19 | 4.88 | 22.36 | 4.79 | 21.58 | 9.77 | 24.12 | 9.67 | 24.61 | 9.67
32.768 | 1.95 | 0.1 | 14.75 | 0.1 | 16.99 | 0.1 | 16.02 | 4.79 | 9.77 | 4.59 | 4.88 | 4.79 | 4.79 | 10.16 | 9.86 | 21.49 | 9.77 | 18.95 | 9.67 | 7.52 | 0.1 | 9.08 | 0.1 | 4.49 | 0.88 | 4.88 | 2.34 | 4.79 | 21.29 | 4.79 | 24.41 | 9.77 | 26.56 | 9.77 | 27.05 | 9.77
32.8 | 1.76 | 0.1 | 14.75 | 0.1 | 16.41 | 0.1 | 16.5 | 4.79 | 10.45 | 4.59 | 4.88 | 4.79 | 4.79 | 9.96 | 9.86 | 21.58 | 9.77 | 19.34 | 9.67 | 8.4 | 0.1 | 9.96 | 0.1 | 5.96 | 0.68 | 4.88 | 3.22 | 4.88 | 21.29 | 4.79 | 24.9 | 9.77 | 27.05 | 9.77 | 27.64 | 9.77
32.832 | 1.46 | 0.1 | 14.55 | 0.1 | 14.06 | 0.1 | 16.9 | 4.79 | 11.04 | 4.59 | 4.88 | 4.69 | 4.79 | 9.86 | 9.86 | 21.58 | 9.77 | 19.82 | 9.67 | 9.08 | 0 | 10.74 | 0.1 | 6.93 | 0.68 | 4.88 | 3.91 | 4.88 | 21.29 | 4.79 | 25.39 | 9.77 | 27.54 | 9.77 | 28.13 | 9.77
32.864 | 1.37 | 0 | 14.45 | 0 | 14.55 | 0 | 16.6 | 4.79 | 11.62 | 4.59 | 4.88 | 4.59 | 4.79 | 9.86 | 9.77 | 21.49 | 9.77 | 20.12 | 9.67 | 10.06 | 0 | 11.43 | 0.1 | 7.81 | 0.59 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 25.88 | 9.77 | 28.22 | 9.77 | 28.61 | 9.67
32.896 | 1.17 | 0.1 | 14.55 | 0 | 14.65 | 0 | 14.65 | 4.79 | 12.21 | 4.59 | 4.88 | 4.59 | 4.79 | 9.86 | 9.86 | 21.49 | 9.77 | 20.61 | 9.67 | 10.74 | 0.1 | 12.31 | 0.1 | 8.59 | 0.49 | 4.88 | 6.35 | 4.88 | 21.19 | 4.88 | 25.2 | 9.77 | 28.52 | 9.77 | 29.1 | 9.77
32.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 14.84 | 4.88 | 12.79 | 4.69 | 4.88 | 4.59 | 4.79 | 9.86 | 9.86 | 21.49 | 9.77 | 21 | 9.67 | 11.52 | 0 | 12.99 | 0.1 | 9.38 | 0.49 | 4.88 | 5.96 | 4.88 | 21.19 | 4.79 | 21.68 | 9.86 | 28.91 | 9.77 | 29.49 | 9.77
32.96 | 0.98 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 14.75 | 4.79 | 13.38 | 4.69 | 4.88 | 4.59 | 4.79 | 9.77 | 9.86 | 21.49 | 9.77 | 21.39 | 9.67 | 12.21 | 0 | 13.77 | 0.1 | 10.16 | 0.39 | 4.88 | 5.66 | 4.88 | 21.19 | 4.88 | 21.58 | 9.77 | 29.49 | 9.77 | 25.98 | 9.77
32.992 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 14.65 | 4.88 | 13.87 | 4.69 | 4.98 | 4.59 | 4.79 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 9.67 | 12.99 | 0 | 14.55 | 0.1 | 10.94 | 0.29 | 4.88 | 5.18 | 4.88 | 21.19 | 4.79 | 21.49 | 9.77 | 26.27 | 9.77 | 25 | 9.77
33.024 | 0.78 | 0 | 14.36 | 0 | 14.45 | 0 | 14.55 | 4.79 | 14.36 | 4.69 | 5.47 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.77 | 22.17 | 9.67 | 13.77 | 0 | 15.33 | 0.1 | 11.72 | 0.2 | 4.88 | 5.08 | 4.88 | 7.91 | 4.79 | 21.29 | 9.77 | 25.2 | 9.86 | 24.71 | 9.77
33.056 | 0.68 | 0.1 | 14.36 | 0 | 14.45 | 0 | 14.55 | 4.79 | 14.94 | 4.69 | 6.54 | 4.59 | 4.79 | 9.67 | 9.77 | 21.49 | 9.77 | 22.56 | 9.77 | 14.36 | 0.1 | 16.02 | 0.1 | 12.5 | 0.2 | 4.88 | 5.08 | 4.88 | 2.44 | 4.88 | 21.29 | 9.86 | 24.81 | 9.77 | 24.61 | 9.77
33.088 | 0.59 | 0.1 | 9.38 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 15.43 | 4.79 | 7.81 | 4.49 | 4.79 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 9.77 | 15.14 | 0.1 | 16.7 | 0.1 | 13.28 | 0.2 | 5.08 | 5.08 | 4.88 | 3.32 | 4.88 | 21.29 | 9.86 | 24.61 | 9.86 | 24.51 | 9.77
33.12 | 0.59 | 0.1 | 2.64 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 15.92 | 4.79 | 8.89 | 4.59 | 4.79 | 9.67 | 9.86 | 19.04 | 9.86 | 23.34 | 9.77 | 15.82 | 0 | 17.29 | 0.1 | 13.87 | 0.2 | 6.05 | 5.08 | 4.79 | 3.81 | 4.88 | 21.29 | 9.86 | 24.61 | 9.86 | 24.41 | 9.86
33.152 | 0.49 | 0.1 | 1.95 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 16.41 | 4.79 | 9.67 | 4.49 | 4.79 | 9.67 | 9.86 | 8.2 | 9.77 | 23.73 | 9.77 | 16.5 | 0 | 18.07 | 0.1 | 14.75 | 0.1 | 7.62 | 4.98 | 4.88 | 4.98 | 4.88 | 21.29 | 9.86 | 24.41 | 9.86 | 24.32 | 9.86
33.184 | 0.29 | 0.1 | 1.76 | 0 | 14.36 | 0 | 14.45 | 4.79 | 16.9 | 4.79 | 10.35 | 4.49 | 5.47 | 9.67 | 9.77 | 7.32 | 9.86 | 24.12 | 9.77 | 17.19 | 0 | 18.56 | 0.1 | 15.43 | 0.1 | 9.96 | 4.88 | 4.88 | 6.15 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.32 | 9.77
33.216 | 0.29 | 0.1 | 1.46 | 0 | 14.36 | 0 | 14.45 | 4.79 | 17.29 | 4.79 | 11.04 | 4.59 | 5.96 | 9.57 | 9.77 | 10.35 | 9.77 | 24.51 | 9.77 | 17.97 | 0.1 | 19.24 | 0.1 | 15.92 | 0.1 | 11.23 | 4.88 | 4.88 | 6.05 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.77
33.248 | 0.29 | 0.1 | 1.46 | 0.1 | 7.13 | 0.1 | 14.45 | 4.88 | 17.77 | 4.79 | 11.52 | 4.59 | 7.03 | 9.67 | 9.86 | 11.52 | 9.86 | 22.36 | 9.77 | 18.56 | 0 | 19.92 | 0.1 | 16.7 | 0.1 | 12.4 | 4.79 | 4.79 | 5.66 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
33.28 | 0.29 | 0.1 | 1.27 | 0.1 | 1.95 | 0.1 | 14.45 | 4.88 | 18.16 | 4.79 | 12.11 | 4.59 | 8.4 | 9.67 | 9.86 | 10.64 | 9.86 | 22.17 | 9.77 | 19.24 | 0.1 | 20.61 | 0.1 | 17.38 | 0.1 | 13.18 | 4.79 | 4.79 | 5.27 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
33.312 | 0.2 | 0.1 | 1.17 | 0.1 | 1.95 | 0.1 | 7.72 | 4.88 | 18.75 | 4.88 | 12.79 | 4.59 | 9.28 | 9.67 | 9.86 | 9.96 | 9.86 | 21.78 | 9.77 | 19.82 | 0 | 21 | 0.1 | 18.07 | 0.1 | 14.06 | 4.69 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
33.344 | 0.1 | 0.1 | 0.98 | 0 | 1.76 | 0.1 | 1.95 | 4.88 | 19.14 | 4.79 | 13.18 | 4.69 | 9.96 | 9.57 | 9.77 | 9.96 | 9.86 | 21.68 | 9.77 | 20.51 | 0 | 21.68 | 0.1 | 18.75 | 0 | 14.84 | 4.69 | 4.88 | 5.08 | 4.79 | 7.72 | 9.77 | 24.12 | 9.86 | 24.22 | 9.86
33.376 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0 | 1.95 | 4.88 | 19.53 | 4.79 | 13.77 | 4.69 | 10.55 | 9.57 | 9.77 | 10.16 | 9.86 | 21.78 | 9.77 | 21 | 0.1 | 22.17 | 0.1 | 19.34 | 0 | 15.53 | 4.59 | 4.88 | 4.98 | 4.88 | 2.44 | 9.86 | 24.22 | 9.86 | 14.45 | 9.77
33.408 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 1.76 | 4.88 | 20.02 | 4.88 | 14.36 | 4.69 | 11.23 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 9.77 | 21.68 | 0.1 | 22.66 | 0.1 | 20.02 | 0.1 | 16.31 | 4.59 | 4.79 | 5.08 | 4.88 | 3.22 | 9.86 | 15.43 | 9.86 | 6.54 | 9.86
33.44 | 0.1 | 0.1 | 0.68 | 0.1 | 1.37 | 0.1 | 1.56 | 4.88 | 20.41 | 4.88 | 14.84 | 4.69 | 11.82 | 9.67 | 10.64 | 9.96 | 9.86 | 21.58 | 9.86 | 22.17 | 0.1 | 19.24 | 0.1 | 20.61 | 0 | 16.9 | 4.59 | 4.79 | 4.98 | 4.88 | 3.91 | 9.86 | 6.54 | 9.86 | 7.52 | 9.86
33.472 | 0 | 0.88 | 0.68 | 0.1 | 1.17 | 0.1 | 1.37 | 4.88 | 20.8 | 4.88 | 15.33 | 4.69 | 12.31 | 9.67 | 11.04 | 9.96 | 9.86 | 21.49 | 9.86 | 20.51 | 0 | 18.26 | 0.1 | 21.29 | 0 | 17.58 | 4.59 | 4.79 | 4.88 | 4.88 | 5.08 | 9.86 | 7.32 | 9.86 | 11.52 | 9.86
33.504 | 0 | 1.76 | 0.59 | 0 | 1.07 | 0.1 | 1.27 | 4.88 | 17.77 | 4.79 | 15.72 | 4.69 | 12.99 | 9.67 | 12.11 | 9.86 | 9.86 | 21.49 | 9.86 | 16.9 | 0 | 18.16 | 0 | 21.78 | 0.1 | 18.26 | 4.59 | 4.88 | 4.88 | 4.79 | 6.15 | 9.77 | 11.33 | 9.86 | 11.82 | 9.77
33.536 | 0 | 3.03 | 0.49 | 0.1 | 0.88 | 0 | 1.17 | 4.88 | 17.58 | 4.88 | 16.21 | 4.79 | 13.38 | 9.67 | 13.38 | 9.86 | 9.86 | 21.49 | 9.86 | 17.68 | 0 | 17.97 | 0.1 | 22.46 | 0 | 18.85 | 4.59 | 4.79 | 4.69 | 4.88 | 5.96 | 9.77 | 11.82 | 9.86 | 10.45 | 9.77
33.568 | 0.1 | 4 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 17.87 | 4.88 | 16.7 | 4.79 | 13.97 | 9.67 | 14.16 | 9.86 | 9.86 | 21.49 | 9.86 | 17.68 | 0 | 17.97 | 0.1 | 19.53 | 0.1 | 19.53 | 4.59 | 5.96 | 4.69 | 4.88 | 5.66 | 9.86 | 10.55 | 9.86 | 10.06 | 9.86
33.6 | 0.1 | 4.69 | 0.39 | 0.1 | 0.68 | 0.1 | 0.88 | 4.88 | 18.16 | 4.88 | 17.19 | 4.79 | 14.45 | 9.67 | 14.75 | 9.77 | 9.86 | 21.49 | 9.86 | 17.68 | 0 | 17.87 | 0.1 | 17.09 | 0.1 | 20.12 | 4.59 | 6.84 | 4.79 | 4.88 | 5.27 | 9.86 | 9.86 | 9.86 | 10.16 | 9.86
33.632 | 0.1 | 5.27 | 0.29 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 17.97 | 4.88 | 17.68 | 4.79 | 14.94 | 9.77 | 15.33 | 9.67 | 9.86 | 21.39 | 9.86 | 17.58 | 0 | 17.77 | 0.1 | 17.68 | 0.1 | 20.8 | 4.59 | 8.69 | 4.69 | 4.88 | 5.18 | 9.86 | 10.16 | 9.86 | 10.45 | 9.86
33.664 | 0 | 5.86 | 0.2 | 0 | 0.59 | 0.1 | 0.68 | 4.88 | 18.07 | 4.79 | 18.16 | 4.79 | 15.53 | 9.67 | 15.82 | 9.67 | 9.86 | 13.38 | 9.86 | 17.58 | 0 | 17.77 | 0 | 17.68 | 0 | 21.39 | 4.59 | 10.45 | 4.59 | 4.79 | 5.18 | 9.77 | 10.35 | 9.86 | 10.35 | 9.77
33.696 | 0 | 6.35 | 0.2 | 0 | 0.49 | 0 | 0.78 | 4.88 | 18.07 | 4.88 | 18.65 | 4.79 | 16.02 | 9.67 | 16.31 | 9.67 | 9.86 | 7.23 | 9.86 | 17.58 | 0 | 17.77 | 0.1 | 17.68 | 0 | 21.97 | 4.59 | 11.72 | 4.49 | 4.88 | 5.08 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86
33.728 | 0 | 7.03 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 16.41 | 9.77 | 16.8 | 9.67 | 9.86 | 8.11 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 22.56 | 4.59 | 12.79 | 4.59 | 4.88 | 5.18 | 9.86 | 10.16 | 9.86 | 10.16 | 9.86
33.76 | 0.1 | 7.72 | 0.2 | 0.1 | 0.39 | 0.1 | 0.59 | 4.88 | 18.07 | 4.88 | 19.53 | 4.88 | 16.99 | 9.77 | 17.29 | 9.67 | 9.86 | 11.04 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 23.05 | 4.59 | 13.57 | 4.59 | 4.88 | 4.98 | 9.86 | 10.16 | 9.86 | 10.25 | 9.86
33.792 | 0 | 8.3 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 17.38 | 9.77 | 17.68 | 9.67 | 9.86 | 11.23 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 23.54 | 4.69 | 14.36 | 4.59 | 4.79 | 4.98 | 9.86 | 10.06 | 9.86 | 10.06 | 9.86
33.824 | 0 | 8.89 | 0 | 0 | 0.2 | 0 | 0.39 | 4.79 | 18.07 | 4.88 | 20.22 | 4.79 | 17.87 | 9.77 | 18.26 | 9.67 | 9.86 | 10.45 | 9.86 | 17.48 | 0 | 17.68 | 0.1 | 17.68 | 0 | 24.22 | 4.69 | 15.04 | 4.59 | 4.79 | 4.88 | 9.77 | 9.96 | 9.86 | 9.96 | 9.77
33.856 | 0 | 9.47 | 0 | 0.2 | 0.2 | 0 | 0.29 | 4.88 | 18.07 | 4.88 | 20.7 | 4.88 | 18.26 | 9.77 | 18.65 | 9.67 | 9.86 | 9.86 | 9.86 | 17.58 | 0 | 8.98 | 0.1 | 17.68 | 0 | 24.61 | 4.69 | 15.82 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 9.96 | 9.77
33.888 | 0.1 | 10.06 | 0.1 | 1.17 | 0.2 | 0.2 | 0.29 | 4.88 | 18.07 | 4.88 | 20.31 | 4.88 | 18.75 | 9.77 | 19.24 | 9.67 | 9.86 | 10.06 | 9.96 | 17.48 | 0 | 2.25 | 0.1 | 17.68 | 0.1 | 25.2 | 4.69 | 16.6 | 4.59 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
33.92 | 0 | 10.64 | 0.1 | 2.25 | 0.1 | 1.17 | 0.2 | 4.88 | 18.07 | 4.88 | 18.26 | 4.88 | 19.24 | 9.77 | 19.53 | 9.67 | 9.86 | 10.25 | 9.96 | 7.13 | 0 | 2.44 | 0.1 | 17.68 | 0 | 25.59 | 4.69 | 17.29 | 4.49 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 9.96 | 9.86
33.952 | 0.1 | 11.23 | 0.1 | 3.61 | 0.1 | 2.05 | 0.2 | 4.88 | 6.35 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 20.02 | 9.57 | 9.86 | 10.16 | 9.96 | 1.86 | 0 | 2.05 | 0.1 | 17.68 | 0.1 | 26.17 | 4.79 | 17.87 | 4.49 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
33.984 | 0 | 11.82 | 0.1 | 4.49 | 0 | 3.52 | 0.1 | 4.79 | 2.54 | 4.88 | 18.36 | 4.88 | 20.12 | 9.77 | 20.41 | 9.57 | 9.86 | 9.96 | 9.86 | 2.15 | 0 | 1.76 | 0.1 | 17.68 | 0 | 21.49 | 4.69 | 18.46 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 9.77 | 9.77
34.016 | 0 | 12.31 | 0.1 | 4.98 | 0.1 | 4.3 | 0.1 | 4.79 | 3.42 | 4.88 | 18.16 | 4.88 | 20.51 | 9.77 | 20.8 | 9.67 | 10.64 | 9.86 | 9.86 | 1.95 | 0 | 1.86 | 0.1 | 6.74 | 0 | 20.31 | 4.69 | 19.14 | 4.59 | 5.76 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
34.048 | 0.1 | 12.99 | 0.1 | 5.47 | 0.1 | 4.98 | 0.1 | 4.88 | 4.39 | 4.88 | 18.16 | 4.88 | 20.9 | 9.86 | 21.19 | 9.67 | 11.33 | 9.86 | 9.96 | 1.66 | 0 | 1.66 | 0.1 | 1.66 | 0.1 | 21 | 4.79 | 19.82 | 4.59 | 6.74 | 4.59 | 10.84 | 9.77 | 9.86 | 9.77 | 9.86
34.08 | 0 | 13.48 | 0.1 | 6.05 | 0.1 | 5.47 | 0.1 | 4.88 | 5.47 | 4.88 | 18.16 | 4.88 | 17.87 | 9.86 | 21.68 | 9.67 | 12.6 | 9.86 | 9.96 | 1.66 | 0 | 1.46 | 0.1 | 2.44 | 0 | 21.19 | 4.79 | 20.41 | 4.69 | 8.5 | 4.59 | 11.43 | 9.77 | 10.74 | 9.77 | 9.86
34.112 | 0.1 | 13.97 | 0.1 | 6.64 | 0 | 5.96 | 0.1 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 18.46 | 9.86 | 22.07 | 9.67 | 13.67 | 9.77 | 9.96 | 1.56 | 0 | 1.46 | 0.1 | 2.15 | 0.1 | 21.29 | 4.79 | 21 | 4.69 | 10.35 | 4.59 | 12.99 | 9.67 | 11.82 | 9.67 | 11.04
34.144 | 0 | 14.55 | 0 | 7.23 | 0 | 6.64 | 0 | 4.79 | 5.47 | 4.88 | 18.07 | 4.79 | 18.16 | 9.86 | 22.56 | 9.67 | 14.45 | 9.77 | 9.86 | 1.37 | 0 | 1.17 | 0.1 | 1.86 | 0 | 21.29 | 4.79 | 21.58 | 4.69 | 11.62 | 4.49 | 15.14 | 9.67 | 13.48 | 9.67 | 12.11
34.176 | 0 | 14.94 | 0.1 | 7.91 | 0.1 | 7.13 | 0.1 | 4.79 | 5.18 | 4.88 | 18.07 | 4.88 | 18.07 | 9.77 | 22.85 | 9.67 | 14.94 | 9.77 | 9.86 | 1.37 | 0 | 1.17 | 0.1 | 1.66 | 0 | 21.29 | 4.79 | 22.17 | 4.69 | 12.7 | 4.49 | 16.31 | 9.67 | 15.23 | 9.67 | 13.77
34.208 | 0 | 15.43 | 0.1 | 8.5 | 0.1 | 7.81 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.86 | 23.24 | 9.67 | 15.53 | 9.77 | 9.96 | 1.17 | 0 | 1.07 | 0.1 | 1.56 | 0.1 | 21.39 | 4.79 | 22.75 | 4.69 | 13.57 | 4.49 | 17.19 | 9.67 | 16.5 | 9.67 | 15.33
34.24 | 0 | 16.02 | 0.1 | 9.08 | 0.1 | 8.4 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 18.07 | 9.86 | 23.54 | 9.67 | 16.02 | 9.67 | 9.96 | 1.07 | 0 | 0.98 | 0.1 | 1.46 | 0.1 | 21.29 | 4.79 | 23.24 | 4.79 | 14.36 | 4.59 | 17.97 | 9.67 | 17.38 | 9.67 | 16.5
34.272 | 0 | 16.5 | 0.1 | 9.67 | 0 | 9.08 | 0.1 | 5.08 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 9.86 | 23.93 | 9.67 | 16.5 | 9.67 | 9.96 | 0.98 | 0 | 0.88 | 0.1 | 1.27 | 0 | 21.29 | 4.88 | 23.83 | 4.79 | 15.14 | 4.59 | 18.65 | 9.67 | 18.16 | 9.67 | 17.48
34.304 | 0 | 16.99 | 0.1 | 10.25 | 0 | 9.67 | 0.1 | 6.05 | 4.79 | 4.79 | 18.07 | 4.79 | 18.07 | 9.86 | 24.32 | 9.67 | 17.09 | 9.67 | 9.96 | 0.88 | 0 | 0.68 | 0 | 1.17 | 0 | 21.29 | 4.79 | 24.32 | 4.79 | 15.63 | 4.59 | 19.34 | 9.57 | 18.95 | 9.67 | 18.26
34.336 | 0 | 16.7 | 0.1 | 10.74 | 0.1 | 10.25 | 0.1 | 6.93 | 4.79 | 4.88 | 8.3 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 17.58 | 9.67 | 9.96 | 0.88 | 0 | 0.68 | 0.1 | 1.17 | 0 | 21.29 | 4.79 | 24.81 | 4.79 | 16.5 | 4.59 | 20.02 | 9.57 | 19.43 | 9.67 | 18.85
34.368 | 0.1 | 14.75 | 0.1 | 11.43 | 0.1 | 10.74 | 0.1 | 8.3 | 4.69 | 4.88 | 2.83 | 4.88 | 18.07 | 9.86 | 21.78 | 9.77 | 17.97 | 9.67 | 10.74 | 0.78 | 0 | 0.68 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.39 | 4.79 | 17.09 | 4.59 | 20.7 | 9.67 | 20.12 | 9.67 | 19.53
34.4 | 0 | 14.65 | 0.1 | 12.01 | 0.1 | 11.43 | 0.1 | 9.28 | 4.79 | 4.88 | 3.32 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 18.46 | 9.67 | 11.72 | 0.68 | 0 | 0.59 | 0.1 | 1.07 | 0.1 | 21.19 | 4.88 | 25.78 | 4.79 | 17.77 | 4.69 | 21.29 | 9.57 | 20.8 | 9.67 | 20.22
34.432 | 0 | 14.75 | 0.1 | 12.6 | 0 | 12.01 | 0.1 | 9.96 | 4.59 | 4.88 | 4.3 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 18.85 | 9.57 | 13.18 | 0.59 | 0 | 0.49 | 0.1 | 0.88 | 0 | 5.96 | 4.88 | 26.37 | 4.79 | 18.46 | 4.69 | 21.88 | 9.57 | 21.29 | 9.67 | 20.7
34.464 | 0 | 14.45 | 0 | 13.18 | 0 | 12.5 | 0.1 | 10.55 | 4.59 | 4.79 | 5.47 | 4.79 | 17.97 | 9.86 | 21.58 | 9.77 | 19.34 | 9.57 | 13.97 | 0.49 | 0 | 0.39 | 0.1 | 0.78 | 0 | 2.83 | 4.88 | 23.05 | 4.79 | 19.14 | 4.69 | 22.46 | 9.57 | 21.97 | 9.67 | 21.39
34.496 | 0 | 14.55 | 0.1 | 13.57 | 0.1 | 13.09 | 0.1 | 11.13 | 4.59 | 4.79 | 5.96 | 4.79 | 15.43 | 9.77 | 21.49 | 9.77 | 19.82 | 9.67 | 14.65 | 0.49 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 3.32 | 4.79 | 22.27 | 4.79 | 19.82 | 4.69 | 22.95 | 9.57 | 22.46 | 9.67 | 21.97
34.528 | 0 | 14.55 | 0.1 | 14.16 | 0.1 | 13.57 | 0.1 | 11.82 | 4.59 | 4.88 | 5.57 | 4.88 | 4.39 | 9.86 | 21.49 | 9.77 | 20.22 | 9.67 | 15.23 | 0.39 | 1.07 | 0.39 | 0.1 | 0.59 | 0.1 | 4.3 | 4.88 | 21.68 | 4.79 | 20.31 | 4.79 | 23.44 | 9.67 | 23.05 | 9.67 | 22.56
34.56 | 0 | 14.45 | 0.1 | 14.65 | 0.1 | 14.06 | 0.1 | 12.31 | 4.59 | 4.88 | 5.27 | 4.88 | 3.03 | 9.86 | 21.49 | 9.77 | 20.61 | 9.67 | 15.72 | 0.39 | 2.05 | 0.29 | 0.68 | 0.49 | 0 | 5.76 | 4.88 | 21.58 | 4.79 | 20.9 | 4.79 | 24.02 | 9.67 | 23.54 | 9.67 | 23.15
34.592 | 0 | 14.45 | 0.1 | 15.23 | 0 | 14.65 | 0.1 | 12.89 | 4.49 | 4.88 | 4.98 | 4.88 | 3.52 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 16.31 | 0.29 | 4 | 0.2 | 1.66 | 0.49 | 0.1 | 6.35 | 4.88 | 21.58 | 4.79 | 21.49 | 4.79 | 24.51 | 9.67 | 24.02 | 9.67 | 23.63
34.624 | 0 | 14.45 | 0 | 15.72 | 0 | 15.14 | 0 | 13.48 | 4.49 | 4.79 | 4.88 | 4.79 | 4.79 | 9.86 | 21.39 | 9.77 | 21.49 | 9.67 | 16.8 | 0.2 | 5.37 | 0.2 | 3.32 | 0.29 | 0 | 6.05 | 4.88 | 21.49 | 4.79 | 22.07 | 4.79 | 25.1 | 9.67 | 24.61 | 9.67 | 24.22
34.656 | 0 | 14.45 | 0.1 | 16.11 | 0.1 | 15.63 | 0.1 | 13.97 | 4.49 | 4.88 | 4.98 | 4.88 | 5.86 | 9.86 | 21.39 | 9.86 | 21.78 | 9.67 | 17.29 | 0.2 | 6.74 | 0.2 | 5.27 | 0.39 | 0 | 5.57 | 4.79 | 21.29 | 4.79 | 22.56 | 4.79 | 25.68 | 9.67 | 25 | 9.77 | 24.71
34.688 | 0 | 14.45 | 0.1 | 16.7 | 0.1 | 16.11 | 0.1 | 14.55 | 4.59 | 5.57 | 4.88 | 4.88 | 5.76 | 9.86 | 21.39 | 9.86 | 22.17 | 9.67 | 17.68 | 0.2 | 7.42 | 0.2 | 6.45 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.88 | 23.24 | 4.79 | 26.08 | 9.67 | 25.59 | 9.77 | 25.2
34.72 | 0 | 14.45 | 0.1 | 17.09 | 0.1 | 16.6 | 0.1 | 15.04 | 4.49 | 6.54 | 4.88 | 4.88 | 5.27 | 9.86 | 21.39 | 9.86 | 22.56 | 9.67 | 18.16 | 0.2 | 8.2 | 0.2 | 7.23 | 0.29 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 23.73 | 4.79 | 26.47 | 9.67 | 26.08 | 9.77 | 25.68
34.752 | 0 | 14.45 | 0.1 | 13.77 | 0 | 17.09 | 0.1 | 15.53 | 4.49 | 7.81 | 4.79 | 4.88 | 5.08 | 9.86 | 18.85 | 9.86 | 22.95 | 9.67 | 18.56 | 0.2 | 8.98 | 0.1 | 8.01 | 0.2 | 0.1 | 5.18 | 4.88 | 21.29 | 4.88 | 24.22 | 4.79 | 27.05 | 9.67 | 26.47 | 9.77 | 26.17
34.784 | 0 | 6.74 | 0 | 14.26 | 0 | 13.67 | 0.1 | 16.02 | 4.49 | 8.89 | 4.69 | 4.79 | 4.88 | 9.86 | 8.11 | 9.86 | 23.34 | 9.67 | 19.04 | 0.1 | 9.77 | 0.1 | 8.79 | 0.1 | 0 | 5.18 | 4.79 | 21.29 | 4.79 | 24.71 | 4.79 | 27.54 | 9.67 | 27.05 | 9.77 | 26.56
34.816 | 0 | 2.05 | 0 | 14.26 | 0.1 | 13.97 | 0.1 | 16.5 | 4.59 | 9.67 | 4.69 | 4.88 | 4.88 | 9.86 | 7.32 | 9.86 | 23.63 | 9.67 | 19.43 | 0.1 | 10.64 | 0.1 | 9.67 | 0.1 | 0 | 5.08 | 4.79 | 21.19 | 4.79 | 25.2 | 4.79 | 28.13 | 9.67 | 27.54 | 9.77 | 27.05
34.848 | 0 | 2.05 | 0.1 | 14.45 | 0.1 | 14.16 | 0.1 | 16.99 | 4.59 | 10.45 | 4.59 | 4.88 | 4.88 | 9.86 | 10.25 | 9.86 | 24.12 | 9.77 | 19.92 | 0.1 | 11.52 | 0.1 | 10.35 | 0.1 | 0.1 | 4.98 | 4.88 | 21.19 | 4.88 | 25.78 | 4.79 | 28.52 | 9.77 | 28.03 | 9.77 | 27.54
34.88 | 0 | 1.86 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 17.38 | 4.59 | 11.04 | 4.59 | 4.79 | 4.79 | 9.86 | 11.52 | 9.86 | 24.41 | 9.67 | 20.31 | 0.1 | 12.21 | 0.1 | 11.23 | 0.1 | 0 | 4.98 | 4.88 | 11.04 | 4.79 | 26.27 | 4.79 | 28.91 | 9.77 | 28.52 | 9.77 | 28.03
34.912 | 0 | 1.37 | 0.1 | 14.45 | 0 | 14.36 | 0.1 | 17.97 | 4.59 | 11.62 | 4.59 | 4.79 | 4.79 | 9.86 | 10.64 | 9.86 | 21.58 | 9.67 | 20.7 | 0.1 | 12.99 | 0.1 | 12.01 | 0.1 | 0.1 | 4.79 | 4.88 | 3.03 | 4.79 | 23.24 | 4.79 | 29.49 | 9.77 | 28.91 | 9.77 | 28.52
34.944 | 0 | 1.37 | 0 | 14.45 | 0 | 14.36 | 0 | 18.36 | 4.69 | 12.21 | 4.49 | 4.79 | 4.69 | 9.86 | 10.06 | 9.86 | 21.88 | 9.77 | 21.19 | 0.1 | 13.77 | 0.1 | 12.79 | 0.1 | 0 | 4.79 | 4.79 | 3.13 | 4.79 | 22.27 | 4.79 | 25.49 | 9.77 | 29.49 | 9.77 | 29.01
34.976 | 0 | 1.27 | 0 | 14.45 | 0.1 | 14.45 | 0.1 | 18.75 | 4.69 | 12.79 | 4.49 | 4.79 | 4.69 | 9.86 | 9.96 | 9.86 | 21.78 | 9.77 | 21.58 | 0.1 | 14.45 | 0.1 | 13.48 | 0.1 | 0 | 4.69 | 4.79 | 3.52 | 4.79 | 21.78 | 4.79 | 24.41 | 9.77 | 26.08 | 9.86 | 29.4
35.008 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.24 | 4.69 | 13.28 | 4.49 | 4.79 | 4.59 | 9.86 | 10.16 | 9.86 | 21.68 | 9.77 | 21.97 | 0.1 | 15.23 | 0.1 | 14.26 | 0.1 | 0.39 | 4.69 | 4.88 | 4.79 | 4.88 | 21.58 | 4.88 | 24.51 | 9.77 | 25.2 | 9.86 | 25.78
35.04 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 19.73 | 4.79 | 13.77 | 4.49 | 4.88 | 4.59 | 9.86 | 10.16 | 9.86 | 21.68 | 9.77 | 22.36 | 0.1 | 15.82 | 0.1 | 14.94 | 0.1 | 1.46 | 4.69 | 4.88 | 5.96 | 4.88 | 21.58 | 4.88 | 24.41 | 9.77 | 24.81 | 9.86 | 25
35.072 | 0 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 20.12 | 4.69 | 14.36 | 4.49 | 5.47 | 4.59 | 9.86 | 10.06 | 9.86 | 21.58 | 9.77 | 22.75 | 0.1 | 16.6 | 0.1 | 15.72 | 0.1 | 3.03 | 4.59 | 4.88 | 6.05 | 4.88 | 21.49 | 4.88 | 24.51 | 9.77 | 24.61 | 9.86 | 24.61
35.104 | 0 | 0.78 | 0 | 14.36 | 0 | 14.45 | 0 | 20.51 | 4.79 | 14.84 | 4.49 | 6.45 | 4.49 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 23.15 | 0.1 | 17.29 | 0.1 | 16.41 | 0 | 5.18 | 4.59 | 4.88 | 5.57 | 4.79 | 21.39 | 4.79 | 24.41 | 9.77 | 24.61 | 9.86 | 24.41
35.136 | 0 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 20.9 | 4.79 | 15.33 | 4.59 | 7.72 | 4.49 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 23.44 | 0.1 | 17.87 | 0.1 | 17.09 | 0.1 | 6.25 | 4.59 | 4.79 | 5.27 | 4.79 | 21.39 | 4.79 | 24.32 | 9.77 | 24.41 | 9.86 | 24.41
35.168 | 0 | 0.68 | 0.1 | 14.16 | 0.1 | 14.36 | 0.1 | 17.58 | 4.79 | 15.82 | 4.59 | 8.79 | 4.49 | 10.64 | 9.96 | 9.86 | 21.49 | 9.77 | 23.83 | 0.1 | 18.56 | 0.1 | 17.68 | 0.1 | 7.13 | 4.59 | 4.88 | 5.27 | 4.88 | 21.29 | 4.79 | 24.41 | 9.77 | 24.32 | 9.86 | 24.41
35.2 | 0.1 | 0.59 | 0.1 | 4.39 | 0.1 | 14.45 | 0.1 | 17.77 | 4.79 | 16.41 | 4.59 | 9.67 | 4.49 | 11.72 | 9.77 | 9.86 | 21.39 | 9.77 | 24.22 | 0.1 | 19.24 | 0.1 | 18.46 | 0.1 | 8.01 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
35.232 | 0 | 0.59 | 0.1 | 1.76 | 0.1 | 5.37 | 0.1 | 17.97 | 4.79 | 16.8 | 4.59 | 10.25 | 4.49 | 13.09 | 9.77 | 9.86 | 21.39 | 9.77 | 21.19 | 0.1 | 19.73 | 0.1 | 19.14 | 0.1 | 8.79 | 4.59 | 4.88 | 5.08 | 4.88 | 21.29 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
35.264 | 0 | 0.39 | 0.1 | 1.86 | 0 | 1.86 | 0.1 | 18.16 | 4.79 | 17.29 | 4.59 | 10.94 | 4.49 | 13.97 | 9.77 | 9.86 | 21.39 | 9.77 | 20.9 | 0.1 | 20.31 | 0.1 | 19.53 | 0 | 9.57 | 4.59 | 4.79 | 5.08 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
35.296 | 0 | 0.39 | 0.1 | 1.46 | 0.1 | 1.95 | 0.1 | 18.07 | 4.88 | 17.77 | 4.69 | 11.52 | 4.49 | 14.55 | 9.67 | 9.86 | 21.39 | 9.86 | 21.19 | 0.1 | 21.09 | 0.1 | 20.22 | 0 | 10.25 | 4.59 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
35.328 | 0 | 0.39 | 0.1 | 1.56 | 0.1 | 1.76 | 0.1 | 18.16 | 4.79 | 18.16 | 4.69 | 12.11 | 4.49 | 15.23 | 9.67 | 9.86 | 13.18 | 9.86 | 21.49 | 0.1 | 21.58 | 0.1 | 20.9 | 0.1 | 11.13 | 4.59 | 4.79 | 4.98 | 4.88 | 13.28 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.12
35.36 | 0 | 0.29 | 0.1 | 1.27 | 0 | 1.46 | 0.1 | 18.16 | 4.88 | 18.65 | 4.69 | 12.7 | 4.59 | 15.72 | 9.67 | 9.86 | 6.93 | 9.86 | 21.39 | 0.1 | 22.07 | 0.1 | 21.39 | 0.1 | 11.91 | 4.59 | 4.88 | 4.88 | 4.79 | 3.22 | 4.88 | 21 | 9.77 | 24.22 | 9.86 | 24.12
35.392 | 0.78 | 0.29 | 0.1 | 1.17 | 0 | 1.37 | 0.1 | 18.16 | 4.79 | 19.14 | 4.69 | 13.18 | 4.59 | 16.21 | 9.67 | 9.86 | 8.11 | 9.86 | 21.39 | 0.1 | 21.97 | 0.1 | 21.97 | 0.1 | 12.6 | 4.59 | 4.88 | 4.88 | 4.79 | 3.22 | 4.88 | 7.81 | 9.77 | 16.9 | 9.86 | 24.12
35.424 | 1.66 | 0.2 | 0 | 0.98 | 0 | 1.07 | 0 | 18.07 | 4.79 | 19.53 | 4.69 | 13.67 | 4.59 | 16.8 | 9.67 | 9.86 | 11.04 | 9.86 | 21.49 | 0.1 | 19.04 | 0.1 | 22.66 | 0 | 13.38 | 4.59 | 5.76 | 4.69 | 4.79 | 3.42 | 4.79 | 6.74 | 9.77 | 6.54 | 9.86 | 13.28
35.456 | 2.93 | 0.1 | 0 | 0.98 | 0.1 | 1.07 | 0.1 | 18.16 | 4.88 | 20.02 | 4.79 | 14.26 | 4.59 | 17.19 | 9.57 | 9.86 | 11.23 | 9.86 | 21.49 | 0.1 | 18.46 | 0.1 | 19.14 | 0.1 | 14.16 | 4.69 | 7.03 | 4.69 | 4.79 | 4.39 | 4.88 | 10.45 | 9.77 | 7.03 | 9.86 | 6.25
35.488 | 3.91 | 0.2 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 18.07 | 4.88 | 20.31 | 4.79 | 14.84 | 4.59 | 17.68 | 9.67 | 9.86 | 10.35 | 9.86 | 21.49 | 0.1 | 18.07 | 0.1 | 17.68 | 0.1 | 14.94 | 4.69 | 8.98 | 4.69 | 4.79 | 5.96 | 4.88 | 12.31 | 9.77 | 11.13 | 9.86 | 7.52
35.52 | 4.69 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 18.07 | 4.88 | 20.8 | 4.79 | 15.23 | 4.69 | 18.16 | 9.57 | 9.86 | 9.96 | 9.86 | 21.49 | 0.1 | 17.97 | 0.1 | 17.87 | 0.1 | 15.63 | 4.69 | 10.55 | 4.59 | 4.79 | 6.25 | 4.88 | 11.04 | 9.77 | 12.01 | 9.86 | 11.43
35.552 | 5.18 | 0.1 | 0.1 | 0.68 | 0 | 0.88 | 0.1 | 18.07 | 4.79 | 19.53 | 4.79 | 15.72 | 4.69 | 18.65 | 9.67 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.77 | 0.1 | 16.31 | 4.79 | 11.91 | 4.59 | 4.79 | 5.66 | 4.88 | 9.86 | 9.77 | 10.64 | 9.86 | 11.82
35.584 | 5.76 | 0 | 0 | 0.59 | 0 | 0.68 | 0 | 17.38 | 4.79 | 17.68 | 4.79 | 16.31 | 4.69 | 19.04 | 9.57 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.87 | 0 | 16.99 | 4.79 | 12.89 | 4.59 | 4.79 | 5.47 | 4.79 | 10.16 | 9.77 | 10.06 | 9.86 | 10.45
35.616 | 6.35 | 0 | 0 | 0.49 | 0.1 | 0.59 | 0.1 | 5.08 | 4.88 | 18.07 | 4.88 | 16.7 | 4.69 | 19.43 | 9.57 | 9.86 | 10.06 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 17.58 | 4.79 | 13.67 | 4.59 | 4.79 | 5.27 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 9.86
35.648 | 6.93 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.83 | 4.88 | 18.26 | 4.88 | 17.19 | 4.79 | 19.92 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 18.16 | 4.79 | 14.45 | 4.49 | 4.79 | 5.18 | 4.88 | 10.45 | 9.77 | 10.35 | 9.86 | 10.16
35.68 | 7.52 | 0.1 | 0.1 | 0.39 | 0.1 | 0.59 | 0.1 | 3.61 | 4.88 | 18.16 | 4.79 | 17.68 | 4.79 | 20.31 | 9.67 | 9.86 | 9.86 | 9.86 | 11.04 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 18.95 | 4.88 | 15.14 | 4.59 | 4.79 | 5.18 | 4.88 | 10.25 | 9.77 | 10.45 | 9.86 | 10.45
35.712 | 8.11 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 0.1 | 4.49 | 4.79 | 18.16 | 4.88 | 18.07 | 4.79 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 6.64 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 19.63 | 4.88 | 15.92 | 4.49 | 5.86 | 5.08 | 4.88 | 10.16 | 9.77 | 10.25 | 9.86 | 10.45
35.744 | 8.79 | 0.1 | 0 | 0.29 | 0 | 0.39 | 0.1 | 5.66 | 4.79 | 18.07 | 4.79 | 18.65 | 4.79 | 21.19 | 9.67 | 9.86 | 9.96 | 9.86 | 9.28 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 20.12 | 4.79 | 16.7 | 4.59 | 7.13 | 4.98 | 4.79 | 10.16 | 9.77 | 9.96 | 9.86 | 10.06
35.776 | 9.38 | 0.1 | 0.68 | 0.29 | 0.1 | 0.29 | 0.1 | 5.86 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 21.49 | 9.57 | 9.86 | 9.77 | 9.96 | 11.43 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 20.8 | 4.88 | 17.29 | 4.59 | 9.08 | 4.98 | 4.79 | 10.16 | 9.77 | 10.06 | 9.86 | 10.06
35.808 | 9.96 | 0.1 | 1.37 | 0.2 | 0.1 | 0.29 | 0.1 | 5.47 | 4.88 | 18.16 | 4.88 | 19.34 | 4.79 | 21.97 | 9.67 | 10.16 | 9.67 | 9.96 | 11.04 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.39 | 4.88 | 17.97 | 4.49 | 10.64 | 4.88 | 4.88 | 10.06 | 9.77 | 10.16 | 9.86 | 10.16
35.84 | 10.45 | 0.1 | 2.54 | 0.2 | 0.1 | 0.2 | 0.1 | 5.18 | 4.88 | 18.16 | 4.88 | 19.92 | 4.79 | 22.36 | 9.67 | 11.04 | 9.77 | 9.96 | 10.25 | 0.1 | 8.89 | 0.1 | 17.68 | 0.1 | 21.97 | 4.88 | 18.65 | 4.59 | 11.91 | 4.88 | 4.88 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16
35.872 | 11.13 | 0.1 | 3.81 | 0.1 | 0 | 0.2 | 0.1 | 4.98 | 4.79 | 18.16 | 4.88 | 20.22 | 4.79 | 22.75 | 9.67 | 11.91 | 9.67 | 9.86 | 9.77 | 0.1 | 1.95 | 0.1 | 15.53 | 0.1 | 22.46 | 4.88 | 19.24 | 4.59 | 12.79 | 4.79 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
35.904 | 11.72 | 0.1 | 4.49 | 0.1 | 0 | 0.2 | 0.1 | 4.88 | 4.79 | 18.07 | 4.88 | 20.7 | 4.79 | 23.15 | 9.67 | 13.28 | 9.67 | 9.86 | 9.96 | 0 | 2.34 | 0.1 | 3.32 | 0 | 18.75 | 4.88 | 19.82 | 4.59 | 13.77 | 4.79 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.77
35.936 | 12.31 | 0.1 | 5.08 | 0 | 0.78 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 23.54 | 9.67 | 14.16 | 9.67 | 9.96 | 10.16 | 0.1 | 2.05 | 0.1 | 2.15 | 0 | 16.99 | 4.88 | 20.51 | 4.59 | 14.36 | 4.69 | 4.88 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
35.968 | 12.79 | 0.1 | 5.57 | 0.1 | 1.56 | 0.1 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 18.46 | 4.88 | 23.93 | 9.67 | 14.75 | 9.67 | 9.96 | 10.16 | 0.1 | 1.66 | 0.1 | 2.25 | 0.1 | 17.58 | 4.88 | 21 | 4.59 | 15.23 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86
36 | 13.48 | 0.1 | 6.15 | 0.1 | 2.83 | 0.1 | 0.78 | 4.79 | 4.88 | 7.72 | 4.88 | 18.36 | 4.79 | 24.22 | 9.67 | 15.33 | 9.57 | 9.96 | 9.86 | 0.1 | 1.76 | 0.1 | 1.86 | 0.1 | 17.68 | 4.88 | 21.58 | 4.69 | 15.92 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77
36.032 | 13.87 | 0.1 | 6.84 | 0.1 | 4 | 0.1 | 1.56 | 4.69 | 4.79 | 2.73 | 4.88 | 18.36 | 4.79 | 21.19 | 9.77 | 15.82 | 9.67 | 9.86 | 9.96 | 0.1 | 1.46 | 0.1 | 1.66 | 0.1 | 17.68 | 4.88 | 22.27 | 4.69 | 16.6 | 4.59 | 5.96 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77
36.064 | 14.36 | 0.1 | 7.42 | 0 | 4.59 | 0 | 2.93 | 4.69 | 4.79 | 3.42 | 4.88 | 18.07 | 4.79 | 20.9 | 9.67 | 16.31 | 9.57 | 9.96 | 9.96 | 0 | 1.37 | 0.1 | 1.56 | 0 | 17.77 | 4.88 | 22.85 | 4.69 | 17.29 | 4.59 | 7.03 | 9.67 | 9.77 | 9.77 | 10.84 | 9.77
36.096 | 14.94 | 0.1 | 8.01 | 0.1 | 5.18 | 0 | 3.91 | 4.59 | 4.88 | 4.3 | 4.88 | 18.16 | 4.79 | 21.19 | 9.67 | 16.8 | 9.57 | 9.96 | 9.86 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 17.68 | 4.88 | 23.24 | 4.69 | 17.97 | 4.59 | 8.89 | 9.67 | 10.64 | 9.67 | 11.91 | 9.67
36.128 | 15.43 | 0.1 | 8.69 | 0.1 | 5.76 | 0.1 | 4.59 | 4.59 | 4.79 | 5.47 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 17.29 | 9.57 | 9.96 | 9.77 | 0.1 | 1.17 | 0.1 | 1.46 | 0.1 | 17.77 | 4.88 | 23.93 | 4.69 | 18.56 | 4.59 | 10.55 | 9.67 | 11.62 | 9.67 | 13.48 | 9.77
36.16 | 15.92 | 0.1 | 9.28 | 0.1 | 6.35 | 0.1 | 5.18 | 4.59 | 4.88 | 5.96 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 17.77 | 9.57 | 9.96 | 9.77 | 0.1 | 1.07 | 0.1 | 1.37 | 0.1 | 17.68 | 4.88 | 24.41 | 4.79 | 19.24 | 4.49 | 11.91 | 9.67 | 13.28 | 9.67 | 15.23 | 9.67
36.192 | 16.41 | 0.1 | 9.86 | 0.1 | 6.93 | 0.1 | 5.76 | 4.59 | 4.79 | 5.66 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 18.26 | 9.57 | 9.96 | 9.77 | 0.1 | 0.98 | 0.1 | 1.07 | 0.1 | 17.68 | 4.88 | 25 | 4.79 | 19.82 | 4.59 | 12.89 | 9.67 | 15.14 | 9.67 | 16.5 | 9.67
36.224 | 16.9 | 0.1 | 10.45 | 0 | 7.52 | 0.1 | 6.35 | 4.49 | 4.79 | 5.18 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 18.65 | 9.57 | 9.96 | 9.67 | 0.1 | 0.88 | 0.1 | 0.98 | 0 | 17.68 | 4.88 | 25.49 | 4.79 | 20.41 | 4.59 | 13.77 | 9.67 | 16.41 | 9.67 | 17.29 | 9.67
36.256 | 16.9 | 0.1 | 11.04 | 0 | 8.11 | 0.1 | 6.93 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 19.14 | 9.57 | 9.96 | 9.67 | 0.1 | 0.88 | 0.1 | 0.88 | 0 | 17.68 | 4.88 | 25.98 | 4.79 | 21 | 4.59 | 14.55 | 9.67 | 17.29 | 9.67 | 18.07 | 9.67
36.288 | 14.94 | 0.1 | 11.52 | 0.1 | 8.69 | 0.1 | 7.52 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 19.43 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 25.78 | 4.79 | 21.68 | 4.59 | 15.23 | 9.67 | 18.07 | 9.67 | 18.75 | 9.67
36.32 | 14.84 | 0.1 | 12.21 | 0.1 | 9.38 | 0.1 | 8.2 | 4.59 | 4.79 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 19.92 | 9.67 | 9.96 | 9.67 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 17.48 | 4.88 | 22.27 | 4.79 | 22.27 | 4.59 | 15.92 | 9.67 | 18.65 | 9.67 | 19.43 | 9.67
36.352 | 14.75 | 0.1 | 12.7 | 0.1 | 9.96 | 0.1 | 8.69 | 4.59 | 4.79 | 4.88 | 4.88 | 17.97 | 4.88 | 21.49 | 9.77 | 20.41 | 9.67 | 10.64 | 9.67 | 0.1 | 0.49 | 0.1 | 0.78 | 0.1 | 4.49 | 4.88 | 21.88 | 4.79 | 22.75 | 4.69 | 16.7 | 9.67 | 19.34 | 9.67 | 20.22 | 9.67
36.384 | 14.45 | 0.1 | 13.28 | 0 | 10.45 | 0.1 | 9.38 | 4.59 | 4.79 | 4.79 | 4.88 | 7.81 | 4.88 | 21.49 | 9.77 | 20.8 | 9.67 | 11.33 | 9.57 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 2.05 | 4.88 | 21.49 | 4.79 | 23.34 | 4.69 | 17.38 | 9.67 | 20.12 | 9.67 | 20.7 | 9.67
36.416 | 14.55 | 0.1 | 13.77 | 0.1 | 11.13 | 0.1 | 9.96 | 4.59 | 4.88 | 4.69 | 4.88 | 2.83 | 4.79 | 21.39 | 9.77 | 21.19 | 9.67 | 12.6 | 9.57 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 2.15 | 4.88 | 21.49 | 4.79 | 23.83 | 4.69 | 17.87 | 9.67 | 20.61 | 9.57 | 21.29 | 9.67
36.448 | 14.45 | 0.1 | 14.26 | 0.1 | 11.72 | 0.1 | 10.55 | 4.59 | 4.79 | 4.79 | 4.88 | 3.22 | 4.88 | 21.49 | 9.77 | 21.58 | 9.67 | 13.67 | 9.57 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 1.86 | 4.88 | 21.49 | 4.79 | 24.32 | 4.69 | 18.56 | 9.67 | 21.29 | 9.67 | 21.97 | 9.67
36.48 | 14.45 | 0.1 | 14.84 | 0.1 | 12.21 | 0.1 | 11.13 | 4.59 | 4.79 | 4.69 | 4.88 | 4.2 | 4.88 | 11.13 | 9.77 | 22.07 | 9.67 | 14.45 | 9.57 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 1.86 | 4.88 | 21.39 | 4.79 | 24.9 | 4.69 | 19.24 | 9.67 | 21.78 | 9.67 | 22.46 | 9.67
36.512 | 14.45 | 0.1 | 15.33 | 0.1 | 12.79 | 0.1 | 11.72 | 4.69 | 4.79 | 4.59 | 4.88 | 5.47 | 4.88 | 6.64 | 9.77 | 22.36 | 9.67 | 15.04 | 9.57 | 0.1 | 0.2 | 0.1 | 0.39 | 0.1 | 1.56 | 4.88 | 21.29 | 4.79 | 25.39 | 4.79 | 19.92 | 9.67 | 22.36 | 9.67 | 23.05 | 9.67
36.544 | 14.36 | 0.1 | 15.82 | 0 | 13.38 | 0.1 | 12.31 | 4.69 | 4.79 | 4.59 | 4.88 | 5.96 | 4.79 | 9.77 | 9.77 | 22.75 | 9.67 | 15.63 | 9.57 | 0 | 0.2 | 0.88 | 0.29 | 0 | 1.37 | 4.88 | 21.29 | 4.88 | 25.98 | 4.79 | 20.51 | 9.67 | 23.05 | 9.67 | 23.54 | 9.67
36.576 | 14.36 | 0.1 | 16.21 | 0.1 | 13.97 | 0 | 12.79 | 4.69 | 4.79 | 4.59 | 4.88 | 5.57 | 4.79 | 11.62 | 9.77 | 23.15 | 9.67 | 16.11 | 9.57 | 0.2 | 0.2 | 1.95 | 0.29 | 0 | 1.27 | 4.88 | 21.29 | 4.88 | 25.1 | 4.79 | 21 | 9.67 | 23.34 | 9.67 | 24.22 | 9.67
36.608 | 14.36 | 0.1 | 16.8 | 0.1 | 14.45 | 0.1 | 13.38 | 4.69 | 4.88 | 4.59 | 4.88 | 5.27 | 4.88 | 10.94 | 9.86 | 23.54 | 9.67 | 16.6 | 9.67 | 1.07 | 0.2 | 3.71 | 0.2 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 22.17 | 4.79 | 21.58 | 9.77 | 24.02 | 9.67 | 24.61 | 9.67
36.64 | 14.36 | 0.1 | 17.29 | 0.1 | 15.04 | 0.1 | 13.97 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 4.88 | 10.16 | 9.77 | 23.93 | 9.67 | 17.09 | 9.57 | 2.25 | 0.1 | 5.37 | 0.2 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 21.78 | 4.79 | 22.17 | 9.77 | 24.51 | 9.67 | 25.2 | 9.67
36.672 | 14.36 | 0.1 | 14.94 | 0.1 | 15.43 | 0.1 | 14.36 | 4.79 | 6.54 | 4.59 | 4.88 | 4.98 | 4.88 | 9.86 | 9.77 | 24.32 | 9.67 | 17.58 | 9.67 | 4.1 | 0.1 | 6.64 | 0.2 | 0.1 | 1.07 | 4.88 | 21.19 | 4.88 | 21.58 | 4.79 | 22.75 | 9.77 | 25.1 | 9.67 | 25.68 | 9.67
36.704 | 7.23 | 0.1 | 15.33 | 0.1 | 15.82 | 0.1 | 14.94 | 4.79 | 8.01 | 4.59 | 4.88 | 4.88 | 4.79 | 10.16 | 9.77 | 20.9 | 9.77 | 17.97 | 9.67 | 5.66 | 0.1 | 7.32 | 0.1 | 1.07 | 0.98 | 4.88 | 21.19 | 4.88 | 21.39 | 4.79 | 23.24 | 9.77 | 25.59 | 9.77 | 26.27 | 9.67
36.736 | 2.15 | 0.1 | 14.75 | 0.1 | 16.41 | 0 | 15.43 | 4.79 | 8.98 | 4.49 | 4.88 | 4.88 | 4.88 | 10.25 | 9.77 | 21.09 | 9.67 | 18.46 | 9.67 | 6.64 | 0.1 | 8.2 | 0.1 | 2.34 | 0.78 | 4.88 | 9.28 | 4.88 | 21.39 | 4.79 | 23.83 | 9.77 | 26.08 | 9.77 | 26.56 | 9.67
36.768 | 1.86 | 0.1 | 14.65 | 0.1 | 16.9 | 0.1 | 15.92 | 4.79 | 9.77 | 4.59 | 4.88 | 4.79 | 4.79 | 10.25 | 9.86 | 21.29 | 9.77 | 18.95 | 9.67 | 7.42 | 0.1 | 8.98 | 0.1 | 4.2 | 0.78 | 4.88 | 2.54 | 4.79 | 21.39 | 4.79 | 24.41 | 9.77 | 26.66 | 9.77 | 27.05 | 9.67
36.8 | 1.66 | 0.1 | 14.65 | 0.1 | 17.29 | 0.1 | 16.41 | 4.79 | 10.45 | 4.59 | 4.88 | 4.79 | 4.88 | 9.96 | 9.86 | 21.39 | 9.77 | 19.34 | 9.67 | 8.2 | 0.1 | 9.77 | 0.1 | 5.66 | 0.68 | 4.88 | 3.22 | 4.79 | 21.29 | 4.79 | 24.9 | 9.77 | 27.05 | 9.77 | 27.64 | 9.77
36.832 | 1.37 | 0.1 | 14.55 | 0.1 | 14.55 | 0.1 | 16.9 | 4.79 | 11.04 | 4.59 | 4.88 | 4.69 | 4.88 | 9.96 | 9.86 | 21.49 | 9.77 | 19.82 | 9.67 | 9.18 | 0 | 10.55 | 0.1 | 6.84 | 0.59 | 4.88 | 3.71 | 4.88 | 21.29 | 4.79 | 25.39 | 9.77 | 27.54 | 9.77 | 28.13 | 9.77
36.864 | 1.46 | 0.1 | 14.55 | 0 | 15.04 | 0.1 | 16.99 | 4.79 | 11.52 | 4.59 | 4.88 | 4.59 | 4.79 | 9.96 | 9.77 | 21.39 | 9.77 | 20.12 | 9.67 | 9.86 | 0 | 11.33 | 0.1 | 7.72 | 0.49 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 25.88 | 9.77 | 28.03 | 9.77 | 28.42 | 9.67
36.896 | 1.27 | 0.1 | 14.55 | 0 | 14.84 | 0 | 15.04 | 4.79 | 12.21 | 4.59 | 4.88 | 4.59 | 4.79 | 9.77 | 9.77 | 21.49 | 9.77 | 20.7 | 9.67 | 10.55 | 0.1 | 12.21 | 0.1 | 8.4 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.79 | 25.88 | 9.77 | 28.52 | 9.77 | 28.91 | 9.67
36.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 15.14 | 4.88 | 12.7 | 4.69 | 4.88 | 4.59 | 4.88 | 9.86 | 9.86 | 21.49 | 9.77 | 21 | 9.67 | 11.43 | 0 | 12.89 | 0.1 | 9.28 | 0.49 | 4.88 | 6.25 | 4.79 | 21.19 | 4.79 | 22.36 | 9.86 | 29.01 | 9.77 | 29.49 | 9.77
36.96 | 0.88 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 14.84 | 4.88 | 13.28 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.39 | 9.77 | 21.39 | 9.67 | 12.21 | 0 | 13.67 | 0.1 | 10.06 | 0.39 | 4.88 | 5.66 | 4.79 | 21.19 | 4.79 | 21.97 | 9.77 | 29.49 | 9.77 | 26.17 | 9.77
36.992 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 14.65 | 4.88 | 13.87 | 4.69 | 4.88 | 4.59 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 9.67 | 12.99 | 0 | 14.36 | 0.1 | 10.84 | 0.29 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 21.68 | 9.77 | 25.88 | 9.77 | 25.2 | 9.77
37.024 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 14.26 | 4.69 | 5.57 | 4.49 | 4.79 | 9.77 | 9.77 | 21.39 | 9.77 | 22.17 | 9.67 | 13.67 | 0 | 15.04 | 0.1 | 11.72 | 0.2 | 4.88 | 5.08 | 4.88 | 8.5 | 4.79 | 21.49 | 9.77 | 25.2 | 9.77 | 24.71 | 9.77
37.056 | 0.68 | 0.1 | 14.36 | 0 | 14.55 | 0.1 | 14.45 | 4.88 | 14.84 | 4.69 | 6.25 | 4.59 | 4.79 | 9.67 | 9.77 | 21.39 | 9.77 | 22.56 | 9.67 | 14.36 | 0 | 15.92 | 0.1 | 12.4 | 0.2 | 4.88 | 5.18 | 4.88 | 2.44 | 4.88 | 21.49 | 9.77 | 24.71 | 9.77 | 24.51 | 9.77
37.088 | 0.59 | 0.1 | 10.94 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 15.33 | 4.79 | 7.62 | 4.49 | 4.88 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 9.67 | 15.14 | 0.1 | 16.6 | 0.1 | 13.09 | 0.2 | 5.08 | 5.08 | 4.79 | 3.22 | 4.88 | 21.39 | 9.86 | 24.51 | 9.86 | 24.51 | 9.77
37.12 | 0.59 | 0.1 | 2.93 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 15.82 | 4.79 | 8.69 | 4.49 | 4.88 | 9.67 | 9.86 | 21.19 | 9.77 | 23.34 | 9.77 | 15.92 | 0.1 | 17.29 | 0.1 | 13.87 | 0.2 | 6.05 | 5.08 | 4.79 | 3.71 | 4.88 | 21.29 | 9.86 | 24.51 | 9.86 | 24.41 | 9.77
37.152 | 0.49 | 0.1 | 2.05 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 16.31 | 4.79 | 9.57 | 4.49 | 4.88 | 9.67 | 9.86 | 8.69 | 9.86 | 23.73 | 9.77 | 16.5 | 0 | 17.87 | 0.1 | 14.55 | 0.1 | 7.62 | 4.98 | 4.88 | 4.98 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.32 | 9.77
37.184 | 0.39 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 16.7 | 4.79 | 10.25 | 4.49 | 5.66 | 9.67 | 9.77 | 6.93 | 9.86 | 24.12 | 9.77 | 17.09 | 0 | 18.46 | 0 | 15.43 | 0.1 | 9.96 | 4.88 | 4.88 | 6.15 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.32 | 9.77
37.216 | 0.29 | 0.1 | 1.46 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.29 | 4.79 | 10.94 | 4.59 | 5.96 | 9.67 | 9.77 | 9.96 | 9.77 | 24.51 | 9.77 | 17.77 | 0 | 19.14 | 0.1 | 16.11 | 0.1 | 11.23 | 4.79 | 4.88 | 6.05 | 4.88 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
37.248 | 0.29 | 0.1 | 1.46 | 0.1 | 8.01 | 0.1 | 14.45 | 4.88 | 17.68 | 4.88 | 11.52 | 4.59 | 6.84 | 9.67 | 9.86 | 11.72 | 9.86 | 22.46 | 9.77 | 18.46 | 0 | 19.73 | 0.1 | 16.8 | 0.1 | 12.31 | 4.79 | 4.79 | 5.57 | 4.88 | 21.29 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
37.28 | 0.2 | 0.1 | 1.37 | 0.1 | 2.25 | 0.1 | 14.36 | 4.88 | 18.16 | 4.88 | 12.01 | 4.59 | 8.3 | 9.67 | 9.86 | 10.84 | 9.86 | 22.27 | 9.77 | 19.24 | 0 | 20.31 | 0.1 | 17.48 | 0.1 | 13.18 | 4.79 | 4.79 | 5.27 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.77
37.312 | 0.2 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 9.08 | 4.88 | 18.56 | 4.88 | 12.7 | 4.59 | 9.28 | 9.67 | 9.86 | 9.96 | 9.86 | 21.88 | 9.77 | 19.73 | 0 | 20.9 | 0.1 | 18.07 | 0.1 | 13.97 | 4.69 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
37.344 | 0.1 | 0 | 1.07 | 0 | 1.86 | 0 | 3.91 | 4.88 | 19.14 | 4.79 | 13.18 | 4.59 | 9.96 | 9.67 | 9.77 | 9.86 | 9.86 | 21.78 | 9.77 | 20.31 | 0 | 21.58 | 0.1 | 18.65 | 0 | 14.84 | 4.59 | 4.88 | 5.08 | 4.79 | 8.69 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77
37.376 | 0.1 | 0.1 | 0.88 | 0 | 1.37 | 0.1 | 3.71 | 4.88 | 19.53 | 4.88 | 13.67 | 4.69 | 10.45 | 9.67 | 9.77 | 10.16 | 9.86 | 21.78 | 9.77 | 20.9 | 0 | 22.17 | 0.1 | 19.24 | 0 | 15.53 | 4.59 | 4.88 | 5.08 | 4.88 | 2.54 | 9.86 | 24.22 | 9.86 | 16.41 | 9.77
37.408 | 0.1 | 0.1 | 0.88 | 0.1 | 1.37 | 0.1 | 4.39 | 4.88 | 19.92 | 4.88 | 14.16 | 4.69 | 11.13 | 9.67 | 9.86 | 10.35 | 9.86 | 21.58 | 9.77 | 21.58 | 0 | 22.75 | 0.1 | 20.02 | 0.1 | 16.21 | 4.59 | 4.88 | 4.98 | 4.88 | 3.22 | 9.86 | 18.65 | 9.86 | 6.74 | 9.86
37.44 | 0.1 | 0.1 | 0.78 | 0.1 | 1.17 | 0.1 | 4.98 | 4.88 | 20.31 | 4.88 | 14.75 | 4.69 | 11.72 | 9.67 | 10.06 | 10.06 | 9.86 | 21.58 | 9.77 | 22.17 | 0 | 19.73 | 0.1 | 20.61 | 0.1 | 16.8 | 4.59 | 4.88 | 4.98 | 4.88 | 3.71 | 9.86 | 7.13 | 9.86 | 7.03 | 9.86
37.472 | 0 | 0.78 | 0.78 | 0.1 | 1.07 | 0.1 | 5.27 | 4.88 | 20.8 | 4.88 | 15.23 | 4.79 | 12.31 | 9.67 | 10.94 | 9.86 | 9.86 | 21.49 | 9.77 | 21.78 | 0 | 18.65 | 0.1 | 21.19 | 0.1 | 17.58 | 4.59 | 4.88 | 4.88 | 4.88 | 4.98 | 9.86 | 6.84 | 9.86 | 11.43 | 9.86
37.504 | 0 | 1.76 | 0.59 | 0 | 0.88 | 0.1 | 4.98 | 4.88 | 17.68 | 4.88 | 15.72 | 4.79 | 12.89 | 9.67 | 11.91 | 9.86 | 9.86 | 21.49 | 9.77 | 17.29 | 0 | 18.26 | 0 | 21.68 | 0.1 | 18.16 | 4.59 | 4.88 | 4.88 | 4.79 | 6.15 | 9.77 | 11.13 | 9.86 | 11.91 | 9.77
37.536 | 0 | 3.03 | 0.49 | 0 | 0.88 | 0.1 | 4.88 | 4.88 | 17.58 | 4.88 | 16.31 | 4.79 | 13.38 | 9.67 | 13.18 | 9.86 | 9.86 | 21.49 | 9.77 | 17.97 | 0 | 18.07 | 0.1 | 22.27 | 0 | 18.85 | 4.59 | 4.88 | 4.79 | 4.88 | 5.96 | 9.86 | 12.11 | 9.86 | 10.55 | 9.86
37.568 | 0 | 4 | 0.49 | 0.1 | 0.78 | 0.1 | 4.79 | 4.88 | 17.77 | 4.88 | 16.7 | 4.79 | 13.87 | 9.77 | 14.16 | 9.86 | 9.86 | 21.49 | 9.86 | 17.87 | 0 | 18.07 | 0.1 | 21.58 | 0.1 | 19.53 | 4.59 | 5.86 | 4.69 | 4.88 | 5.57 | 9.86 | 10.55 | 9.86 | 9.96 | 9.86
37.6 | 0.1 | 4.69 | 0.49 | 0.1 | 0.68 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 17.19 | 4.79 | 14.45 | 9.77 | 14.75 | 9.77 | 9.86 | 21.49 | 9.86 | 17.77 | 0 | 17.97 | 0.1 | 18.07 | 0.1 | 20.12 | 4.49 | 6.84 | 4.69 | 4.88 | 5.27 | 9.86 | 9.96 | 9.86 | 10.06 | 9.86
37.632 | 0.1 | 5.27 | 0.29 | 0.1 | 0.68 | 0.1 | 4.79 | 4.88 | 17.97 | 4.88 | 17.58 | 4.79 | 14.94 | 9.77 | 15.23 | 9.77 | 9.86 | 21.49 | 9.86 | 17.68 | 0 | 17.77 | 0.1 | 18.07 | 0.1 | 20.7 | 4.49 | 8.5 | 4.69 | 4.79 | 5.18 | 9.86 | 10.35 | 9.86 | 10.35 | 9.86
37.664 | 0 | 5.86 | 0.29 | 0.1 | 0.49 | 0.1 | 4.59 | 4.88 | 17.97 | 4.88 | 18.07 | 4.79 | 15.53 | 9.77 | 15.72 | 9.67 | 9.86 | 13.97 | 9.86 | 17.58 | 0 | 17.77 | 0 | 17.87 | 0 | 21.39 | 4.49 | 10.35 | 4.59 | 4.79 | 5.08 | 9.77 | 10.35 | 9.86 | 10.45 | 9.77
37.696 | 0.1 | 6.45 | 0.2 | 0.1 | 0.59 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 18.46 | 4.79 | 15.92 | 9.77 | 16.31 | 9.67 | 9.86 | 7.42 | 9.77 | 17.58 | 0 | 17.77 | 0.1 | 17.87 | 0 | 21.97 | 4.59 | 11.72 | 4.59 | 4.88 | 4.98 | 9.86 | 10.45 | 9.86 | 10.25 | 9.77
37.728 | 0 | 7.03 | 0.2 | 0.1 | 0.49 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 18.95 | 4.88 | 16.41 | 9.77 | 16.8 | 9.67 | 9.86 | 8.01 | 9.86 | 17.58 | 0 | 17.77 | 0.1 | 17.87 | 0.1 | 22.46 | 4.59 | 12.7 | 4.59 | 4.79 | 5.08 | 9.86 | 10.16 | 9.86 | 10.06 | 9.86
37.76 | 0 | 7.62 | 0.2 | 0.1 | 0.39 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 19.43 | 4.88 | 16.9 | 9.77 | 17.29 | 9.67 | 9.86 | 10.84 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.87 | 0.1 | 23.05 | 4.59 | 13.48 | 4.49 | 4.79 | 4.98 | 9.86 | 10.06 | 9.86 | 10.16 | 9.86
37.792 | 0 | 8.2 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 17.97 | 4.88 | 19.82 | 4.88 | 17.38 | 9.77 | 17.68 | 9.67 | 9.86 | 11.43 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 23.54 | 4.69 | 14.36 | 4.59 | 4.79 | 5.08 | 9.86 | 10.16 | 9.86 | 10.16 | 9.86
37.824 | 0 | 8.89 | 0.1 | 0 | 0.2 | 0.1 | 4.49 | 4.88 | 17.97 | 4.88 | 20.22 | 4.79 | 17.77 | 9.77 | 18.16 | 9.67 | 9.86 | 10.45 | 9.86 | 17.58 | 0 | 17.68 | 0 | 17.77 | 0 | 24.12 | 4.69 | 15.14 | 4.49 | 4.79 | 4.88 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
37.856 | 0 | 9.47 | 0 | 0.2 | 0.2 | 0 | 4.49 | 4.88 | 18.07 | 4.88 | 20.7 | 4.88 | 18.16 | 9.77 | 18.65 | 9.67 | 9.86 | 9.86 | 9.86 | 17.58 | 0 | 10.55 | 0.1 | 17.77 | 0 | 24.61 | 4.69 | 15.92 | 4.49 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.77
37.888 | 0.1 | 10.06 | 0.1 | 1.17 | 0.2 | 0.2 | 4.49 | 4.88 | 17.97 | 4.88 | 21.09 | 4.88 | 18.75 | 9.86 | 19.04 | 9.57 | 9.86 | 9.96 | 9.86 | 17.58 | 0 | 2.64 | 0.1 | 17.68 | 0.1 | 25.1 | 4.69 | 16.5 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
37.92 | 0 | 10.64 | 0.1 | 2.15 | 0.1 | 1.17 | 4.49 | 4.88 | 17.97 | 4.88 | 18.95 | 4.88 | 19.24 | 9.86 | 19.53 | 9.57 | 9.86 | 10.16 | 9.86 | 8.01 | 0 | 2.44 | 0.1 | 17.68 | 0 | 25.49 | 4.69 | 17.19 | 4.59 | 4.79 | 4.79 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
37.952 | 0 | 11.23 | 0.1 | 3.52 | 0 | 1.95 | 4.49 | 4.88 | 6.54 | 4.88 | 18.46 | 4.88 | 19.53 | 9.86 | 19.92 | 9.57 | 9.86 | 10.16 | 9.86 | 1.95 | 0 | 1.95 | 0.1 | 17.68 | 0 | 26.08 | 4.69 | 17.77 | 4.49 | 4.79 | 4.69 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
37.984 | 0 | 11.82 | 0.1 | 4.39 | 0 | 3.42 | 4.49 | 4.88 | 2.44 | 4.88 | 18.46 | 4.88 | 20.12 | 9.86 | 20.41 | 9.57 | 9.86 | 9.96 | 9.86 | 2.25 | 0 | 1.66 | 0.1 | 17.68 | 0 | 21.58 | 4.69 | 18.46 | 4.59 | 4.88 | 4.69 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77
38.016 | 0 | 12.31 | 0 | 4.98 | 0.1 | 4.2 | 4.49 | 4.88 | 3.32 | 4.88 | 18.36 | 4.88 | 20.51 | 9.86 | 20.8 | 9.67 | 10.55 | 9.96 | 9.86 | 1.86 | 0.1 | 1.86 | 0.1 | 8.01 | 0 | 20.31 | 4.69 | 19.14 | 4.49 | 5.76 | 4.59 | 9.86 | 9.77 | 9.86 | 9.77 | 9.77
38.048 | 0 | 12.89 | 0.1 | 5.47 | 0 | 4.88 | 4.59 | 4.88 | 4.3 | 4.88 | 18.26 | 4.88 | 20.9 | 9.86 | 21.19 | 9.67 | 11.33 | 9.96 | 9.86 | 1.76 | 0 | 1.66 | 0.1 | 1.95 | 0.1 | 21 | 4.79 | 19.82 | 4.59 | 6.54 | 4.59 | 10.16 | 9.77 | 9.86 | 9.77 | 9.86
38.08 | 0 | 13.38 | 0.1 | 6.05 | 0.1 | 5.37 | 4.59 | 4.88 | 5.47 | 4.88 | 18.16 | 4.88 | 17.58 | 9.86 | 21.58 | 9.67 | 12.5 | 9.86 | 9.86 | 1.66 | 0 | 1.56 | 0.1 | 2.34 | 0.1 | 21.19 | 4.79 | 20.22 | 4.69 | 8.2 | 4.59 | 11.23 | 9.77 | 10.74 | 9.77 | 9.86
38.112 | 0 | 13.97 | 0.1 | 6.64 | 0 | 5.96 | 4.59 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 17.77 | 9.86 | 21.97 | 9.67 | 13.57 | 9.86 | 9.86 | 1.46 | 0 | 1.27 | 0.1 | 1.95 | 0 | 21.29 | 4.79 | 20.9 | 4.69 | 10.25 | 4.59 | 12.79 | 9.77 | 11.72 | 9.67 | 10.74
38.144 | 0 | 14.45 | 0 | 7.23 | 0 | 6.54 | 4.59 | 4.88 | 5.47 | 4.79 | 18.16 | 4.88 | 17.97 | 9.86 | 22.46 | 9.67 | 14.36 | 9.77 | 9.86 | 1.46 | 0 | 1.27 | 0.1 | 1.76 | 0 | 21.29 | 4.79 | 21.58 | 4.69 | 11.52 | 4.59 | 14.84 | 9.67 | 13.48 | 9.67 | 11.91
38.176 | 0 | 14.94 | 0 | 7.91 | 0.1 | 7.13 | 4.69 | 4.88 | 5.18 | 4.79 | 18.07 | 4.88 | 17.97 | 9.86 | 22.75 | 9.67 | 14.84 | 9.77 | 9.86 | 1.27 | 0 | 1.17 | 0.1 | 1.66 | 0 | 21.29 | 4.79 | 22.17 | 4.69 | 12.6 | 4.49 | 16.21 | 9.67 | 15.14 | 9.67 | 13.38
38.208 | 0.1 | 15.53 | 0.1 | 8.5 | 0.1 | 7.72 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.07 | 9.86 | 23.05 | 9.67 | 15.53 | 9.77 | 9.86 | 1.27 | 0 | 1.07 | 0.1 | 1.46 | 0.1 | 21.29 | 4.79 | 22.75 | 4.69 | 13.38 | 4.59 | 17.09 | 9.67 | 16.5 | 9.67 | 15.23
38.24 | 0 | 16.02 | 0.1 | 9.08 | 0.1 | 8.4 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.07 | 9.86 | 23.54 | 9.67 | 15.92 | 9.77 | 9.96 | 1.17 | 0 | 1.07 | 0.1 | 1.46 | 0.1 | 21.29 | 4.79 | 23.15 | 4.69 | 14.16 | 4.59 | 17.97 | 9.67 | 17.29 | 9.67 | 16.5
38.272 | 0 | 16.5 | 0.1 | 9.67 | 0 | 8.98 | 4.79 | 4.98 | 4.79 | 4.88 | 18.07 | 4.88 | 18.07 | 9.86 | 23.93 | 9.67 | 16.5 | 9.67 | 9.86 | 0.98 | 0 | 0.78 | 0.1 | 1.37 | 0.1 | 21.29 | 4.79 | 23.73 | 4.79 | 14.94 | 4.49 | 18.65 | 9.67 | 18.07 | 9.67 | 17.38
38.304 | 0 | 16.99 | 0 | 10.25 | 0 | 9.67 | 4.79 | 5.86 | 4.79 | 4.79 | 18.07 | 4.88 | 18.07 | 9.86 | 24.22 | 9.67 | 17.09 | 9.67 | 9.96 | 0.98 | 0 | 0.78 | 0.1 | 1.07 | 0 | 21.29 | 4.79 | 24.32 | 4.79 | 15.72 | 4.59 | 19.34 | 9.67 | 18.85 | 9.67 | 18.16
38.336 | 0 | 16.99 | 0.1 | 10.74 | 0.1 | 10.16 | 4.79 | 6.74 | 4.79 | 4.88 | 9.38 | 4.88 | 18.07 | 9.86 | 20.9 | 9.67 | 17.38 | 9.67 | 9.96 | 0.78 | 0 | 0.78 | 0.1 | 1.07 | 0 | 21.29 | 4.79 | 24.81 | 4.69 | 16.41 | 4.59 | 20.02 | 9.57 | 19.43 | 9.67 | 18.85
38.368 | 0 | 14.94 | 0.1 | 11.43 | 0.1 | 10.74 | 4.79 | 8.2 | 4.69 | 4.88 | 3.03 | 4.88 | 18.07 | 9.86 | 21 | 9.67 | 17.97 | 9.67 | 10.55 | 0.78 | 0 | 0.68 | 0.1 | 1.07 | 0 | 21.29 | 4.88 | 25.29 | 4.79 | 16.99 | 4.59 | 20.61 | 9.67 | 20.12 | 9.67 | 19.43
38.4 | 0 | 14.84 | 0.1 | 12.01 | 0.1 | 11.43 | 4.79 | 9.18 | 4.69 | 4.88 | 3.13 | 4.88 | 18.07 | 9.86 | 21.29 | 9.77 | 18.46 | 9.67 | 11.62 | 0.68 | 0 | 0.59 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 17.68 | 4.69 | 21.09 | 9.67 | 20.61 | 9.67 | 20.12
38.432 | 0 | 14.84 | 0.1 | 12.5 | 0.1 | 11.91 | 4.79 | 9.86 | 4.59 | 4.88 | 4 | 4.88 | 18.07 | 9.86 | 21.39 | 9.67 | 18.75 | 9.57 | 12.99 | 0.59 | 0 | 0.59 | 0.1 | 0.78 | 0 | 6.54 | 4.88 | 26.27 | 4.79 | 18.26 | 4.69 | 21.78 | 9.67 | 21.19 | 9.67 | 20.8
38.464 | 0 | 14.55 | 0 | 13.09 | 0 | 12.5 | 4.79 | 10.55 | 4.59 | 4.79 | 5.18 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 19.24 | 9.57 | 13.87 | 0.49 | 0 | 0.49 | 0.1 | 0.78 | 0 | 2.64 | 4.88 | 23.24 | 4.79 | 18.95 | 4.69 | 22.36 | 9.67 | 21.88 | 9.67 | 21.29
38.496 | 0 | 14.65 | 0.1 | 13.57 | 0.1 | 12.99 | 4.88 | 11.13 | 4.59 | 4.88 | 5.86 | 4.88 | 17.77 | 9.86 | 21.39 | 9.77 | 19.73 | 9.67 | 14.45 | 0.49 | 0.1 | 0.49 | 0.1 | 0.68 | 0 | 3.22 | 4.88 | 22.27 | 4.79 | 19.63 | 4.69 | 22.85 | 9.67 | 22.36 | 9.67 | 21.88
38.528 | 0 | 14.55 | 0.1 | 14.16 | 0.1 | 13.57 | 4.88 | 11.72 | 4.49 | 4.88 | 5.57 | 4.88 | 10.45 | 9.86 | 21.49 | 9.77 | 20.12 | 9.67 | 15.14 | 0.39 | 0.88 | 0.39 | 0.1 | 0.59 | 0.1 | 4.2 | 4.88 | 21.78 | 4.79 | 20.22 | 4.69 | 23.34 | 9.67 | 22.95 | 9.67 | 22.46
38.56 | 0 | 14.55 | 0.1 | 14.65 | 0.1 | 14.06 | 4.88 | 12.31 | 4.49 | 4.88 | 5.27 | 4.88 | 8.3 | 9.86 | 21.39 | 9.77 | 20.61 | 9.67 | 15.63 | 0.39 | 1.86 | 0.29 | 0.59 | 0.59 | 0.1 | 5.57 | 4.88 | 21.58 | 4.79 | 20.8 | 4.79 | 23.83 | 9.67 | 23.54 | 9.67 | 22.95
38.592 | 0 | 14.55 | 0.1 | 15.23 | 0 | 14.65 | 4.88 | 12.89 | 4.49 | 4.88 | 4.98 | 4.88 | 9.77 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 16.21 | 0.29 | 3.71 | 0.2 | 1.66 | 0.49 | 0.1 | 6.15 | 4.88 | 21.58 | 4.79 | 21.49 | 4.79 | 24.51 | 9.67 | 24.02 | 9.67 | 23.54
38.624 | 0 | 14.45 | 0.1 | 15.63 | 0 | 15.14 | 4.88 | 13.38 | 4.49 | 4.79 | 4.98 | 4.88 | 10.64 | 9.86 | 21.39 | 9.77 | 21.39 | 9.67 | 16.7 | 0.2 | 5.37 | 0.2 | 3.13 | 0.29 | 0 | 5.86 | 4.79 | 21.49 | 4.79 | 21.97 | 4.69 | 25 | 9.67 | 24.61 | 9.67 | 24.02
38.656 | 0 | 14.45 | 0.1 | 16.11 | 0.1 | 15.63 | 4.88 | 13.87 | 4.49 | 4.88 | 4.98 | 4.88 | 10.45 | 9.86 | 21.39 | 9.77 | 21.78 | 9.67 | 17.19 | 0.2 | 6.45 | 0.2 | 5.18 | 0.29 | 0 | 5.47 | 4.79 | 21.39 | 4.79 | 22.56 | 4.79 | 25.49 | 9.67 | 25 | 9.77 | 24.61
38.688 | 0 | 14.45 | 0.1 | 16.6 | 0 | 16.11 | 4.88 | 14.45 | 4.49 | 5.57 | 4.88 | 4.88 | 9.96 | 9.86 | 21.39 | 9.77 | 22.17 | 9.67 | 17.58 | 0.2 | 7.23 | 0.2 | 6.45 | 0.29 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 23.15 | 4.79 | 26.08 | 9.67 | 25.59 | 9.77 | 25.1
38.72 | 0 | 14.45 | 0.1 | 17.09 | 0.1 | 16.5 | 4.88 | 14.94 | 4.49 | 6.45 | 4.79 | 4.88 | 9.77 | 9.86 | 21.39 | 9.77 | 22.56 | 9.67 | 18.07 | 0.2 | 8.11 | 0.2 | 7.23 | 0.29 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 23.63 | 4.79 | 26.56 | 9.77 | 26.17 | 9.77 | 25.59
38.752 | 0 | 14.45 | 0.1 | 13.77 | 0 | 17.09 | 4.88 | 15.43 | 4.49 | 7.81 | 4.79 | 4.88 | 9.86 | 9.86 | 21.19 | 9.77 | 22.95 | 9.67 | 18.56 | 0.2 | 8.89 | 0.1 | 7.91 | 0.2 | 0 | 5.18 | 4.88 | 21.29 | 4.79 | 24.12 | 4.79 | 26.86 | 9.77 | 26.56 | 9.77 | 26.08
38.784 | 0 | 7.32 | 0 | 14.16 | 0 | 14.06 | 4.88 | 16.02 | 4.49 | 8.79 | 4.69 | 4.88 | 9.86 | 9.86 | 8.79 | 9.86 | 23.34 | 9.67 | 18.95 | 0.1 | 9.57 | 0.1 | 8.79 | 0.2 | 0 | 5.08 | 4.88 | 21.29 | 4.79 | 24.71 | 4.79 | 27.34 | 9.67 | 27.15 | 9.77 | 26.56
38.816 | 0 | 2.15 | 0 | 14.16 | 0.1 | 14.06 | 4.88 | 16.41 | 4.59 | 9.57 | 4.69 | 4.88 | 9.77 | 9.86 | 6.93 | 9.86 | 23.73 | 9.77 | 19.34 | 0.1 | 10.45 | 0.1 | 9.57 | 0.1 | 0 | 5.08 | 4.79 | 21.19 | 4.79 | 25.2 | 4.79 | 28.03 | 9.77 | 27.54 | 9.77 | 27.15
38.848 | 0 | 2.05 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 16.9 | 4.59 | 10.35 | 4.59 | 4.88 | 9.77 | 9.86 | 10.06 | 9.86 | 24.02 | 9.77 | 19.82 | 0.1 | 11.23 | 0.1 | 10.35 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 25.68 | 4.79 | 28.42 | 9.77 | 28.03 | 9.77 | 27.54
38.88 | 0 | 1.76 | 0.1 | 14.36 | 0.1 | 14.55 | 4.98 | 17.38 | 4.59 | 10.94 | 4.59 | 4.88 | 9.77 | 9.86 | 11.52 | 9.86 | 24.51 | 9.77 | 20.22 | 0.1 | 12.21 | 0.1 | 11.13 | 0.1 | 0.1 | 4.88 | 4.88 | 12.01 | 4.88 | 26.17 | 4.79 | 28.91 | 9.77 | 28.52 | 9.77 | 28.03
38.912 | 0 | 1.46 | 0.1 | 14.45 | 0 | 14.45 | 4.98 | 17.87 | 4.59 | 11.52 | 4.59 | 4.88 | 9.67 | 9.86 | 10.64 | 9.86 | 22.46 | 9.77 | 20.7 | 0.1 | 12.89 | 0.1 | 11.91 | 0.1 | 0.1 | 4.88 | 4.88 | 2.93 | 4.88 | 22.75 | 4.79 | 29.4 | 9.77 | 28.91 | 9.77 | 28.52
38.944 | 0 | 1.37 | 0 | 14.45 | 0 | 14.45 | 4.88 | 18.26 | 4.59 | 12.11 | 4.49 | 4.88 | 9.67 | 9.86 | 10.06 | 9.86 | 22.07 | 9.77 | 21.09 | 0.1 | 13.67 | 0.1 | 12.7 | 0.1 | 0 | 4.79 | 4.88 | 3.13 | 4.79 | 21.78 | 4.79 | 25 | 9.77 | 29.49 | 9.77 | 29.01
38.976 | 0 | 1.37 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 18.75 | 4.69 | 12.6 | 4.59 | 4.88 | 9.67 | 9.86 | 9.86 | 9.86 | 21.88 | 9.77 | 21.49 | 0.1 | 14.45 | 0.1 | 13.38 | 0.1 | 0 | 4.79 | 4.79 | 3.61 | 4.79 | 21.58 | 4.79 | 23.54 | 9.77 | 25.88 | 9.86 | 29.49
39.008 | 0 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.98 | 19.24 | 4.69 | 13.18 | 4.59 | 4.88 | 9.67 | 9.86 | 9.96 | 9.86 | 21.68 | 9.77 | 21.97 | 0.1 | 15.04 | 0.1 | 14.06 | 0.1 | 0.29 | 4.69 | 4.88 | 4.59 | 4.79 | 21.49 | 4.88 | 24.22 | 9.77 | 25.1 | 9.86 | 26.17
39.04 | 0 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.63 | 4.69 | 13.77 | 4.49 | 4.88 | 9.67 | 9.86 | 10.25 | 9.86 | 21.68 | 9.77 | 22.27 | 0.1 | 15.72 | 0.1 | 14.84 | 0.1 | 1.46 | 4.69 | 4.79 | 6.05 | 4.79 | 21.49 | 4.79 | 24.22 | 9.77 | 24.71 | 9.77 | 25.2
39.072 | 0 | 0.98 | 0.1 | 14.36 | 0 | 14.45 | 4.98 | 20.02 | 4.69 | 14.26 | 4.49 | 5.57 | 9.57 | 9.86 | 10.06 | 9.86 | 21.58 | 9.77 | 22.66 | 0.1 | 16.5 | 0.1 | 15.53 | 0.1 | 2.83 | 4.69 | 4.88 | 6.15 | 4.79 | 21.49 | 4.88 | 24.32 | 9.77 | 24.51 | 9.86 | 24.81
39.104 | 0 | 0.98 | 0 | 14.36 | 0 | 14.45 | 4.98 | 20.51 | 4.69 | 14.75 | 4.49 | 6.45 | 9.57 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 23.05 | 0.1 | 17.19 | 0.1 | 16.31 | 0 | 5.08 | 4.59 | 4.79 | 5.66 | 4.79 | 21.29 | 4.79 | 24.32 | 9.77 | 24.61 | 9.86 | 24.51
39.136 | 0 | 0.68 | 0 | 14.36 | 0.1 | 14.45 | 4.98 | 20.9 | 4.79 | 15.23 | 4.59 | 7.72 | 9.57 | 9.96 | 9.86 | 9.86 | 21.49 | 9.86 | 23.34 | 0.1 | 17.87 | 0.1 | 16.9 | 0 | 6.15 | 4.59 | 4.79 | 5.27 | 4.79 | 21.39 | 4.79 | 24.32 | 9.77 | 24.41 | 9.86 | 24.51
39.168 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.98 | 17.58 | 4.79 | 15.72 | 4.59 | 8.79 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 9.86 | 23.83 | 0.1 | 18.46 | 0.1 | 17.58 | 0.1 | 7.03 | 4.59 | 4.88 | 5.18 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.41
39.2 | 0 | 0.68 | 0.1 | 4.88 | 0.1 | 14.45 | 4.98 | 17.58 | 4.79 | 16.31 | 4.59 | 9.67 | 9.57 | 11.62 | 9.86 | 9.86 | 21.49 | 9.86 | 24.12 | 0.1 | 19.14 | 0.1 | 18.26 | 0.1 | 8.01 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
39.232 | 0 | 0.59 | 0.1 | 1.95 | 0.1 | 6.15 | 4.98 | 17.87 | 4.79 | 16.8 | 4.59 | 10.25 | 9.57 | 12.89 | 9.77 | 9.86 | 21.39 | 9.77 | 22.95 | 0.1 | 19.73 | 0.1 | 18.95 | 0.1 | 8.59 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.32
39.264 | 0 | 0.49 | 0 | 1.76 | 0 | 1.76 | 4.88 | 18.07 | 4.79 | 17.19 | 4.59 | 10.94 | 9.57 | 13.97 | 9.77 | 9.86 | 21.39 | 9.86 | 21 | 0.1 | 20.31 | 0.1 | 19.53 | 0 | 9.57 | 4.59 | 4.79 | 5.18 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 24.32 | 9.86 | 24.22
39.296 | 0 | 0.39 | 0.1 | 1.76 | 0.1 | 1.95 | 4.98 | 18.07 | 4.79 | 17.68 | 4.69 | 11.43 | 9.57 | 14.55 | 9.77 | 9.86 | 21.39 | 9.86 | 21.49 | 0.1 | 21 | 0.1 | 20.12 | 0 | 10.16 | 4.59 | 4.79 | 4.98 | 4.79 | 21.19 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
39.328 | 0.1 | 0.39 | 0.1 | 1.46 | 0 | 1.76 | 4.98 | 18.07 | 4.79 | 18.16 | 4.69 | 12.01 | 9.67 | 15.14 | 9.67 | 9.86 | 14.06 | 9.86 | 21.58 | 0.1 | 21.49 | 0.1 | 20.8 | 0.1 | 11.13 | 4.59 | 4.88 | 4.79 | 4.79 | 16.02 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
39.36 | 0 | 0.29 | 0.1 | 1.37 | 0 | 1.37 | 4.98 | 18.16 | 4.79 | 18.56 | 4.69 | 12.7 | 9.67 | 15.63 | 9.67 | 9.86 | 7.13 | 9.86 | 21.49 | 0.1 | 21.97 | 0.1 | 21.39 | 0.1 | 11.82 | 4.59 | 4.79 | 4.79 | 4.88 | 3.71 | 4.88 | 24.02 | 9.77 | 24.22 | 9.86 | 24.12
39.392 | 0.78 | 0.29 | 0.1 | 1.17 | 0 | 1.46 | 4.98 | 18.16 | 4.79 | 19.14 | 4.69 | 13.18 | 9.57 | 16.21 | 9.67 | 9.86 | 8.01 | 9.86 | 21.49 | 0.1 | 22.56 | 0.1 | 21.97 | 0.1 | 12.6 | 4.59 | 4.88 | 4.79 | 4.79 | 3.22 | 4.88 | 8.5 | 9.77 | 17.97 | 9.86 | 24.12
39.424 | 1.56 | 0.2 | 0 | 1.07 | 0 | 1.17 | 4.88 | 18.16 | 4.79 | 19.53 | 4.69 | 13.67 | 9.67 | 16.7 | 9.67 | 9.86 | 11.04 | 9.86 | 21.49 | 0.1 | 19.53 | 0.1 | 22.56 | 0 | 13.38 | 4.59 | 5.96 | 4.69 | 4.79 | 3.71 | 4.88 | 6.64 | 9.77 | 6.64 | 9.86 | 14.55
39.456 | 2.73 | 0.2 | 0 | 0.88 | 0.1 | 1.17 | 4.98 | 18.16 | 4.79 | 19.92 | 4.79 | 14.26 | 9.67 | 17.19 | 9.67 | 9.86 | 11.23 | 9.86 | 21.49 | 0.1 | 18.65 | 0.1 | 18.85 | 0 | 14.06 | 4.59 | 6.93 | 4.79 | 4.79 | 4.3 | 4.88 | 10.25 | 9.77 | 6.84 | 9.86 | 6.64
39.488 | 3.81 | 0.2 | 0.1 | 0.88 | 0 | 1.07 | 4.98 | 18.07 | 4.88 | 20.22 | 4.79 | 14.75 | 9.67 | 17.68 | 9.67 | 9.86 | 10.45 | 9.86 | 21.49 | 0.1 | 18.16 | 0.1 | 17.38 | 0.1 | 14.75 | 4.69 | 8.69 | 4.69 | 4.79 | 5.76 | 4.88 | 12.31 | 9.77 | 10.74 | 9.86 | 7.23
39.52 | 4.59 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.98 | 18.16 | 4.79 | 20.8 | 4.79 | 15.23 | 9.67 | 18.07 | 9.67 | 9.86 | 9.86 | 9.96 | 21.49 | 0.1 | 17.97 | 0.1 | 17.77 | 0.1 | 15.53 | 4.69 | 10.45 | 4.59 | 4.79 | 6.25 | 4.88 | 11.04 | 9.77 | 12.11 | 9.86 | 11.13
39.552 | 5.18 | 0.1 | 0.1 | 0.78 | 0 | 0.78 | 4.98 | 18.07 | 4.79 | 20.7 | 4.79 | 15.72 | 9.67 | 18.56 | 9.67 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 16.31 | 4.69 | 11.82 | 4.59 | 4.79 | 5.76 | 4.88 | 9.96 | 9.77 | 10.64 | 9.86 | 11.91
39.584 | 5.76 | 0.1 | 0.1 | 0.59 | 0 | 0.68 | 4.88 | 17.97 | 4.79 | 18.16 | 4.79 | 16.21 | 9.67 | 19.14 | 9.57 | 9.86 | 10.16 | 9.96 | 21.39 | 0.1 | 17.77 | 0.1 | 17.87 | 0 | 16.9 | 4.69 | 12.79 | 4.59 | 4.79 | 5.37 | 4.79 | 10.25 | 9.77 | 9.96 | 9.86 | 10.55
39.616 | 6.25 | 0.1 | 0 | 0.49 | 0.1 | 0.59 | 4.88 | 6.35 | 4.88 | 18.46 | 4.79 | 16.8 | 9.67 | 19.34 | 9.57 | 9.86 | 10.06 | 9.96 | 21.39 | 0.1 | 17.77 | 0.1 | 17.77 | 0 | 17.38 | 4.79 | 13.57 | 4.59 | 4.79 | 5.18 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 9.86
39.648 | 6.84 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.98 | 2.73 | 4.88 | 18.46 | 4.88 | 17.19 | 9.67 | 19.92 | 9.67 | 9.86 | 9.86 | 9.96 | 21.39 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 18.07 | 4.79 | 14.36 | 4.59 | 4.79 | 5.08 | 4.88 | 10.55 | 9.77 | 10.35 | 9.86 | 9.96
39.68 | 7.52 | 0.1 | 0.1 | 0.49 | 0 | 0.59 | 4.88 | 3.52 | 4.88 | 18.26 | 4.88 | 17.68 | 9.67 | 20.31 | 9.67 | 9.86 | 9.86 | 9.96 | 11.91 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 18.85 | 4.79 | 15.04 | 4.59 | 4.79 | 5.08 | 4.88 | 10.35 | 9.77 | 10.45 | 9.86 | 10.45
39.712 | 8.11 | 0.1 | 0.1 | 0.29 | 0 | 0.49 | 4.98 | 4.49 | 4.79 | 18.26 | 4.88 | 18.07 | 9.67 | 20.7 | 9.67 | 9.86 | 9.96 | 9.86 | 6.74 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 19.43 | 4.79 | 15.82 | 4.49 | 5.86 | 5.18 | 4.88 | 10.06 | 9.77 | 10.16 | 9.86 | 10.55
39.744 | 8.69 | 0.1 | 0 | 0.29 | 0 | 0.29 | 4.88 | 5.57 | 4.79 | 18.16 | 4.79 | 18.56 | 9.67 | 21.29 | 9.67 | 9.86 | 9.86 | 9.96 | 8.79 | 0 | 17.58 | 0.1 | 17.68 | 0 | 20.22 | 4.79 | 16.6 | 4.59 | 6.84 | 5.08 | 4.79 | 10.25 | 9.77 | 10.06 | 9.86 | 10.06
39.776 | 9.38 | 0.1 | 0.68 | 0.2 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 18.95 | 9.67 | 21.58 | 9.67 | 9.86 | 9.77 | 9.96 | 11.33 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 20.7 | 4.79 | 17.19 | 4.59 | 8.79 | 4.98 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 9.96
39.808 | 9.86 | 0.1 | 1.37 | 0.2 | 0.1 | 0.29 | 4.98 | 5.47 | 4.88 | 18.16 | 4.88 | 19.43 | 9.77 | 21.97 | 9.67 | 9.96 | 9.77 | 9.96 | 11.23 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 21.29 | 4.79 | 17.87 | 4.49 | 10.45 | 4.88 | 4.88 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16
39.84 | 10.45 | 0.1 | 2.34 | 0.2 | 0.1 | 0.29 | 4.88 | 5.27 | 4.88 | 18.16 | 4.88 | 19.82 | 9.77 | 22.36 | 9.67 | 10.74 | 9.77 | 9.96 | 10.25 | 0.1 | 8.89 | 0.1 | 17.68 | 0.1 | 21.97 | 4.79 | 18.46 | 4.59 | 11.62 | 4.88 | 4.88 | 9.86 | 9.77 | 10.06 | 9.86 | 10.16
39.872 | 11.13 | 0.1 | 3.71 | 0.1 | 0 | 0.2 | 4.98 | 4.98 | 4.79 | 18.07 | 4.88 | 20.22 | 9.77 | 22.75 | 9.67 | 11.82 | 9.67 | 9.96 | 9.77 | 0.1 | 2.44 | 0.1 | 16.9 | 0.1 | 22.46 | 4.88 | 19.24 | 4.59 | 12.7 | 4.79 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
39.904 | 11.72 | 0.1 | 4.49 | 0.1 | 0 | 0.2 | 4.88 | 4.88 | 4.79 | 18.07 | 4.88 | 20.7 | 9.77 | 23.15 | 9.67 | 13.18 | 9.67 | 9.96 | 9.86 | 0.1 | 2.34 | 0.1 | 3.91 | 0 | 18.07 | 4.79 | 19.82 | 4.59 | 13.57 | 4.69 | 4.79 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
39.936 | 12.21 | 0.1 | 4.98 | 0 | 0.68 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 9.77 | 23.54 | 9.67 | 14.06 | 9.67 | 9.96 | 10.16 | 0.1 | 1.86 | 0.1 | 1.86 | 0 | 16.8 | 4.88 | 20.31 | 4.69 | 14.36 | 4.69 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 9.77
39.968 | 12.7 | 0.1 | 5.57 | 0.1 | 1.46 | 0.1 | 4.98 | 4.88 | 4.88 | 18.07 | 4.88 | 18.56 | 9.77 | 23.93 | 9.67 | 14.65 | 9.67 | 9.96 | 10.25 | 0.1 | 1.76 | 0.1 | 2.25 | 0.1 | 17.48 | 4.88 | 21 | 4.59 | 15.23 | 4.69 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
40 | 13.28 | 0.1 | 6.15 | 0 | 2.64 | 0.1 | 5.57 | 4.88 | 4.88 | 8.89 | 4.88 | 18.46 | 9.77 | 24.22 | 9.67 | 15.14 | 9.67 | 9.96 | 9.86 | 0.1 | 1.76 | 0.1 | 1.76 | 0.1 | 17.68 | 4.88 | 21.58 | 4.69 | 15.82 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77
40.032 | 13.87 | 0.1 | 6.74 | 0.1 | 3.91 | 0.1 | 6.35 | 4.79 | 4.79 | 2.93 | 4.88 | 18.46 | 9.77 | 22.46 | 9.67 | 15.82 | 9.57 | 9.96 | 9.86 | 0.1 | 1.46 | 0.1 | 1.86 | 0.1 | 17.68 | 4.88 | 22.17 | 4.69 | 16.5 | 4.59 | 5.96 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77
40.064 | 14.45 | 0.1 | 7.32 | 0 | 4.49 | 0.1 | 7.52 | 4.69 | 4.79 | 3.22 | 4.88 | 18.16 | 9.77 | 21 | 9.67 | 16.21 | 9.57 | 9.96 | 9.96 | 0.1 | 1.46 | 0.1 | 1.46 | 0 | 17.68 | 4.88 | 22.75 | 4.69 | 17.19 | 4.59 | 7.03 | 9.67 | 9.77 | 9.77 | 10.64 | 9.77
40.096 | 14.84 | 0.1 | 7.91 | 0 | 5.18 | 0.1 | 8.79 | 4.69 | 4.88 | 4.1 | 4.88 | 18.26 | 9.77 | 21.39 | 9.67 | 16.8 | 9.57 | 9.96 | 9.86 | 0.1 | 1.27 | 0.1 | 1.56 | 0 | 17.68 | 4.88 | 23.34 | 4.69 | 17.77 | 4.59 | 8.89 | 9.67 | 10.64 | 9.67 | 11.52 | 9.67
40.128 | 15.33 | 0.1 | 8.59 | 0.1 | 5.76 | 0.1 | 9.57 | 4.59 | 4.88 | 5.37 | 4.88 | 18.16 | 9.77 | 21.58 | 9.77 | 17.19 | 9.57 | 9.96 | 9.77 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 17.77 | 4.88 | 23.83 | 4.69 | 18.46 | 4.59 | 10.55 | 9.67 | 11.52 | 9.67 | 13.18 | 9.67
40.16 | 15.92 | 0.1 | 9.08 | 0.1 | 6.25 | 0.1 | 10.16 | 4.59 | 4.88 | 6.15 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 17.68 | 9.57 | 9.96 | 9.77 | 0.1 | 1.07 | 0.1 | 1.37 | 0.1 | 17.68 | 4.88 | 24.32 | 4.69 | 19.24 | 4.49 | 11.82 | 9.67 | 13.09 | 9.67 | 15.14 | 9.67
40.192 | 16.31 | 0.1 | 9.77 | 0.1 | 6.84 | 0.1 | 10.74 | 4.59 | 4.79 | 5.66 | 4.88 | 18.16 | 9.77 | 21.49 | 9.77 | 18.16 | 9.57 | 9.96 | 9.77 | 0.1 | 0.98 | 0.1 | 1.07 | 0.1 | 17.77 | 4.88 | 24.81 | 4.79 | 19.82 | 4.49 | 12.79 | 9.67 | 15.04 | 9.67 | 16.5 | 9.67
40.224 | 16.8 | 0.1 | 10.35 | 0.1 | 7.42 | 0.1 | 11.43 | 4.49 | 4.79 | 5.37 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.56 | 9.67 | 9.96 | 9.67 | 0.1 | 0.88 | 0.1 | 1.07 | 0 | 17.68 | 4.88 | 25.29 | 4.79 | 20.31 | 4.59 | 13.77 | 9.67 | 16.31 | 9.67 | 17.19 | 9.67
40.256 | 17.29 | 0.1 | 10.94 | 0 | 8.11 | 0.1 | 12.01 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.95 | 9.57 | 9.96 | 9.67 | 0.1 | 0.78 | 0.1 | 0.98 | 0.1 | 17.68 | 4.88 | 25.78 | 4.79 | 20.9 | 4.59 | 14.45 | 9.57 | 17.19 | 9.67 | 18.07 | 9.57
40.288 | 15.33 | 0.1 | 11.52 | 0.1 | 8.69 | 0.1 | 12.6 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 19.43 | 9.57 | 9.96 | 9.67 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 17.68 | 4.88 | 26.37 | 4.79 | 21.58 | 4.59 | 15.23 | 9.67 | 18.07 | 9.67 | 18.75 | 9.67
40.32 | 15.04 | 0.1 | 12.11 | 0.1 | 9.38 | 0.1 | 13.18 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 19.92 | 9.67 | 9.96 | 9.67 | 0.1 | 0.59 | 0.1 | 0.78 | 0.1 | 17.68 | 4.88 | 23.05 | 4.79 | 22.17 | 4.59 | 15.82 | 9.67 | 18.65 | 9.67 | 19.43 | 9.57
40.352 | 14.84 | 0.1 | 12.7 | 0.1 | 9.86 | 0.1 | 13.67 | 4.49 | 4.79 | 4.88 | 4.88 | 17.97 | 9.77 | 21.49 | 9.77 | 20.31 | 9.67 | 10.64 | 9.67 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 6.15 | 4.88 | 22.27 | 4.79 | 22.66 | 4.59 | 16.7 | 9.67 | 19.34 | 9.67 | 20.02 | 9.67
40.384 | 14.55 | 0.1 | 13.18 | 0 | 10.45 | 0.1 | 14.26 | 4.49 | 4.79 | 4.79 | 4.88 | 8.4 | 9.77 | 21.49 | 9.77 | 20.8 | 9.67 | 11.33 | 9.57 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 2.44 | 4.88 | 21.68 | 4.79 | 23.24 | 4.69 | 16.99 | 9.67 | 20.02 | 9.67 | 20.61 | 9.57
40.416 | 14.65 | 0.1 | 13.67 | 0 | 11.13 | 0.1 | 14.75 | 4.59 | 4.88 | 4.69 | 4.88 | 3.03 | 9.77 | 21.39 | 9.77 | 21.09 | 9.67 | 12.5 | 9.57 | 0.1 | 0.49 | 0.1 | 0.49 | 0 | 2.83 | 4.88 | 21.58 | 4.88 | 23.73 | 4.69 | 17.87 | 9.67 | 20.61 | 9.57 | 21.19 | 9.57
40.448 | 14.45 | 0.1 | 14.26 | 0.1 | 11.62 | 0.1 | 15.23 | 4.59 | 4.88 | 4.69 | 4.88 | 3.22 | 9.77 | 21.39 | 9.77 | 21.58 | 9.67 | 13.57 | 9.57 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 2.93 | 4.88 | 21.58 | 4.79 | 24.32 | 4.69 | 18.56 | 9.67 | 21.19 | 9.67 | 21.88 | 9.67
40.48 | 14.45 | 0.1 | 14.75 | 0.1 | 12.21 | 0.1 | 15.72 | 4.59 | 4.79 | 4.79 | 4.88 | 4.1 | 9.77 | 11.43 | 9.77 | 21.97 | 9.67 | 14.45 | 9.57 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 3.13 | 4.88 | 21.49 | 4.79 | 24.71 | 4.69 | 19.14 | 9.67 | 21.78 | 9.67 | 22.36 | 9.57
40.512 | 14.45 | 0.1 | 15.23 | 0.1 | 12.7 | 0.1 | 16.21 | 4.59 | 4.79 | 4.69 | 4.88 | 5.37 | 9.77 | 6.54 | 9.77 | 22.36 | 9.67 | 14.94 | 9.57 | 0.1 | 0.2 | 0.1 | 0.39 | 0.1 | 3.52 | 4.88 | 21.29 | 4.88 | 25.29 | 4.79 | 19.82 | 9.67 | 22.36 | 9.67 | 22.85 | 9.67
40.544 | 14.36 | 0.1 | 15.72 | 0 | 13.28 | 0.1 | 16.8 | 4.59 | 4.79 | 4.59 | 4.88 | 5.76 | 9.77 | 9.57 | 9.77 | 22.75 | 9.67 | 15.53 | 9.57 | 0 | 0.2 | 0.88 | 0.29 | 0 | 3.71 | 4.88 | 21.29 | 4.79 | 25.68 | 4.69 | 20.51 | 9.67 | 22.85 | 9.67 | 23.44 | 9.67
40.576 | 14.36 | 0.1 | 16.21 | 0 | 13.77 | 0.1 | 17.19 | 4.69 | 4.88 | 4.59 | 4.88 | 5.57 | 9.77 | 11.52 | 9.77 | 23.15 | 9.67 | 16.02 | 9.57 | 0.1 | 0.2 | 1.86 | 0.29 | 0 | 3.81 | 4.88 | 21.29 | 4.79 | 26.17 | 4.79 | 20.9 | 9.67 | 23.34 | 9.67 | 24.02 | 9.67
40.608 | 14.36 | 0.1 | 16.7 | 0.1 | 14.36 | 0.1 | 17.68 | 4.69 | 4.88 | 4.59 | 4.88 | 5.27 | 9.77 | 11.04 | 9.86 | 23.54 | 9.67 | 16.6 | 9.57 | 1.17 | 0.2 | 3.52 | 0.29 | 0.1 | 3.61 | 4.88 | 21.29 | 4.79 | 23.05 | 4.79 | 21.68 | 9.77 | 24.02 | 9.67 | 24.41 | 9.67
40.64 | 14.36 | 0.1 | 17.19 | 0.1 | 14.84 | 0.1 | 18.16 | 4.69 | 5.37 | 4.59 | 4.88 | 4.98 | 9.77 | 10.35 | 9.86 | 23.93 | 9.67 | 16.99 | 9.67 | 2.25 | 0.1 | 5.37 | 0.2 | 0.1 | 3.42 | 4.88 | 21.29 | 4.79 | 22.17 | 4.79 | 22.17 | 9.77 | 24.51 | 9.67 | 25.1 | 9.67
40.672 | 14.36 | 0.1 | 14.16 | 0.1 | 15.43 | 0.1 | 18.56 | 4.69 | 6.45 | 4.59 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 24.22 | 9.67 | 17.58 | 9.67 | 4 | 0.1 | 6.54 | 0.2 | 0.1 | 3.22 | 4.88 | 21.19 | 4.88 | 21.78 | 4.79 | 22.66 | 9.77 | 25.1 | 9.67 | 25.59 | 9.67
40.704 | 7.52 | 0.1 | 15.14 | 0 | 15.82 | 0.1 | 19.04 | 4.69 | 7.72 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.77 | 20.8 | 9.67 | 17.97 | 9.67 | 5.57 | 0.1 | 7.42 | 0.1 | 1.07 | 3.13 | 4.88 | 21.19 | 4.88 | 21.49 | 4.79 | 23.24 | 9.77 | 25.49 | 9.67 | 26.08 | 9.67
40.736 | 2.25 | 0.1 | 14.55 | 0 | 16.41 | 0.1 | 19.43 | 4.69 | 8.79 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.77 | 20.9 | 9.67 | 18.46 | 9.67 | 6.64 | 0.1 | 8.2 | 0.1 | 2.15 | 3.03 | 4.88 | 10.55 | 4.79 | 21.49 | 4.79 | 23.83 | 9.77 | 25.88 | 9.67 | 26.56 | 9.67
40.768 | 1.95 | 0.1 | 14.55 | 0.1 | 16.8 | 0.1 | 19.92 | 4.79 | 9.67 | 4.59 | 4.88 | 4.79 | 9.77 | 10.16 | 9.86 | 21.19 | 9.67 | 18.85 | 9.67 | 7.42 | 0.1 | 8.89 | 0.1 | 3.91 | 2.93 | 4.88 | 2.83 | 4.79 | 21.39 | 4.79 | 24.41 | 9.77 | 26.47 | 9.77 | 26.95 | 9.67
40.8 | 1.76 | 0.1 | 14.55 | 0.1 | 17.29 | 0.1 | 20.31 | 4.79 | 10.25 | 4.59 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.39 | 9.77 | 19.34 | 9.67 | 8.3 | 0.1 | 9.77 | 0.1 | 5.57 | 2.93 | 4.88 | 3.13 | 4.79 | 21.29 | 4.79 | 24.9 | 9.77 | 26.95 | 9.77 | 27.44 | 9.67
40.832 | 1.37 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 20.7 | 4.79 | 10.94 | 4.59 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 21.39 | 9.77 | 19.82 | 9.67 | 9.08 | 0 | 10.64 | 0.1 | 6.64 | 2.83 | 4.88 | 3.61 | 4.88 | 21.29 | 4.79 | 25.39 | 9.77 | 27.44 | 9.77 | 28.03 | 9.77
40.864 | 1.37 | 0.1 | 14.55 | 0 | 15.23 | 0.1 | 21.19 | 4.79 | 11.52 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.77 | 21.39 | 9.77 | 20.12 | 9.67 | 9.77 | 0 | 11.43 | 0.1 | 7.42 | 2.73 | 4.79 | 4.79 | 4.88 | 21.19 | 4.79 | 25.88 | 9.77 | 27.93 | 9.77 | 28.42 | 9.77
40.896 | 1.27 | 0.1 | 14.45 | 0 | 14.94 | 0.1 | 19.24 | 4.79 | 12.11 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.77 | 21.49 | 9.67 | 20.61 | 9.67 | 10.64 | 0.1 | 12.21 | 0.1 | 8.2 | 2.64 | 4.88 | 6.05 | 4.88 | 21.19 | 4.88 | 26.17 | 9.77 | 28.42 | 9.77 | 28.91 | 9.77
40.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.75 | 0.1 | 19.04 | 4.88 | 12.6 | 4.69 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 21 | 9.67 | 11.52 | 0 | 12.89 | 0.1 | 9.08 | 2.64 | 4.88 | 6.05 | 4.79 | 21.19 | 4.88 | 22.66 | 9.77 | 28.91 | 9.77 | 29.49 | 9.77
40.96 | 0.98 | 0.1 | 14.45 | 0.1 | 14.75 | 0.1 | 18.65 | 4.88 | 13.18 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.39 | 9.67 | 12.11 | 0 | 13.67 | 0.1 | 9.86 | 2.64 | 4.88 | 5.66 | 4.79 | 21.19 | 4.88 | 22.07 | 9.77 | 29.4 | 9.77 | 25.39 | 9.77
40.992 | 0.78 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.77 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 9.67 | 12.99 | 0 | 14.55 | 0.1 | 10.64 | 2.54 | 4.88 | 5.27 | 4.88 | 21.19 | 4.88 | 21.68 | 9.77 | 25.2 | 9.77 | 25 | 9.77
41.024 | 0.78 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 18.46 | 4.79 | 14.26 | 4.69 | 5.47 | 4.49 | 9.77 | 9.77 | 9.77 | 21.39 | 9.77 | 22.17 | 9.67 | 13.77 | 0 | 15.14 | 0.1 | 11.52 | 2.54 | 4.88 | 5.18 | 4.88 | 9.57 | 4.79 | 21.49 | 9.77 | 24.9 | 9.77 | 24.51 | 9.77
41.056 | 0.78 | 0.1 | 14.36 | 0 | 14.55 | 0.1 | 18.26 | 4.88 | 14.84 | 4.69 | 6.35 | 4.59 | 9.77 | 9.67 | 9.77 | 21.49 | 9.77 | 22.56 | 9.67 | 14.45 | 0.1 | 15.82 | 0.1 | 12.21 | 2.54 | 4.79 | 5.08 | 4.88 | 2.93 | 4.88 | 21.49 | 9.77 | 24.51 | 9.77 | 24.41 | 9.77
41.088 | 0.59 | 0.1 | 12.01 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.33 | 4.79 | 7.42 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 9.77 | 15.14 | 0.1 | 16.41 | 0.1 | 13.09 | 2.54 | 4.88 | 5.18 | 4.79 | 3.22 | 4.88 | 21.39 | 9.77 | 24.41 | 9.77 | 24.41 | 9.77
41.12 | 0.59 | 0.1 | 4 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.82 | 4.79 | 8.59 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 23.34 | 9.77 | 15.82 | 0.1 | 17.19 | 0.1 | 13.67 | 2.54 | 5.86 | 5.18 | 4.79 | 3.71 | 4.88 | 21.29 | 9.77 | 24.41 | 9.77 | 24.32 | 9.77
41.152 | 0.49 | 0.1 | 2.05 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 16.31 | 4.79 | 9.57 | 4.49 | 9.77 | 9.67 | 9.86 | 9.77 | 9.77 | 23.73 | 9.77 | 16.5 | 0 | 17.87 | 0.1 | 14.55 | 2.54 | 7.32 | 4.98 | 4.88 | 4.79 | 4.88 | 21.29 | 9.77 | 24.41 | 9.77 | 24.32 | 9.77
41.184 | 0.39 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 18.16 | 4.79 | 16.8 | 4.79 | 10.16 | 4.49 | 9.86 | 9.67 | 9.77 | 6.64 | 9.77 | 24.12 | 9.77 | 17.19 | 0 | 18.46 | 0 | 15.23 | 2.54 | 9.67 | 4.88 | 4.88 | 6.05 | 4.88 | 21.19 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77
41.216 | 0.29 | 0.1 | 1.66 | 0 | 14.36 | 0.1 | 18.16 | 4.79 | 17.19 | 4.79 | 10.74 | 4.59 | 10.55 | 9.67 | 9.77 | 9.57 | 9.77 | 24.41 | 9.77 | 17.87 | 0 | 19.24 | 0.1 | 16.02 | 2.54 | 11.04 | 4.88 | 4.88 | 6.05 | 4.88 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
41.248 | 0.29 | 0.1 | 1.46 | 0.1 | 9.57 | 0.1 | 18.16 | 4.88 | 17.68 | 4.88 | 11.43 | 4.59 | 11.62 | 9.67 | 9.86 | 11.62 | 9.86 | 22.36 | 9.77 | 18.46 | 0.1 | 19.82 | 0.1 | 16.8 | 2.54 | 12.21 | 4.79 | 4.79 | 5.57 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
41.28 | 0.29 | 0.1 | 1.27 | 0.1 | 2.73 | 0.1 | 18.07 | 4.88 | 18.16 | 4.88 | 12.01 | 4.59 | 12.99 | 9.67 | 9.86 | 10.94 | 9.86 | 22.36 | 9.77 | 19.14 | 0 | 20.41 | 0.1 | 17.19 | 2.64 | 13.09 | 4.79 | 4.88 | 5.27 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
41.312 | 0.2 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 9.47 | 4.88 | 18.56 | 4.88 | 12.6 | 4.69 | 13.87 | 9.67 | 9.86 | 10.16 | 9.86 | 21.97 | 9.77 | 19.82 | 0 | 20.9 | 0.1 | 17.97 | 2.64 | 13.97 | 4.69 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
41.344 | 0.1 | 0.1 | 0.98 | 0 | 1.86 | 0.1 | 3.13 | 4.88 | 18.95 | 4.88 | 13.09 | 4.69 | 14.55 | 9.67 | 9.86 | 9.86 | 9.86 | 21.78 | 9.77 | 20.31 | 0 | 21.58 | 0.1 | 18.65 | 2.64 | 14.55 | 4.69 | 4.88 | 5.08 | 4.79 | 12.99 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
41.376 | 0.1 | 0.1 | 1.07 | 0 | 1.46 | 0.1 | 3.22 | 4.88 | 19.43 | 4.88 | 13.67 | 4.69 | 15.04 | 9.67 | 9.77 | 10.06 | 9.77 | 21.78 | 9.77 | 20.9 | 0.1 | 22.07 | 0.1 | 19.24 | 2.64 | 15.33 | 4.59 | 4.88 | 5.08 | 4.88 | 7.62 | 9.77 | 24.12 | 9.77 | 19.43 | 9.77
41.408 | 0.1 | 0.1 | 0.78 | 0.1 | 1.46 | 0.1 | 3.71 | 4.88 | 19.92 | 4.88 | 14.16 | 4.69 | 15.63 | 9.67 | 9.86 | 10.25 | 9.86 | 21.68 | 9.86 | 21.58 | 0 | 22.66 | 0.1 | 19.82 | 2.73 | 16.11 | 4.59 | 4.79 | 5.08 | 4.88 | 8.69 | 9.86 | 19.73 | 9.86 | 7.42 | 9.86
41.44 | 0 | 0.1 | 0.88 | 0.1 | 1.27 | 0.1 | 4.69 | 4.88 | 20.31 | 4.88 | 14.75 | 4.69 | 16.11 | 9.67 | 9.96 | 10.16 | 9.86 | 21.58 | 9.77 | 22.17 | 0 | 19.53 | 0.1 | 20.51 | 2.73 | 16.7 | 4.59 | 4.79 | 4.98 | 4.88 | 11.04 | 9.77 | 7.42 | 9.86 | 6.84 | 9.86
41.472 | 0 | 0.78 | 0.68 | 0.1 | 1.07 | 0.1 | 5.86 | 4.88 | 20.7 | 4.88 | 15.23 | 4.79 | 16.7 | 9.67 | 10.64 | 9.96 | 9.86 | 21.58 | 9.86 | 21.97 | 0 | 18.75 | 0.1 | 21 | 2.73 | 17.48 | 4.59 | 4.88 | 4.88 | 4.88 | 11.04 | 9.77 | 6.74 | 9.86 | 10.94 | 9.86
41.504 | 0 | 1.56 | 0.59 | 0 | 0.88 | 0.1 | 5.76 | 4.88 | 20.31 | 4.88 | 15.72 | 4.79 | 17.19 | 9.67 | 11.82 | 9.96 | 9.86 | 21.49 | 9.77 | 17.77 | 0 | 18.26 | 0.1 | 21.68 | 2.73 | 18.16 | 4.59 | 4.88 | 4.88 | 4.88 | 10.25 | 9.77 | 10.74 | 9.86 | 12.01 | 9.77
41.536 | 0 | 2.83 | 0.49 | 0.1 | 0.88 | 0.1 | 5.37 | 4.88 | 18.36 | 4.88 | 16.21 | 4.79 | 17.58 | 9.67 | 13.09 | 9.86 | 9.86 | 21.49 | 9.77 | 18.07 | 0.1 | 18.07 | 0.1 | 22.27 | 2.73 | 18.75 | 4.49 | 4.88 | 4.79 | 4.88 | 9.96 | 9.77 | 12.01 | 9.77 | 10.74 | 9.77
41.568 | 0 | 3.91 | 0.49 | 0.1 | 0.88 | 0.1 | 5.08 | 4.88 | 18.26 | 4.88 | 16.7 | 4.79 | 18.07 | 9.67 | 13.97 | 9.86 | 9.86 | 21.49 | 9.86 | 17.87 | 0.1 | 18.07 | 0.1 | 22.66 | 2.73 | 19.43 | 4.49 | 5.76 | 4.79 | 4.88 | 10.16 | 9.77 | 10.74 | 9.86 | 9.96 | 9.86
41.6 | 0 | 4.59 | 0.49 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.36 | 4.88 | 17.19 | 4.79 | 18.56 | 9.77 | 14.65 | 9.77 | 9.86 | 21.49 | 9.86 | 17.87 | 0 | 17.97 | 0.1 | 19.24 | 2.73 | 20.12 | 4.59 | 6.84 | 4.69 | 4.88 | 10.25 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
41.632 | 0.1 | 5.18 | 0.39 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 17.58 | 4.79 | 18.95 | 9.77 | 15.14 | 9.67 | 9.86 | 21.49 | 9.86 | 17.77 | 0 | 17.77 | 0.1 | 18.56 | 2.83 | 20.7 | 4.59 | 8.4 | 4.69 | 4.79 | 10.25 | 9.77 | 10.16 | 9.86 | 10.45 | 9.86
41.664 | 0 | 5.76 | 0.29 | 0 | 0.59 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 17.97 | 4.79 | 19.43 | 9.77 | 15.72 | 9.77 | 9.86 | 15.82 | 9.86 | 17.58 | 0 | 17.77 | 0.1 | 18.07 | 2.73 | 21.29 | 4.59 | 10.35 | 4.59 | 4.79 | 9.96 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86
41.696 | 0 | 6.35 | 0.2 | 0 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.46 | 4.79 | 19.82 | 9.77 | 16.11 | 9.67 | 9.77 | 7.72 | 9.86 | 17.68 | 0 | 17.77 | 0.1 | 17.97 | 2.83 | 21.78 | 4.59 | 11.62 | 4.59 | 4.88 | 9.86 | 9.77 | 10.45 | 9.77 | 10.25 | 9.86
41.728 | 0.1 | 6.93 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 18.95 | 4.88 | 20.31 | 9.77 | 16.7 | 9.67 | 9.86 | 7.62 | 9.86 | 17.58 | 0 | 17.77 | 0.1 | 17.97 | 2.83 | 22.36 | 4.59 | 12.7 | 4.59 | 4.88 | 9.86 | 9.77 | 10.35 | 9.86 | 10.06 | 9.86
41.76 | 0 | 7.52 | 0.2 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.43 | 4.88 | 20.7 | 9.77 | 17.19 | 9.67 | 9.86 | 10.64 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.87 | 2.83 | 22.95 | 4.59 | 13.57 | 4.59 | 4.88 | 9.86 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
41.792 | 0 | 8.11 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 21.19 | 9.77 | 17.68 | 9.67 | 9.86 | 11.43 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 2.83 | 23.44 | 4.59 | 14.26 | 4.59 | 4.88 | 9.86 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
41.824 | 0 | 8.79 | 0.1 | 0 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.22 | 4.79 | 21.58 | 9.77 | 18.16 | 9.67 | 9.86 | 10.64 | 9.86 | 17.58 | 0 | 17.68 | 0.1 | 17.68 | 2.83 | 24.02 | 4.69 | 15.14 | 4.49 | 4.79 | 9.77 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
41.856 | 0 | 9.38 | 0.1 | 0.1 | 0.2 | 0.1 | 4.49 | 4.88 | 18.07 | 4.88 | 20.7 | 4.88 | 21.97 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 9.86 | 17.58 | 0 | 12.01 | 0.1 | 17.68 | 2.83 | 24.51 | 4.69 | 15.82 | 4.49 | 4.88 | 9.77 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
41.888 | 0.1 | 9.96 | 0.1 | 1.17 | 0.2 | 0.1 | 4.49 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.36 | 9.86 | 19.04 | 9.57 | 9.86 | 9.96 | 9.86 | 17.58 | 0 | 2.73 | 0.1 | 17.68 | 2.93 | 25 | 4.69 | 16.5 | 4.49 | 4.88 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
41.92 | 0 | 10.55 | 0.1 | 1.95 | 0.1 | 1.17 | 4.59 | 4.88 | 18.07 | 4.88 | 19.14 | 4.88 | 22.75 | 9.86 | 19.34 | 9.57 | 9.86 | 10.16 | 9.86 | 8.4 | 0 | 2.34 | 0.1 | 17.68 | 2.93 | 25.49 | 4.69 | 17.09 | 4.59 | 4.88 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
41.952 | 0 | 11.13 | 0.1 | 3.32 | 0.1 | 1.95 | 4.49 | 4.88 | 7.81 | 4.88 | 18.85 | 4.88 | 23.05 | 9.86 | 19.82 | 9.57 | 9.86 | 10.16 | 9.86 | 2.15 | 0 | 2.25 | 0.1 | 17.68 | 2.93 | 25.98 | 4.79 | 17.87 | 4.59 | 4.88 | 9.67 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
41.984 | 0 | 11.72 | 0.1 | 4.3 | 0 | 3.32 | 4.49 | 4.88 | 2.73 | 4.88 | 18.56 | 4.79 | 23.44 | 9.86 | 20.22 | 9.57 | 9.86 | 9.96 | 9.86 | 2.15 | 0 | 1.66 | 0.1 | 17.68 | 2.83 | 21.88 | 4.79 | 18.46 | 4.59 | 4.88 | 9.67 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
42.016 | 0 | 12.21 | 0 | 4.88 | 0.1 | 4.1 | 4.49 | 4.88 | 3.13 | 4.88 | 18.46 | 4.88 | 23.83 | 9.86 | 20.7 | 9.67 | 10.45 | 9.96 | 9.86 | 1.95 | 0 | 1.86 | 0.1 | 8.89 | 2.83 | 20.61 | 4.79 | 19.04 | 4.59 | 5.76 | 9.67 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
42.048 | 0 | 12.89 | 0.1 | 5.37 | 0 | 4.79 | 4.49 | 4.88 | 4.3 | 4.88 | 18.46 | 4.88 | 24.22 | 9.86 | 21.19 | 9.67 | 11.33 | 9.96 | 9.96 | 1.76 | 0 | 1.56 | 0.1 | 2.15 | 2.93 | 21.19 | 4.79 | 19.82 | 4.59 | 6.35 | 9.67 | 9.96 | 9.77 | 9.86 | 9.77 | 9.86
42.08 | 0 | 13.38 | 0.1 | 5.96 | 0.1 | 5.37 | 4.49 | 4.88 | 5.47 | 4.88 | 18.26 | 4.88 | 20.8 | 9.86 | 21.49 | 9.67 | 12.4 | 9.96 | 9.96 | 1.56 | 0 | 1.46 | 0.1 | 2.34 | 2.93 | 21.29 | 4.79 | 20.31 | 4.59 | 8.01 | 9.67 | 10.94 | 9.77 | 10.84 | 9.77 | 9.86
42.112 | 0 | 13.87 | 0.1 | 6.54 | 0 | 5.96 | 4.49 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 20.9 | 9.86 | 21.88 | 9.67 | 13.57 | 9.77 | 9.96 | 1.56 | 0 | 1.37 | 0.1 | 2.25 | 2.93 | 21.29 | 4.79 | 20.9 | 4.69 | 10.25 | 9.67 | 12.5 | 9.77 | 11.52 | 9.67 | 10.74
42.144 | 0 | 14.45 | 0 | 7.23 | 0 | 6.54 | 4.49 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.09 | 9.86 | 22.36 | 9.67 | 14.36 | 9.77 | 9.86 | 1.37 | 0 | 1.17 | 0 | 1.66 | 2.83 | 21.19 | 4.79 | 21.58 | 4.69 | 11.52 | 9.67 | 14.65 | 9.67 | 13.18 | 9.67 | 11.91
42.176 | 0 | 14.94 | 0.1 | 7.81 | 0.1 | 7.03 | 4.59 | 4.88 | 5.18 | 4.88 | 18.16 | 4.88 | 21.29 | 9.86 | 22.66 | 9.67 | 14.94 | 9.77 | 9.86 | 1.37 | 0 | 1.27 | 0.1 | 1.76 | 2.83 | 21.29 | 4.79 | 22.17 | 4.69 | 12.5 | 9.57 | 16.02 | 9.67 | 15.04 | 9.67 | 13.38
42.208 | 0.1 | 15.43 | 0.1 | 8.4 | 0.1 | 7.72 | 4.59 | 4.88 | 4.98 | 4.98 | 18.16 | 4.88 | 21.29 | 9.86 | 23.05 | 9.67 | 15.43 | 9.77 | 9.96 | 1.17 | 0 | 1.07 | 0.1 | 1.56 | 2.93 | 21.29 | 4.79 | 22.75 | 4.69 | 13.28 | 9.67 | 16.99 | 9.67 | 16.31 | 9.67 | 15.14
42.24 | 0 | 15.92 | 0.1 | 8.98 | 0.1 | 8.3 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.29 | 9.86 | 23.44 | 9.67 | 15.92 | 9.77 | 9.86 | 1.07 | 0 | 0.88 | 0.1 | 1.46 | 2.93 | 21.29 | 4.79 | 23.24 | 4.69 | 14.16 | 9.67 | 17.77 | 9.67 | 17.19 | 9.67 | 16.41
42.272 | 0 | 16.41 | 0.1 | 9.57 | 0 | 8.89 | 4.69 | 4.98 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 23.83 | 9.67 | 16.41 | 9.67 | 9.96 | 0.98 | 0 | 0.88 | 0.1 | 1.27 | 2.93 | 21.29 | 4.88 | 23.73 | 4.79 | 14.94 | 9.67 | 18.46 | 9.67 | 17.97 | 9.67 | 17.29
42.304 | 0 | 16.9 | 0.1 | 10.16 | 0 | 9.67 | 4.69 | 5.66 | 4.88 | 4.88 | 18.07 | 4.88 | 21.29 | 9.86 | 24.22 | 9.67 | 16.99 | 9.67 | 9.86 | 0.88 | 0 | 0.78 | 0.1 | 1.17 | 2.93 | 21.19 | 4.79 | 24.22 | 4.69 | 15.72 | 9.67 | 19.24 | 9.67 | 18.75 | 9.67 | 18.16
42.336 | 0 | 17.38 | 0.1 | 10.64 | 0.1 | 10.06 | 4.69 | 6.64 | 4.79 | 4.88 | 10.84 | 4.88 | 21.29 | 9.86 | 20.8 | 9.67 | 17.38 | 9.67 | 9.96 | 0.88 | 0 | 0.78 | 0.1 | 1.07 | 2.83 | 21.29 | 4.79 | 24.81 | 4.69 | 16.31 | 9.67 | 19.82 | 9.67 | 19.34 | 9.67 | 18.75
42.368 | 0 | 15.43 | 0.1 | 11.33 | 0.1 | 10.74 | 4.69 | 8.11 | 4.79 | 4.98 | 3.32 | 4.88 | 21.39 | 9.86 | 20.9 | 9.67 | 17.87 | 9.67 | 10.45 | 0.78 | 0 | 0.59 | 0.1 | 1.07 | 2.93 | 21.19 | 4.88 | 25.29 | 4.79 | 17.09 | 9.67 | 20.51 | 9.67 | 20.02 | 9.67 | 19.34
42.4 | 0 | 15.23 | 0.1 | 11.91 | 0.1 | 11.23 | 4.79 | 9.08 | 4.69 | 4.88 | 3.13 | 4.88 | 21.29 | 9.86 | 21.09 | 9.67 | 18.26 | 9.67 | 11.52 | 0.68 | 0 | 0.59 | 0.1 | 0.88 | 2.93 | 21.19 | 4.88 | 25.78 | 4.79 | 17.68 | 9.67 | 21.09 | 9.67 | 20.61 | 9.67 | 20.12
42.432 | 0 | 14.94 | 0.1 | 12.5 | 0.1 | 11.82 | 4.79 | 9.77 | 4.59 | 4.88 | 4 | 4.88 | 21.29 | 9.86 | 21.39 | 9.67 | 18.75 | 9.67 | 12.79 | 0.59 | 0 | 0.49 | 0.1 | 0.88 | 2.93 | 7.42 | 4.88 | 26.27 | 4.79 | 18.36 | 9.67 | 21.68 | 9.67 | 21.19 | 9.67 | 20.7
42.464 | 0 | 14.65 | 0 | 12.99 | 0 | 12.4 | 4.79 | 10.45 | 4.59 | 4.88 | 4.98 | 4.88 | 21.29 | 9.86 | 21.29 | 9.67 | 19.24 | 9.67 | 13.87 | 0.49 | 0 | 0.49 | 0 | 0.59 | 2.93 | 2.34 | 4.88 | 23.15 | 4.79 | 18.95 | 9.67 | 22.27 | 9.67 | 21.78 | 9.67 | 21.39
42.496 | 0 | 14.75 | 0.1 | 13.48 | 0.1 | 12.89 | 4.79 | 11.04 | 4.59 | 4.88 | 5.96 | 4.88 | 21.29 | 9.86 | 21.39 | 9.77 | 19.53 | 9.67 | 14.45 | 0.49 | 0 | 0.49 | 0.1 | 0.68 | 2.93 | 3.22 | 4.88 | 22.27 | 4.79 | 19.63 | 9.67 | 22.75 | 9.67 | 22.36 | 9.67 | 21.88
42.528 | 0 | 14.65 | 0.1 | 14.06 | 0.1 | 13.48 | 4.88 | 11.72 | 4.59 | 4.88 | 5.76 | 4.88 | 10.64 | 9.86 | 21.49 | 9.77 | 20.12 | 9.67 | 15.04 | 0.49 | 0.78 | 0.29 | 0.1 | 0.59 | 2.93 | 4 | 4.88 | 21.78 | 4.79 | 20.12 | 9.77 | 23.34 | 9.67 | 22.85 | 9.67 | 22.46
42.56 | 0 | 14.55 | 0.1 | 14.55 | 0.1 | 14.06 | 4.88 | 12.31 | 4.59 | 4.88 | 5.27 | 4.88 | 6.54 | 9.86 | 21.39 | 9.77 | 20.51 | 9.67 | 15.63 | 0.29 | 1.76 | 0.29 | 0.29 | 0.49 | 2.93 | 5.18 | 4.88 | 21.58 | 4.79 | 20.9 | 9.77 | 23.93 | 9.67 | 23.44 | 9.67 | 23.05
42.592 | 0.1 | 14.55 | 0.1 | 15.04 | 0 | 14.55 | 4.88 | 12.79 | 4.49 | 4.88 | 4.98 | 4.88 | 9.47 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 16.11 | 0.29 | 3.42 | 0.2 | 1.56 | 0.49 | 2.93 | 6.25 | 4.88 | 21.58 | 4.79 | 21.39 | 9.77 | 24.51 | 9.67 | 24.02 | 9.67 | 23.44
42.624 | 0 | 14.45 | 0 | 15.63 | 0 | 15.04 | 4.88 | 13.28 | 4.49 | 4.88 | 4.88 | 4.88 | 11.62 | 9.86 | 21.39 | 9.77 | 21.39 | 9.67 | 16.7 | 0.2 | 5.18 | 0.2 | 3.03 | 0.29 | 2.93 | 5.96 | 4.79 | 21.49 | 4.79 | 21.97 | 9.77 | 24.9 | 9.67 | 24.61 | 9.67 | 24.02
42.656 | 0 | 14.45 | 0.1 | 16.02 | 0.1 | 15.53 | 4.88 | 13.87 | 4.59 | 4.88 | 4.88 | 4.88 | 10.64 | 9.86 | 21.39 | 9.77 | 21.68 | 9.67 | 17.09 | 0.2 | 6.45 | 0.2 | 5.08 | 0.29 | 2.93 | 5.57 | 4.79 | 21.39 | 4.79 | 22.56 | 9.77 | 25.39 | 9.67 | 25.1 | 9.67 | 24.51
42.688 | 0 | 14.45 | 0.1 | 16.6 | 0.1 | 16.02 | 4.88 | 14.45 | 4.49 | 5.66 | 4.88 | 4.88 | 10.16 | 9.86 | 21.39 | 9.77 | 22.17 | 9.67 | 17.58 | 0.2 | 7.32 | 0.2 | 6.35 | 0.29 | 2.93 | 5.27 | 4.88 | 21.29 | 4.79 | 23.15 | 9.77 | 25.98 | 9.67 | 25.59 | 9.67 | 25.1
42.72 | 0 | 14.45 | 0.1 | 17.09 | 0 | 16.5 | 4.88 | 14.94 | 4.49 | 6.35 | 4.79 | 4.88 | 9.86 | 9.86 | 21.39 | 9.77 | 22.56 | 9.67 | 18.07 | 0.2 | 8.01 | 0.2 | 7.13 | 0.2 | 2.93 | 5.08 | 4.88 | 21.29 | 4.79 | 23.63 | 9.77 | 26.37 | 9.77 | 25.98 | 9.77 | 25.59
42.752 | 0 | 14.36 | 0.1 | 14.06 | 0 | 17.09 | 4.88 | 15.43 | 4.49 | 7.62 | 4.69 | 4.88 | 10.16 | 9.86 | 21.39 | 9.77 | 22.85 | 9.67 | 18.46 | 0.1 | 8.79 | 0.1 | 8.01 | 0.2 | 2.93 | 5.08 | 4.88 | 21.29 | 4.79 | 24.12 | 9.77 | 26.86 | 9.77 | 26.47 | 9.67 | 26.08
42.784 | 0 | 8.2 | 0.1 | 13.97 | 0 | 15.33 | 4.88 | 15.92 | 4.49 | 8.79 | 4.69 | 4.79 | 10.25 | 9.86 | 9.77 | 9.77 | 23.34 | 9.67 | 18.95 | 0.1 | 9.67 | 0.1 | 8.69 | 0.1 | 2.93 | 4.98 | 4.88 | 21.29 | 4.79 | 24.71 | 9.77 | 27.44 | 9.67 | 27.05 | 9.67 | 26.66
42.816 | 0 | 2.64 | 0 | 14.16 | 0.1 | 14.16 | 4.88 | 16.31 | 4.59 | 9.57 | 4.69 | 4.88 | 9.96 | 9.86 | 6.64 | 9.86 | 23.63 | 9.77 | 19.34 | 0.1 | 10.45 | 0.1 | 9.47 | 0.1 | 2.93 | 5.08 | 4.79 | 21.19 | 4.79 | 25.2 | 9.77 | 27.93 | 9.67 | 27.44 | 9.77 | 27.05
42.848 | 0 | 2.05 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 16.9 | 4.59 | 10.35 | 4.69 | 4.88 | 9.86 | 9.86 | 9.47 | 9.77 | 24.12 | 9.77 | 19.82 | 0.1 | 11.23 | 0.1 | 10.35 | 0.1 | 2.93 | 4.98 | 4.88 | 21.19 | 4.79 | 25.68 | 9.77 | 28.32 | 9.77 | 28.03 | 9.77 | 27.44
42.88 | 0.1 | 1.86 | 0.1 | 14.36 | 0.1 | 14.65 | 4.88 | 17.38 | 4.59 | 10.94 | 4.59 | 4.88 | 9.86 | 9.86 | 11.62 | 9.86 | 24.41 | 9.77 | 20.22 | 0.1 | 11.91 | 0.1 | 11.13 | 0.1 | 2.93 | 4.98 | 4.88 | 11.82 | 4.79 | 26.17 | 9.77 | 28.91 | 9.77 | 28.52 | 9.77 | 28.03
42.912 | 0 | 1.46 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 17.77 | 4.59 | 11.43 | 4.59 | 4.88 | 9.86 | 9.86 | 10.94 | 9.86 | 22.46 | 9.77 | 20.7 | 0.1 | 12.89 | 0.1 | 11.82 | 0.1 | 2.93 | 4.79 | 4.88 | 3.22 | 4.79 | 22.17 | 9.77 | 29.3 | 9.77 | 28.91 | 9.77 | 28.42
42.944 | 0 | 1.46 | 0 | 14.45 | 0 | 14.45 | 4.88 | 18.26 | 4.59 | 12.11 | 4.49 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 22.17 | 9.77 | 21.09 | 0.1 | 13.57 | 0.1 | 12.7 | 0.1 | 2.93 | 4.79 | 4.88 | 3.13 | 4.79 | 21.19 | 9.77 | 24.61 | 9.77 | 29.4 | 9.77 | 28.91
42.976 | 0 | 1.27 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 18.65 | 4.69 | 12.6 | 4.59 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 21.88 | 9.77 | 21.49 | 0.1 | 14.16 | 0.1 | 13.38 | 0 | 2.93 | 4.69 | 4.79 | 3.52 | 4.79 | 21.39 | 9.77 | 23.34 | 9.77 | 25.1 | 9.77 | 29.4
43.008 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.98 | 19.24 | 4.69 | 13.18 | 4.59 | 4.88 | 9.67 | 9.86 | 10.16 | 9.86 | 21.68 | 9.77 | 21.88 | 0.1 | 14.94 | 0.1 | 14.06 | 0.1 | 3.03 | 4.79 | 4.88 | 4.49 | 4.79 | 21.29 | 9.86 | 24.02 | 9.77 | 24.41 | 9.77 | 25.88
43.04 | 0 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.98 | 19.53 | 4.69 | 13.67 | 4.49 | 4.88 | 9.77 | 9.86 | 10.16 | 9.86 | 21.68 | 9.77 | 22.27 | 0.1 | 15.72 | 0.1 | 14.84 | 0.1 | 4.1 | 4.69 | 4.88 | 5.96 | 4.79 | 21.39 | 9.77 | 24.22 | 9.77 | 24.41 | 9.77 | 25
43.072 | 0 | 0.88 | 0.1 | 14.36 | 0 | 14.45 | 4.98 | 20.02 | 4.69 | 14.26 | 4.49 | 5.57 | 9.67 | 9.86 | 9.96 | 9.86 | 21.58 | 9.77 | 22.56 | 0.1 | 16.41 | 0.1 | 15.63 | 0.1 | 5.57 | 4.69 | 4.88 | 6.15 | 4.79 | 21.39 | 9.77 | 24.22 | 9.77 | 24.32 | 9.77 | 24.61
43.104 | 0 | 0.78 | 0 | 14.36 | 0 | 14.45 | 4.88 | 20.51 | 4.69 | 14.84 | 4.49 | 6.25 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 23.05 | 0 | 16.99 | 0.1 | 16.31 | 0 | 8.01 | 4.59 | 4.88 | 5.66 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.41 | 9.77 | 24.51
43.136 | 0 | 0.78 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 20.9 | 4.79 | 15.23 | 4.59 | 7.52 | 9.57 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 23.34 | 0.1 | 17.68 | 0.1 | 16.9 | 0 | 9.28 | 4.59 | 4.79 | 5.57 | 4.79 | 21.39 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77 | 24.41
43.168 | 0 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.98 | 17.97 | 4.79 | 15.72 | 4.59 | 8.69 | 9.67 | 10.45 | 9.86 | 9.86 | 21.39 | 9.77 | 23.73 | 0.1 | 18.36 | 0.1 | 17.58 | 0.1 | 10.45 | 4.59 | 4.88 | 5.27 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77 | 24.41
43.2 | 0 | 0.68 | 0.1 | 5.76 | 0.1 | 14.36 | 4.98 | 17.68 | 4.79 | 16.31 | 4.59 | 9.57 | 9.67 | 11.52 | 9.86 | 9.86 | 21.39 | 9.77 | 24.12 | 0.1 | 18.85 | 0.1 | 18.26 | 0.1 | 11.33 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86 | 24.22
43.232 | 0 | 0.59 | 0.1 | 1.95 | 0.1 | 6.74 | 4.98 | 17.97 | 4.79 | 16.7 | 4.59 | 10.25 | 9.57 | 12.79 | 9.77 | 9.86 | 21.39 | 9.77 | 23.93 | 0.1 | 19.53 | 0.1 | 18.85 | 0.1 | 12.21 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
43.264 | 0 | 0.49 | 0 | 1.95 | 0 | 1.76 | 4.88 | 18.07 | 4.79 | 17.19 | 4.59 | 10.84 | 9.57 | 13.77 | 9.67 | 9.86 | 21.39 | 9.86 | 21.29 | 0.1 | 20.22 | 0.1 | 19.53 | 0 | 12.99 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22
43.296 | 0 | 0.39 | 0 | 1.66 | 0.1 | 1.95 | 4.98 | 18.07 | 4.79 | 17.68 | 4.69 | 11.43 | 9.57 | 14.45 | 9.67 | 9.86 | 21.39 | 9.86 | 21.68 | 0.1 | 20.8 | 0.1 | 20.12 | 0 | 13.67 | 4.59 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
43.328 | 0 | 0.39 | 0.1 | 1.37 | 0.1 | 1.76 | 4.98 | 18.07 | 4.79 | 18.16 | 4.69 | 12.01 | 9.57 | 15.04 | 9.67 | 9.86 | 14.55 | 9.86 | 21.68 | 0.1 | 21.39 | 0.1 | 20.7 | 0.1 | 14.36 | 4.59 | 4.88 | 4.98 | 4.79 | 18.07 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
43.36 | 0 | 0.29 | 0.1 | 1.46 | 0.1 | 1.46 | 4.98 | 18.16 | 4.79 | 18.46 | 4.69 | 12.5 | 9.57 | 15.63 | 9.67 | 9.86 | 7.23 | 9.86 | 21.58 | 0.1 | 22.07 | 0.1 | 21.39 | 0.1 | 15.23 | 4.59 | 4.88 | 4.88 | 4.79 | 4.3 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
43.392 | 0.68 | 0.29 | 0.1 | 1.17 | 0 | 1.37 | 4.98 | 18.07 | 4.79 | 18.95 | 4.69 | 13.09 | 9.57 | 16.11 | 9.67 | 9.86 | 8.01 | 9.86 | 21.58 | 0.1 | 22.56 | 0.1 | 21.97 | 0.1 | 15.92 | 4.59 | 4.88 | 4.79 | 4.79 | 3.03 | 9.77 | 10.74 | 9.86 | 21 | 9.86 | 24.12
43.424 | 1.56 | 0.2 | 0.1 | 1.07 | 0 | 1.27 | 4.88 | 18.07 | 4.79 | 19.43 | 4.69 | 13.67 | 9.57 | 16.6 | 9.67 | 9.86 | 10.84 | 9.86 | 21.49 | 0.1 | 19.14 | 0.1 | 22.46 | 0 | 16.5 | 4.59 | 5.76 | 4.79 | 4.79 | 3.32 | 9.77 | 5.96 | 9.77 | 7.03 | 9.77 | 16.6
43.456 | 2.73 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.98 | 18.07 | 4.79 | 19.92 | 4.79 | 14.16 | 9.57 | 17.09 | 9.67 | 9.86 | 11.33 | 9.86 | 21.49 | 0.1 | 18.46 | 0.1 | 18.16 | 0 | 17.19 | 4.59 | 6.74 | 4.69 | 4.79 | 4.3 | 9.77 | 9.28 | 9.77 | 6.54 | 9.86 | 7.03
43.488 | 3.81 | 0.2 | 0.1 | 1.07 | 0 | 1.07 | 4.98 | 18.07 | 4.79 | 20.22 | 4.79 | 14.75 | 9.67 | 17.58 | 9.67 | 9.86 | 10.45 | 9.86 | 21.49 | 0.1 | 18.07 | 0.1 | 16.8 | 0.1 | 17.87 | 4.69 | 8.4 | 4.69 | 4.79 | 5.66 | 9.86 | 12.31 | 9.86 | 10.45 | 9.86 | 7.03
43.52 | 4.59 | 0.1 | 0.1 | 0.78 | 0 | 0.88 | 4.98 | 18.07 | 4.88 | 20.7 | 4.79 | 15.14 | 9.67 | 17.97 | 9.57 | 9.86 | 9.96 | 9.86 | 21.49 | 0.1 | 17.87 | 0.1 | 17.58 | 0.1 | 18.56 | 4.69 | 10.25 | 4.59 | 4.79 | 6.25 | 9.77 | 11.23 | 9.86 | 12.21 | 9.86 | 10.74
43.552 | 5.08 | 0.1 | 0.1 | 0.68 | 0 | 0.88 | 4.98 | 18.07 | 4.79 | 21 | 4.79 | 15.72 | 9.67 | 18.46 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 17.97 | 0.1 | 17.68 | 0.1 | 19.24 | 4.69 | 11.72 | 4.59 | 4.79 | 5.86 | 9.77 | 10.16 | 9.86 | 10.74 | 9.86 | 12.01
43.584 | 5.76 | 0.1 | 0 | 0.59 | 0 | 0.78 | 4.88 | 18.07 | 4.79 | 18.95 | 4.79 | 16.11 | 9.67 | 18.95 | 9.57 | 9.86 | 10.06 | 9.86 | 21.39 | 0 | 17.77 | 0.1 | 17.68 | 0 | 19.82 | 4.69 | 12.7 | 4.59 | 4.79 | 5.37 | 9.77 | 10.16 | 9.77 | 9.96 | 9.77 | 10.94
43.616 | 6.25 | 0.1 | 0.1 | 0.59 | 0.1 | 0.59 | 4.98 | 6.93 | 4.88 | 18.85 | 4.88 | 16.6 | 9.67 | 19.34 | 9.57 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.68 | 0 | 20.31 | 4.79 | 13.48 | 4.59 | 4.79 | 5.27 | 9.77 | 10.35 | 9.77 | 10.06 | 9.86 | 9.86
43.648 | 6.84 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 4.98 | 2.54 | 4.88 | 18.56 | 4.79 | 17.09 | 9.67 | 19.92 | 9.67 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 21 | 4.79 | 14.26 | 4.49 | 4.79 | 5.08 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86 | 9.96
43.68 | 7.52 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.98 | 3.42 | 4.79 | 18.46 | 4.88 | 17.58 | 9.67 | 20.22 | 9.67 | 9.86 | 9.77 | 9.96 | 8.01 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 21.58 | 4.79 | 15.04 | 4.49 | 4.79 | 5.18 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86 | 10.35
43.712 | 8.01 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.98 | 4.39 | 4.79 | 18.46 | 4.88 | 18.07 | 9.67 | 20.7 | 9.67 | 9.86 | 9.86 | 9.86 | 1.46 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 22.17 | 4.79 | 15.72 | 4.49 | 5.66 | 5.08 | 9.77 | 10.06 | 9.86 | 10.25 | 9.86 | 10.45
43.744 | 8.69 | 0.1 | 0 | 0.29 | 0 | 0.39 | 4.88 | 5.57 | 4.79 | 18.16 | 4.79 | 18.46 | 9.67 | 21.09 | 9.67 | 9.86 | 9.86 | 9.96 | 2.44 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 22.85 | 4.79 | 16.5 | 4.49 | 6.74 | 4.98 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86 | 10.25
43.776 | 9.38 | 0.1 | 0.49 | 0.2 | 0 | 0.29 | 4.88 | 5.76 | 4.79 | 18.16 | 4.88 | 18.95 | 9.67 | 21.49 | 9.67 | 9.86 | 9.77 | 9.96 | 2.15 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 23.24 | 4.79 | 17.09 | 4.59 | 8.5 | 4.98 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86 | 9.96
43.808 | 9.86 | 0.1 | 1.27 | 0.2 | 0.1 | 0.29 | 4.98 | 5.47 | 4.79 | 18.26 | 4.88 | 19.34 | 9.77 | 21.88 | 9.67 | 9.96 | 9.77 | 9.96 | 2.05 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 23.73 | 4.79 | 17.87 | 4.59 | 10.35 | 4.98 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86 | 9.96
43.84 | 10.45 | 0.1 | 2.25 | 0.2 | 0.1 | 0.2 | 4.98 | 5.18 | 4.79 | 18.16 | 4.88 | 19.82 | 9.77 | 22.27 | 9.67 | 10.64 | 9.77 | 9.96 | 1.95 | 0.1 | 12.6 | 0.1 | 17.68 | 0.1 | 24.32 | 4.79 | 18.46 | 4.59 | 11.62 | 4.79 | 9.86 | 9.96 | 9.77 | 10.06 | 9.86 | 10.16
43.872 | 11.13 | 0.1 | 3.71 | 0.1 | 0 | 0.2 | 4.88 | 4.98 | 4.79 | 18.16 | 4.88 | 20.22 | 9.77 | 22.66 | 9.67 | 11.72 | 9.67 | 9.96 | 1.76 | 0.1 | 2.73 | 0.1 | 17.58 | 0.1 | 24.9 | 4.88 | 19.14 | 4.59 | 12.7 | 4.79 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
43.904 | 11.62 | 0.1 | 4.49 | 0.1 | 0 | 0.2 | 4.88 | 4.98 | 4.79 | 18.07 | 4.79 | 20.7 | 9.77 | 23.05 | 9.67 | 13.09 | 9.67 | 9.96 | 1.76 | 0 | 2.44 | 0.1 | 4.39 | 0 | 20.31 | 4.79 | 19.92 | 4.59 | 13.57 | 4.69 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
43.936 | 12.21 | 0.1 | 4.98 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 4.79 | 18.07 | 4.88 | 21 | 9.77 | 23.44 | 9.67 | 13.97 | 9.67 | 9.96 | 1.37 | 0.1 | 1.95 | 0.1 | 1.95 | 0 | 19.14 | 4.88 | 20.22 | 4.59 | 14.26 | 4.69 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86
43.968 | 12.79 | 0.1 | 5.57 | 0.1 | 1.27 | 0.1 | 4.98 | 4.88 | 4.79 | 18.07 | 4.88 | 19.04 | 9.77 | 23.83 | 9.67 | 14.65 | 9.67 | 9.96 | 1.46 | 0.1 | 1.66 | 0.1 | 2.25 | 0.1 | 19.73 | 4.88 | 20.9 | 4.59 | 15.04 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
44 | 13.28 | 0.1 | 6.15 | 0.1 | 2.54 | 0.1 | 5.57 | 4.79 | 4.79 | 10.16 | 4.88 | 18.85 | 9.77 | 24.22 | 9.67 | 15.14 | 9.57 | 9.96 | 1.27 | 0.1 | 1.86 | 0.1 | 1.86 | 0.1 | 20.02 | 4.88 | 21.58 | 4.69 | 15.82 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
44.032 | 13.77 | 0.1 | 6.74 | 0.1 | 3.81 | 0.1 | 6.35 | 4.79 | 4.79 | 3.32 | 4.79 | 18.56 | 9.77 | 23.05 | 9.67 | 15.72 | 9.57 | 9.96 | 1.17 | 0.1 | 1.56 | 0.1 | 1.66 | 0.1 | 20.02 | 4.88 | 22.07 | 4.69 | 16.41 | 4.59 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
44.064 | 14.36 | 0.1 | 7.32 | 0.1 | 4.49 | 0.1 | 7.52 | 4.69 | 4.79 | 3.22 | 4.79 | 18.26 | 9.77 | 21 | 9.67 | 16.21 | 9.57 | 9.96 | 1.37 | 0.1 | 1.37 | 0.1 | 1.46 | 0 | 20.02 | 4.79 | 22.75 | 4.69 | 17.09 | 4.59 | 11.82 | 9.67 | 9.77 | 9.77 | 10.74 | 9.77
44.096 | 14.94 | 0.1 | 7.91 | 0 | 5.08 | 0.1 | 8.69 | 4.59 | 4.79 | 3.81 | 4.88 | 18.36 | 9.77 | 21.49 | 9.67 | 16.7 | 9.57 | 9.96 | 0.88 | 0 | 1.27 | 0.1 | 1.37 | 0 | 20.02 | 4.88 | 23.15 | 4.69 | 17.77 | 4.59 | 13.48 | 9.77 | 10.64 | 9.77 | 11.43 | 9.77
44.128 | 15.33 | 0.1 | 8.5 | 0.1 | 5.57 | 0.1 | 9.57 | 4.59 | 4.79 | 5.27 | 4.88 | 18.16 | 9.77 | 21.58 | 9.77 | 17.19 | 9.57 | 9.96 | 1.07 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 20.12 | 4.88 | 23.73 | 4.69 | 18.46 | 4.59 | 15.14 | 9.67 | 11.43 | 9.77 | 12.89 | 9.77
44.16 | 15.92 | 0.1 | 9.08 | 0.1 | 6.15 | 0.1 | 10.16 | 4.59 | 4.88 | 6.05 | 4.88 | 18.16 | 9.77 | 21.49 | 9.77 | 17.68 | 9.57 | 9.96 | 0.88 | 0.1 | 1.17 | 0.1 | 1.27 | 0.1 | 20.12 | 4.88 | 24.22 | 4.69 | 19.14 | 4.49 | 16.41 | 9.67 | 12.99 | 9.67 | 15.04 | 9.67
44.192 | 16.31 | 0.1 | 9.67 | 0.1 | 6.84 | 0.1 | 10.84 | 4.59 | 4.79 | 5.66 | 4.88 | 18.16 | 9.77 | 21.49 | 9.77 | 18.16 | 9.57 | 9.96 | 0.78 | 0.1 | 0.88 | 0.1 | 1.27 | 0.1 | 20.12 | 4.88 | 24.81 | 4.79 | 19.73 | 4.49 | 17.29 | 9.67 | 15.04 | 9.67 | 16.31 | 9.67
44.224 | 16.8 | 0.1 | 10.25 | 0 | 7.42 | 0.1 | 11.52 | 4.49 | 4.79 | 5.27 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.56 | 9.57 | 9.96 | 0.78 | 0 | 0.88 | 0.1 | 1.07 | 0 | 20.02 | 4.79 | 25.29 | 4.79 | 20.31 | 4.49 | 18.16 | 9.67 | 16.31 | 9.67 | 17.19 | 9.67
44.256 | 17.29 | 0.1 | 10.84 | 0 | 8.01 | 0.1 | 12.01 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 19.04 | 9.57 | 9.96 | 0.68 | 0 | 0.78 | 0.1 | 0.88 | 0 | 20.02 | 4.88 | 25.78 | 4.79 | 20.9 | 4.59 | 18.85 | 9.67 | 17.19 | 9.57 | 17.97 | 9.67
44.288 | 15.33 | 0.1 | 11.52 | 0.1 | 8.59 | 0.1 | 12.6 | 4.59 | 4.79 | 4.98 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 19.34 | 9.67 | 9.96 | 0.59 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 20.02 | 4.88 | 26.37 | 4.79 | 21.49 | 4.59 | 19.43 | 9.67 | 17.97 | 9.67 | 18.65 | 9.67
44.32 | 15.14 | 0.1 | 12.01 | 0.1 | 9.28 | 0.1 | 13.18 | 4.59 | 4.79 | 4.98 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 19.92 | 9.67 | 9.96 | 0.59 | 0.1 | 0.59 | 0.1 | 0.88 | 0.1 | 20.02 | 4.88 | 23.05 | 4.79 | 22.07 | 4.59 | 20.12 | 9.67 | 18.75 | 9.67 | 19.34 | 9.67
44.352 | 14.84 | 0.1 | 12.6 | 0.1 | 9.77 | 0.1 | 13.67 | 4.59 | 4.79 | 4.98 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 20.31 | 9.67 | 10.55 | 0.49 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 6.64 | 4.88 | 22.27 | 4.79 | 22.66 | 4.59 | 20.7 | 9.67 | 19.34 | 9.67 | 20.02 | 9.67
44.384 | 14.65 | 0.1 | 13.18 | 0 | 10.35 | 0.1 | 14.16 | 4.49 | 4.79 | 4.88 | 4.88 | 9.38 | 9.77 | 21.39 | 9.77 | 20.7 | 9.67 | 11.23 | 0.49 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 2.64 | 4.88 | 21.78 | 4.79 | 23.24 | 4.59 | 21.39 | 9.67 | 20.02 | 9.67 | 20.61 | 9.57
44.416 | 14.65 | 0.1 | 13.67 | 0.1 | 11.04 | 0.1 | 14.75 | 4.59 | 4.79 | 4.79 | 4.88 | 3.22 | 9.77 | 21.39 | 9.77 | 21.19 | 9.67 | 12.31 | 0.39 | 0 | 0.49 | 0.1 | 0.59 | 0 | 3.32 | 4.88 | 21.58 | 4.79 | 23.73 | 4.69 | 21.97 | 9.67 | 20.51 | 9.57 | 21.19 | 9.57
44.448 | 14.55 | 0.1 | 14.16 | 0.1 | 11.62 | 0.1 | 15.23 | 4.59 | 4.79 | 4.69 | 4.88 | 3.22 | 9.77 | 21.39 | 9.77 | 21.49 | 9.67 | 13.48 | 0.39 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4 | 4.88 | 21.58 | 4.79 | 24.22 | 4.69 | 22.56 | 9.67 | 21.19 | 9.67 | 21.68 | 9.57
44.48 | 14.45 | 0.1 | 14.75 | 0.1 | 12.11 | 0.1 | 15.72 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 9.77 | 12.11 | 9.86 | 21.88 | 9.67 | 14.36 | 0.29 | 0.1 | 0.29 | 0.1 | 0.49 | 0.1 | 5.27 | 4.88 | 21.49 | 4.79 | 24.71 | 4.69 | 23.15 | 9.67 | 21.78 | 9.67 | 22.27 | 9.67
44.512 | 14.45 | 0.1 | 15.23 | 0.1 | 12.7 | 0.1 | 16.21 | 4.59 | 4.79 | 4.69 | 4.88 | 5.08 | 9.77 | 6.74 | 9.86 | 22.36 | 9.67 | 14.94 | 0.29 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 6.05 | 4.88 | 21.39 | 4.79 | 25.29 | 4.69 | 23.54 | 9.67 | 22.17 | 9.67 | 22.95 | 9.67
44.544 | 14.36 | 0.1 | 15.72 | 0 | 13.18 | 0.1 | 16.8 | 4.59 | 4.79 | 4.59 | 4.79 | 5.86 | 9.77 | 9.28 | 9.77 | 22.66 | 9.67 | 15.43 | 0.2 | 0 | 0.2 | 0.88 | 0.29 | 0 | 5.86 | 4.88 | 21.39 | 4.79 | 25.78 | 4.69 | 24.12 | 9.67 | 22.95 | 9.67 | 23.44 | 9.67
44.576 | 14.36 | 0.1 | 16.21 | 0.1 | 13.77 | 0.1 | 17.29 | 4.69 | 4.88 | 4.59 | 4.88 | 5.66 | 9.77 | 11.43 | 9.77 | 23.05 | 9.67 | 16.02 | 0.2 | 0.1 | 0.2 | 1.86 | 0.29 | 0 | 5.37 | 4.88 | 21.29 | 4.79 | 26.27 | 4.69 | 24.61 | 9.67 | 23.34 | 9.67 | 23.93 | 9.67
44.608 | 14.36 | 0.1 | 16.7 | 0.1 | 14.26 | 0.1 | 17.68 | 4.69 | 4.79 | 4.59 | 4.88 | 5.27 | 9.77 | 11.04 | 9.86 | 23.54 | 9.67 | 16.5 | 0.2 | 1.07 | 0.2 | 3.42 | 0.29 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 23.05 | 4.79 | 25.2 | 9.67 | 24.02 | 9.67 | 24.41 | 9.67
44.64 | 14.36 | 0.1 | 17.19 | 0.1 | 14.84 | 0.1 | 18.16 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 9.77 | 10.25 | 9.86 | 23.83 | 9.67 | 16.99 | 0.1 | 2.05 | 0.2 | 5.27 | 0.2 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 22.17 | 4.79 | 25.68 | 9.67 | 24.61 | 9.67 | 25.1 | 9.67
44.672 | 14.36 | 0.1 | 13.97 | 0.1 | 15.23 | 0.1 | 18.56 | 4.69 | 6.45 | 4.59 | 4.88 | 4.88 | 9.77 | 9.86 | 9.86 | 24.22 | 9.67 | 17.48 | 0.1 | 3.71 | 0.1 | 6.45 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 21.78 | 4.69 | 26.17 | 9.67 | 25 | 9.67 | 25.49 | 9.67
44.704 | 8.01 | 0.1 | 14.65 | 0 | 15.72 | 0.1 | 19.04 | 4.69 | 7.62 | 4.49 | 4.79 | 4.88 | 9.77 | 10.06 | 9.77 | 21.19 | 9.77 | 17.87 | 0.1 | 5.47 | 0.1 | 7.32 | 0.1 | 1.07 | 5.08 | 4.88 | 21.19 | 4.79 | 21.58 | 4.69 | 26.76 | 9.67 | 25.49 | 9.67 | 26.08 | 9.67
44.736 | 2.44 | 0.1 | 14.36 | 0 | 16.31 | 0.1 | 19.43 | 4.79 | 8.69 | 4.49 | 4.88 | 4.88 | 9.77 | 10.25 | 9.77 | 21 | 9.67 | 18.46 | 0.1 | 6.54 | 0.1 | 8.11 | 0.1 | 2.05 | 4.88 | 4.88 | 13.09 | 4.79 | 21.49 | 4.79 | 27.25 | 9.77 | 25.98 | 9.67 | 26.47 | 9.67
44.768 | 1.95 | 0.1 | 14.45 | 0.1 | 16.7 | 0.1 | 20.02 | 4.79 | 9.67 | 4.59 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 21.19 | 9.77 | 18.85 | 0.1 | 7.32 | 0.1 | 8.79 | 0.2 | 3.71 | 4.79 | 4.88 | 3.13 | 4.79 | 21.49 | 4.79 | 27.54 | 9.77 | 26.47 | 9.77 | 26.95 | 9.77
44.8 | 1.66 | 0.1 | 14.45 | 0.1 | 17.19 | 0.1 | 20.31 | 4.79 | 10.25 | 4.59 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 19.24 | 0.1 | 8.11 | 0.1 | 9.77 | 0.1 | 5.37 | 4.79 | 4.88 | 3.32 | 4.79 | 21.29 | 4.79 | 28.13 | 9.77 | 26.95 | 9.77 | 27.54 | 9.77
44.832 | 1.37 | 0.1 | 14.45 | 0.1 | 15.43 | 0.1 | 20.7 | 4.79 | 10.94 | 4.59 | 4.88 | 4.79 | 9.77 | 9.86 | 9.86 | 21.39 | 9.77 | 19.73 | 0.1 | 8.79 | 0 | 10.55 | 0.1 | 6.54 | 4.79 | 4.88 | 3.42 | 4.88 | 21.39 | 4.79 | 28.61 | 9.77 | 27.54 | 9.77 | 27.83 | 9.77
44.864 | 1.37 | 0.1 | 14.55 | 0 | 15.23 | 0.1 | 21.19 | 4.79 | 11.43 | 4.59 | 4.79 | 4.59 | 9.77 | 9.96 | 9.77 | 21.39 | 9.77 | 20.12 | 0.1 | 9.67 | 0 | 11.33 | 0.1 | 7.32 | 4.59 | 4.88 | 4.49 | 4.88 | 21.29 | 4.79 | 29.2 | 9.77 | 28.03 | 9.77 | 28.32 | 9.77
44.896 | 1.17 | 0.1 | 14.45 | 0 | 14.75 | 0.1 | 19.14 | 4.79 | 12.11 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.77 | 21.49 | 9.77 | 20.61 | 0 | 10.45 | 0.1 | 12.01 | 0.1 | 8.11 | 4.59 | 4.88 | 6.25 | 4.88 | 21.29 | 4.79 | 29.59 | 9.77 | 28.42 | 9.77 | 28.91 | 9.77
44.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 19.14 | 4.88 | 12.7 | 4.69 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.9 | 0.1 | 11.33 | 0.1 | 12.89 | 0.1 | 8.98 | 4.59 | 4.88 | 6.25 | 4.79 | 21.19 | 4.79 | 26.37 | 9.77 | 28.91 | 9.77 | 29.4 | 9.77
44.96 | 0.98 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 18.65 | 4.88 | 13.18 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.29 | 0.1 | 12.21 | 0.1 | 13.57 | 0.1 | 9.96 | 4.59 | 4.88 | 5.86 | 4.79 | 21.19 | 4.79 | 25.39 | 9.77 | 29.4 | 9.77 | 25 | 9.77
44.992 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.67 | 4.69 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 12.79 | 0 | 14.36 | 0.1 | 10.74 | 4.59 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 24.9 | 9.77 | 25.1 | 9.77 | 24.71 | 9.77
45.024 | 0.78 | 0.1 | 14.45 | 0 | 14.55 | 0.1 | 18.46 | 4.88 | 14.16 | 4.69 | 5.66 | 4.49 | 9.77 | 9.77 | 9.77 | 21.39 | 9.77 | 22.17 | 0.1 | 13.57 | 0 | 15.04 | 0.1 | 11.43 | 4.49 | 4.88 | 5.18 | 4.88 | 11.62 | 4.79 | 24.71 | 9.77 | 24.81 | 9.77 | 24.32 | 9.77
45.056 | 0.68 | 0.1 | 14.36 | 0 | 14.55 | 0.1 | 18.36 | 4.88 | 14.84 | 4.79 | 6.25 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.56 | 0 | 14.26 | 0.1 | 15.82 | 0.1 | 12.21 | 4.49 | 4.88 | 5.18 | 4.79 | 3.13 | 4.79 | 24.71 | 9.77 | 24.41 | 9.77 | 24.22 | 9.77
45.088 | 0.59 | 0.1 | 14.16 | 0.1 | 14.45 | 0.1 | 18.36 | 4.88 | 15.23 | 4.79 | 7.42 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 0.1 | 15.04 | 0 | 16.5 | 0.1 | 12.89 | 4.49 | 4.88 | 5.08 | 4.79 | 3.13 | 4.79 | 24.61 | 9.77 | 24.32 | 9.86 | 24.32 | 9.77
45.12 | 0.59 | 0.1 | 3.91 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.72 | 4.79 | 8.59 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 23.34 | 0.1 | 15.72 | 0.1 | 17.19 | 0.1 | 13.77 | 4.49 | 5.76 | 5.08 | 4.79 | 3.52 | 4.79 | 24.41 | 9.77 | 24.41 | 9.86 | 24.32 | 9.77
45.152 | 0.49 | 0.1 | 1.95 | 0.1 | 14.36 | 0.1 | 18.26 | 4.88 | 16.21 | 4.79 | 9.47 | 4.49 | 9.77 | 9.67 | 9.86 | 10.64 | 9.77 | 23.73 | 0.1 | 16.41 | 0 | 17.77 | 0.1 | 14.45 | 4.49 | 7.13 | 4.98 | 4.88 | 4.49 | 4.79 | 24.41 | 9.77 | 24.41 | 9.86 | 24.32 | 9.77
45.184 | 0.29 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 18.16 | 4.88 | 16.7 | 4.79 | 10.06 | 4.49 | 9.86 | 9.67 | 9.86 | 6.54 | 9.86 | 24.02 | 0.1 | 17.09 | 0 | 18.46 | 0.1 | 15.23 | 4.49 | 9.38 | 4.88 | 4.88 | 6.05 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.32 | 9.86
45.216 | 0.29 | 0.1 | 1.46 | 0 | 14.36 | 0.1 | 18.16 | 4.88 | 17.19 | 4.79 | 10.74 | 4.59 | 10.55 | 9.67 | 9.77 | 9.08 | 9.77 | 24.41 | 0 | 17.68 | 0 | 19.14 | 0.1 | 16.02 | 4.49 | 10.74 | 4.88 | 4.79 | 6.15 | 4.79 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
45.248 | 0.29 | 0.1 | 1.46 | 0.1 | 11.33 | 0.1 | 18.16 | 4.88 | 17.68 | 4.88 | 11.33 | 4.59 | 11.62 | 9.67 | 9.86 | 11.43 | 9.86 | 21.97 | 0.1 | 18.46 | 0 | 19.73 | 0.1 | 16.6 | 4.49 | 12.01 | 4.79 | 4.79 | 5.76 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
45.28 | 0.29 | 0.1 | 1.27 | 0.1 | 3.13 | 0.1 | 18.16 | 4.88 | 18.07 | 4.88 | 11.91 | 4.59 | 12.89 | 9.67 | 9.86 | 11.04 | 9.86 | 22.27 | 0.1 | 19.04 | 0.1 | 20.31 | 0.1 | 17.19 | 4.49 | 12.89 | 4.79 | 4.79 | 5.27 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
45.312 | 0.2 | 0.1 | 1.17 | 0.1 | 2.05 | 0.1 | 11.91 | 4.88 | 18.56 | 4.88 | 12.6 | 4.59 | 13.87 | 9.67 | 9.86 | 10.16 | 9.86 | 21.88 | 0.1 | 19.53 | 0 | 20.9 | 0.1 | 17.97 | 4.59 | 13.77 | 4.79 | 4.79 | 5.27 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
45.344 | 0.1 | 0.1 | 1.07 | 0 | 1.76 | 0.1 | 3.32 | 4.88 | 18.95 | 4.88 | 12.99 | 4.69 | 14.65 | 9.67 | 9.86 | 9.86 | 9.86 | 21.78 | 0.1 | 20.31 | 0 | 21.58 | 0.1 | 18.56 | 4.59 | 14.55 | 4.69 | 4.79 | 5.08 | 4.79 | 14.16 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
45.376 | 0.1 | 0.1 | 0.98 | 0 | 1.56 | 0.1 | 3.32 | 4.88 | 19.43 | 4.88 | 13.67 | 4.69 | 15.14 | 9.67 | 9.77 | 10.16 | 9.77 | 21.68 | 0 | 20.8 | 0 | 22.07 | 0.1 | 19.14 | 4.59 | 15.23 | 4.69 | 4.88 | 5.08 | 4.79 | 6.54 | 9.77 | 24.22 | 9.86 | 21.39 | 9.86
45.408 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.91 | 4.88 | 19.92 | 4.88 | 14.16 | 4.69 | 15.63 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 0.1 | 21.39 | 0 | 22.66 | 0.1 | 19.82 | 4.69 | 16.02 | 4.69 | 4.79 | 5.18 | 4.79 | 7.42 | 9.77 | 20.9 | 9.86 | 8.11 | 9.86
45.44 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 4.69 | 4.88 | 20.31 | 4.88 | 14.65 | 4.69 | 16.11 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 0.1 | 21.97 | 0 | 19.24 | 0.1 | 20.31 | 4.69 | 16.7 | 4.59 | 4.88 | 4.98 | 4.79 | 11.43 | 9.77 | 7.81 | 9.86 | 6.54 | 9.86
45.472 | 0 | 0.68 | 0.68 | 0.1 | 1.07 | 0.1 | 5.76 | 4.88 | 20.7 | 4.88 | 15.23 | 4.79 | 16.7 | 9.67 | 10.55 | 9.86 | 9.86 | 21.58 | 0.1 | 22.56 | 0 | 18.65 | 0.1 | 20.9 | 4.69 | 17.48 | 4.59 | 4.88 | 4.98 | 4.79 | 11.72 | 9.77 | 6.64 | 9.86 | 10.64 | 9.86
45.504 | 0 | 1.46 | 0.59 | 0 | 0.88 | 0.1 | 5.76 | 4.88 | 20.61 | 4.88 | 15.63 | 4.79 | 17.19 | 9.67 | 11.62 | 9.86 | 9.86 | 21.49 | 0.1 | 18.95 | 0 | 18.16 | 0.1 | 21.58 | 4.69 | 17.97 | 4.59 | 4.79 | 4.88 | 4.79 | 10.45 | 9.77 | 10.55 | 9.86 | 12.11 | 9.86
45.536 | 0 | 2.64 | 0.59 | 0 | 0.98 | 0.1 | 5.27 | 4.88 | 18.56 | 4.88 | 16.21 | 4.79 | 17.68 | 9.67 | 12.89 | 10.06 | 9.86 | 21.49 | 0 | 18.46 | 0.1 | 18.07 | 0.1 | 22.07 | 4.69 | 18.65 | 4.59 | 4.88 | 4.79 | 4.79 | 10.06 | 9.77 | 12.21 | 9.77 | 10.94 | 9.86
45.568 | 0 | 3.91 | 0.49 | 0.1 | 0.78 | 0.1 | 5.08 | 4.88 | 18.46 | 4.88 | 16.6 | 4.79 | 18.07 | 9.67 | 13.77 | 9.86 | 9.86 | 21.49 | 0.1 | 18.07 | 0 | 17.97 | 0.1 | 22.75 | 4.79 | 19.34 | 4.49 | 5.76 | 4.79 | 4.79 | 10.25 | 9.77 | 11.13 | 9.86 | 9.96 | 9.86
45.6 | 0 | 4.59 | 0.49 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.46 | 4.88 | 17.09 | 4.79 | 18.56 | 9.77 | 14.45 | 9.77 | 9.86 | 21.49 | 0.1 | 17.97 | 0.1 | 17.87 | 0.1 | 19.53 | 4.79 | 19.92 | 4.59 | 6.54 | 4.69 | 4.79 | 10.45 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
45.632 | 0.1 | 5.18 | 0.29 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 17.58 | 4.79 | 18.95 | 9.77 | 15.04 | 9.67 | 9.86 | 21.39 | 0.1 | 17.87 | 0 | 17.77 | 0.1 | 18.75 | 4.79 | 20.61 | 4.49 | 8.2 | 4.69 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 10.35 | 9.86
45.664 | 0 | 5.66 | 0.29 | 0 | 0.59 | 0.1 | 4.88 | 4.88 | 18.26 | 4.88 | 17.97 | 4.79 | 19.34 | 9.77 | 15.63 | 9.77 | 9.86 | 17.09 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 18.16 | 4.79 | 21.19 | 4.59 | 10.25 | 4.69 | 4.79 | 10.16 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86
45.696 | 0.1 | 6.25 | 0.2 | 0 | 0.59 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.56 | 4.88 | 19.82 | 9.77 | 16.11 | 9.67 | 9.86 | 7.91 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 18.07 | 4.79 | 21.78 | 4.59 | 11.52 | 4.59 | 4.79 | 10.16 | 9.77 | 10.45 | 9.86 | 10.25 | 9.86
45.728 | 0.1 | 6.84 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.95 | 4.88 | 20.22 | 9.77 | 16.6 | 9.67 | 9.86 | 7.52 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 18.07 | 4.88 | 22.36 | 4.59 | 12.6 | 4.59 | 4.79 | 10.16 | 9.77 | 10.25 | 9.86 | 10.06 | 9.86
45.76 | 0.1 | 7.52 | 0.2 | 0.1 | 0.39 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.34 | 4.88 | 20.8 | 9.77 | 17.09 | 9.67 | 9.86 | 10.64 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.87 | 4.88 | 22.85 | 4.59 | 13.38 | 4.59 | 4.79 | 10.16 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
45.792 | 0 | 8.11 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 21.19 | 9.77 | 17.58 | 9.67 | 9.86 | 11.43 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.88 | 23.44 | 4.59 | 14.16 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
45.824 | 0 | 8.69 | 0.1 | 0 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 21.58 | 9.77 | 18.07 | 9.67 | 9.86 | 10.64 | 0 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.88 | 23.93 | 4.59 | 14.94 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
45.856 | 0 | 9.28 | 0.1 | 0.1 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.7 | 4.88 | 21.88 | 9.77 | 18.56 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0.1 | 12.89 | 0.1 | 17.68 | 4.88 | 24.41 | 4.69 | 15.63 | 4.49 | 4.79 | 9.96 | 9.77 | 10.06 | 9.86 | 9.86 | 9.86
45.888 | 0.1 | 9.86 | 0.1 | 1.07 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.36 | 9.77 | 18.95 | 9.57 | 9.86 | 9.96 | 0.1 | 17.58 | 0.1 | 3.03 | 0.1 | 17.77 | 4.88 | 24.9 | 4.69 | 16.41 | 4.49 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
45.92 | 0 | 10.45 | 0.1 | 1.86 | 0.1 | 0.88 | 4.59 | 4.88 | 18.07 | 4.88 | 19.14 | 4.88 | 22.75 | 9.86 | 19.43 | 9.67 | 9.86 | 10.16 | 0.1 | 9.67 | 0.1 | 2.25 | 0.1 | 17.68 | 4.88 | 25.49 | 4.69 | 17.09 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
45.952 | 0.1 | 11.04 | 0.1 | 3.13 | 0.1 | 1.86 | 4.59 | 4.88 | 8.4 | 4.88 | 18.85 | 4.88 | 23.05 | 9.86 | 19.82 | 9.57 | 9.86 | 10.16 | 0.1 | 2.64 | 0 | 1.95 | 0.1 | 17.68 | 4.88 | 25.98 | 4.69 | 17.68 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.77 | 9.86
45.984 | 0 | 11.62 | 0 | 4.2 | 0 | 3.22 | 4.49 | 4.88 | 2.83 | 4.88 | 18.56 | 4.88 | 23.54 | 9.77 | 20.22 | 9.57 | 9.86 | 9.96 | 0.1 | 2.15 | 0 | 1.76 | 0.1 | 17.68 | 4.88 | 24.22 | 4.69 | 18.36 | 4.59 | 4.79 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
46.016 | 0 | 12.11 | 0.1 | 4.79 | 0 | 4 | 4.49 | 4.88 | 3.13 | 4.88 | 18.36 | 4.88 | 23.83 | 9.77 | 20.61 | 9.67 | 10.45 | 9.96 | 0 | 1.95 | 0.1 | 1.86 | 0.1 | 9.77 | 4.88 | 21.68 | 4.69 | 18.95 | 4.59 | 5.27 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
46.048 | 0.1 | 12.79 | 0.1 | 5.37 | 0 | 4.79 | 4.49 | 4.88 | 4.1 | 4.88 | 18.46 | 4.88 | 24.22 | 9.86 | 21.09 | 9.67 | 11.13 | 9.96 | 0.1 | 1.95 | 0 | 1.66 | 0.1 | 2.34 | 4.98 | 21.58 | 4.79 | 19.63 | 4.59 | 6.35 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
46.08 | 0 | 13.28 | 0.1 | 5.86 | 0.1 | 5.27 | 4.49 | 4.88 | 5.18 | 4.88 | 18.26 | 4.88 | 21.19 | 9.86 | 21.49 | 9.67 | 12.31 | 9.96 | 0.1 | 1.76 | 0 | 1.37 | 0.1 | 2.25 | 4.88 | 21.49 | 4.79 | 20.22 | 4.69 | 7.91 | 9.67 | 10.94 | 9.77 | 10.16 | 9.77 | 9.86
46.112 | 0.1 | 13.77 | 0.1 | 6.45 | 0 | 5.86 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21 | 9.86 | 21.88 | 9.67 | 13.48 | 9.77 | 0.1 | 1.56 | 0 | 1.37 | 0.1 | 1.95 | 4.98 | 21.39 | 4.79 | 20.9 | 4.69 | 10.06 | 9.67 | 12.21 | 9.77 | 11.23 | 9.67 | 10.84
46.144 | 0 | 14.36 | 0.1 | 7.03 | 0 | 6.45 | 4.49 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 21.09 | 9.86 | 22.36 | 9.67 | 14.26 | 9.77 | 0.1 | 1.37 | 0 | 1.27 | 0.1 | 1.56 | 4.88 | 21.29 | 4.79 | 21.39 | 4.69 | 11.33 | 9.67 | 14.45 | 9.77 | 12.89 | 9.67 | 11.72
46.176 | 0 | 14.84 | 0.1 | 7.72 | 0.1 | 7.03 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 22.66 | 9.67 | 14.84 | 9.77 | 0 | 1.27 | 0 | 1.17 | 0.1 | 1.76 | 4.88 | 21.39 | 4.79 | 21.88 | 4.69 | 12.4 | 9.67 | 15.92 | 9.67 | 14.84 | 9.67 | 12.99
46.208 | 0.1 | 15.43 | 0.1 | 8.3 | 0.1 | 7.72 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 23.05 | 9.67 | 15.43 | 9.77 | 0.1 | 1.27 | 0.1 | 0.98 | 0.1 | 1.56 | 4.98 | 21.29 | 4.79 | 22.56 | 4.69 | 13.38 | 9.67 | 16.9 | 9.67 | 16.11 | 9.67 | 15.04
46.24 | 0 | 15.82 | 0.1 | 8.89 | 0.1 | 8.2 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 23.44 | 9.67 | 15.92 | 9.67 | 0.1 | 1.17 | 0 | 0.88 | 0.1 | 1.37 | 4.98 | 21.29 | 4.88 | 23.15 | 4.79 | 14.16 | 9.67 | 17.68 | 9.67 | 17.09 | 9.67 | 16.31
46.272 | 0 | 16.31 | 0.1 | 9.47 | 0 | 8.89 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 23.73 | 9.67 | 16.41 | 9.67 | 0.1 | 0.98 | 0 | 0.78 | 0.1 | 1.27 | 4.98 | 21.29 | 4.79 | 23.63 | 4.79 | 14.84 | 9.57 | 18.46 | 9.67 | 17.87 | 9.67 | 17.19
46.304 | 0 | 16.8 | 0 | 10.06 | 0 | 9.57 | 4.59 | 5.66 | 4.79 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 24.22 | 9.67 | 16.9 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0.1 | 1.17 | 4.98 | 21.19 | 4.79 | 24.22 | 4.69 | 15.53 | 9.67 | 19.24 | 9.67 | 18.75 | 9.67 | 18.07
46.336 | 0 | 17.29 | 0.1 | 10.64 | 0.1 | 9.96 | 4.69 | 6.54 | 4.79 | 4.88 | 10.94 | 4.88 | 21.39 | 9.86 | 22.27 | 9.77 | 17.38 | 9.67 | 0.1 | 0.78 | 0 | 0.78 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.61 | 4.79 | 16.21 | 9.67 | 19.73 | 9.57 | 19.24 | 9.67 | 18.65
46.368 | 0.1 | 15.23 | 0.1 | 11.23 | 0.1 | 10.64 | 4.69 | 8.01 | 4.79 | 4.88 | 3.32 | 4.88 | 21.39 | 9.86 | 21 | 9.77 | 17.87 | 9.67 | 1.17 | 0.78 | 0.1 | 0.59 | 0.1 | 0.88 | 4.98 | 21.19 | 4.88 | 25.2 | 4.79 | 16.99 | 9.67 | 20.31 | 9.67 | 19.82 | 9.67 | 19.34
46.4 | 0 | 15.33 | 0.1 | 11.82 | 0.1 | 11.23 | 4.79 | 8.98 | 4.79 | 4.88 | 3.22 | 4.88 | 21.39 | 9.86 | 21.29 | 9.77 | 18.36 | 9.67 | 2.34 | 0.68 | 0 | 0.59 | 0.1 | 0.88 | 4.98 | 21.19 | 4.88 | 25.68 | 4.79 | 17.58 | 9.67 | 21 | 9.67 | 20.51 | 9.67 | 19.92
46.432 | 0 | 14.84 | 0.1 | 12.4 | 0 | 11.82 | 4.79 | 9.77 | 4.59 | 4.88 | 3.91 | 4.88 | 21.39 | 9.86 | 21.49 | 9.77 | 18.65 | 9.67 | 4.1 | 0.59 | 0 | 0.49 | 0.1 | 0.78 | 4.98 | 8.5 | 4.88 | 26.17 | 4.79 | 18.26 | 9.67 | 21.58 | 9.67 | 21.19 | 9.67 | 20.51
46.464 | 0 | 14.75 | 0.1 | 12.89 | 0 | 12.4 | 4.79 | 10.45 | 4.59 | 4.88 | 4.98 | 4.88 | 21.39 | 9.86 | 21.29 | 9.77 | 19.24 | 9.67 | 5.76 | 0.49 | 0 | 0.49 | 0 | 0.68 | 4.88 | 2.25 | 4.79 | 22.36 | 4.79 | 18.85 | 9.67 | 22.17 | 9.67 | 21.88 | 9.67 | 21.19
46.496 | 0 | 14.75 | 0.1 | 13.48 | 0.1 | 12.89 | 4.79 | 11.04 | 4.59 | 4.88 | 6.05 | 4.88 | 21.39 | 9.77 | 21.39 | 9.77 | 19.53 | 9.57 | 6.84 | 0.49 | 0 | 0.39 | 0.1 | 0.59 | 4.88 | 3.22 | 4.88 | 21.78 | 4.79 | 19.53 | 9.67 | 22.75 | 9.67 | 22.27 | 9.67 | 21.78
46.528 | 0 | 14.65 | 0.1 | 13.97 | 0.1 | 13.48 | 4.79 | 11.62 | 4.59 | 4.88 | 5.86 | 4.88 | 11.13 | 9.86 | 21.49 | 9.77 | 20.12 | 9.57 | 7.72 | 0.39 | 0.78 | 0.29 | 0.1 | 0.59 | 4.98 | 3.91 | 4.88 | 21.68 | 4.79 | 20.12 | 9.67 | 23.15 | 9.67 | 22.85 | 9.67 | 22.36
46.56 | 0 | 14.55 | 0.1 | 14.55 | 0.1 | 13.97 | 4.79 | 12.21 | 4.49 | 4.88 | 5.27 | 4.88 | 6.64 | 9.86 | 21.39 | 9.77 | 20.51 | 9.67 | 8.5 | 0.39 | 1.66 | 0.29 | 0.2 | 0.59 | 4.98 | 4.98 | 4.88 | 21.39 | 4.88 | 20.7 | 9.67 | 23.83 | 9.67 | 23.34 | 9.67 | 22.85
46.592 | 0 | 14.55 | 0.1 | 15.04 | 0 | 14.55 | 4.88 | 12.79 | 4.49 | 4.88 | 4.98 | 4.88 | 9.08 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 9.28 | 0.29 | 3.32 | 0.2 | 1.46 | 0.49 | 4.98 | 6.15 | 4.88 | 21.49 | 4.79 | 21.29 | 9.77 | 24.32 | 9.67 | 23.93 | 9.67 | 23.44
46.624 | 0 | 14.45 | 0 | 15.53 | 0 | 15.04 | 4.88 | 13.28 | 4.49 | 4.88 | 4.88 | 4.88 | 11.62 | 9.86 | 21.39 | 9.77 | 21.39 | 9.67 | 10.16 | 0.29 | 5.18 | 0.2 | 2.73 | 0.29 | 4.88 | 5.96 | 4.88 | 21.39 | 4.79 | 21.88 | 9.77 | 24.81 | 9.67 | 24.51 | 9.67 | 23.93
46.656 | 0 | 14.45 | 0 | 16.02 | 0.1 | 15.53 | 4.88 | 13.77 | 4.49 | 4.88 | 4.98 | 4.88 | 10.94 | 9.86 | 21.39 | 9.86 | 21.68 | 9.67 | 10.94 | 0.2 | 6.35 | 0.2 | 4.79 | 0.29 | 4.88 | 5.57 | 4.79 | 21.29 | 4.79 | 22.46 | 9.77 | 25.29 | 9.67 | 24.9 | 9.67 | 24.41
46.688 | 0 | 14.45 | 0.1 | 16.5 | 0.1 | 16.02 | 4.88 | 14.36 | 4.49 | 5.57 | 4.88 | 4.88 | 10.16 | 9.86 | 21.39 | 9.86 | 22.07 | 9.67 | 11.62 | 0.2 | 7.23 | 0.2 | 6.05 | 0.29 | 4.98 | 5.27 | 4.88 | 21.29 | 4.88 | 23.05 | 9.77 | 25.78 | 9.77 | 25.39 | 9.67 | 25
46.72 | 0 | 14.45 | 0.1 | 16.99 | 0.1 | 16.41 | 4.88 | 14.94 | 4.49 | 6.25 | 4.88 | 4.88 | 9.86 | 9.86 | 21.39 | 9.86 | 22.46 | 9.67 | 12.5 | 0.2 | 8.01 | 0.1 | 7.03 | 0.29 | 4.98 | 5.18 | 4.88 | 21.19 | 4.88 | 23.54 | 9.77 | 26.37 | 9.77 | 25.88 | 9.77 | 25.39
46.752 | 0 | 14.45 | 0.1 | 15.82 | 0 | 16.99 | 4.88 | 15.43 | 4.49 | 7.42 | 4.69 | 4.88 | 10.16 | 9.86 | 21.39 | 9.77 | 22.95 | 9.67 | 13.28 | 0.1 | 8.79 | 0.1 | 7.81 | 0.2 | 4.98 | 5.18 | 4.88 | 21.29 | 4.88 | 24.02 | 9.77 | 26.86 | 9.77 | 26.47 | 9.67 | 25.98
46.784 | 0 | 9.67 | 0 | 14.06 | 0 | 15.82 | 4.88 | 15.92 | 4.49 | 8.59 | 4.69 | 4.79 | 10.25 | 9.86 | 11.04 | 9.86 | 23.34 | 9.67 | 14.06 | 0.1 | 9.67 | 0.1 | 8.69 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 24.61 | 9.77 | 27.34 | 9.77 | 26.95 | 9.77 | 26.47
46.816 | 0 | 3.03 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 16.41 | 4.59 | 9.47 | 4.69 | 4.88 | 9.96 | 9.77 | 6.54 | 9.86 | 23.63 | 9.67 | 14.75 | 0.1 | 10.45 | 0.1 | 9.47 | 0.1 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.1 | 9.77 | 27.74 | 9.77 | 27.34 | 9.77 | 27.05
46.848 | 0 | 2.05 | 0.1 | 14.65 | 0.1 | 14.65 | 4.88 | 16.9 | 4.59 | 10.16 | 4.59 | 4.88 | 9.86 | 9.86 | 8.98 | 9.86 | 24.02 | 9.67 | 15.43 | 0.1 | 11.23 | 0.1 | 10.16 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.59 | 9.77 | 28.22 | 9.77 | 28.03 | 9.77 | 27.44
46.88 | 0 | 1.76 | 0.1 | 14.45 | 0.1 | 14.75 | 4.88 | 17.29 | 4.59 | 10.74 | 4.59 | 4.88 | 9.96 | 9.86 | 11.43 | 9.86 | 24.41 | 9.67 | 16.21 | 0.1 | 11.91 | 0.1 | 11.04 | 0.1 | 4.98 | 4.88 | 4.88 | 16.21 | 4.88 | 26.08 | 9.77 | 28.71 | 9.86 | 28.32 | 9.77 | 28.03
46.912 | 0.1 | 1.46 | 0.1 | 14.45 | 0 | 14.55 | 4.88 | 17.77 | 4.59 | 11.43 | 4.59 | 4.88 | 9.96 | 9.86 | 11.04 | 9.86 | 21.97 | 9.67 | 16.8 | 0.1 | 12.6 | 0.1 | 11.91 | 0.1 | 4.98 | 4.79 | 4.88 | 3.81 | 4.79 | 22.36 | 9.77 | 29.2 | 9.86 | 28.91 | 9.77 | 28.42
46.944 | 0 | 1.46 | 0 | 14.45 | 0 | 14.55 | 4.88 | 18.26 | 4.59 | 11.91 | 4.49 | 4.79 | 9.77 | 9.86 | 10.16 | 9.86 | 22.27 | 9.67 | 17.58 | 0.1 | 13.38 | 0.1 | 12.6 | 0 | 4.88 | 4.69 | 4.79 | 3.13 | 4.79 | 20.9 | 9.77 | 24.51 | 9.86 | 29.4 | 9.77 | 28.91
46.976 | 0 | 1.27 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 18.65 | 4.69 | 12.6 | 4.59 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 18.16 | 0.1 | 14.26 | 0.1 | 13.38 | 0 | 4.88 | 4.69 | 4.79 | 3.42 | 4.88 | 21.29 | 9.77 | 23.54 | 9.77 | 24.61 | 9.77 | 29.3
47.008 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.14 | 4.69 | 13.09 | 4.59 | 4.88 | 9.77 | 9.86 | 10.06 | 9.86 | 21.68 | 9.77 | 18.75 | 0.1 | 14.94 | 0.1 | 14.06 | 0.1 | 4.98 | 4.69 | 4.88 | 4.39 | 4.88 | 21.19 | 9.77 | 24.12 | 9.86 | 23.54 | 9.77 | 25.1
47.04 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.69 | 13.67 | 4.49 | 4.88 | 9.77 | 9.86 | 10.25 | 9.86 | 21.68 | 9.77 | 19.43 | 0.1 | 15.72 | 0.1 | 14.75 | 0.1 | 5.96 | 4.69 | 4.88 | 5.86 | 4.88 | 21.29 | 9.86 | 24.22 | 9.86 | 24.12 | 9.77 | 24.32
47.072 | 0 | 0.98 | 0.1 | 14.36 | 0 | 14.45 | 4.88 | 20.02 | 4.69 | 14.16 | 4.49 | 5.47 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 9.77 | 20.12 | 0.1 | 16.41 | 0.1 | 15.53 | 0.1 | 7.32 | 4.59 | 4.88 | 6.15 | 4.88 | 21.29 | 9.86 | 24.22 | 9.86 | 24.12 | 9.77 | 24.41
47.104 | 0 | 0.88 | 0.1 | 14.36 | 0 | 14.45 | 4.88 | 20.41 | 4.69 | 14.65 | 4.49 | 6.25 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.8 | 0 | 16.9 | 0 | 16.11 | 0 | 9.67 | 4.59 | 4.88 | 5.66 | 4.79 | 21.29 | 9.77 | 24.22 | 9.86 | 24.32 | 9.86 | 24.32
47.136 | 0 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.8 | 4.79 | 15.14 | 4.59 | 7.32 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 21.39 | 0.1 | 17.68 | 0.1 | 16.9 | 0 | 11.04 | 4.59 | 4.79 | 5.37 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
47.168 | 0 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 18.75 | 4.79 | 15.72 | 4.59 | 8.59 | 9.67 | 10.45 | 9.96 | 9.86 | 21.39 | 9.77 | 21.97 | 0.1 | 18.36 | 0.1 | 17.58 | 0.1 | 12.21 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86 | 24.32
47.2 | 0 | 0.59 | 0.1 | 6.54 | 0.1 | 14.45 | 4.88 | 17.77 | 4.79 | 16.11 | 4.59 | 9.47 | 9.67 | 11.43 | 9.86 | 9.86 | 21.49 | 9.77 | 22.46 | 0.1 | 19.04 | 0.1 | 18.26 | 0.1 | 13.09 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 9.86 | 24.22 | 9.86 | 24.32 | 9.86 | 24.22
47.232 | 0 | 0.59 | 0.1 | 1.86 | 0.1 | 7.23 | 4.88 | 17.97 | 4.79 | 16.6 | 4.59 | 10.16 | 9.67 | 12.7 | 9.77 | 9.86 | 21.39 | 9.77 | 22.46 | 0.1 | 19.53 | 0.1 | 18.85 | 0.1 | 13.97 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.32
47.264 | 0 | 0.49 | 0 | 1.95 | 0 | 1.86 | 4.88 | 18.07 | 4.79 | 17.09 | 4.59 | 10.74 | 9.57 | 13.77 | 9.67 | 9.86 | 21.39 | 9.77 | 18.07 | 0.1 | 20.12 | 0.1 | 19.43 | 0 | 14.75 | 4.49 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
47.296 | 0 | 0.39 | 0 | 1.76 | 0.1 | 1.95 | 4.88 | 18.07 | 4.79 | 17.58 | 4.69 | 11.33 | 9.57 | 14.45 | 9.67 | 9.96 | 21.39 | 9.77 | 18.26 | 0.1 | 20.8 | 0.1 | 20.12 | 0 | 15.33 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
47.328 | 0 | 0.39 | 0.1 | 1.46 | 0.1 | 1.95 | 4.88 | 18.16 | 4.88 | 17.97 | 4.69 | 11.91 | 9.57 | 15.04 | 9.67 | 9.96 | 16.6 | 9.77 | 18.07 | 0.1 | 21.29 | 0.1 | 20.7 | 0.1 | 16.11 | 4.59 | 4.88 | 4.88 | 4.79 | 20.12 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
47.36 | 0 | 0.29 | 0.1 | 1.46 | 0.1 | 1.46 | 4.88 | 18.16 | 4.79 | 18.46 | 4.69 | 12.5 | 9.57 | 15.53 | 9.67 | 9.86 | 7.81 | 9.77 | 17.97 | 0.1 | 21.97 | 0.1 | 21.29 | 0.1 | 16.9 | 4.59 | 4.88 | 4.88 | 4.79 | 4.49 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
47.392 | 0.68 | 0.29 | 0.1 | 1.17 | 0 | 1.37 | 4.88 | 18.07 | 4.79 | 18.95 | 4.69 | 13.09 | 9.67 | 16.11 | 9.67 | 9.86 | 7.62 | 9.86 | 17.87 | 0.1 | 22.46 | 0.1 | 21.88 | 0 | 17.58 | 4.59 | 4.88 | 4.79 | 4.79 | 3.03 | 9.86 | 11.13 | 9.86 | 23.93 | 9.86 | 24.12
47.424 | 1.37 | 0.2 | 0 | 1.07 | 0 | 1.27 | 4.88 | 18.07 | 4.88 | 19.34 | 4.69 | 13.57 | 9.57 | 16.7 | 9.67 | 9.86 | 10.94 | 9.86 | 17.77 | 0 | 18.46 | 0.1 | 22.46 | 0 | 18.16 | 4.59 | 5.66 | 4.79 | 4.79 | 3.52 | 9.77 | 5.86 | 9.86 | 7.62 | 9.77 | 17.68
47.456 | 2.54 | 0.2 | 0 | 0.98 | 0.1 | 1.27 | 4.88 | 18.07 | 4.88 | 19.82 | 4.79 | 14.16 | 9.57 | 17.09 | 9.67 | 9.96 | 11.33 | 9.86 | 17.77 | 0.1 | 18.46 | 0.1 | 17.87 | 0 | 18.85 | 4.59 | 6.64 | 4.69 | 4.79 | 4.3 | 9.77 | 8.89 | 9.86 | 6.54 | 9.86 | 7.62
47.488 | 3.81 | 0.2 | 0.1 | 1.07 | 0.1 | 0.88 | 4.98 | 18.07 | 4.88 | 20.22 | 4.79 | 14.65 | 9.67 | 17.58 | 9.67 | 9.96 | 10.45 | 9.86 | 17.77 | 0.1 | 17.87 | 0.1 | 16.9 | 0.1 | 19.43 | 4.59 | 8.3 | 4.69 | 4.79 | 5.66 | 9.86 | 12.21 | 9.86 | 9.96 | 9.86 | 6.64
47.52 | 4.59 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.7 | 4.79 | 15.14 | 9.67 | 17.97 | 9.67 | 9.96 | 9.86 | 9.86 | 17.68 | 0.1 | 17.77 | 0.1 | 17.58 | 0.1 | 20.12 | 4.69 | 10.25 | 4.59 | 4.79 | 6.15 | 9.77 | 11.43 | 9.86 | 12.31 | 9.86 | 10.55
47.552 | 5.08 | 0.1 | 0.1 | 0.68 | 0 | 0.78 | 4.88 | 18.07 | 4.79 | 21 | 4.79 | 15.72 | 9.67 | 18.46 | 9.67 | 9.86 | 9.86 | 9.86 | 17.68 | 0.1 | 17.87 | 0 | 17.68 | 0 | 20.8 | 4.69 | 11.52 | 4.59 | 4.79 | 5.76 | 9.77 | 10.25 | 9.86 | 11.04 | 9.86 | 12.21
47.584 | 5.66 | 0.1 | 0 | 0.59 | 0 | 0.78 | 4.88 | 18.07 | 4.79 | 19.14 | 4.79 | 16.11 | 9.67 | 18.95 | 9.67 | 9.96 | 9.96 | 9.86 | 17.68 | 0 | 17.68 | 0.1 | 17.68 | 0 | 21.39 | 4.69 | 12.6 | 4.59 | 4.79 | 5.37 | 9.77 | 9.96 | 9.77 | 10.06 | 9.86 | 10.74
47.616 | 6.25 | 0.1 | 0 | 0.49 | 0.1 | 0.59 | 4.88 | 7.23 | 4.88 | 18.95 | 4.79 | 16.6 | 9.67 | 19.34 | 9.57 | 9.96 | 9.96 | 9.86 | 17.68 | 0 | 17.68 | 0.1 | 17.68 | 0 | 21.88 | 4.69 | 13.38 | 4.59 | 4.79 | 5.27 | 9.77 | 10.45 | 9.86 | 9.96 | 9.86 | 9.96
47.648 | 6.84 | 0.1 | 0.1 | 0.49 | 0 | 0.68 | 4.88 | 2.93 | 4.88 | 18.56 | 4.88 | 17.09 | 9.67 | 19.82 | 9.67 | 9.96 | 9.96 | 9.86 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.46 | 4.79 | 14.26 | 4.49 | 4.79 | 5.18 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86 | 9.96
47.68 | 7.42 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.22 | 4.88 | 18.46 | 4.88 | 17.58 | 9.67 | 20.12 | 9.67 | 9.96 | 9.86 | 9.86 | 8.59 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 23.15 | 4.79 | 15.04 | 4.49 | 4.88 | 5.18 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86 | 10.35
47.712 | 8.01 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.88 | 4.2 | 4.79 | 18.46 | 4.79 | 17.97 | 9.67 | 20.61 | 9.67 | 9.86 | 9.86 | 9.86 | 2.15 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 23.63 | 4.79 | 15.72 | 4.49 | 5.66 | 5.08 | 9.77 | 10.06 | 9.86 | 10.25 | 9.86 | 10.45
47.744 | 8.59 | 0.1 | 0.1 | 0.29 | 0 | 0.29 | 4.88 | 5.47 | 4.79 | 18.26 | 4.79 | 18.46 | 9.67 | 21.09 | 9.67 | 9.86 | 9.77 | 9.86 | 2.25 | 0 | 17.58 | 0.1 | 17.68 | 0 | 24.12 | 4.79 | 16.41 | 4.59 | 6.74 | 4.98 | 9.77 | 10.16 | 9.86 | 10.25 | 9.77 | 10.16
47.776 | 9.18 | 0.1 | 0.49 | 0.2 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.26 | 4.88 | 18.85 | 9.67 | 21.49 | 9.67 | 9.86 | 9.67 | 9.86 | 1.95 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 24.51 | 4.79 | 17.09 | 4.59 | 8.4 | 4.98 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 9.96
47.808 | 9.86 | 0.1 | 1.27 | 0.2 | 0.1 | 0.29 | 4.88 | 5.66 | 4.88 | 18.26 | 4.88 | 19.34 | 9.77 | 21.88 | 9.67 | 9.96 | 9.77 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.1 | 4.79 | 17.77 | 4.59 | 10.25 | 4.88 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86 | 10.25
47.84 | 10.35 | 0.1 | 2.25 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.16 | 4.88 | 19.82 | 9.77 | 22.27 | 9.67 | 10.64 | 9.77 | 9.86 | 1.66 | 0.1 | 13.97 | 0.1 | 17.68 | 0.1 | 25.59 | 4.79 | 18.46 | 4.59 | 11.52 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
47.872 | 11.04 | 0.1 | 3.61 | 0.1 | 0 | 0.2 | 4.88 | 4.98 | 4.79 | 18.16 | 4.79 | 20.12 | 9.77 | 22.66 | 9.67 | 11.62 | 9.67 | 9.86 | 1.46 | 0.1 | 3.03 | 0.1 | 17.68 | 0.1 | 26.08 | 4.79 | 18.95 | 4.59 | 12.6 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
47.904 | 11.62 | 0.1 | 4.39 | 0.1 | 0 | 0.2 | 4.88 | 4.88 | 4.79 | 18.07 | 4.79 | 20.61 | 9.77 | 23.05 | 9.67 | 12.99 | 9.67 | 9.86 | 1.46 | 0.1 | 2.05 | 0.1 | 5.66 | 0 | 21.58 | 4.79 | 19.73 | 4.59 | 13.48 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
47.936 | 12.11 | 0.1 | 4.98 | 0 | 0.59 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21 | 9.77 | 23.44 | 9.67 | 13.97 | 9.67 | 9.86 | 1.27 | 0.1 | 2.34 | 0.1 | 1.66 | 0 | 20.7 | 4.79 | 20.22 | 4.59 | 14.16 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86
47.968 | 12.7 | 0.1 | 5.57 | 0.1 | 1.37 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.95 | 9.77 | 23.83 | 9.77 | 14.55 | 9.67 | 9.96 | 1.27 | 0.1 | 1.76 | 0.1 | 2.25 | 0.1 | 21.29 | 4.88 | 20.9 | 4.59 | 15.04 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
48 | 13.18 | 0.1 | 6.05 | 0.1 | 2.44 | 0.1 | 5.57 | 4.79 | 4.88 | 12.5 | 4.88 | 18.95 | 9.77 | 24.12 | 9.77 | 15.14 | 9.67 | 9.96 | 1.17 | 0.1 | 1.86 | 0.1 | 1.95 | 0.1 | 21.39 | 4.79 | 21.39 | 4.69 | 15.72 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
48.032 | 13.77 | 0.1 | 6.64 | 0.1 | 3.71 | 0.1 | 6.25 | 4.79 | 4.79 | 3.61 | 4.88 | 18.46 | 9.77 | 23.73 | 9.77 | 15.63 | 9.67 | 9.96 | 0.98 | 0.1 | 1.56 | 0.1 | 1.76 | 0 | 21.39 | 4.88 | 21.97 | 4.69 | 16.5 | 4.59 | 10.64 | 9.77 | 9.86 | 9.86 | 9.77 | 9.77
48.064 | 14.26 | 0.1 | 7.23 | 0 | 4.49 | 0 | 7.32 | 4.69 | 4.79 | 3.22 | 4.79 | 18.36 | 9.77 | 21.19 | 9.77 | 16.31 | 9.57 | 9.86 | 0.88 | 0 | 1.37 | 0.1 | 1.66 | 0 | 21.29 | 4.79 | 22.56 | 4.69 | 17.09 | 4.59 | 11.82 | 9.77 | 9.86 | 9.77 | 10.84 | 9.77
48.096 | 14.84 | 0.1 | 7.81 | 0 | 5.08 | 0.1 | 8.59 | 4.59 | 4.88 | 3.71 | 4.88 | 18.46 | 9.77 | 21.68 | 9.77 | 16.7 | 9.57 | 9.96 | 0.88 | 0.1 | 1.37 | 0.1 | 1.46 | 0 | 21.29 | 4.79 | 23.15 | 4.69 | 17.77 | 4.59 | 13.28 | 9.67 | 10.45 | 9.77 | 11.33 | 9.67
48.128 | 15.33 | 0.1 | 8.5 | 0.1 | 5.57 | 0.1 | 9.57 | 4.59 | 4.88 | 4.88 | 4.88 | 18.26 | 9.77 | 21.68 | 9.77 | 17.19 | 9.57 | 9.96 | 0.78 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 21.39 | 4.88 | 23.63 | 4.69 | 18.46 | 4.59 | 15.14 | 9.67 | 11.23 | 9.77 | 12.89 | 9.67
48.16 | 15.82 | 0.1 | 9.08 | 0.1 | 6.15 | 0.1 | 10.16 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 9.77 | 21.58 | 9.77 | 17.58 | 9.57 | 9.96 | 0.78 | 0.1 | 1.17 | 0.1 | 1.27 | 0.1 | 21.39 | 4.88 | 24.22 | 4.69 | 19.14 | 4.59 | 16.41 | 9.67 | 12.79 | 9.67 | 14.94 | 9.67
48.192 | 16.31 | 0.1 | 9.67 | 0.1 | 6.74 | 0.1 | 10.64 | 4.59 | 4.79 | 5.76 | 4.88 | 18.16 | 9.77 | 21.58 | 9.77 | 18.07 | 9.57 | 9.86 | 0.59 | 0 | 0.88 | 0 | 1.17 | 0.1 | 21.39 | 4.88 | 24.71 | 4.79 | 19.53 | 4.59 | 17.29 | 9.67 | 14.94 | 9.67 | 16.31 | 9.67
48.224 | 16.8 | 0.1 | 10.25 | 0 | 7.32 | 0.1 | 11.43 | 4.49 | 4.79 | 5.37 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 0 | 21.29 | 4.79 | 25.29 | 4.79 | 20.31 | 4.49 | 18.07 | 9.67 | 16.31 | 9.67 | 17.09 | 9.67
48.256 | 17.29 | 0.1 | 10.84 | 0 | 8.01 | 0.1 | 11.91 | 4.49 | 4.88 | 5.18 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.95 | 9.57 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 0.98 | 0 | 21.29 | 4.88 | 25.68 | 4.79 | 20.8 | 4.59 | 18.85 | 9.67 | 17.09 | 9.67 | 17.87 | 9.57
48.288 | 15.23 | 0.1 | 11.43 | 0.1 | 8.59 | 0.1 | 12.5 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.49 | 9.86 | 19.43 | 9.57 | 9.96 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 26.17 | 4.79 | 21.39 | 4.59 | 19.43 | 9.57 | 17.87 | 9.67 | 18.56 | 9.67
48.32 | 15.23 | 0.1 | 12.01 | 0.1 | 9.18 | 0.1 | 13.09 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 19.92 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 22.36 | 4.79 | 22.07 | 4.59 | 20.12 | 9.57 | 18.56 | 9.67 | 19.34 | 9.67
48.352 | 14.84 | 0.1 | 12.6 | 0.1 | 9.67 | 0.1 | 13.57 | 4.49 | 4.79 | 4.98 | 4.79 | 18.07 | 9.77 | 21.39 | 9.86 | 20.31 | 9.67 | 10.55 | 0.39 | 0.1 | 0.59 | 0 | 0.78 | 0.1 | 7.62 | 4.88 | 21.88 | 4.79 | 22.66 | 4.59 | 20.7 | 9.67 | 19.24 | 9.57 | 20.02 | 9.67
48.384 | 14.65 | 0.1 | 13.09 | 0 | 10.35 | 0.1 | 14.16 | 4.49 | 4.79 | 4.79 | 4.88 | 10.74 | 9.77 | 21.39 | 9.77 | 20.7 | 9.67 | 11.23 | 0.29 | 0 | 0.49 | 0.1 | 0.59 | 0 | 2.15 | 4.79 | 21.68 | 4.79 | 23.24 | 4.69 | 21.39 | 9.57 | 20.02 | 9.57 | 20.51 | 9.57
48.416 | 14.65 | 0.1 | 13.67 | 0 | 11.04 | 0.1 | 14.75 | 4.49 | 4.88 | 4.79 | 4.88 | 3.42 | 9.77 | 21.39 | 9.77 | 21.19 | 9.67 | 12.21 | 0.2 | 0 | 0.49 | 0.1 | 0.59 | 0 | 3.32 | 4.79 | 21.49 | 4.79 | 23.63 | 4.69 | 21.88 | 9.67 | 20.51 | 9.67 | 21.09 | 9.57
48.448 | 14.55 | 0.1 | 14.16 | 0.1 | 11.52 | 0.1 | 15.23 | 4.59 | 4.88 | 4.79 | 4.88 | 3.22 | 9.77 | 21.39 | 9.86 | 21.58 | 9.67 | 13.48 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.81 | 4.88 | 21.49 | 4.79 | 24.12 | 4.69 | 22.46 | 9.67 | 21.19 | 9.67 | 21.68 | 9.57
48.48 | 14.55 | 0.1 | 14.65 | 0.1 | 12.11 | 0.1 | 15.72 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 9.77 | 12.5 | 9.86 | 21.88 | 9.67 | 14.26 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.98 | 4.88 | 21.39 | 4.79 | 24.71 | 4.69 | 22.95 | 9.67 | 21.78 | 9.67 | 22.27 | 9.67
48.512 | 14.45 | 0.1 | 15.14 | 0.1 | 12.7 | 0.1 | 16.21 | 4.59 | 4.79 | 4.59 | 4.88 | 4.88 | 9.77 | 6.84 | 9.86 | 22.36 | 9.67 | 14.84 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 6.15 | 4.88 | 21.39 | 4.79 | 25.29 | 4.69 | 23.44 | 9.67 | 22.36 | 9.67 | 22.95 | 9.67
48.544 | 14.36 | 0.1 | 15.72 | 0 | 13.18 | 0.1 | 16.7 | 4.59 | 4.79 | 4.59 | 4.88 | 5.86 | 9.77 | 8.98 | 9.77 | 22.75 | 9.67 | 15.53 | 0.1 | 0 | 0.2 | 0.78 | 0.29 | 0 | 5.96 | 4.79 | 21.39 | 4.79 | 25.88 | 4.69 | 24.02 | 9.67 | 22.95 | 9.67 | 23.34 | 9.67
48.576 | 14.45 | 0.1 | 16.21 | 0 | 13.67 | 0.1 | 17.09 | 4.69 | 4.88 | 4.49 | 4.88 | 5.76 | 9.77 | 11.33 | 9.86 | 23.15 | 9.67 | 15.92 | 0.1 | 0 | 0.2 | 1.76 | 0.29 | 0 | 5.86 | 4.79 | 21.29 | 4.79 | 26.27 | 4.79 | 24.61 | 9.67 | 23.34 | 9.67 | 23.93 | 9.67
48.608 | 14.36 | 0.1 | 16.7 | 0.1 | 14.26 | 0.1 | 17.58 | 4.69 | 4.88 | 4.59 | 4.88 | 5.47 | 9.77 | 11.13 | 9.86 | 23.44 | 9.67 | 16.5 | 0.1 | 0.88 | 0.2 | 3.22 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.79 | 22.66 | 4.79 | 25.2 | 9.67 | 23.93 | 9.67 | 24.41 | 9.67
48.64 | 14.36 | 0.1 | 17.19 | 0.1 | 14.75 | 0.1 | 18.07 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 9.77 | 10.35 | 9.86 | 23.83 | 9.67 | 16.9 | 0.1 | 1.86 | 0.2 | 5.18 | 0.2 | 0 | 4.98 | 4.88 | 21.29 | 4.79 | 21.97 | 4.79 | 25.68 | 9.67 | 24.41 | 9.67 | 24.9 | 9.67
48.672 | 14.45 | 0.1 | 13.87 | 0.1 | 15.23 | 0.1 | 18.56 | 4.79 | 6.25 | 4.59 | 4.88 | 4.98 | 9.77 | 9.96 | 9.86 | 24.22 | 9.77 | 17.38 | 0.1 | 3.52 | 0.1 | 6.45 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 21.68 | 4.79 | 26.17 | 9.67 | 25 | 9.67 | 25.49 | 9.67
48.704 | 9.08 | 0.1 | 14.45 | 0 | 15.72 | 0.1 | 18.95 | 4.69 | 7.42 | 4.49 | 4.88 | 4.98 | 9.77 | 10.06 | 9.86 | 22.07 | 9.77 | 17.87 | 0.1 | 5.27 | 0.1 | 7.13 | 0.2 | 0.98 | 5.08 | 4.79 | 21.19 | 4.79 | 21.58 | 4.79 | 26.56 | 9.67 | 25.49 | 9.67 | 25.98 | 9.67
48.736 | 2.83 | 0.1 | 14.26 | 0 | 16.21 | 0 | 19.34 | 4.79 | 8.59 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 21 | 9.67 | 18.36 | 0.1 | 6.45 | 0.1 | 8.01 | 0.1 | 1.95 | 4.98 | 4.88 | 16.11 | 4.79 | 21.49 | 4.79 | 27.15 | 9.77 | 25.98 | 9.67 | 26.37 | 9.67
48.768 | 2.05 | 0.1 | 14.45 | 0.1 | 16.8 | 0.1 | 19.82 | 4.79 | 9.47 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 21.29 | 9.77 | 18.75 | 0.1 | 7.23 | 0.1 | 8.79 | 0.2 | 3.71 | 4.88 | 4.88 | 3.81 | 4.79 | 21.49 | 4.79 | 27.54 | 9.77 | 26.47 | 9.77 | 26.86 | 9.67
48.8 | 1.86 | 0.1 | 14.45 | 0.1 | 17.19 | 0.1 | 20.22 | 4.79 | 10.25 | 4.59 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 19.24 | 0.1 | 8.11 | 0.1 | 9.67 | 0.1 | 5.37 | 4.98 | 4.88 | 3.22 | 4.79 | 21.39 | 4.79 | 28.13 | 9.77 | 26.95 | 9.77 | 27.34 | 9.77
48.832 | 1.37 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 20.7 | 4.79 | 10.84 | 4.59 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 21.39 | 9.77 | 19.73 | 0.1 | 8.89 | 0 | 10.45 | 0.1 | 6.54 | 4.88 | 4.88 | 3.52 | 4.79 | 21.39 | 4.79 | 28.52 | 9.77 | 27.54 | 9.77 | 28.03 | 9.77
48.864 | 1.46 | 0.1 | 14.45 | 0 | 15.04 | 0 | 21 | 4.79 | 11.43 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.12 | 0.1 | 9.67 | 0 | 11.23 | 0.1 | 7.42 | 4.79 | 4.88 | 4.39 | 4.79 | 21.29 | 4.79 | 29.1 | 9.77 | 28.03 | 9.77 | 28.42 | 9.67
48.896 | 1.17 | 0.1 | 14.45 | 0 | 14.55 | 0 | 18.95 | 4.79 | 12.01 | 4.59 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.51 | 0 | 10.45 | 0.1 | 12.01 | 0.1 | 8.2 | 4.69 | 4.88 | 5.96 | 4.79 | 21.29 | 4.88 | 29.49 | 9.77 | 28.42 | 9.77 | 28.91 | 9.67
48.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.95 | 4.79 | 12.6 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 20.9 | 0.1 | 11.23 | 0.1 | 12.79 | 0.1 | 8.89 | 4.69 | 4.88 | 6.25 | 4.79 | 21.29 | 4.88 | 25.98 | 9.77 | 28.81 | 9.86 | 29.2 | 9.77
48.96 | 0.88 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 18.46 | 4.88 | 13.18 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 12.01 | 0.1 | 13.48 | 0.1 | 9.67 | 4.69 | 4.88 | 5.86 | 4.79 | 21.29 | 4.88 | 25.2 | 9.77 | 29.4 | 9.77 | 24.81 | 9.77
48.992 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.67 | 4.69 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.86 | 21.68 | 0.1 | 12.7 | 0 | 14.26 | 0.1 | 10.55 | 4.59 | 4.88 | 5.47 | 4.79 | 21.19 | 4.79 | 24.81 | 9.77 | 24.71 | 9.77 | 24.32 | 9.77
49.024 | 0.78 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 18.46 | 4.79 | 14.26 | 4.69 | 5.47 | 4.49 | 9.77 | 9.77 | 9.86 | 21.39 | 9.77 | 22.07 | 0.1 | 13.48 | 0 | 14.94 | 0.1 | 11.33 | 4.59 | 4.79 | 5.18 | 4.79 | 13.38 | 4.79 | 24.71 | 9.77 | 23.93 | 9.86 | 24.22 | 9.77
49.056 | 0.68 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 18.26 | 4.79 | 14.75 | 4.69 | 6.15 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.46 | 0.1 | 14.16 | 0.1 | 15.63 | 0.1 | 12.21 | 4.49 | 4.88 | 5.08 | 4.79 | 3.42 | 4.88 | 24.71 | 9.77 | 24.12 | 9.77 | 24.22 | 9.77
49.088 | 0.59 | 0.1 | 14.16 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.23 | 4.69 | 7.13 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 0.1 | 14.94 | 0.1 | 16.41 | 0.1 | 12.99 | 4.59 | 4.88 | 5.18 | 4.79 | 3.22 | 4.88 | 24.61 | 9.77 | 24.22 | 9.86 | 24.32 | 9.77
49.12 | 0.59 | 0.1 | 4.2 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.72 | 4.79 | 8.59 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.24 | 0.1 | 15.63 | 0.1 | 17.09 | 0.1 | 13.67 | 4.59 | 5.86 | 5.18 | 4.79 | 3.42 | 4.88 | 24.41 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
49.152 | 0.49 | 0.1 | 1.86 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 16.21 | 4.79 | 9.47 | 4.49 | 9.77 | 9.67 | 9.86 | 11.04 | 9.86 | 23.63 | 0.1 | 16.5 | 0 | 17.68 | 0.1 | 14.36 | 4.59 | 7.03 | 4.98 | 4.79 | 4.49 | 4.88 | 24.51 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
49.184 | 0.49 | 0.1 | 1.86 | 0 | 14.36 | 0.1 | 18.16 | 4.79 | 16.7 | 4.79 | 10.06 | 4.49 | 9.86 | 9.67 | 9.86 | 6.64 | 9.86 | 24.02 | 0.1 | 16.99 | 0 | 18.46 | 0.1 | 15.14 | 4.49 | 9.28 | 4.88 | 4.79 | 5.96 | 4.88 | 24.41 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
49.216 | 0.39 | 0.1 | 1.56 | 0 | 14.36 | 0.1 | 18.16 | 4.88 | 17.19 | 4.79 | 10.64 | 4.49 | 10.45 | 9.67 | 9.86 | 8.89 | 9.86 | 24.41 | 0.1 | 17.68 | 0.1 | 19.14 | 0.1 | 15.82 | 4.49 | 10.55 | 4.88 | 4.79 | 6.25 | 4.88 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
49.248 | 0.29 | 0.1 | 1.46 | 0.1 | 11.82 | 0.1 | 18.16 | 4.88 | 17.68 | 4.79 | 11.23 | 4.59 | 11.52 | 9.67 | 9.86 | 11.52 | 9.86 | 21.58 | 0.1 | 18.26 | 0.1 | 19.73 | 0.1 | 16.6 | 4.59 | 11.82 | 4.88 | 4.79 | 5.76 | 4.88 | 24.41 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
49.28 | 0.29 | 0.1 | 1.37 | 0.1 | 3.32 | 0.1 | 18.16 | 4.88 | 18.07 | 4.79 | 11.91 | 4.59 | 12.89 | 9.57 | 9.86 | 11.04 | 9.86 | 22.07 | 0.1 | 18.95 | 0.1 | 20.31 | 0.1 | 17.29 | 4.59 | 12.79 | 4.79 | 4.79 | 5.37 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.12 | 9.77
49.312 | 0.2 | 0.1 | 1.17 | 0.1 | 2.05 | 0.1 | 13.48 | 4.88 | 18.46 | 4.79 | 12.5 | 4.59 | 13.87 | 9.67 | 9.86 | 10.16 | 9.86 | 21.78 | 0.1 | 19.63 | 0 | 20.9 | 0.1 | 17.87 | 4.59 | 13.67 | 4.69 | 4.79 | 5.18 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
49.344 | 0.1 | 0.1 | 1.07 | 0 | 1.86 | 0.1 | 3.61 | 4.79 | 18.95 | 4.79 | 12.99 | 4.59 | 14.55 | 9.67 | 9.86 | 9.86 | 9.86 | 21.68 | 0 | 20.22 | 0 | 21.39 | 0.1 | 18.56 | 4.49 | 14.55 | 4.69 | 4.79 | 5.08 | 4.79 | 16.5 | 9.77 | 24.32 | 9.86 | 24.12 | 9.77
49.376 | 0.1 | 0.1 | 0.88 | 0 | 1.56 | 0.1 | 3.32 | 4.88 | 19.34 | 4.88 | 13.57 | 4.69 | 15.04 | 9.67 | 9.86 | 9.96 | 9.86 | 21.58 | 0 | 20.9 | 0.1 | 21.97 | 0.1 | 19.24 | 4.59 | 15.23 | 4.59 | 4.79 | 5.18 | 4.88 | 6.93 | 9.77 | 24.22 | 9.77 | 23.34 | 9.77
49.408 | 0.1 | 0.1 | 0.88 | 0.1 | 1.56 | 0.1 | 3.61 | 4.88 | 19.82 | 4.88 | 14.16 | 4.69 | 15.63 | 9.67 | 9.86 | 10.25 | 9.86 | 21.58 | 0.1 | 21.49 | 0.1 | 22.56 | 0.1 | 19.92 | 4.69 | 15.92 | 4.59 | 4.79 | 5.08 | 4.88 | 6.84 | 9.77 | 24.02 | 9.86 | 8.4 | 9.86
49.44 | 0.1 | 0.1 | 0.78 | 0.1 | 1.37 | 0.1 | 4.69 | 4.88 | 20.22 | 4.88 | 14.55 | 4.69 | 16.11 | 9.67 | 9.86 | 9.96 | 9.86 | 21.58 | 0.1 | 21.97 | 0.1 | 18.46 | 0.1 | 20.41 | 4.69 | 16.6 | 4.59 | 4.79 | 4.98 | 4.88 | 11.23 | 9.77 | 8.89 | 9.86 | 6.45 | 9.86
49.472 | 0 | 0.68 | 0.68 | 0.1 | 1.07 | 0.1 | 5.76 | 4.88 | 20.7 | 4.88 | 15.14 | 4.79 | 16.7 | 9.67 | 10.45 | 9.86 | 9.86 | 21.49 | 0.1 | 22.56 | 0 | 18.26 | 0.1 | 21 | 4.69 | 17.29 | 4.59 | 4.79 | 4.98 | 4.88 | 12.01 | 9.77 | 6.35 | 9.86 | 10.55 | 9.86
49.504 | 0 | 1.37 | 0.59 | 0 | 0.98 | 0.1 | 5.76 | 4.88 | 21.09 | 4.88 | 15.53 | 4.69 | 17.19 | 9.67 | 11.52 | 9.96 | 9.86 | 21.49 | 0.1 | 19.43 | 0 | 17.87 | 0.1 | 21.68 | 4.69 | 17.97 | 4.59 | 4.79 | 4.88 | 4.79 | 10.55 | 9.77 | 10.25 | 9.86 | 12.21 | 9.86
49.536 | 0 | 2.44 | 0.59 | 0 | 0.98 | 0 | 5.37 | 4.88 | 19.14 | 4.88 | 16.11 | 4.79 | 17.58 | 9.67 | 12.79 | 9.96 | 9.86 | 21.49 | 0 | 18.46 | 0.1 | 17.77 | 0.1 | 22.07 | 4.69 | 18.65 | 4.49 | 4.79 | 4.79 | 4.88 | 9.96 | 9.77 | 12.4 | 9.77 | 10.94 | 9.86
49.568 | 0 | 3.81 | 0.49 | 0.1 | 0.88 | 0.1 | 5.18 | 4.88 | 18.85 | 4.88 | 16.5 | 4.79 | 18.07 | 9.67 | 13.77 | 9.96 | 9.86 | 21.49 | 0.1 | 18.07 | 0.1 | 17.87 | 0.1 | 22.75 | 4.79 | 19.34 | 4.49 | 5.76 | 4.79 | 4.88 | 10.16 | 9.77 | 11.13 | 9.86 | 9.96 | 9.86
49.6 | 0.1 | 4.49 | 0.39 | 0.1 | 0.68 | 0.1 | 4.98 | 4.88 | 18.56 | 4.88 | 17.09 | 4.79 | 18.46 | 9.67 | 14.45 | 9.77 | 9.86 | 21.39 | 0.1 | 17.87 | 0.1 | 17.77 | 0.1 | 19.63 | 4.79 | 19.92 | 4.59 | 6.35 | 4.79 | 4.88 | 10.45 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
49.632 | 0.1 | 5.18 | 0.39 | 0.1 | 0.68 | 0.1 | 4.98 | 4.88 | 18.26 | 4.88 | 17.58 | 4.79 | 18.95 | 9.67 | 15.04 | 9.77 | 9.86 | 21.39 | 0.1 | 17.87 | 0 | 17.77 | 0.1 | 18.75 | 4.79 | 20.51 | 4.59 | 7.91 | 4.79 | 4.88 | 10.45 | 9.77 | 10.16 | 9.86 | 10.35 | 9.86
49.664 | 0 | 5.66 | 0.29 | 0 | 0.49 | 0 | 4.88 | 4.88 | 18.36 | 4.88 | 17.97 | 4.79 | 19.43 | 9.67 | 15.63 | 9.77 | 9.86 | 17.77 | 0.1 | 17.77 | 0 | 17.77 | 0.1 | 18.16 | 4.79 | 21 | 4.59 | 10.06 | 4.59 | 4.79 | 10.16 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86
49.696 | 0.1 | 6.15 | 0.2 | 0 | 0.59 | 0 | 4.79 | 4.88 | 18.26 | 4.88 | 18.46 | 4.79 | 19.82 | 9.67 | 16.11 | 9.67 | 9.86 | 8.2 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 18.07 | 4.79 | 21.68 | 4.59 | 11.43 | 4.59 | 4.88 | 10.25 | 9.77 | 10.35 | 9.77 | 10.35 | 9.86
49.728 | 0.1 | 6.84 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.85 | 4.88 | 20.22 | 9.77 | 16.7 | 9.67 | 9.86 | 7.52 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 18.07 | 4.88 | 22.36 | 4.59 | 12.4 | 4.59 | 4.88 | 10.25 | 9.77 | 10.35 | 9.86 | 10.06 | 9.86
49.76 | 0 | 7.42 | 0.2 | 0.1 | 0.49 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.34 | 4.88 | 20.7 | 9.77 | 17.19 | 9.67 | 9.86 | 10.45 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.97 | 4.88 | 22.75 | 4.59 | 13.28 | 4.59 | 4.88 | 10.06 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
49.792 | 0.1 | 8.01 | 0.1 | 0.1 | 0.29 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 19.73 | 4.88 | 21.09 | 9.77 | 17.58 | 9.67 | 9.86 | 11.52 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.88 | 23.44 | 4.59 | 14.16 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
49.824 | 0 | 8.59 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 21.49 | 9.77 | 18.07 | 9.67 | 9.86 | 10.64 | 0.1 | 17.58 | 0 | 17.68 | 0 | 17.77 | 4.88 | 23.93 | 4.69 | 14.84 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.06 | 9.77
49.856 | 0 | 9.18 | 0 | 0.1 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.61 | 4.88 | 21.97 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0.1 | 15.23 | 0.1 | 17.68 | 4.88 | 24.51 | 4.69 | 15.53 | 4.49 | 4.88 | 9.96 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
49.888 | 0.1 | 9.77 | 0.1 | 0.78 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.36 | 9.77 | 18.95 | 9.57 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 3.91 | 0.1 | 17.77 | 4.88 | 25 | 4.69 | 16.21 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
49.92 | 0.1 | 10.35 | 0.1 | 1.66 | 0.1 | 0.78 | 4.59 | 4.88 | 18.07 | 4.88 | 18.16 | 4.88 | 22.66 | 9.77 | 19.34 | 9.57 | 9.86 | 10.25 | 0.1 | 12.01 | 0.1 | 2.05 | 0.1 | 17.68 | 4.88 | 25.49 | 4.69 | 16.99 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
49.952 | 0.1 | 11.04 | 0.1 | 2.93 | 0.1 | 1.56 | 4.59 | 4.88 | 9.08 | 4.88 | 18.85 | 4.88 | 23.05 | 9.77 | 19.82 | 9.67 | 9.86 | 10.16 | 0.1 | 2.93 | 0 | 2.05 | 0.1 | 17.68 | 4.88 | 25.88 | 4.79 | 17.68 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
49.984 | 0 | 11.62 | 0 | 4 | 0 | 2.93 | 4.49 | 4.88 | 3.22 | 4.88 | 18.26 | 4.88 | 23.54 | 9.77 | 20.22 | 9.67 | 9.86 | 9.86 | 0.1 | 2.25 | 0 | 1.76 | 0.1 | 17.68 | 4.88 | 25.49 | 4.69 | 18.26 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
50.016 | 0 | 12.11 | 0 | 4.59 | 0.1 | 3.91 | 4.49 | 4.88 | 3.13 | 4.88 | 18.26 | 4.88 | 23.83 | 9.77 | 20.61 | 9.67 | 10.64 | 9.86 | 0 | 2.05 | 0.1 | 1.76 | 0.1 | 11.13 | 4.88 | 22.46 | 4.79 | 18.85 | 4.59 | 5.08 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
50.048 | 0.1 | 12.7 | 0.1 | 5.27 | 0.1 | 4.69 | 4.49 | 4.88 | 3.91 | 4.88 | 18.26 | 4.88 | 24.22 | 9.77 | 21 | 9.67 | 11.13 | 9.96 | 0.1 | 1.76 | 0.1 | 1.56 | 0.1 | 2.64 | 4.88 | 21.97 | 4.79 | 19.43 | 4.59 | 6.05 | 9.67 | 9.86 | 9.86 | 9.86 | 9.77 | 9.86
50.08 | 0.1 | 13.28 | 0.1 | 5.76 | 0.1 | 5.18 | 4.59 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 22.17 | 9.77 | 21.39 | 9.67 | 12.11 | 9.86 | 0.1 | 1.66 | 0.1 | 1.37 | 0.1 | 2.44 | 4.88 | 21.68 | 4.79 | 20.12 | 4.69 | 7.72 | 9.67 | 10.74 | 9.86 | 9.96 | 9.77 | 9.86
50.112 | 0.1 | 13.77 | 0.1 | 6.35 | 0 | 5.76 | 4.59 | 4.88 | 5.86 | 4.88 | 18.26 | 4.88 | 21 | 9.77 | 21.88 | 9.67 | 13.48 | 9.77 | 0.1 | 1.56 | 0 | 1.37 | 0.1 | 2.05 | 4.88 | 21.58 | 4.79 | 20.9 | 4.69 | 9.86 | 9.67 | 12.21 | 9.77 | 10.94 | 9.67 | 10.35
50.144 | 0 | 14.26 | 0.1 | 7.03 | 0 | 6.35 | 4.49 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 21.19 | 9.86 | 22.36 | 9.67 | 14.26 | 9.77 | 0.1 | 1.37 | 0 | 1.27 | 0.1 | 1.66 | 4.88 | 21.39 | 4.79 | 21.49 | 4.69 | 11.23 | 9.67 | 14.36 | 9.77 | 12.6 | 9.67 | 11.52
50.176 | 0 | 14.84 | 0.1 | 7.62 | 0.1 | 6.93 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 22.66 | 9.67 | 14.84 | 9.77 | 0 | 1.37 | 0.1 | 1.07 | 0.1 | 1.76 | 4.88 | 21.39 | 4.79 | 21.97 | 4.69 | 12.31 | 9.67 | 15.63 | 9.67 | 14.55 | 9.67 | 12.79
50.208 | 0.1 | 15.33 | 0.1 | 8.2 | 0.1 | 7.52 | 4.59 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 21.39 | 9.86 | 23.05 | 9.67 | 15.43 | 9.77 | 0.1 | 1.17 | 0.1 | 1.07 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 22.56 | 4.79 | 13.28 | 9.67 | 16.8 | 9.67 | 16.02 | 9.67 | 14.94
50.24 | 0.1 | 15.82 | 0.1 | 8.79 | 0.1 | 8.11 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 23.44 | 9.67 | 15.92 | 9.67 | 0.1 | 1.07 | 0.1 | 0.98 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 23.05 | 4.79 | 14.06 | 9.67 | 17.68 | 9.67 | 16.99 | 9.67 | 16.21
50.272 | 0.1 | 16.21 | 0.1 | 9.38 | 0 | 8.79 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 23.83 | 9.67 | 16.41 | 9.67 | 0.1 | 1.07 | 0 | 0.88 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 23.54 | 4.79 | 14.84 | 9.67 | 18.36 | 9.67 | 17.87 | 9.67 | 17.19
50.304 | 0 | 16.8 | 0 | 10.06 | 0 | 9.47 | 4.69 | 5.47 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 24.22 | 9.67 | 16.9 | 9.67 | 0 | 0.88 | 0 | 0.78 | 0.1 | 1.17 | 4.88 | 21.19 | 4.88 | 24.12 | 4.79 | 15.53 | 9.67 | 19.14 | 9.67 | 18.56 | 9.67 | 17.87
50.336 | 0 | 17.29 | 0.1 | 10.55 | 0.1 | 9.86 | 4.69 | 6.45 | 4.88 | 4.88 | 14.55 | 4.88 | 21.39 | 9.77 | 23.15 | 9.67 | 17.29 | 9.67 | 0 | 0.78 | 0.1 | 0.78 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.71 | 4.79 | 16.21 | 9.67 | 19.53 | 9.67 | 19.24 | 9.67 | 18.65
50.368 | 0.1 | 14.75 | 0.1 | 11.23 | 0.1 | 10.55 | 4.69 | 7.81 | 4.79 | 4.88 | 4.2 | 4.88 | 21.39 | 9.77 | 21.29 | 9.67 | 17.77 | 9.67 | 0.88 | 0.88 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 21.29 | 4.88 | 25.29 | 4.79 | 16.9 | 9.67 | 20.31 | 9.67 | 19.82 | 9.67 | 19.24
50.4 | 0.1 | 15.23 | 0.1 | 11.82 | 0.1 | 11.13 | 4.79 | 8.79 | 4.69 | 4.88 | 3.32 | 4.88 | 21.39 | 9.77 | 21.58 | 9.77 | 18.26 | 9.67 | 1.86 | 0.68 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 21.19 | 4.88 | 25.68 | 4.79 | 17.58 | 9.67 | 20.9 | 9.67 | 20.51 | 9.67 | 19.92
50.432 | 0.1 | 14.75 | 0.1 | 12.31 | 0.1 | 11.72 | 4.79 | 9.67 | 4.59 | 4.88 | 3.71 | 4.88 | 21.39 | 9.77 | 21.58 | 9.67 | 18.75 | 9.57 | 3.52 | 0.68 | 0 | 0.49 | 0.1 | 0.88 | 4.88 | 9.38 | 4.88 | 26.17 | 4.79 | 18.26 | 9.67 | 21.49 | 9.67 | 20.9 | 9.67 | 20.61
50.464 | 0 | 14.65 | 0.1 | 12.89 | 0 | 12.31 | 4.79 | 10.35 | 4.59 | 4.88 | 4.69 | 4.88 | 21.39 | 9.77 | 21.39 | 9.77 | 19.14 | 9.57 | 5.27 | 0.49 | 0 | 0.49 | 0.1 | 0.68 | 4.88 | 2.34 | 4.79 | 21.78 | 4.79 | 18.85 | 9.67 | 22.07 | 9.67 | 21.68 | 9.67 | 21.19
50.496 | 0.1 | 14.65 | 0.1 | 13.38 | 0.1 | 12.79 | 4.79 | 10.94 | 4.59 | 4.88 | 5.76 | 4.88 | 21.39 | 9.77 | 21.49 | 9.77 | 19.53 | 9.67 | 6.35 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 4.88 | 3.22 | 4.88 | 20.8 | 4.79 | 19.43 | 9.67 | 22.66 | 9.67 | 22.17 | 9.67 | 21.78
50.528 | 0 | 14.55 | 0.1 | 13.97 | 0.1 | 13.38 | 4.79 | 11.52 | 4.59 | 4.88 | 5.86 | 4.88 | 11.43 | 9.86 | 21.49 | 9.77 | 20.12 | 9.67 | 7.23 | 0.49 | 0.68 | 0.29 | 0.1 | 0.59 | 4.88 | 3.71 | 4.88 | 21.39 | 4.88 | 20.12 | 9.67 | 23.24 | 9.67 | 22.75 | 9.67 | 22.27
50.56 | 0.1 | 14.55 | 0.1 | 14.45 | 0.1 | 13.87 | 4.88 | 12.11 | 4.59 | 4.88 | 5.47 | 4.88 | 6.54 | 9.86 | 21.49 | 9.77 | 20.51 | 9.67 | 8.01 | 0.49 | 1.76 | 0.29 | 0.1 | 0.59 | 4.88 | 4.88 | 4.88 | 21.29 | 4.88 | 20.61 | 9.67 | 23.73 | 9.67 | 23.34 | 9.67 | 22.85
50.592 | 0 | 14.55 | 0.1 | 14.94 | 0 | 14.36 | 4.88 | 12.7 | 4.49 | 4.88 | 5.08 | 4.88 | 8.89 | 9.77 | 21.49 | 9.77 | 20.8 | 9.57 | 8.89 | 0.29 | 3.22 | 0.2 | 1.27 | 0.49 | 4.88 | 6.15 | 4.88 | 21.29 | 4.88 | 21.29 | 9.77 | 24.32 | 9.67 | 23.83 | 9.67 | 23.44
50.624 | 0 | 14.45 | 0 | 15.43 | 0 | 14.94 | 4.79 | 13.18 | 4.49 | 4.88 | 4.88 | 4.88 | 11.52 | 9.86 | 21.39 | 9.77 | 21.29 | 9.67 | 9.57 | 0.2 | 5.08 | 0.2 | 2.64 | 0.29 | 4.88 | 6.05 | 4.88 | 21.29 | 4.79 | 21.88 | 9.67 | 24.9 | 9.67 | 24.41 | 9.67 | 23.83
50.656 | 0 | 14.45 | 0.1 | 15.92 | 0.1 | 15.43 | 4.88 | 13.77 | 4.49 | 4.88 | 4.98 | 4.88 | 10.94 | 9.77 | 21.39 | 9.77 | 21.68 | 9.67 | 10.35 | 0.2 | 6.35 | 0.2 | 4.59 | 0.29 | 4.88 | 5.57 | 4.79 | 21.29 | 4.79 | 22.46 | 9.77 | 25.29 | 9.67 | 24.9 | 9.77 | 24.41
50.688 | 0.1 | 14.45 | 0.1 | 16.5 | 0.1 | 15.92 | 4.88 | 14.26 | 4.49 | 5.57 | 4.88 | 4.88 | 10.16 | 9.77 | 21.39 | 9.77 | 22.07 | 9.67 | 11.13 | 0.2 | 7.23 | 0.2 | 5.96 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.88 | 23.05 | 9.77 | 25.88 | 9.77 | 25.39 | 9.77 | 24.9
50.72 | 0.1 | 14.45 | 0.1 | 16.9 | 0.1 | 16.41 | 4.88 | 14.84 | 4.49 | 6.15 | 4.88 | 4.88 | 9.86 | 9.77 | 21.39 | 9.77 | 22.46 | 9.67 | 12.01 | 0.2 | 8.01 | 0.1 | 6.93 | 0.29 | 4.88 | 5.18 | 4.88 | 21.19 | 4.88 | 23.54 | 9.77 | 26.37 | 9.77 | 25.98 | 9.77 | 25.29
50.752 | 0.1 | 14.36 | 0.1 | 16.9 | 0 | 16.8 | 4.88 | 15.23 | 4.49 | 7.23 | 4.79 | 4.88 | 10.16 | 9.77 | 21.39 | 9.77 | 22.75 | 9.67 | 12.7 | 0.2 | 8.79 | 0.1 | 7.72 | 0.2 | 4.88 | 5.08 | 4.88 | 21.29 | 4.88 | 24.12 | 9.77 | 26.76 | 9.77 | 26.37 | 9.77 | 25.88
50.784 | 0 | 9.57 | 0 | 14.55 | 0 | 16.9 | 4.88 | 15.82 | 4.49 | 8.5 | 4.69 | 4.88 | 10.25 | 9.77 | 11.43 | 9.77 | 23.24 | 9.67 | 13.48 | 0.1 | 9.57 | 0.1 | 8.59 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 24.51 | 9.77 | 27.34 | 9.77 | 26.95 | 9.77 | 26.56
50.816 | 0 | 3.03 | 0.1 | 14.84 | 0.1 | 15.23 | 4.88 | 16.41 | 4.59 | 9.38 | 4.69 | 4.88 | 10.16 | 9.77 | 6.64 | 9.86 | 23.54 | 9.67 | 14.16 | 0.1 | 10.35 | 0.1 | 9.38 | 0.1 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25 | 9.77 | 27.74 | 9.77 | 27.34 | 9.77 | 26.86
50.848 | 0.1 | 1.95 | 0.1 | 14.75 | 0.1 | 15.04 | 4.88 | 16.8 | 4.59 | 10.16 | 4.59 | 4.88 | 9.86 | 9.77 | 8.69 | 9.86 | 23.93 | 9.67 | 14.84 | 0.1 | 11.13 | 0.1 | 10.16 | 0.2 | 4.88 | 5.18 | 4.88 | 21.19 | 4.88 | 25.59 | 9.77 | 28.22 | 9.77 | 27.83 | 9.77 | 27.34
50.88 | 0.1 | 1.76 | 0.1 | 14.65 | 0.1 | 14.84 | 4.88 | 17.29 | 4.59 | 10.74 | 4.59 | 4.88 | 9.86 | 9.77 | 11.52 | 9.86 | 24.32 | 9.67 | 15.53 | 0.1 | 11.82 | 0.1 | 10.84 | 0.1 | 4.88 | 4.98 | 4.88 | 19.73 | 4.88 | 26.08 | 9.77 | 28.81 | 9.86 | 28.32 | 9.77 | 27.83
50.912 | 0.1 | 1.56 | 0.1 | 14.65 | 0 | 14.65 | 4.88 | 17.68 | 4.59 | 11.33 | 4.59 | 4.88 | 9.96 | 9.77 | 11.04 | 9.86 | 21.29 | 9.67 | 16.21 | 0.1 | 12.6 | 0.1 | 11.72 | 0.1 | 4.88 | 4.79 | 4.88 | 4.39 | 4.88 | 21.49 | 9.77 | 29.2 | 9.86 | 28.71 | 9.77 | 28.32
50.944 | 0 | 1.46 | 0 | 14.45 | 0 | 14.55 | 4.88 | 18.16 | 4.59 | 11.91 | 4.49 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 22.07 | 9.67 | 16.9 | 0.1 | 13.38 | 0.1 | 12.5 | 0 | 4.88 | 4.79 | 4.88 | 3.03 | 4.79 | 20.22 | 9.77 | 25 | 9.77 | 29.4 | 9.77 | 28.91
50.976 | 0 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 18.65 | 4.69 | 12.4 | 4.49 | 4.88 | 9.77 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 17.58 | 0.1 | 14.26 | 0.1 | 13.38 | 0 | 4.88 | 4.69 | 4.88 | 3.52 | 4.88 | 20.9 | 9.77 | 23.63 | 9.77 | 24.61 | 9.77 | 29.3
51.008 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.14 | 4.69 | 13.09 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.49 | 9.77 | 18.16 | 0.1 | 14.94 | 0.1 | 13.97 | 0.1 | 4.88 | 4.69 | 4.88 | 4.3 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 23.24 | 9.86 | 24.81
51.04 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.69 | 13.57 | 4.49 | 4.88 | 9.77 | 9.77 | 10.25 | 9.86 | 21.58 | 9.77 | 18.85 | 0.1 | 15.63 | 0.1 | 14.65 | 0.1 | 5.86 | 4.69 | 4.88 | 5.57 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 23.93 | 9.86 | 23.63
51.072 | 0.1 | 0.98 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 20.02 | 4.69 | 14.06 | 4.49 | 5.57 | 9.77 | 9.77 | 10.16 | 9.86 | 21.49 | 9.77 | 19.53 | 0.1 | 16.21 | 0 | 15.43 | 0 | 7.23 | 4.59 | 4.88 | 6.15 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
51.104 | 0 | 0.78 | 0 | 14.36 | 0 | 14.45 | 4.88 | 20.31 | 4.69 | 14.65 | 4.49 | 6.25 | 9.67 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.02 | 0 | 16.9 | 0.1 | 16.21 | 0 | 9.38 | 4.59 | 4.88 | 5.96 | 4.79 | 21.19 | 9.77 | 24.12 | 9.77 | 24.12 | 9.86 | 24.22
51.136 | 0 | 0.78 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 20.7 | 4.79 | 15.14 | 4.59 | 7.23 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.7 | 0.1 | 17.58 | 0.1 | 16.8 | 0 | 10.74 | 4.59 | 4.79 | 5.37 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.12 | 9.86 | 24.22
51.168 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 20.31 | 4.79 | 15.63 | 4.59 | 8.59 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 9.77 | 21.29 | 0.1 | 18.36 | 0.1 | 17.38 | 0.1 | 12.01 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
51.2 | 0.1 | 0.59 | 0.1 | 7.52 | 0.1 | 14.36 | 4.88 | 18.26 | 4.79 | 16.11 | 4.59 | 9.47 | 9.67 | 11.23 | 9.96 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 18.85 | 0.1 | 18.07 | 0.1 | 12.89 | 4.59 | 4.88 | 5.08 | 4.88 | 21.29 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
51.232 | 0.1 | 0.59 | 0.1 | 2.25 | 0.1 | 7.91 | 4.88 | 18.36 | 4.79 | 16.6 | 4.59 | 10.16 | 9.67 | 12.5 | 9.77 | 9.86 | 21.39 | 9.77 | 22.36 | 0.1 | 19.53 | 0 | 18.75 | 0.1 | 13.87 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
51.264 | 0 | 0.49 | 0.1 | 1.86 | 0 | 1.95 | 4.88 | 18.26 | 4.79 | 17.09 | 4.59 | 10.74 | 9.57 | 13.57 | 9.77 | 9.86 | 21.39 | 9.77 | 19.24 | 0 | 20.12 | 0.1 | 19.34 | 0 | 14.65 | 4.49 | 4.88 | 4.98 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
51.296 | 0 | 0.49 | 0 | 1.76 | 0.1 | 1.95 | 4.88 | 18.26 | 4.79 | 17.48 | 4.69 | 11.33 | 9.57 | 14.26 | 9.77 | 9.86 | 21.39 | 9.77 | 18.36 | 0.1 | 20.7 | 0.1 | 20.02 | 0 | 15.33 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
51.328 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.76 | 4.88 | 18.26 | 4.79 | 18.07 | 4.69 | 11.91 | 9.57 | 14.94 | 9.77 | 9.86 | 18.75 | 9.77 | 17.97 | 0.1 | 21.29 | 0.1 | 20.61 | 0.1 | 16.02 | 4.59 | 4.88 | 4.88 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
51.36 | 0.1 | 0.29 | 0.1 | 1.37 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.46 | 4.69 | 12.5 | 9.57 | 15.43 | 9.67 | 9.86 | 8.4 | 9.77 | 17.77 | 0.1 | 21.88 | 0.1 | 21.19 | 0.1 | 16.8 | 4.59 | 4.88 | 4.88 | 4.88 | 5.66 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.12
51.392 | 0.68 | 0.2 | 0.1 | 1.17 | 0 | 1.46 | 4.88 | 18.16 | 4.79 | 18.85 | 4.69 | 12.99 | 9.67 | 16.02 | 9.67 | 9.86 | 7.23 | 9.77 | 17.68 | 0.1 | 22.46 | 0 | 21.88 | 0.1 | 17.58 | 4.59 | 4.88 | 4.79 | 4.88 | 2.93 | 9.86 | 11.52 | 9.86 | 24.12 | 9.86 | 24.12
51.424 | 1.27 | 0.2 | 0.1 | 1.07 | 0 | 1.17 | 4.88 | 18.16 | 4.79 | 19.34 | 4.69 | 13.57 | 9.57 | 16.5 | 9.67 | 9.86 | 10.45 | 9.86 | 17.58 | 0.1 | 18.75 | 0.1 | 22.27 | 0 | 18.16 | 4.59 | 5.76 | 4.79 | 4.79 | 3.32 | 9.77 | 5.57 | 9.86 | 8.69 | 9.86 | 21.09
51.456 | 2.44 | 0.1 | 0 | 0.98 | 0.1 | 1.17 | 4.88 | 18.07 | 4.88 | 19.73 | 4.79 | 14.06 | 9.67 | 16.9 | 9.67 | 9.86 | 11.52 | 9.86 | 17.58 | 0.1 | 17.87 | 0.1 | 18.26 | 0 | 18.75 | 4.59 | 6.35 | 4.69 | 4.79 | 4.2 | 9.77 | 8.5 | 9.86 | 6.05 | 9.86 | 8.2
51.488 | 3.71 | 0.2 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.07 | 4.79 | 20.12 | 4.79 | 14.65 | 9.67 | 17.48 | 9.67 | 9.86 | 10.64 | 9.86 | 17.58 | 0.1 | 17.77 | 0.1 | 17.19 | 0.1 | 19.34 | 4.59 | 8.01 | 4.69 | 4.88 | 5.47 | 9.86 | 12.31 | 9.86 | 9.08 | 9.86 | 6.35
51.52 | 4.49 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 15.04 | 9.67 | 17.97 | 9.67 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.68 | 0 | 20.02 | 4.69 | 10.06 | 4.59 | 4.88 | 6.25 | 9.86 | 11.43 | 9.86 | 12.21 | 9.86 | 10.16
51.552 | 5.08 | 0.1 | 0.1 | 0.78 | 0 | 0.88 | 4.88 | 18.07 | 4.79 | 21 | 4.79 | 15.63 | 9.67 | 18.36 | 9.57 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.87 | 0 | 17.77 | 0.1 | 20.7 | 4.69 | 11.43 | 4.59 | 4.88 | 5.86 | 9.86 | 10.16 | 9.86 | 11.23 | 9.86 | 12.21
51.584 | 5.66 | 0 | 0.1 | 0.59 | 0 | 0.68 | 4.88 | 18.07 | 4.79 | 18.85 | 4.79 | 16.11 | 9.67 | 18.85 | 9.57 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.68 | 0 | 17.68 | 0 | 21.29 | 4.69 | 12.4 | 4.59 | 4.79 | 5.47 | 9.77 | 10.06 | 9.86 | 10.25 | 9.86 | 10.94
51.616 | 6.15 | 0 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 8.4 | 4.88 | 18.85 | 4.79 | 16.6 | 9.67 | 19.24 | 9.57 | 9.86 | 10.16 | 9.86 | 17.29 | 0 | 17.68 | 0.1 | 17.68 | 0 | 21.78 | 4.69 | 13.28 | 4.59 | 4.79 | 5.18 | 9.86 | 10.35 | 9.86 | 9.96 | 9.86 | 10.16
51.648 | 6.74 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 2.83 | 4.88 | 18.46 | 4.79 | 17.09 | 9.67 | 19.73 | 9.67 | 9.86 | 10.06 | 9.86 | 17.38 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 22.27 | 4.79 | 14.06 | 4.49 | 4.88 | 5.08 | 9.86 | 10.55 | 9.86 | 10.35 | 9.86 | 9.96
51.68 | 7.32 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.42 | 4.88 | 18.36 | 4.79 | 17.48 | 9.67 | 20.12 | 9.67 | 9.86 | 9.96 | 9.86 | 10.74 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 22.95 | 4.79 | 14.94 | 4.49 | 4.88 | 5.18 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86 | 10.25
51.712 | 8.01 | 0.1 | 0.1 | 0.39 | 0 | 0.49 | 4.88 | 4 | 4.79 | 18.36 | 4.79 | 17.87 | 9.67 | 20.61 | 9.67 | 9.86 | 9.86 | 9.86 | 2.73 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 23.54 | 4.79 | 15.53 | 4.49 | 5.76 | 5.08 | 9.77 | 9.96 | 9.86 | 10.25 | 9.86 | 10.45
51.744 | 8.59 | 0.1 | 0.1 | 0.29 | 0 | 0.39 | 4.88 | 5.37 | 4.79 | 18.16 | 4.79 | 18.36 | 9.67 | 21 | 9.67 | 9.86 | 9.86 | 9.86 | 2.05 | 0 | 17.58 | 0.1 | 17.68 | 0 | 24.02 | 4.79 | 16.31 | 4.59 | 6.64 | 4.98 | 9.77 | 10.06 | 9.86 | 10.06 | 9.86 | 10.25
51.776 | 9.18 | 0.1 | 0.2 | 0.2 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 18.85 | 9.67 | 21.39 | 9.67 | 9.86 | 9.77 | 9.86 | 2.05 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 24.51 | 4.79 | 16.99 | 4.49 | 8.3 | 4.98 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 9.96
51.808 | 9.77 | 0.1 | 1.17 | 0.2 | 0.1 | 0.29 | 4.88 | 5.57 | 4.88 | 18.16 | 4.88 | 19.24 | 9.77 | 21.78 | 9.67 | 9.86 | 9.77 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25 | 4.79 | 17.68 | 4.49 | 10.25 | 5.08 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86 | 10.16
51.84 | 10.45 | 0.1 | 2.05 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.73 | 9.77 | 22.17 | 9.67 | 10.45 | 9.77 | 9.86 | 1.66 | 0.1 | 15.04 | 0.1 | 17.68 | 0.1 | 25.49 | 4.79 | 18.36 | 4.49 | 11.52 | 4.88 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
51.872 | 10.94 | 0.1 | 3.52 | 0.1 | 0 | 0.2 | 4.88 | 4.98 | 4.79 | 18.16 | 4.79 | 20.12 | 9.77 | 22.56 | 9.67 | 11.52 | 9.77 | 9.86 | 1.66 | 0.1 | 3.42 | 0 | 17.68 | 0.1 | 26.08 | 4.79 | 18.95 | 4.59 | 12.5 | 4.79 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86 | 10.06
51.904 | 11.62 | 0 | 4.39 | 0.1 | 0 | 0.1 | 4.88 | 4.79 | 4.79 | 18.07 | 4.79 | 20.61 | 9.77 | 22.95 | 9.67 | 12.79 | 9.67 | 9.86 | 1.37 | 0 | 2.15 | 0.1 | 6.45 | 0 | 23.63 | 4.79 | 19.63 | 4.59 | 13.48 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86
51.936 | 12.11 | 0.1 | 4.88 | 0 | 0.59 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 9.77 | 23.34 | 9.67 | 13.77 | 9.67 | 9.96 | 1.27 | 0 | 2.15 | 0.1 | 1.46 | 0 | 21.78 | 4.79 | 20.22 | 4.69 | 14.16 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
51.968 | 12.6 | 0.1 | 5.47 | 0.1 | 1.37 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.85 | 9.77 | 23.63 | 9.77 | 14.45 | 9.67 | 9.86 | 1.17 | 0.1 | 1.95 | 0.1 | 2.44 | 0.1 | 21.68 | 4.79 | 20.8 | 4.59 | 14.94 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
52 | 13.18 | 0.1 | 6.05 | 0.1 | 2.34 | 0.1 | 5.66 | 4.79 | 4.88 | 13.57 | 4.88 | 18.85 | 9.77 | 24.12 | 9.77 | 15.04 | 9.67 | 9.96 | 1.17 | 0.1 | 1.76 | 0.1 | 2.05 | 0.1 | 21.58 | 4.79 | 21.39 | 4.69 | 15.63 | 4.69 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
52.032 | 13.77 | 0.1 | 6.54 | 0.1 | 3.81 | 0.1 | 6.15 | 4.79 | 4.79 | 3.81 | 4.79 | 18.46 | 9.77 | 24.41 | 9.77 | 15.63 | 9.67 | 9.86 | 0.88 | 0.1 | 1.56 | 0.1 | 1.76 | 0.1 | 21.49 | 4.88 | 21.97 | 4.69 | 16.41 | 4.59 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
52.064 | 14.26 | 0.1 | 7.13 | 0 | 4.49 | 0 | 7.23 | 4.69 | 4.79 | 3.22 | 4.79 | 18.36 | 9.77 | 21.97 | 9.77 | 16.11 | 9.57 | 9.96 | 0.88 | 0.1 | 1.46 | 0 | 1.56 | 0 | 21.39 | 4.79 | 22.66 | 4.69 | 17.09 | 4.59 | 11.52 | 9.67 | 9.86 | 9.77 | 10.06 | 9.77
52.096 | 14.84 | 0.1 | 7.81 | 0 | 5.08 | 0 | 8.59 | 4.69 | 4.88 | 3.81 | 4.88 | 18.26 | 9.77 | 21.97 | 9.77 | 16.7 | 9.57 | 9.96 | 0.78 | 0 | 1.46 | 0.1 | 1.46 | 0 | 21.39 | 4.79 | 23.15 | 4.69 | 17.77 | 4.59 | 13.09 | 9.67 | 10.16 | 9.77 | 11.13 | 9.77
52.128 | 15.23 | 0.1 | 8.4 | 0.1 | 5.57 | 0.1 | 9.47 | 4.59 | 4.88 | 4.98 | 4.88 | 18.26 | 9.77 | 21.78 | 9.77 | 17.09 | 9.57 | 9.96 | 0.78 | 0.1 | 1.27 | 0.1 | 1.37 | 0 | 21.39 | 4.88 | 23.54 | 4.69 | 18.36 | 4.59 | 15.14 | 9.67 | 11.04 | 9.67 | 12.7 | 9.77
52.16 | 15.82 | 0.1 | 8.98 | 0.1 | 6.15 | 0.1 | 10.06 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 9.77 | 21.68 | 9.77 | 17.58 | 9.57 | 9.96 | 0.68 | 0.1 | 1.07 | 0.1 | 1.27 | 0.1 | 21.39 | 4.88 | 24.12 | 4.69 | 18.95 | 4.59 | 16.31 | 9.67 | 12.5 | 9.77 | 14.75 | 9.67
52.192 | 16.31 | 0.1 | 9.67 | 0.1 | 6.84 | 0.1 | 10.64 | 4.59 | 4.79 | 5.76 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.97 | 9.57 | 9.86 | 0.59 | 0.1 | 0.88 | 0.1 | 1.17 | 0.1 | 21.39 | 4.88 | 24.71 | 4.79 | 19.63 | 4.59 | 17.29 | 9.67 | 14.55 | 9.67 | 16.21 | 9.67
52.224 | 16.7 | 0.1 | 10.25 | 0 | 7.32 | 0 | 11.43 | 4.49 | 4.79 | 5.27 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 0 | 21.29 | 4.79 | 25.2 | 4.79 | 20.31 | 4.49 | 18.07 | 9.57 | 16.11 | 9.67 | 17.09 | 9.67
52.256 | 17.19 | 0.1 | 10.74 | 0 | 8.01 | 0 | 11.91 | 4.49 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.95 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 0.98 | 0 | 21.29 | 4.79 | 25.68 | 4.79 | 20.9 | 4.59 | 18.65 | 9.67 | 16.99 | 9.67 | 17.87 | 9.67
52.288 | 15.14 | 0.1 | 11.33 | 0.1 | 8.5 | 0.1 | 12.5 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.49 | 9.86 | 19.34 | 9.57 | 9.96 | 0.39 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 26.27 | 4.79 | 21.49 | 4.59 | 19.43 | 9.67 | 17.87 | 9.67 | 18.65 | 9.67
52.32 | 15.23 | 0.1 | 11.91 | 0 | 9.08 | 0.1 | 13.09 | 4.49 | 4.79 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 19.82 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 21.88 | 4.79 | 21.97 | 4.59 | 20.12 | 9.57 | 18.56 | 9.67 | 19.24 | 9.67
52.352 | 14.75 | 0.1 | 12.5 | 0.1 | 9.77 | 0.1 | 13.57 | 4.49 | 4.79 | 4.98 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 20.22 | 9.67 | 10.74 | 0.29 | 0.1 | 0.59 | 0 | 0.68 | 0 | 8.59 | 4.88 | 20.9 | 4.79 | 22.56 | 4.59 | 20.61 | 9.57 | 19.24 | 9.67 | 19.92 | 9.67
52.384 | 14.65 | 0 | 13.09 | 0 | 10.35 | 0 | 14.16 | 4.49 | 4.79 | 4.88 | 4.88 | 12.21 | 9.77 | 21.39 | 9.77 | 20.7 | 9.67 | 11.13 | 0.29 | 0 | 0.49 | 0.1 | 0.68 | 0 | 2.34 | 4.79 | 21.49 | 4.79 | 23.15 | 4.59 | 21.39 | 9.57 | 19.92 | 9.67 | 20.41 | 9.57
52.416 | 14.65 | 0 | 13.57 | 0 | 10.94 | 0 | 14.65 | 4.49 | 4.88 | 4.79 | 4.88 | 3.71 | 9.77 | 21.39 | 9.86 | 21.09 | 9.67 | 12.11 | 0.2 | 0 | 0.49 | 0.1 | 0.68 | 0 | 3.22 | 4.88 | 21.29 | 4.79 | 23.63 | 4.69 | 21.88 | 9.67 | 20.51 | 9.67 | 21 | 9.57
52.448 | 14.55 | 0.1 | 14.16 | 0.1 | 11.52 | 0.1 | 15.14 | 4.59 | 4.79 | 4.69 | 4.88 | 3.13 | 9.86 | 21.39 | 9.86 | 21.49 | 9.67 | 13.38 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.81 | 4.88 | 21.39 | 4.79 | 24.22 | 4.69 | 22.46 | 9.67 | 21.09 | 9.67 | 21.68 | 9.57
52.48 | 14.55 | 0.1 | 14.55 | 0.1 | 12.11 | 0.1 | 15.63 | 4.59 | 4.79 | 4.69 | 4.88 | 3.81 | 9.77 | 13.48 | 9.86 | 21.88 | 9.67 | 14.26 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.88 | 4.88 | 21.29 | 4.79 | 24.71 | 4.69 | 23.05 | 9.67 | 21.68 | 9.67 | 22.27 | 9.67
52.512 | 14.55 | 0.1 | 15.14 | 0.1 | 12.6 | 0.1 | 16.11 | 4.59 | 4.79 | 4.69 | 4.88 | 4.79 | 9.86 | 7.23 | 9.86 | 22.27 | 9.67 | 14.84 | 0.2 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 6.15 | 4.88 | 21.29 | 4.88 | 25.39 | 4.69 | 23.44 | 9.67 | 22.27 | 9.67 | 22.75 | 9.67
52.544 | 14.36 | 0 | 15.63 | 0 | 13.18 | 0 | 16.7 | 4.59 | 4.79 | 4.59 | 4.88 | 5.86 | 9.77 | 8.4 | 9.86 | 22.66 | 9.67 | 15.43 | 0.1 | 0 | 0.29 | 0.29 | 0.29 | 0 | 6.05 | 4.79 | 21.39 | 4.79 | 25.78 | 4.69 | 24.12 | 9.67 | 22.75 | 9.67 | 23.24 | 9.67
52.576 | 14.45 | 0.1 | 16.11 | 0 | 13.77 | 0 | 17.09 | 4.59 | 4.79 | 4.59 | 4.88 | 5.76 | 9.77 | 11.04 | 9.86 | 22.95 | 9.67 | 15.82 | 0.1 | 0 | 0.2 | 1.46 | 0.29 | 0 | 5.57 | 4.79 | 21.29 | 4.88 | 26.27 | 4.79 | 24.61 | 9.67 | 23.34 | 9.67 | 23.83 | 9.67
52.608 | 14.45 | 0.1 | 16.6 | 0.1 | 14.26 | 0.1 | 17.58 | 4.69 | 4.79 | 4.59 | 4.88 | 5.37 | 9.86 | 11.13 | 9.86 | 23.34 | 9.77 | 16.41 | 0.1 | 0.78 | 0.2 | 3.03 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.79 | 22.46 | 4.79 | 25.1 | 9.67 | 23.93 | 9.67 | 24.32 | 9.67
52.64 | 14.45 | 0.1 | 17.09 | 0 | 14.75 | 0.1 | 18.07 | 4.69 | 5.47 | 4.59 | 4.88 | 4.98 | 9.86 | 10.45 | 9.86 | 23.73 | 9.77 | 16.9 | 0.1 | 1.76 | 0.2 | 4.98 | 0.29 | 0 | 5.08 | 4.88 | 21.29 | 4.79 | 21.88 | 4.69 | 25.49 | 9.67 | 24.41 | 9.67 | 24.9 | 9.67
52.672 | 14.36 | 0.1 | 13.77 | 0.1 | 15.23 | 0.1 | 18.46 | 4.69 | 6.25 | 4.49 | 4.88 | 4.98 | 9.86 | 9.86 | 9.86 | 24.12 | 9.77 | 17.38 | 0.1 | 3.22 | 0.1 | 6.35 | 0.2 | 0 | 5.18 | 4.88 | 21.29 | 4.79 | 21.58 | 4.79 | 26.17 | 9.67 | 25 | 9.77 | 25.29 | 9.67
52.704 | 10.25 | 0.1 | 13.97 | 0 | 15.72 | 0.1 | 18.95 | 4.69 | 7.32 | 4.49 | 4.88 | 4.88 | 9.77 | 10.06 | 9.86 | 23.63 | 9.77 | 17.87 | 0.1 | 5.08 | 0.1 | 7.13 | 0.2 | 0.88 | 5.08 | 4.79 | 21.19 | 4.79 | 21.58 | 4.79 | 26.56 | 9.67 | 25.49 | 9.67 | 25.88 | 9.67
52.736 | 2.83 | 0 | 14.06 | 0 | 16.21 | 0 | 19.43 | 4.79 | 8.59 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 18.26 | 0 | 6.35 | 0.1 | 7.81 | 0.1 | 1.76 | 4.98 | 4.88 | 19.53 | 4.79 | 21.49 | 4.79 | 27.05 | 9.67 | 25.88 | 9.67 | 26.47 | 9.67
52.768 | 2.05 | 0.1 | 14.45 | 0.1 | 16.7 | 0.1 | 19.82 | 4.79 | 9.47 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 21.68 | 9.77 | 18.65 | 0.1 | 7.23 | 0.1 | 8.69 | 0.2 | 3.42 | 4.98 | 4.88 | 4.39 | 4.79 | 21.49 | 4.79 | 27.64 | 9.77 | 26.37 | 9.77 | 26.95 | 9.67
52.8 | 1.86 | 0.1 | 14.36 | 0.1 | 17.19 | 0.1 | 20.22 | 4.79 | 10.16 | 4.59 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.68 | 9.77 | 19.24 | 0.1 | 8.01 | 0.1 | 9.47 | 0.1 | 5.27 | 4.98 | 4.88 | 3.22 | 4.79 | 21.39 | 4.79 | 28.13 | 9.77 | 26.86 | 9.77 | 27.44 | 9.67
52.832 | 1.46 | 0.1 | 14.45 | 0.1 | 15.04 | 0.1 | 20.7 | 4.79 | 10.64 | 4.49 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 19.63 | 0.1 | 8.79 | 0 | 10.45 | 0.1 | 6.45 | 4.79 | 4.88 | 3.52 | 4.88 | 21.39 | 4.79 | 28.61 | 9.77 | 27.44 | 9.77 | 27.83 | 9.77
52.864 | 1.37 | 0.1 | 14.45 | 0 | 14.84 | 0.1 | 21.09 | 4.79 | 11.23 | 4.59 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.02 | 0.1 | 9.57 | 0 | 11.13 | 0.1 | 7.32 | 4.69 | 4.79 | 4.3 | 4.79 | 21.29 | 4.79 | 29.01 | 9.77 | 27.83 | 9.77 | 28.32 | 9.67
52.896 | 1.37 | 0 | 14.45 | 0 | 14.55 | 0 | 18.75 | 4.79 | 11.91 | 4.59 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.51 | 0 | 10.25 | 0.1 | 11.82 | 0.1 | 8.11 | 4.69 | 4.88 | 5.96 | 4.79 | 21.29 | 4.79 | 29.49 | 9.77 | 28.32 | 9.77 | 28.91 | 9.67
52.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.85 | 4.88 | 12.5 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.9 | 0.1 | 11.23 | 0.1 | 12.7 | 0.1 | 8.79 | 4.79 | 4.88 | 6.25 | 4.79 | 21.19 | 4.79 | 25.49 | 9.77 | 28.91 | 9.77 | 29.3 | 9.77
52.96 | 0.98 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.09 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.29 | 0.1 | 11.91 | 0.1 | 13.38 | 0.1 | 9.67 | 4.69 | 4.88 | 5.86 | 4.79 | 21.19 | 4.79 | 24.81 | 9.77 | 29.3 | 9.77 | 24.61 | 9.77
52.992 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.46 | 4.88 | 13.57 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.39 | 9.86 | 21.68 | 0.1 | 12.6 | 0.1 | 14.16 | 0.1 | 10.45 | 4.59 | 4.79 | 5.57 | 4.79 | 21.19 | 4.79 | 24.71 | 9.77 | 24.71 | 9.86 | 23.54 | 9.77
53.024 | 0.88 | 0.1 | 14.45 | 0 | 14.55 | 0 | 18.46 | 4.79 | 13.97 | 4.69 | 5.66 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.07 | 0.1 | 13.48 | 0 | 14.94 | 0.1 | 11.23 | 4.59 | 4.79 | 5.37 | 4.79 | 14.45 | 4.79 | 24.61 | 9.77 | 23.34 | 9.86 | 24.02 | 9.77
53.056 | 0.78 | 0 | 14.36 | 0 | 14.45 | 0 | 18.26 | 4.88 | 14.55 | 4.69 | 6.05 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.46 | 0.1 | 14.06 | 0.1 | 15.63 | 0.1 | 12.01 | 4.59 | 4.79 | 4.98 | 4.79 | 3.81 | 4.79 | 24.61 | 9.77 | 24.02 | 9.77 | 24.12 | 9.77
53.088 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.14 | 4.79 | 6.93 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 0.1 | 14.84 | 0.1 | 16.31 | 0.1 | 12.7 | 4.59 | 4.88 | 5.18 | 4.79 | 3.22 | 4.79 | 24.51 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77
53.12 | 0.59 | 0.1 | 5.27 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.63 | 4.79 | 8.4 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 23.24 | 0.1 | 15.53 | 0.1 | 16.9 | 0.1 | 13.57 | 4.59 | 5.96 | 5.08 | 4.79 | 3.61 | 4.79 | 24.41 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77
53.152 | 0.49 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 18.16 | 4.88 | 16.11 | 4.79 | 9.38 | 4.49 | 9.86 | 9.67 | 9.86 | 11.91 | 9.86 | 23.54 | 0.1 | 16.31 | 0 | 17.68 | 0.1 | 14.26 | 4.59 | 6.93 | 4.98 | 4.88 | 4.39 | 4.79 | 24.51 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
53.184 | 0.49 | 0 | 1.95 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.5 | 4.79 | 9.96 | 4.49 | 9.86 | 9.67 | 9.86 | 6.74 | 9.86 | 23.93 | 0.1 | 16.99 | 0.1 | 18.16 | 0.1 | 15.14 | 4.49 | 9.08 | 4.98 | 4.79 | 5.86 | 4.79 | 24.41 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
53.216 | 0.39 | 0 | 1.66 | 0 | 14.45 | 0 | 18.16 | 4.88 | 17.09 | 4.79 | 10.64 | 4.59 | 10.45 | 9.57 | 9.86 | 8.5 | 9.86 | 24.32 | 0 | 17.68 | 0.1 | 18.95 | 0.1 | 15.72 | 4.59 | 10.55 | 4.88 | 4.79 | 6.15 | 4.79 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
53.248 | 0.29 | 0.1 | 1.46 | 0.1 | 12.7 | 0.1 | 18.16 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 11.43 | 9.67 | 9.86 | 11.23 | 9.86 | 21.29 | 0.1 | 18.26 | 0.1 | 19.53 | 0.1 | 16.41 | 4.59 | 11.82 | 4.88 | 4.79 | 5.86 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
53.28 | 0.29 | 0.1 | 1.37 | 0.1 | 3.91 | 0.1 | 18.16 | 4.88 | 17.97 | 4.79 | 11.82 | 4.59 | 12.7 | 9.67 | 9.86 | 11.13 | 9.86 | 21.78 | 0.1 | 18.95 | 0.1 | 20.12 | 0.1 | 16.99 | 4.59 | 12.79 | 4.79 | 4.79 | 5.37 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
53.312 | 0.2 | 0.1 | 1.17 | 0 | 1.95 | 0.1 | 14.26 | 4.88 | 18.46 | 4.79 | 12.4 | 4.59 | 13.67 | 9.57 | 9.86 | 10.25 | 9.86 | 21.58 | 0.1 | 19.43 | 0 | 20.7 | 0.1 | 17.77 | 4.59 | 13.67 | 4.69 | 4.79 | 5.18 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
53.344 | 0.2 | 0 | 1.17 | 0 | 1.76 | 0 | 3.61 | 4.88 | 18.85 | 4.79 | 12.89 | 4.59 | 14.45 | 9.67 | 9.86 | 9.86 | 9.86 | 21.58 | 0.1 | 20.12 | 0 | 21.49 | 0.1 | 18.46 | 4.59 | 14.55 | 4.69 | 4.79 | 5.18 | 4.79 | 18.36 | 9.77 | 24.32 | 9.86 | 24.12 | 9.77
53.376 | 0.1 | 0.1 | 0.88 | 0 | 1.56 | 0 | 3.13 | 4.88 | 19.34 | 4.79 | 13.48 | 4.69 | 15.04 | 9.57 | 9.86 | 9.96 | 9.86 | 21.58 | 0 | 20.7 | 0.1 | 21.88 | 0.1 | 19.14 | 4.59 | 15.23 | 4.69 | 4.79 | 5.18 | 4.79 | 7.52 | 9.77 | 24.22 | 9.77 | 24.12 | 9.77
53.408 | 0.1 | 0.1 | 0.98 | 0.1 | 1.46 | 0.1 | 3.52 | 4.88 | 19.73 | 4.88 | 13.97 | 4.69 | 15.53 | 9.67 | 9.86 | 10.25 | 9.86 | 21.49 | 0.1 | 21.39 | 0.1 | 22.56 | 0.1 | 19.73 | 4.69 | 15.92 | 4.59 | 4.79 | 5.18 | 4.79 | 6.74 | 9.77 | 24.22 | 9.86 | 9.57 | 9.86
53.44 | 0.1 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 4.59 | 4.88 | 20.12 | 4.88 | 14.45 | 4.69 | 16.11 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 21.97 | 0.1 | 17.87 | 0.1 | 20.22 | 4.69 | 16.6 | 4.59 | 4.79 | 4.98 | 4.79 | 10.74 | 9.77 | 10.45 | 9.86 | 6.05 | 9.86
53.472 | 0.1 | 0.68 | 0.78 | 0 | 1.17 | 0.1 | 5.76 | 4.88 | 20.51 | 4.88 | 15.04 | 4.69 | 16.6 | 9.67 | 10.45 | 9.86 | 9.86 | 21.49 | 0.1 | 22.36 | 0 | 17.29 | 0.1 | 20.9 | 4.69 | 17.29 | 4.59 | 4.79 | 4.98 | 4.79 | 12.01 | 9.77 | 5.96 | 9.86 | 9.77 | 9.86
53.504 | 0 | 1.27 | 0.59 | 0 | 0.98 | 0.1 | 5.76 | 4.88 | 20.9 | 4.79 | 15.53 | 4.69 | 16.99 | 9.67 | 11.52 | 9.86 | 9.86 | 21.49 | 0.1 | 19.14 | 0 | 17.58 | 0.1 | 21.49 | 4.69 | 17.97 | 4.59 | 4.79 | 4.88 | 4.79 | 10.64 | 9.77 | 9.67 | 9.86 | 12.31 | 9.77
53.536 | 0 | 2.44 | 0.59 | 0 | 0.98 | 0 | 5.37 | 4.88 | 18.85 | 4.88 | 16.02 | 4.69 | 17.58 | 9.67 | 12.79 | 9.96 | 9.86 | 21.49 | 0 | 18.07 | 0.1 | 17.68 | 0.1 | 22.07 | 4.69 | 18.56 | 4.49 | 4.79 | 4.88 | 4.79 | 10.06 | 9.77 | 12.31 | 9.77 | 11.23 | 9.86
53.568 | 0.1 | 3.81 | 0.59 | 0.1 | 0.88 | 0.1 | 5.08 | 4.88 | 18.95 | 4.88 | 16.41 | 4.79 | 18.07 | 9.67 | 13.77 | 9.86 | 9.96 | 21.39 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 22.66 | 4.79 | 19.34 | 4.49 | 5.76 | 4.79 | 4.79 | 10.25 | 9.77 | 11.33 | 9.86 | 10.25 | 9.86
53.6 | 0.1 | 4.49 | 0.49 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.46 | 4.88 | 16.9 | 4.79 | 18.46 | 9.67 | 14.45 | 9.77 | 9.86 | 21.49 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 19.24 | 4.79 | 19.82 | 4.59 | 6.15 | 4.79 | 4.79 | 10.35 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86
53.632 | 0.1 | 5.08 | 0.39 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.26 | 4.88 | 17.38 | 4.79 | 18.95 | 9.67 | 15.04 | 9.67 | 9.86 | 21.39 | 0.1 | 17.77 | 0 | 17.77 | 0.1 | 18.56 | 4.79 | 20.51 | 4.59 | 7.72 | 4.69 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 10.35 | 9.86
53.664 | 0.1 | 5.57 | 0.29 | 0 | 0.49 | 0 | 4.88 | 4.79 | 18.36 | 4.88 | 17.77 | 4.79 | 19.34 | 9.67 | 15.63 | 9.67 | 9.86 | 19.43 | 0.1 | 17.77 | 0 | 17.68 | 0.1 | 18.16 | 4.79 | 21.09 | 4.49 | 9.77 | 4.69 | 4.79 | 10.25 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86
53.696 | 0.1 | 6.15 | 0.2 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.16 | 4.88 | 18.36 | 4.79 | 19.82 | 9.67 | 16.11 | 9.67 | 9.86 | 8.69 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 17.97 | 4.79 | 21.68 | 4.59 | 11.33 | 4.59 | 4.79 | 10.06 | 9.77 | 10.45 | 9.86 | 10.45 | 9.77
53.728 | 0.1 | 6.74 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.65 | 4.79 | 20.22 | 9.77 | 16.6 | 9.67 | 9.86 | 7.23 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.97 | 4.88 | 22.17 | 4.59 | 12.5 | 4.59 | 4.79 | 10.06 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86
53.76 | 0.1 | 7.42 | 0.2 | 0.1 | 0.39 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 20.7 | 9.77 | 17.09 | 9.67 | 9.86 | 10.25 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 22.75 | 4.59 | 13.18 | 4.59 | 4.79 | 10.06 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86
53.792 | 0.1 | 8.01 | 0.2 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.73 | 4.79 | 21.09 | 9.77 | 17.58 | 9.67 | 9.96 | 11.52 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.88 | 23.34 | 4.59 | 13.97 | 4.59 | 4.79 | 10.16 | 9.77 | 10.06 | 9.86 | 10.25 | 9.86
53.824 | 0 | 8.59 | 0.1 | 0 | 0.2 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 21.58 | 9.77 | 18.07 | 9.67 | 9.86 | 10.84 | 0 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 4.79 | 24.02 | 4.69 | 14.84 | 4.59 | 4.79 | 9.96 | 9.77 | 10.25 | 9.86 | 9.96 | 9.77
53.856 | 0.1 | 9.18 | 0.1 | 0.1 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.51 | 4.88 | 21.97 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0.1 | 17.38 | 0.1 | 17.68 | 4.79 | 24.41 | 4.69 | 15.53 | 4.49 | 4.79 | 9.86 | 9.77 | 10.06 | 9.77 | 9.86 | 9.86
53.888 | 0.1 | 9.77 | 0.1 | 0.78 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 22.36 | 9.77 | 18.95 | 9.67 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 4.49 | 0.1 | 17.77 | 4.88 | 24.9 | 4.69 | 16.21 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
53.92 | 0.1 | 10.45 | 0.1 | 1.56 | 0.2 | 0.68 | 4.59 | 4.88 | 18.07 | 4.88 | 17.68 | 4.88 | 22.75 | 9.77 | 19.34 | 9.57 | 9.86 | 10.06 | 0.1 | 13.77 | 0.1 | 2.05 | 0.1 | 17.68 | 4.88 | 25.39 | 4.69 | 16.8 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
53.952 | 0.1 | 11.04 | 0.1 | 2.83 | 0.1 | 1.46 | 4.59 | 4.88 | 11.23 | 4.88 | 18.26 | 4.88 | 23.05 | 9.77 | 19.82 | 9.57 | 9.96 | 10.16 | 0.1 | 3.42 | 0 | 2.15 | 0.1 | 17.68 | 4.88 | 25.98 | 4.69 | 17.58 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
53.984 | 0.1 | 11.52 | 0.1 | 3.91 | 0 | 2.73 | 4.59 | 4.88 | 3.61 | 4.88 | 17.97 | 4.88 | 23.44 | 9.77 | 20.22 | 9.57 | 9.86 | 9.96 | 0.1 | 2.25 | 0 | 1.95 | 0.1 | 17.68 | 4.88 | 25.68 | 4.69 | 18.16 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
54.016 | 0.1 | 12.11 | 0.1 | 4.59 | 0.1 | 3.81 | 4.49 | 4.79 | 3.22 | 4.88 | 18.07 | 4.88 | 23.83 | 9.77 | 20.61 | 9.67 | 10.25 | 9.96 | 0 | 2.05 | 0.1 | 1.66 | 0.1 | 12.89 | 4.88 | 22.17 | 4.79 | 18.85 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
54.048 | 0.1 | 12.7 | 0.1 | 5.18 | 0.1 | 4.59 | 4.59 | 4.88 | 3.71 | 4.88 | 18.07 | 4.88 | 24.22 | 9.77 | 21 | 9.67 | 11.13 | 9.96 | 0.1 | 1.86 | 0.1 | 1.66 | 0.1 | 3.03 | 4.88 | 22.07 | 4.79 | 19.43 | 4.69 | 5.86 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
54.08 | 0.1 | 13.18 | 0.1 | 5.76 | 0.1 | 5.18 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 4.88 | 23.73 | 9.77 | 21.49 | 9.67 | 12.01 | 9.86 | 0.1 | 1.76 | 0.1 | 1.46 | 0.1 | 2.54 | 4.88 | 21.68 | 4.79 | 20.12 | 4.69 | 7.32 | 9.67 | 10.64 | 9.86 | 9.86 | 9.77 | 9.86
54.112 | 0.1 | 13.67 | 0.1 | 6.35 | 0 | 5.76 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.58 | 9.77 | 21.88 | 9.57 | 13.38 | 9.86 | 0.1 | 1.66 | 0 | 1.27 | 0.1 | 2.05 | 4.88 | 21.58 | 4.79 | 20.7 | 4.69 | 9.57 | 9.67 | 12.01 | 9.77 | 10.84 | 9.67 | 10.16
54.144 | 0 | 14.26 | 0.1 | 6.93 | 0 | 6.25 | 4.49 | 4.79 | 5.76 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 22.27 | 9.67 | 14.16 | 9.77 | 0 | 1.27 | 0 | 1.37 | 0.1 | 1.66 | 4.88 | 21.39 | 4.79 | 21.19 | 4.69 | 11.04 | 9.67 | 14.06 | 9.77 | 12.4 | 9.67 | 11.52
54.176 | 0 | 14.75 | 0 | 7.52 | 0.1 | 6.84 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.68 | 9.77 | 22.66 | 9.67 | 14.75 | 9.77 | 0 | 1.37 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.49 | 4.79 | 21.78 | 4.69 | 12.21 | 9.67 | 15.53 | 9.67 | 14.26 | 9.67 | 12.7
54.208 | 0.1 | 15.33 | 0.1 | 8.11 | 0.1 | 7.52 | 4.59 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 23.05 | 9.67 | 15.33 | 9.67 | 0.1 | 1.27 | 0.1 | 1.07 | 0.1 | 1.56 | 4.88 | 21.29 | 4.88 | 22.46 | 4.79 | 12.99 | 9.67 | 16.6 | 9.67 | 15.72 | 9.67 | 14.84
54.24 | 0.1 | 15.82 | 0.1 | 8.69 | 0.1 | 8.01 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.77 | 23.34 | 9.67 | 15.82 | 9.67 | 0.1 | 1.17 | 0.1 | 0.98 | 0.1 | 1.37 | 4.88 | 21.39 | 4.79 | 22.95 | 4.79 | 13.87 | 9.67 | 17.48 | 9.67 | 16.8 | 9.67 | 16.21
54.272 | 0.1 | 16.21 | 0.1 | 9.38 | 0 | 8.69 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.73 | 9.67 | 16.41 | 9.67 | 0.1 | 0.98 | 0 | 0.88 | 0.1 | 1.27 | 4.88 | 21.29 | 4.88 | 23.54 | 4.79 | 14.75 | 9.67 | 18.26 | 9.67 | 17.68 | 9.67 | 17.19
54.304 | 0 | 16.7 | 0.1 | 9.96 | 0 | 9.38 | 4.69 | 5.57 | 4.79 | 4.88 | 17.97 | 4.79 | 21.49 | 9.77 | 24.12 | 9.67 | 16.9 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.02 | 4.79 | 15.33 | 9.67 | 18.95 | 9.67 | 18.46 | 9.57 | 17.87
54.336 | 0 | 17.19 | 0.1 | 10.45 | 0.1 | 9.86 | 4.69 | 6.25 | 4.88 | 4.88 | 17.58 | 4.88 | 21.49 | 9.77 | 23.54 | 9.67 | 17.29 | 9.67 | 0 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.61 | 4.79 | 16.02 | 9.57 | 19.63 | 9.67 | 18.95 | 9.67 | 18.56
54.368 | 0.1 | 14.55 | 0.1 | 11.13 | 0.1 | 10.45 | 4.69 | 7.62 | 4.79 | 4.88 | 4.59 | 4.88 | 21.49 | 9.77 | 21.58 | 9.67 | 17.77 | 9.67 | 0.88 | 0.78 | 0.1 | 0.59 | 0.1 | 0.98 | 4.88 | 21.19 | 4.88 | 25.1 | 4.79 | 16.7 | 9.67 | 20.22 | 9.67 | 19.73 | 9.67 | 19.24
54.4 | 0.1 | 15.14 | 0.1 | 11.72 | 0.1 | 11.04 | 4.79 | 8.69 | 4.69 | 4.88 | 3.03 | 4.88 | 21.49 | 9.77 | 21.68 | 9.67 | 18.26 | 9.67 | 1.66 | 0.68 | 0.1 | 0.49 | 0.1 | 0.98 | 4.88 | 21.19 | 4.88 | 25.49 | 4.79 | 17.48 | 9.67 | 20.9 | 9.57 | 20.31 | 9.67 | 19.92
54.432 | 0.1 | 14.65 | 0.1 | 12.31 | 0.1 | 11.62 | 4.79 | 9.57 | 4.69 | 4.88 | 3.61 | 4.88 | 21.39 | 9.77 | 21.68 | 9.67 | 18.65 | 9.57 | 3.32 | 0.59 | 0 | 0.49 | 0.1 | 0.88 | 4.88 | 9.77 | 4.88 | 26.08 | 4.79 | 18.16 | 9.67 | 21.39 | 9.67 | 20.9 | 9.67 | 20.41
54.464 | 0 | 14.65 | 0.1 | 12.89 | 0 | 12.21 | 4.79 | 10.25 | 4.59 | 4.79 | 4.69 | 4.88 | 21.39 | 9.77 | 21.49 | 9.77 | 19.24 | 9.57 | 5.27 | 0.49 | 0 | 0.49 | 0 | 0.68 | 4.88 | 2.44 | 4.79 | 21.39 | 4.79 | 18.85 | 9.67 | 22.07 | 9.67 | 21.58 | 9.67 | 21.09
54.496 | 0.1 | 14.65 | 0.1 | 13.28 | 0.1 | 12.79 | 4.79 | 10.84 | 4.59 | 4.88 | 5.86 | 4.79 | 21.39 | 9.77 | 21.49 | 9.77 | 19.53 | 9.57 | 6.45 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 4.88 | 3.32 | 4.88 | 20.31 | 4.79 | 19.34 | 9.67 | 22.56 | 9.67 | 22.07 | 9.67 | 21.78
54.528 | 0.1 | 14.55 | 0.1 | 13.87 | 0.1 | 13.38 | 4.88 | 11.52 | 4.59 | 4.88 | 5.76 | 4.88 | 11.91 | 9.86 | 21.49 | 9.77 | 20.02 | 9.57 | 7.13 | 0.39 | 0.39 | 0.39 | 0.1 | 0.59 | 4.88 | 3.61 | 4.88 | 21.19 | 4.88 | 20.02 | 9.67 | 23.15 | 9.67 | 22.66 | 9.67 | 22.27
54.56 | 0.1 | 14.55 | 0.1 | 14.36 | 0.1 | 13.87 | 4.88 | 12.01 | 4.59 | 4.88 | 5.37 | 4.88 | 6.74 | 9.77 | 21.49 | 9.77 | 20.51 | 9.57 | 8.01 | 0.29 | 1.46 | 0.29 | 0.1 | 0.59 | 4.88 | 4.69 | 4.88 | 21.19 | 4.88 | 20.7 | 9.77 | 23.63 | 9.67 | 23.24 | 9.67 | 22.75
54.592 | 0.1 | 14.45 | 0.1 | 14.94 | 0 | 14.36 | 4.88 | 12.6 | 4.49 | 4.88 | 5.08 | 4.88 | 8.59 | 9.77 | 21.49 | 9.77 | 20.9 | 9.57 | 8.59 | 0.29 | 3.13 | 0.2 | 1.17 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 21.29 | 9.77 | 24.22 | 9.67 | 23.73 | 9.67 | 23.34
54.624 | 0 | 14.45 | 0.1 | 15.43 | 0 | 14.84 | 4.79 | 13.18 | 4.49 | 4.88 | 5.08 | 4.79 | 11.43 | 9.77 | 21.39 | 9.77 | 21.29 | 9.57 | 9.67 | 0.2 | 4.98 | 0.2 | 2.64 | 0.39 | 4.88 | 6.05 | 4.88 | 21.19 | 4.79 | 21.78 | 9.67 | 24.81 | 9.67 | 24.32 | 9.67 | 23.93
54.656 | 0.1 | 14.45 | 0.1 | 15.92 | 0.1 | 15.33 | 4.88 | 13.67 | 4.49 | 4.79 | 4.98 | 4.88 | 11.04 | 9.77 | 21.39 | 9.77 | 21.68 | 9.67 | 10.45 | 0.2 | 6.25 | 0.2 | 4.49 | 0.29 | 4.88 | 5.57 | 4.79 | 21.29 | 4.79 | 22.27 | 9.77 | 25.29 | 9.67 | 24.71 | 9.77 | 24.41
54.688 | 0.1 | 14.45 | 0.1 | 16.41 | 0.1 | 15.82 | 4.88 | 14.26 | 4.49 | 5.66 | 4.88 | 4.88 | 10.25 | 9.77 | 21.49 | 9.77 | 22.07 | 9.67 | 11.04 | 0.2 | 7.13 | 0.2 | 5.86 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.88 | 22.95 | 9.77 | 25.78 | 9.77 | 25.29 | 9.77 | 25
54.72 | 0.1 | 14.45 | 0.1 | 16.9 | 0.1 | 16.31 | 4.88 | 14.75 | 4.59 | 6.15 | 4.88 | 4.88 | 9.86 | 9.77 | 21.39 | 9.77 | 22.46 | 9.67 | 11.82 | 0.2 | 7.91 | 0.1 | 6.93 | 0.29 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 23.34 | 9.77 | 26.27 | 9.77 | 25.68 | 9.77 | 25.49
54.752 | 0.1 | 14.36 | 0.1 | 17.29 | 0 | 16.9 | 4.88 | 15.23 | 4.49 | 7.23 | 4.69 | 4.79 | 9.96 | 9.77 | 21.39 | 9.77 | 22.85 | 9.67 | 12.6 | 0.2 | 8.79 | 0.1 | 7.72 | 0.2 | 4.88 | 5.18 | 4.88 | 21.29 | 4.88 | 23.93 | 9.77 | 26.76 | 9.77 | 26.27 | 9.77 | 25.78
54.784 | 0 | 11.04 | 0 | 14.94 | 0 | 17.29 | 4.88 | 15.72 | 4.49 | 8.59 | 4.69 | 4.79 | 10.16 | 9.77 | 11.82 | 9.77 | 23.24 | 9.67 | 13.48 | 0.1 | 9.47 | 0.1 | 8.59 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 24.51 | 9.77 | 27.25 | 9.77 | 26.86 | 9.77 | 26.47
54.816 | 0.1 | 3.22 | 0.1 | 15.04 | 0.1 | 15.43 | 4.88 | 16.21 | 4.59 | 9.38 | 4.69 | 4.79 | 10.16 | 9.77 | 6.74 | 9.77 | 23.54 | 9.67 | 14.06 | 0.1 | 10.16 | 0.1 | 9.38 | 0.1 | 4.88 | 5.08 | 4.79 | 21.19 | 4.88 | 25 | 9.77 | 27.74 | 9.77 | 27.34 | 9.77 | 26.95
54.848 | 0.1 | 2.05 | 0.1 | 14.84 | 0.1 | 15.33 | 4.88 | 16.7 | 4.59 | 10.06 | 4.69 | 4.88 | 9.86 | 9.77 | 8.59 | 9.86 | 24.02 | 9.67 | 14.84 | 0.1 | 11.13 | 0.1 | 10.16 | 0.1 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 25.49 | 9.77 | 28.13 | 9.86 | 27.74 | 9.77 | 27.44
54.88 | 0.1 | 1.86 | 0.1 | 14.65 | 0.1 | 14.94 | 4.88 | 17.19 | 4.69 | 10.74 | 4.59 | 4.88 | 9.86 | 9.77 | 11.23 | 9.86 | 24.32 | 9.67 | 15.53 | 0.1 | 11.82 | 0.1 | 10.94 | 0.1 | 4.88 | 4.88 | 4.88 | 20.9 | 4.88 | 25.98 | 9.77 | 28.61 | 9.77 | 28.22 | 9.77 | 27.83
54.912 | 0.1 | 1.46 | 0.1 | 14.65 | 0 | 14.75 | 4.88 | 17.68 | 4.59 | 11.33 | 4.59 | 4.79 | 9.96 | 9.77 | 11.23 | 9.86 | 21 | 9.67 | 16.11 | 0.1 | 12.79 | 0.1 | 11.62 | 0.1 | 4.88 | 4.88 | 4.88 | 5.08 | 4.88 | 21.49 | 9.77 | 29.1 | 9.86 | 28.71 | 9.77 | 28.22
54.944 | 0 | 1.56 | 0.1 | 14.45 | 0 | 14.65 | 4.88 | 18.07 | 4.69 | 11.91 | 4.49 | 4.79 | 9.86 | 9.77 | 10.16 | 9.86 | 21.68 | 9.67 | 16.8 | 0.1 | 13.38 | 0.1 | 12.6 | 0 | 4.88 | 4.88 | 4.88 | 2.83 | 4.79 | 20.31 | 9.77 | 27.34 | 9.86 | 29.3 | 9.77 | 28.81
54.976 | 0 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 18.56 | 4.69 | 12.4 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 21.39 | 9.67 | 17.58 | 0 | 14.06 | 0.1 | 13.18 | 0 | 4.88 | 4.69 | 4.79 | 3.32 | 4.88 | 21 | 9.77 | 24.61 | 9.86 | 26.17 | 9.86 | 29.3
55.008 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.04 | 4.79 | 12.99 | 4.59 | 4.88 | 9.77 | 9.77 | 10.06 | 9.86 | 21.49 | 9.67 | 18.16 | 0.1 | 14.84 | 0.1 | 13.97 | 0.1 | 4.88 | 4.69 | 4.88 | 4.1 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 23.54 | 9.86 | 24.61
55.04 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.43 | 4.79 | 13.57 | 4.49 | 4.79 | 9.77 | 9.77 | 10.35 | 9.86 | 21.49 | 9.77 | 18.75 | 0.1 | 15.53 | 0.1 | 14.65 | 0 | 5.96 | 4.69 | 4.88 | 5.47 | 4.88 | 21.19 | 9.77 | 24.51 | 9.86 | 24.12 | 9.86 | 23.44
55.072 | 0.1 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 19.92 | 4.69 | 14.06 | 4.49 | 5.57 | 9.77 | 9.77 | 10.06 | 9.86 | 21.49 | 9.67 | 19.43 | 0.1 | 16.31 | 0.1 | 15.43 | 0 | 7.13 | 4.59 | 4.88 | 6.25 | 4.88 | 21.19 | 9.86 | 24.41 | 9.86 | 24.22 | 9.86 | 24.02
55.104 | 0 | 0.88 | 0 | 14.36 | 0 | 14.45 | 4.88 | 20.22 | 4.79 | 14.55 | 4.49 | 6.15 | 9.67 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.12 | 0 | 16.8 | 0.1 | 16.11 | 0 | 9.28 | 4.59 | 4.88 | 5.86 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
55.136 | 0 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 20.61 | 4.79 | 15.04 | 4.49 | 7.13 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.51 | 0.1 | 17.58 | 0.1 | 16.8 | 0 | 10.64 | 4.59 | 4.88 | 5.37 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
55.168 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.98 | 21.09 | 4.79 | 15.53 | 4.59 | 8.59 | 9.67 | 10.55 | 10.06 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 18.16 | 0.1 | 17.38 | 0 | 11.91 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
55.2 | 0.1 | 0.59 | 0.1 | 8.11 | 0.1 | 14.45 | 4.88 | 18.95 | 4.79 | 16.11 | 4.59 | 9.47 | 9.67 | 11.23 | 9.86 | 9.86 | 21.49 | 9.77 | 21.68 | 0.1 | 18.85 | 0.1 | 18.26 | 0.1 | 12.99 | 4.59 | 4.88 | 5.08 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
55.232 | 0.1 | 0.59 | 0.1 | 2.34 | 0.1 | 9.28 | 4.88 | 18.75 | 4.79 | 16.6 | 4.59 | 10.06 | 9.67 | 12.4 | 9.86 | 9.86 | 21.39 | 9.77 | 22.36 | 0.1 | 19.43 | 0 | 18.75 | 0 | 13.77 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
55.264 | 0 | 0.49 | 0 | 1.86 | 0 | 2.44 | 4.88 | 18.46 | 4.79 | 17.09 | 4.59 | 10.64 | 9.57 | 13.57 | 9.77 | 9.86 | 21.39 | 9.77 | 19.14 | 0 | 20.12 | 0 | 19.43 | 0 | 14.55 | 4.49 | 4.88 | 5.08 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
55.296 | 0.1 | 0.39 | 0.1 | 1.66 | 0.1 | 1.95 | 4.88 | 18.36 | 4.88 | 17.58 | 4.69 | 11.33 | 9.57 | 14.26 | 9.77 | 9.86 | 21.39 | 9.77 | 18.16 | 0 | 20.7 | 0.1 | 20.12 | 0 | 15.33 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
55.328 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.86 | 4.88 | 18.36 | 4.88 | 17.97 | 4.69 | 11.91 | 9.67 | 14.94 | 9.77 | 9.86 | 21 | 9.77 | 17.77 | 0.1 | 21.29 | 0.1 | 20.8 | 0.1 | 16.02 | 4.49 | 4.88 | 4.88 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
55.36 | 0.1 | 0.29 | 0.1 | 1.37 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.36 | 4.69 | 12.4 | 9.57 | 15.43 | 9.67 | 9.86 | 8.69 | 9.77 | 17.68 | 0.1 | 21.97 | 0.1 | 21.29 | 0 | 16.7 | 4.59 | 4.88 | 4.79 | 4.88 | 7.03 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
55.392 | 0.68 | 0.29 | 0.1 | 1.27 | 0 | 1.46 | 4.88 | 18.16 | 4.79 | 18.85 | 4.69 | 12.99 | 9.67 | 15.92 | 9.67 | 9.86 | 7.03 | 9.77 | 17.68 | 0.1 | 22.46 | 0 | 21.88 | 0 | 17.38 | 4.59 | 4.88 | 4.88 | 4.88 | 2.44 | 9.86 | 12.6 | 9.86 | 24.12 | 9.86 | 24.12
55.424 | 1.27 | 0.2 | 0.1 | 1.07 | 0 | 1.27 | 4.88 | 18.16 | 4.88 | 19.24 | 4.69 | 13.57 | 9.57 | 16.5 | 9.67 | 9.86 | 10.25 | 9.77 | 17.68 | 0 | 17.87 | 0.1 | 22.36 | 0 | 17.97 | 4.59 | 5.18 | 4.79 | 4.88 | 3.22 | 9.77 | 5.66 | 9.86 | 10.45 | 9.77 | 23.05
55.456 | 2.44 | 0.1 | 0.1 | 0.98 | 0.1 | 1.07 | 4.88 | 18.07 | 4.88 | 19.73 | 4.79 | 14.06 | 9.57 | 16.9 | 9.67 | 9.86 | 11.52 | 9.86 | 17.58 | 0.1 | 16.9 | 0.1 | 19.63 | 0 | 18.65 | 4.59 | 6.05 | 4.69 | 4.88 | 4.2 | 9.77 | 8.11 | 9.86 | 5.86 | 9.86 | 8.3
55.488 | 3.81 | 0.1 | 0.1 | 0.98 | 0.1 | 1.07 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 14.55 | 9.67 | 17.48 | 9.67 | 9.86 | 10.74 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 17.38 | 0.1 | 19.34 | 4.59 | 7.72 | 4.69 | 4.88 | 5.18 | 9.86 | 12.01 | 9.86 | 8.69 | 9.86 | 6.35
55.52 | 4.49 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 15.04 | 9.67 | 17.87 | 9.67 | 9.86 | 10.16 | 9.86 | 17.48 | 0.1 | 17.58 | 0.1 | 17.87 | 0.1 | 20.02 | 4.69 | 9.86 | 4.59 | 4.88 | 6.25 | 9.86 | 11.62 | 9.86 | 12.01 | 9.86 | 9.96
55.552 | 5.08 | 0.1 | 0.1 | 0.78 | 0 | 0.78 | 4.88 | 18.07 | 4.79 | 20.9 | 4.79 | 15.63 | 9.67 | 18.36 | 9.67 | 9.86 | 9.86 | 9.77 | 17.48 | 0.1 | 17.68 | 0 | 17.87 | 0 | 20.61 | 4.69 | 11.33 | 4.59 | 4.88 | 5.96 | 9.86 | 10.35 | 9.86 | 11.33 | 9.86 | 12.4
55.584 | 5.66 | 0 | 0 | 0.59 | 0 | 0.68 | 4.88 | 18.07 | 4.88 | 18.65 | 4.79 | 16.11 | 9.67 | 18.85 | 9.57 | 9.86 | 10.06 | 9.86 | 17.38 | 0 | 17.68 | 0.1 | 17.68 | 0 | 21.29 | 4.69 | 12.4 | 4.59 | 4.88 | 5.47 | 9.86 | 9.96 | 9.86 | 10.25 | 9.86 | 10.94
55.616 | 6.15 | 0 | 0 | 0.49 | 0.1 | 0.68 | 4.88 | 9.08 | 4.88 | 18.75 | 4.88 | 16.5 | 9.67 | 19.24 | 9.57 | 9.86 | 10.06 | 9.86 | 17.38 | 0 | 17.68 | 0.1 | 17.77 | 0 | 21.78 | 4.69 | 13.18 | 4.59 | 4.79 | 5.18 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86 | 10.16
55.648 | 6.84 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 3.13 | 4.88 | 18.26 | 4.79 | 16.99 | 9.67 | 19.73 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.36 | 4.79 | 14.06 | 4.49 | 4.88 | 5.18 | 9.86 | 10.45 | 9.86 | 10.25 | 9.86 | 9.96
55.68 | 7.32 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.22 | 4.88 | 18.26 | 4.88 | 17.48 | 9.77 | 20.12 | 9.67 | 9.86 | 9.86 | 9.86 | 13.09 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 22.95 | 4.79 | 14.84 | 4.49 | 4.88 | 5.18 | 9.77 | 10.55 | 9.86 | 10.55 | 9.86 | 10.35
55.712 | 8.01 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 4 | 4.79 | 18.26 | 4.88 | 17.87 | 9.77 | 20.51 | 9.67 | 9.86 | 9.86 | 9.86 | 3.32 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 23.44 | 4.79 | 15.63 | 4.49 | 5.37 | 5.08 | 9.86 | 9.96 | 9.86 | 10.35 | 9.86 | 10.55
55.744 | 8.59 | 0.1 | 0.1 | 0.29 | 0 | 0.29 | 4.88 | 5.27 | 4.88 | 18.26 | 4.79 | 18.36 | 9.67 | 21 | 9.67 | 9.86 | 9.86 | 9.86 | 2.15 | 0 | 17.58 | 0.1 | 17.68 | 0 | 23.93 | 4.79 | 16.31 | 4.59 | 6.35 | 5.08 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.25
55.776 | 9.18 | 0 | 0.2 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 18.75 | 9.77 | 21.39 | 9.67 | 9.86 | 9.77 | 9.86 | 2.25 | 0 | 17.68 | 0.1 | 17.68 | 0 | 24.41 | 4.79 | 16.9 | 4.49 | 8.01 | 4.98 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 9.96
55.808 | 9.77 | 0.1 | 1.17 | 0.2 | 0.1 | 0.29 | 4.88 | 5.57 | 4.88 | 18.16 | 4.88 | 19.24 | 9.77 | 21.78 | 9.67 | 9.86 | 9.77 | 9.86 | 1.86 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 24.9 | 4.79 | 17.58 | 4.49 | 10.16 | 4.98 | 9.77 | 10.25 | 9.86 | 10.16 | 9.86 | 10.16
55.84 | 10.45 | 0.1 | 1.95 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 19.73 | 9.77 | 22.17 | 9.67 | 10.64 | 9.77 | 9.86 | 1.66 | 0.1 | 17.38 | 0.1 | 17.68 | 0.1 | 25.49 | 4.79 | 18.26 | 4.49 | 11.52 | 4.88 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 10.16
55.872 | 11.04 | 0.1 | 3.22 | 0.1 | 0 | 0.2 | 4.88 | 4.98 | 4.79 | 18.16 | 4.79 | 20.12 | 9.77 | 22.46 | 9.67 | 11.43 | 9.77 | 9.86 | 1.46 | 0.1 | 4.1 | 0.1 | 17.68 | 0 | 26.08 | 4.79 | 18.85 | 4.59 | 12.5 | 4.79 | 9.86 | 9.86 | 9.86 | 9.96 | 9.86 | 10.06
55.904 | 11.52 | 0.1 | 4.2 | 0.1 | 0 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.61 | 9.77 | 22.95 | 9.67 | 12.6 | 9.67 | 9.86 | 1.46 | 0 | 1.95 | 0 | 6.74 | 0 | 24.61 | 4.79 | 19.53 | 4.59 | 13.38 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
55.936 | 12.11 | 0 | 4.79 | 0 | 0.29 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21 | 9.77 | 23.34 | 9.67 | 13.67 | 9.67 | 9.86 | 1.27 | 0 | 2.15 | 0.1 | 1.76 | 0 | 22.46 | 4.79 | 20.02 | 4.59 | 14.16 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
55.968 | 12.7 | 0.1 | 5.37 | 0 | 1.27 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 9.77 | 23.63 | 9.67 | 14.45 | 9.67 | 9.86 | 1.27 | 0.1 | 1.95 | 0.1 | 2.25 | 0 | 21.97 | 4.88 | 20.7 | 4.59 | 14.94 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
56 | 13.18 | 0.1 | 5.96 | 0 | 2.34 | 0 | 5.27 | 4.88 | 4.88 | 14.26 | 4.88 | 18.75 | 9.77 | 24.02 | 9.67 | 14.94 | 9.67 | 9.86 | 0.98 | 0.1 | 1.66 | 0.1 | 2.15 | 0 | 21.68 | 4.88 | 21.29 | 4.59 | 15.63 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77
56.032 | 13.77 | 0.1 | 6.54 | 0 | 3.71 | 0 | 6.15 | 4.69 | 4.79 | 4.1 | 4.88 | 18.26 | 9.77 | 24.41 | 9.77 | 15.53 | 9.67 | 9.86 | 1.07 | 0.1 | 1.56 | 0 | 1.66 | 0 | 21.58 | 4.79 | 21.78 | 4.69 | 16.41 | 4.59 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
56.064 | 14.26 | 0.1 | 7.13 | 0 | 4.39 | 0 | 7.13 | 4.69 | 4.88 | 3.13 | 4.79 | 18.26 | 9.77 | 22.36 | 9.77 | 16.02 | 9.67 | 9.86 | 0.88 | 0 | 1.46 | 0.1 | 1.66 | 0 | 21.49 | 4.79 | 22.46 | 4.69 | 17.09 | 4.49 | 11.52 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77
56.096 | 14.75 | 0.1 | 7.72 | 0 | 4.98 | 0 | 8.5 | 4.69 | 4.88 | 3.61 | 4.88 | 18.26 | 9.77 | 22.17 | 9.77 | 16.6 | 9.57 | 9.86 | 0.88 | 0 | 1.37 | 0.1 | 1.46 | 0 | 21.49 | 4.88 | 23.05 | 4.69 | 17.58 | 4.59 | 12.99 | 9.67 | 9.96 | 9.67 | 10.94 | 9.77
56.128 | 15.23 | 0.1 | 8.3 | 0.1 | 5.57 | 0.1 | 9.38 | 4.59 | 4.88 | 4.79 | 4.88 | 18.16 | 9.86 | 21.88 | 9.77 | 16.99 | 9.57 | 9.86 | 0.78 | 0.1 | 1.27 | 0.1 | 1.46 | 0 | 21.39 | 4.88 | 23.54 | 4.69 | 18.36 | 4.59 | 15.04 | 9.67 | 10.94 | 9.77 | 12.31 | 9.67
56.16 | 15.72 | 0.1 | 8.89 | 0.1 | 6.05 | 0.1 | 10.06 | 4.59 | 4.88 | 5.96 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.48 | 9.67 | 9.86 | 0.68 | 0.1 | 1.27 | 0.1 | 1.37 | 0 | 21.39 | 4.88 | 24.02 | 4.69 | 18.95 | 4.59 | 16.31 | 9.67 | 12.4 | 9.67 | 14.36 | 9.67
56.192 | 16.31 | 0.1 | 9.57 | 0.1 | 6.64 | 0.1 | 10.64 | 4.59 | 4.79 | 5.76 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.97 | 9.67 | 9.86 | 0.59 | 0.1 | 0.98 | 0 | 1.27 | 0 | 21.39 | 4.88 | 24.61 | 4.79 | 19.63 | 4.49 | 17.19 | 9.67 | 14.36 | 9.67 | 16.11 | 9.67
56.224 | 16.7 | 0 | 10.16 | 0 | 7.23 | 0 | 11.33 | 4.49 | 4.88 | 5.37 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.67 | 9.86 | 0.59 | 0 | 0.88 | 0 | 1.07 | 0 | 21.29 | 4.79 | 25.1 | 4.79 | 20.22 | 4.49 | 18.07 | 9.67 | 16.02 | 9.67 | 16.9 | 9.67
56.256 | 17.09 | 0 | 10.64 | 0 | 7.91 | 0 | 11.82 | 4.49 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.85 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 25.59 | 4.79 | 20.8 | 4.59 | 18.75 | 9.67 | 16.9 | 9.67 | 17.77 | 9.57
56.288 | 14.55 | 0.1 | 11.33 | 0.1 | 8.5 | 0.1 | 12.4 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 19.24 | 9.67 | 9.96 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 26.08 | 4.69 | 21.39 | 4.59 | 19.43 | 9.67 | 17.77 | 9.67 | 18.46 | 9.67
56.32 | 15.14 | 0.1 | 11.82 | 0.1 | 9.08 | 0.1 | 12.99 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 19.73 | 9.67 | 9.96 | 0.29 | 0 | 0.68 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 21.58 | 4.79 | 22.07 | 4.59 | 19.92 | 9.57 | 18.46 | 9.67 | 19.14 | 9.67
56.352 | 14.65 | 0.1 | 12.4 | 0.1 | 9.67 | 0.1 | 13.57 | 4.49 | 4.79 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.77 | 20.12 | 9.67 | 10.74 | 0.29 | 0.1 | 0.59 | 0 | 0.78 | 0.1 | 9.47 | 4.88 | 20.31 | 4.79 | 22.56 | 4.59 | 20.61 | 9.67 | 19.14 | 9.67 | 19.82 | 9.57
56.384 | 14.65 | 0 | 13.09 | 0 | 10.25 | 0 | 14.06 | 4.49 | 4.88 | 4.88 | 4.88 | 13.09 | 9.77 | 21.39 | 9.77 | 20.61 | 9.67 | 11.13 | 0.2 | 0 | 0.59 | 0.1 | 0.59 | 0 | 2.64 | 4.79 | 21.19 | 4.79 | 23.15 | 4.59 | 21.19 | 9.57 | 19.92 | 9.67 | 20.31 | 9.57
56.416 | 14.65 | 0 | 13.57 | 0 | 10.94 | 0 | 14.55 | 4.49 | 4.88 | 4.79 | 4.88 | 4 | 9.77 | 21.39 | 9.77 | 20.9 | 9.67 | 12.01 | 0.2 | 0 | 0.59 | 0.1 | 0.59 | 0 | 3.32 | 4.88 | 21.19 | 4.79 | 23.63 | 4.69 | 21.68 | 9.67 | 20.41 | 9.67 | 21 | 9.57
56.448 | 14.55 | 0.1 | 14.06 | 0.1 | 11.43 | 0.1 | 15.14 | 4.59 | 4.88 | 4.69 | 4.88 | 3.13 | 9.86 | 21.39 | 9.86 | 21.39 | 9.67 | 13.38 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.61 | 4.88 | 21.19 | 4.79 | 24.22 | 4.69 | 22.36 | 9.67 | 21 | 9.67 | 21.58 | 9.67
56.48 | 14.55 | 0.1 | 14.55 | 0.1 | 12.01 | 0.1 | 15.63 | 4.59 | 4.88 | 4.69 | 4.88 | 3.61 | 9.86 | 13.97 | 9.86 | 21.78 | 9.67 | 14.16 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.59 | 4.88 | 21.19 | 4.79 | 24.71 | 4.69 | 22.85 | 9.67 | 21.58 | 9.67 | 22.17 | 9.67
56.512 | 14.45 | 0.1 | 15.04 | 0 | 12.6 | 0.1 | 16.11 | 4.59 | 4.79 | 4.69 | 4.88 | 4.79 | 9.86 | 7.13 | 9.86 | 22.17 | 9.67 | 14.84 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0 | 6.15 | 4.88 | 21.29 | 4.79 | 25.2 | 4.69 | 23.44 | 9.67 | 22.27 | 9.67 | 22.75 | 9.67
56.544 | 14.45 | 0 | 15.53 | 0 | 13.09 | 0 | 16.6 | 4.59 | 4.88 | 4.59 | 4.79 | 5.86 | 9.86 | 8.3 | 9.86 | 22.56 | 9.67 | 15.33 | 0.1 | 0 | 0.29 | 0.2 | 0.39 | 0 | 6.05 | 4.79 | 21.29 | 4.79 | 25.78 | 4.69 | 24.02 | 9.67 | 22.85 | 9.67 | 23.15 | 9.57
56.576 | 14.45 | 0.1 | 16.02 | 0 | 13.67 | 0 | 16.99 | 4.59 | 4.88 | 4.49 | 4.88 | 5.86 | 9.77 | 10.94 | 9.86 | 22.95 | 9.67 | 15.82 | 0.1 | 0 | 0.29 | 1.56 | 0.29 | 0 | 5.76 | 4.88 | 21.19 | 4.79 | 26.17 | 4.79 | 24.51 | 9.67 | 23.24 | 9.67 | 23.73 | 9.67
56.608 | 14.45 | 0.1 | 16.6 | 0.1 | 14.16 | 0.1 | 17.58 | 4.69 | 4.88 | 4.59 | 4.88 | 5.27 | 9.86 | 11.23 | 9.86 | 23.34 | 9.77 | 16.41 | 0.1 | 0.59 | 0.2 | 2.83 | 0.29 | 0.1 | 5.37 | 4.88 | 21.29 | 4.79 | 21.78 | 4.79 | 25.1 | 9.67 | 23.83 | 9.67 | 24.32 | 9.67
56.64 | 14.45 | 0.1 | 16.99 | 0 | 14.65 | 0.1 | 17.97 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 9.86 | 10.45 | 9.86 | 23.73 | 9.77 | 16.9 | 0.1 | 1.66 | 0.2 | 4.69 | 0.2 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 20.9 | 4.79 | 25.49 | 9.67 | 24.32 | 9.67 | 24.71 | 9.67
56.672 | 14.36 | 0.1 | 15.33 | 0.1 | 15.23 | 0.1 | 18.46 | 4.69 | 6.15 | 4.59 | 4.88 | 4.98 | 9.86 | 9.86 | 9.86 | 24.02 | 9.77 | 17.29 | 0.1 | 3.03 | 0.1 | 6.25 | 0.2 | 0 | 5.08 | 4.88 | 21.29 | 4.79 | 21.39 | 4.79 | 26.17 | 9.67 | 24.81 | 9.67 | 25.39 | 9.67
56.704 | 10.94 | 0.1 | 13.97 | 0 | 15.72 | 0.1 | 18.95 | 4.69 | 7.23 | 4.49 | 4.79 | 4.88 | 9.86 | 9.96 | 9.86 | 24.32 | 9.77 | 17.87 | 0.1 | 5.18 | 0.1 | 7.03 | 0.2 | 0.68 | 5.18 | 4.88 | 21.19 | 4.79 | 21.39 | 4.79 | 26.56 | 9.67 | 25.39 | 9.67 | 25.88 | 9.67
56.736 | 3.22 | 0 | 14.26 | 0 | 16.21 | 0 | 19.24 | 4.79 | 8.59 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 22.27 | 9.77 | 18.26 | 0 | 6.25 | 0.2 | 7.81 | 0.1 | 1.66 | 5.08 | 4.79 | 21.19 | 4.79 | 21.29 | 4.79 | 27.05 | 9.67 | 25.88 | 9.67 | 26.37 | 9.67
56.768 | 2.05 | 0.1 | 14.55 | 0.1 | 16.7 | 0.1 | 19.82 | 4.79 | 9.47 | 4.49 | 4.88 | 4.79 | 9.86 | 10.25 | 9.86 | 22.07 | 9.77 | 18.65 | 0.1 | 7.23 | 0.1 | 8.59 | 0.1 | 3.13 | 5.08 | 4.88 | 5.57 | 4.79 | 21.39 | 4.79 | 27.54 | 9.77 | 26.37 | 9.77 | 26.86 | 9.67
56.8 | 1.76 | 0.1 | 14.45 | 0 | 17.19 | 0.1 | 20.22 | 4.79 | 10.16 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.78 | 9.77 | 19.14 | 0.1 | 8.01 | 0.1 | 9.38 | 0.1 | 5.18 | 4.98 | 4.88 | 2.83 | 4.79 | 21.39 | 4.79 | 28.03 | 9.77 | 26.86 | 9.77 | 27.34 | 9.67
56.832 | 1.46 | 0.1 | 14.45 | 0.1 | 14.26 | 0.1 | 20.61 | 4.79 | 10.64 | 4.59 | 4.88 | 4.79 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 19.53 | 0.1 | 8.59 | 0.1 | 10.25 | 0.1 | 6.45 | 4.88 | 4.88 | 3.42 | 4.79 | 21.39 | 4.79 | 28.52 | 9.77 | 27.34 | 9.77 | 27.74 | 9.67
56.864 | 1.56 | 0.1 | 14.45 | 0 | 14.16 | 0 | 21 | 4.79 | 11.33 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.68 | 9.77 | 20.02 | 0.1 | 9.38 | 0.1 | 11.04 | 0.1 | 7.23 | 4.79 | 4.79 | 4.2 | 4.79 | 21.29 | 4.79 | 29.01 | 9.77 | 27.83 | 9.77 | 28.22 | 9.67
56.896 | 1.27 | 0 | 14.45 | 0 | 14.26 | 0 | 18.07 | 4.79 | 11.91 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.31 | 0 | 10.25 | 0.1 | 11.72 | 0.1 | 8.01 | 4.69 | 4.88 | 5.66 | 4.79 | 21.29 | 4.88 | 29.49 | 9.77 | 28.32 | 9.77 | 28.81 | 9.67
56.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.16 | 4.79 | 12.5 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.9 | 0.1 | 11.04 | 0.1 | 12.5 | 0.1 | 8.89 | 4.69 | 4.88 | 6.25 | 4.79 | 21.29 | 4.79 | 25.29 | 9.77 | 28.91 | 9.77 | 29.2 | 9.77
56.96 | 1.07 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 13.09 | 4.69 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.29 | 0.1 | 11.72 | 0.1 | 13.38 | 0.1 | 9.67 | 4.69 | 4.79 | 5.86 | 4.79 | 21.29 | 4.79 | 24.51 | 9.77 | 29.1 | 9.77 | 24.9 | 9.77
56.992 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 13.57 | 4.69 | 4.88 | 4.59 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 21.68 | 0.1 | 12.6 | 0 | 14.06 | 0.1 | 10.45 | 4.69 | 4.88 | 5.37 | 4.79 | 21.29 | 4.79 | 24.61 | 9.77 | 25 | 9.77 | 23.24 | 9.77
57.024 | 0.88 | 0 | 14.36 | 0 | 14.45 | 0 | 18.26 | 4.79 | 14.06 | 4.69 | 5.08 | 4.49 | 9.77 | 9.77 | 9.86 | 21.39 | 9.86 | 21.97 | 0.1 | 13.38 | 0.1 | 14.65 | 0.1 | 11.23 | 4.59 | 4.79 | 5.27 | 4.79 | 17.77 | 4.79 | 24.61 | 9.77 | 23.34 | 9.77 | 24.02 | 9.77
57.056 | 0.78 | 0 | 14.36 | 0 | 14.45 | 0 | 18.16 | 4.79 | 14.65 | 4.69 | 5.96 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.46 | 0.1 | 14.06 | 0.1 | 15.53 | 0.1 | 12.01 | 4.59 | 4.88 | 5.27 | 4.79 | 4.59 | 4.88 | 24.51 | 9.77 | 24.12 | 9.77 | 24.22 | 9.77
57.088 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.14 | 4.69 | 6.93 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 0.1 | 14.75 | 0.1 | 16.11 | 0.1 | 12.7 | 4.59 | 4.88 | 5.08 | 4.79 | 3.13 | 4.79 | 24.51 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
57.12 | 0.68 | 0.1 | 6.05 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.72 | 4.79 | 8.4 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 23.24 | 0.1 | 15.53 | 0.1 | 16.8 | 0.1 | 13.48 | 4.59 | 5.76 | 5.08 | 4.79 | 3.42 | 4.88 | 24.41 | 9.77 | 24.32 | 9.77 | 24.12 | 9.77
57.152 | 0.59 | 0.1 | 1.76 | 0 | 14.36 | 0.1 | 18.16 | 4.88 | 16.11 | 4.79 | 9.38 | 4.59 | 9.86 | 9.67 | 9.86 | 12.79 | 9.86 | 23.54 | 0.1 | 16.21 | 0.1 | 17.58 | 0.1 | 14.26 | 4.59 | 6.84 | 5.08 | 4.79 | 4.3 | 4.79 | 24.41 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
57.184 | 0.49 | 0 | 1.95 | 0 | 14.45 | 0 | 18.16 | 4.79 | 16.5 | 4.79 | 9.96 | 4.49 | 9.77 | 9.67 | 9.86 | 7.03 | 9.86 | 23.93 | 0.1 | 16.9 | 0 | 18.16 | 0.1 | 15.04 | 4.49 | 8.79 | 4.98 | 4.79 | 5.76 | 4.88 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
57.216 | 0.39 | 0 | 1.66 | 0 | 14.45 | 0 | 18.16 | 4.79 | 17.09 | 4.79 | 10.55 | 4.59 | 10.55 | 9.57 | 9.86 | 8.11 | 9.86 | 24.32 | 0 | 17.58 | 0.1 | 18.85 | 0.1 | 15.63 | 4.49 | 10.35 | 4.88 | 4.79 | 6.25 | 4.88 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
57.248 | 0.39 | 0.1 | 1.56 | 0.1 | 14.26 | 0.1 | 18.16 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 11.23 | 9.57 | 9.86 | 11.04 | 9.86 | 21.09 | 0.1 | 18.26 | 0.1 | 19.53 | 0.1 | 16.41 | 4.59 | 11.72 | 4.88 | 4.79 | 5.86 | 4.88 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
57.28 | 0.29 | 0.1 | 1.46 | 0.1 | 4.49 | 0.1 | 18.16 | 4.88 | 18.07 | 4.79 | 11.82 | 4.59 | 12.6 | 9.57 | 9.86 | 11.23 | 9.96 | 21.39 | 0.1 | 18.85 | 0.1 | 20.02 | 0.1 | 16.99 | 4.59 | 12.7 | 4.79 | 4.79 | 5.47 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
57.312 | 0.29 | 0.1 | 1.17 | 0 | 1.76 | 0.1 | 17.09 | 4.88 | 18.46 | 4.79 | 12.4 | 4.59 | 13.57 | 9.67 | 9.86 | 10.35 | 9.96 | 21.49 | 0.1 | 19.43 | 0 | 20.7 | 0.1 | 17.68 | 4.59 | 13.57 | 4.79 | 4.79 | 5.27 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
57.344 | 0.2 | 0 | 1.17 | 0 | 1.86 | 0 | 4.39 | 4.88 | 18.85 | 4.79 | 12.79 | 4.59 | 14.36 | 9.67 | 9.86 | 9.86 | 9.86 | 21.58 | 0.1 | 20.12 | 0.1 | 21.29 | 0.1 | 18.46 | 4.59 | 14.26 | 4.79 | 4.79 | 5.08 | 4.79 | 20.02 | 9.77 | 24.22 | 9.77 | 24.12 | 9.77
57.376 | 0.2 | 0 | 0.98 | 0 | 1.86 | 0 | 3.03 | 4.88 | 19.34 | 4.79 | 13.48 | 4.69 | 14.94 | 9.57 | 9.86 | 9.96 | 9.86 | 21.49 | 0 | 20.61 | 0.1 | 21.88 | 0.1 | 18.85 | 4.59 | 15.04 | 4.59 | 4.79 | 5.18 | 4.88 | 7.91 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
57.408 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.61 | 4.88 | 19.82 | 4.88 | 13.97 | 4.69 | 15.53 | 9.67 | 9.86 | 10.06 | 9.96 | 21.49 | 0.1 | 21.29 | 0.1 | 22.36 | 0.1 | 19.53 | 4.69 | 15.82 | 4.59 | 4.79 | 5.18 | 4.88 | 6.45 | 9.77 | 24.22 | 9.86 | 11.23 | 9.86
57.44 | 0.1 | 0.1 | 0.88 | 0.1 | 1.37 | 0.1 | 4.59 | 4.88 | 20.12 | 4.88 | 14.55 | 4.69 | 16.02 | 9.67 | 9.86 | 10.16 | 9.96 | 21.49 | 0.1 | 21.88 | 0.1 | 17.97 | 0.1 | 20.22 | 4.69 | 16.5 | 4.59 | 4.79 | 5.08 | 4.88 | 10.35 | 9.77 | 11.23 | 9.86 | 5.76 | 9.86
57.472 | 0.1 | 0.39 | 0.68 | 0 | 1.17 | 0.1 | 5.66 | 4.88 | 20.61 | 4.88 | 15.04 | 4.69 | 16.5 | 9.67 | 10.55 | 9.86 | 9.96 | 21.49 | 0.1 | 22.36 | 0.1 | 16.7 | 0.1 | 20.8 | 4.69 | 17.19 | 4.59 | 4.79 | 4.98 | 4.79 | 12.21 | 9.77 | 5.66 | 9.86 | 9.08 | 9.86
57.504 | 0 | 1.37 | 0.68 | 0 | 1.07 | 0.1 | 5.76 | 4.79 | 21 | 4.79 | 15.43 | 4.69 | 16.99 | 9.67 | 11.33 | 9.86 | 9.96 | 21.49 | 0.1 | 18.26 | 0 | 17.38 | 0.1 | 21.49 | 4.79 | 17.77 | 4.59 | 4.79 | 4.98 | 4.79 | 10.64 | 9.77 | 9.28 | 9.77 | 12.31 | 9.77
57.536 | 0 | 2.25 | 0.59 | 0 | 1.07 | 0 | 5.47 | 4.88 | 18.56 | 4.79 | 15.92 | 4.79 | 17.58 | 9.67 | 12.6 | 9.96 | 9.86 | 21.49 | 0 | 17.29 | 0.1 | 17.58 | 0.1 | 21.97 | 4.69 | 18.46 | 4.49 | 4.79 | 4.88 | 4.88 | 9.96 | 9.77 | 12.4 | 9.77 | 11.52 | 9.86
57.568 | 0.1 | 3.71 | 0.49 | 0.1 | 0.88 | 0.1 | 5.18 | 4.88 | 18.95 | 4.88 | 16.41 | 4.79 | 17.87 | 9.67 | 13.67 | 9.86 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 22.46 | 4.79 | 19.14 | 4.59 | 5.27 | 4.79 | 4.88 | 10.16 | 9.77 | 11.52 | 9.86 | 10.25 | 9.86
57.6 | 0.1 | 4.49 | 0.49 | 0 | 0.78 | 0.1 | 4.98 | 4.88 | 18.46 | 4.88 | 16.99 | 4.79 | 18.46 | 9.67 | 14.45 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 18.46 | 4.79 | 19.82 | 4.59 | 6.05 | 4.79 | 4.88 | 10.35 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86
57.632 | 0.1 | 5.08 | 0.39 | 0.1 | 0.68 | 0.1 | 4.79 | 4.88 | 18.36 | 4.88 | 17.58 | 4.79 | 18.85 | 9.67 | 14.94 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 20.41 | 4.49 | 7.62 | 4.79 | 4.88 | 10.45 | 9.77 | 9.96 | 9.86 | 10.35 | 9.86
57.664 | 0.1 | 5.47 | 0.29 | 0 | 0.59 | 0 | 4.79 | 4.79 | 18.36 | 4.79 | 17.87 | 4.79 | 19.34 | 9.67 | 15.53 | 9.77 | 9.96 | 20.9 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.97 | 4.79 | 20.9 | 4.59 | 9.67 | 4.69 | 4.79 | 10.25 | 9.77 | 10.35 | 9.77 | 10.45 | 9.77
57.696 | 0.1 | 6.05 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.26 | 4.79 | 18.36 | 4.79 | 19.82 | 9.67 | 15.92 | 9.67 | 9.86 | 8.89 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 21.49 | 4.59 | 11.23 | 4.59 | 4.88 | 10.16 | 9.77 | 10.45 | 9.77 | 10.45 | 9.86
57.728 | 0.1 | 6.74 | 0.29 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 18.75 | 4.79 | 20.12 | 9.77 | 16.5 | 9.67 | 9.96 | 7.13 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 22.17 | 4.59 | 12.31 | 4.59 | 4.79 | 10.06 | 9.77 | 10.45 | 9.86 | 10.06 | 9.86
57.76 | 0.1 | 7.32 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 20.61 | 9.77 | 16.99 | 9.67 | 9.96 | 10.16 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 22.66 | 4.59 | 13.09 | 4.69 | 4.88 | 10.16 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
57.792 | 0.1 | 8.01 | 0.1 | 0 | 0.29 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.63 | 4.79 | 21 | 9.77 | 17.58 | 9.67 | 9.96 | 11.62 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.24 | 4.59 | 13.97 | 4.59 | 4.88 | 10.16 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
57.824 | 0 | 8.5 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.79 | 18.16 | 4.79 | 20.12 | 4.79 | 21.49 | 9.77 | 17.97 | 9.67 | 9.96 | 10.74 | 0 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 4.88 | 23.83 | 4.59 | 14.75 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 9.96 | 9.77
57.856 | 0.1 | 9.08 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.41 | 4.88 | 21.78 | 9.77 | 18.36 | 9.67 | 9.86 | 10.16 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 4.88 | 24.32 | 4.69 | 15.43 | 4.59 | 4.88 | 9.86 | 9.77 | 10.06 | 9.77 | 9.96 | 9.77
57.888 | 0.1 | 9.77 | 0.1 | 0.78 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 22.27 | 9.77 | 18.85 | 9.57 | 9.96 | 9.86 | 0.1 | 17.58 | 0.1 | 5.86 | 0.1 | 17.77 | 4.88 | 24.81 | 4.69 | 16.21 | 4.59 | 4.88 | 9.86 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86
57.92 | 0.1 | 10.35 | 0.1 | 1.46 | 0.2 | 0.68 | 4.59 | 4.88 | 18.07 | 4.88 | 17.48 | 4.88 | 22.66 | 9.77 | 19.24 | 9.57 | 9.96 | 10.16 | 0.1 | 16.8 | 0.1 | 1.66 | 0.1 | 17.68 | 4.88 | 25.29 | 4.69 | 16.9 | 4.59 | 4.88 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
57.952 | 0.1 | 10.94 | 0.1 | 2.73 | 0.1 | 1.37 | 4.59 | 4.88 | 12.21 | 4.88 | 18.07 | 4.88 | 23.05 | 9.77 | 19.73 | 9.47 | 9.96 | 10.16 | 0.1 | 4.1 | 0.1 | 2.25 | 0.1 | 17.68 | 4.88 | 25.88 | 4.79 | 17.58 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
57.984 | 0 | 11.43 | 0.1 | 3.81 | 0.1 | 2.64 | 4.49 | 4.79 | 3.81 | 4.79 | 17.87 | 4.79 | 23.34 | 9.77 | 20.12 | 9.57 | 9.86 | 10.16 | 0 | 1.95 | 0.1 | 1.95 | 0.1 | 17.58 | 4.88 | 26.27 | 4.69 | 18.16 | 4.59 | 4.88 | 9.77 | 9.77 | 9.96 | 9.86 | 9.77 | 9.77
58.016 | 0 | 12.01 | 0.1 | 4.59 | 0.1 | 3.81 | 4.49 | 4.88 | 3.13 | 4.88 | 18.07 | 4.88 | 23.73 | 9.77 | 20.61 | 9.57 | 10.35 | 9.86 | 0 | 2.25 | 0.1 | 1.76 | 0.1 | 14.84 | 4.88 | 23.15 | 4.69 | 18.75 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
58.048 | 0.1 | 12.6 | 0.1 | 5.08 | 0.1 | 4.49 | 4.49 | 4.88 | 3.91 | 4.88 | 18.07 | 4.88 | 24.12 | 9.77 | 21 | 9.57 | 11.04 | 9.86 | 0.1 | 1.95 | 0.1 | 1.66 | 0.1 | 3.71 | 4.88 | 22.27 | 4.79 | 19.43 | 4.69 | 5.86 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77
58.08 | 0.1 | 13.18 | 0.1 | 5.66 | 0.1 | 5.08 | 4.49 | 4.79 | 4.69 | 4.88 | 18.07 | 4.88 | 24.41 | 9.77 | 21.39 | 9.67 | 11.91 | 9.96 | 0.1 | 1.76 | 0.1 | 1.37 | 0.1 | 2.15 | 4.88 | 21.78 | 4.79 | 20.12 | 4.69 | 7.23 | 9.67 | 10.94 | 9.77 | 9.86 | 9.77 | 9.86
58.112 | 0.1 | 13.67 | 0.1 | 6.25 | 0.1 | 5.66 | 4.49 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 22.36 | 9.77 | 21.78 | 9.57 | 13.28 | 9.77 | 0.1 | 1.76 | 0.1 | 1.37 | 0.1 | 2.15 | 4.88 | 21.58 | 4.79 | 20.7 | 4.69 | 9.47 | 9.67 | 11.91 | 9.77 | 10.74 | 9.67 | 9.96
58.144 | 0 | 14.16 | 0.1 | 6.93 | 0.1 | 6.25 | 4.59 | 4.79 | 5.76 | 4.88 | 18.07 | 4.88 | 21.88 | 9.77 | 22.27 | 9.67 | 14.16 | 9.77 | 0 | 1.46 | 0 | 1.17 | 0 | 1.56 | 4.88 | 21.58 | 4.79 | 21.29 | 4.69 | 10.94 | 9.67 | 13.87 | 9.77 | 12.21 | 9.67 | 11.23
58.176 | 0.1 | 14.65 | 0.1 | 7.52 | 0.1 | 6.74 | 4.59 | 4.79 | 5.27 | 4.88 | 18.07 | 4.88 | 21.78 | 9.77 | 22.56 | 9.67 | 14.75 | 9.77 | 0 | 1.27 | 0.1 | 1.27 | 0.1 | 1.86 | 4.88 | 21.49 | 4.79 | 21.88 | 4.79 | 12.11 | 9.67 | 15.43 | 9.67 | 13.97 | 9.67 | 12.4
58.208 | 0.1 | 15.14 | 0.1 | 8.11 | 0.1 | 7.42 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.78 | 9.77 | 22.95 | 9.67 | 15.33 | 9.77 | 0.1 | 1.27 | 0.1 | 1.07 | 0.1 | 1.56 | 4.98 | 21.39 | 4.88 | 22.46 | 4.79 | 12.99 | 9.67 | 16.5 | 9.67 | 15.63 | 9.67 | 14.55
58.24 | 0.1 | 15.72 | 0.1 | 8.69 | 0.1 | 8.01 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.68 | 9.86 | 23.34 | 9.67 | 15.82 | 9.77 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.46 | 4.98 | 21.39 | 4.79 | 22.95 | 4.79 | 13.87 | 9.67 | 17.38 | 9.67 | 16.6 | 9.67 | 15.92
58.272 | 0.1 | 16.21 | 0.1 | 9.38 | 0.1 | 8.69 | 4.69 | 4.79 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.77 | 23.73 | 9.67 | 16.41 | 9.67 | 0.1 | 1.17 | 0.1 | 0.88 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 23.54 | 4.79 | 14.75 | 9.67 | 18.16 | 9.67 | 17.58 | 9.67 | 16.99
58.304 | 0 | 16.7 | 0.1 | 9.86 | 0.1 | 9.38 | 4.69 | 5.57 | 4.88 | 4.88 | 17.97 | 4.79 | 21.49 | 9.77 | 24.02 | 9.67 | 16.9 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 23.93 | 4.79 | 15.43 | 9.57 | 18.85 | 9.67 | 18.36 | 9.67 | 17.77
58.336 | 0.1 | 17.19 | 0.1 | 10.45 | 0.1 | 9.77 | 4.69 | 6.25 | 4.88 | 4.88 | 17.87 | 4.88 | 21.49 | 9.77 | 24.32 | 9.67 | 17.29 | 9.67 | 0 | 0.88 | 0.1 | 0.78 | 0.1 | 0.98 | 4.88 | 21.29 | 4.79 | 24.51 | 4.79 | 16.11 | 9.57 | 19.63 | 9.67 | 18.95 | 9.67 | 18.46
58.368 | 0.1 | 13.97 | 0.1 | 11.13 | 0.1 | 10.45 | 4.69 | 7.42 | 4.79 | 4.88 | 5.47 | 4.88 | 21.49 | 9.86 | 22.27 | 9.67 | 17.77 | 9.67 | 0.78 | 0.78 | 0.1 | 0.59 | 0.1 | 1.07 | 4.98 | 21.19 | 4.79 | 25.1 | 4.88 | 16.8 | 9.67 | 20.12 | 9.67 | 19.73 | 9.67 | 19.04
58.4 | 0.1 | 14.36 | 0.1 | 11.62 | 0.1 | 11.04 | 4.79 | 8.59 | 4.79 | 4.88 | 2.93 | 4.88 | 21.49 | 9.86 | 22.07 | 9.67 | 18.26 | 9.67 | 1.76 | 0.68 | 0.1 | 0.59 | 0.1 | 0.88 | 4.98 | 21.19 | 4.88 | 25.59 | 4.79 | 17.48 | 9.67 | 20.8 | 9.67 | 20.22 | 9.67 | 19.82
58.432 | 0.1 | 14.45 | 0.1 | 12.21 | 0.1 | 11.62 | 4.79 | 9.47 | 4.69 | 4.88 | 3.61 | 4.88 | 21.49 | 9.86 | 21.78 | 9.67 | 18.65 | 9.57 | 3.22 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 11.62 | 4.88 | 26.08 | 4.79 | 18.16 | 9.67 | 21.39 | 9.67 | 20.9 | 9.67 | 20.41
58.464 | 0.1 | 14.55 | 0 | 12.79 | 0 | 12.11 | 4.79 | 10.16 | 4.59 | 4.88 | 4.59 | 4.88 | 21.39 | 9.77 | 21.58 | 9.67 | 19.24 | 9.57 | 5.18 | 0.49 | 0.1 | 0.49 | 0.1 | 0.68 | 4.88 | 3.03 | 4.79 | 21.49 | 4.79 | 18.65 | 9.67 | 21.97 | 9.67 | 21.58 | 9.67 | 20.9
58.496 | 0.1 | 14.45 | 0.1 | 13.38 | 0.1 | 12.7 | 4.79 | 10.74 | 4.59 | 4.88 | 5.66 | 4.88 | 21.39 | 9.77 | 21.68 | 9.77 | 19.53 | 9.57 | 6.25 | 0.49 | 0.1 | 0.39 | 0.1 | 0.59 | 4.88 | 3.13 | 4.88 | 20.22 | 4.88 | 19.34 | 9.67 | 22.56 | 9.67 | 22.07 | 9.67 | 21.58
58.528 | 0.1 | 14.45 | 0.1 | 13.87 | 0.1 | 13.28 | 4.79 | 11.33 | 4.59 | 4.88 | 5.76 | 4.88 | 13.28 | 9.86 | 21.58 | 9.77 | 20.02 | 9.57 | 7.23 | 0.39 | 0.2 | 0.39 | 0.1 | 0.59 | 4.98 | 3.52 | 4.88 | 21 | 4.88 | 20.02 | 9.67 | 23.15 | 9.67 | 22.75 | 9.67 | 22.17
58.56 | 0.1 | 14.55 | 0.1 | 14.36 | 0.1 | 13.77 | 4.88 | 12.01 | 4.59 | 4.88 | 5.47 | 4.88 | 7.03 | 9.77 | 21.49 | 9.77 | 20.51 | 9.57 | 8.01 | 0.39 | 1.56 | 0.29 | 0.1 | 0.49 | 4.88 | 4.59 | 4.88 | 21.09 | 4.88 | 20.61 | 9.67 | 23.54 | 9.67 | 23.24 | 9.67 | 22.75
58.592 | 0.1 | 14.45 | 0.1 | 14.94 | 0.1 | 14.36 | 4.88 | 12.6 | 4.49 | 4.88 | 5.08 | 4.88 | 8.3 | 9.77 | 21.49 | 9.77 | 20.9 | 9.57 | 8.69 | 0.29 | 2.93 | 0.2 | 1.17 | 0.49 | 4.98 | 5.96 | 4.88 | 21.19 | 4.88 | 21.19 | 9.67 | 24.12 | 9.67 | 23.63 | 9.67 | 23.24
58.624 | 0.1 | 14.45 | 0 | 15.43 | 0.1 | 14.75 | 4.79 | 13.09 | 4.49 | 4.88 | 4.98 | 4.79 | 11.04 | 9.77 | 21.39 | 9.77 | 21.29 | 9.57 | 9.57 | 0.29 | 4.69 | 0.2 | 2.34 | 0.29 | 4.88 | 6.05 | 4.88 | 21.19 | 4.88 | 21.68 | 9.67 | 24.61 | 9.67 | 24.32 | 9.67 | 23.63
58.656 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 15.33 | 4.88 | 13.57 | 4.49 | 4.88 | 4.88 | 4.88 | 11.23 | 9.77 | 21.39 | 9.77 | 21.58 | 9.67 | 10.35 | 0.2 | 6.15 | 0.2 | 4.2 | 0.29 | 4.88 | 5.66 | 4.79 | 21.19 | 4.88 | 22.27 | 9.77 | 25.2 | 9.67 | 24.71 | 9.67 | 24.32
58.688 | 0.1 | 14.36 | 0.1 | 16.41 | 0.1 | 15.82 | 4.88 | 14.16 | 4.49 | 5.27 | 4.88 | 4.88 | 10.45 | 9.86 | 21.39 | 9.77 | 22.07 | 9.67 | 11.04 | 0.2 | 7.13 | 0.2 | 5.57 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.88 | 22.95 | 9.77 | 25.68 | 9.77 | 25.2 | 9.77 | 24.81
58.72 | 0.1 | 14.45 | 0.1 | 16.9 | 0.1 | 16.31 | 4.88 | 14.65 | 4.59 | 6.05 | 4.88 | 4.88 | 9.86 | 9.86 | 21.39 | 9.77 | 22.46 | 9.67 | 11.72 | 0.2 | 8.01 | 0.2 | 6.74 | 0.29 | 4.88 | 5.18 | 4.88 | 21.19 | 4.88 | 23.44 | 9.77 | 26.17 | 9.77 | 25.68 | 9.77 | 25.29
58.752 | 0.1 | 14.36 | 0.1 | 17.38 | 0.1 | 16.8 | 4.88 | 15.14 | 4.49 | 7.03 | 4.79 | 4.88 | 10.06 | 9.77 | 21.39 | 9.77 | 22.85 | 9.67 | 12.6 | 0.2 | 8.69 | 0.1 | 7.62 | 0.2 | 4.88 | 5.18 | 4.88 | 21.29 | 4.88 | 24.02 | 9.77 | 26.56 | 9.77 | 26.17 | 9.77 | 25.78
58.784 | 0 | 12.7 | 0.1 | 15.23 | 0 | 17.29 | 4.88 | 15.72 | 4.59 | 8.5 | 4.69 | 4.79 | 10.25 | 9.77 | 12.89 | 9.77 | 23.15 | 9.67 | 13.28 | 0.1 | 9.38 | 0.1 | 8.4 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.88 | 24.61 | 9.77 | 27.25 | 9.77 | 26.66 | 9.77 | 26.27
58.816 | 0.1 | 3.81 | 0 | 15.14 | 0.1 | 15.23 | 4.88 | 16.11 | 4.59 | 9.38 | 4.69 | 4.88 | 10.16 | 9.77 | 6.93 | 9.86 | 23.54 | 9.67 | 13.97 | 0.1 | 10.25 | 0.1 | 9.18 | 0.1 | 4.88 | 5.08 | 4.79 | 21.19 | 4.88 | 25 | 9.77 | 27.54 | 9.77 | 27.25 | 9.77 | 26.76
58.848 | 0.1 | 1.76 | 0.1 | 14.84 | 0.1 | 15.23 | 4.88 | 16.6 | 4.59 | 9.96 | 4.59 | 4.88 | 9.96 | 9.86 | 8.2 | 9.86 | 23.93 | 9.67 | 14.65 | 0.1 | 11.04 | 0.1 | 9.96 | 0.1 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.49 | 9.77 | 28.13 | 9.77 | 27.64 | 9.77 | 27.34
58.88 | 0.1 | 1.95 | 0.1 | 14.65 | 0.1 | 14.84 | 4.88 | 17.09 | 4.69 | 10.64 | 4.59 | 4.88 | 9.96 | 9.86 | 10.94 | 9.86 | 24.41 | 9.67 | 15.53 | 0.1 | 11.72 | 0.1 | 10.74 | 0.1 | 4.88 | 4.88 | 4.88 | 21.09 | 4.88 | 26.08 | 9.77 | 28.71 | 9.77 | 28.22 | 9.77 | 27.74
58.912 | 0.1 | 1.56 | 0.1 | 14.65 | 0.1 | 14.65 | 4.88 | 17.58 | 4.59 | 11.23 | 4.59 | 4.88 | 9.86 | 9.77 | 11.33 | 9.77 | 20.9 | 9.67 | 16.11 | 0.1 | 12.5 | 0 | 11.62 | 0.1 | 4.88 | 4.88 | 4.88 | 5.76 | 4.88 | 21.88 | 9.77 | 29.1 | 9.86 | 28.71 | 9.77 | 28.22
58.944 | 0.1 | 1.56 | 0.1 | 14.45 | 0.1 | 14.75 | 4.88 | 17.97 | 4.69 | 11.91 | 4.49 | 4.79 | 9.86 | 9.86 | 10.35 | 9.86 | 21.49 | 9.67 | 16.8 | 0 | 13.28 | 0.1 | 12.31 | 0 | 4.88 | 4.79 | 4.88 | 2.64 | 4.88 | 20.7 | 9.77 | 28.91 | 9.86 | 29.2 | 9.77 | 28.71
58.976 | 0.1 | 1.27 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 18.46 | 4.69 | 12.4 | 4.59 | 4.88 | 9.77 | 9.77 | 9.77 | 9.86 | 21.29 | 9.67 | 17.48 | 0 | 14.06 | 0.1 | 12.99 | 0 | 4.88 | 4.69 | 4.79 | 3.42 | 4.88 | 21.29 | 9.77 | 25.88 | 9.77 | 28.32 | 9.86 | 29.1
59.008 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 4.98 | 18.95 | 4.79 | 12.99 | 4.49 | 4.88 | 9.77 | 9.86 | 9.96 | 9.86 | 21.49 | 9.67 | 18.07 | 0.1 | 14.84 | 0.1 | 13.77 | 0 | 4.88 | 4.69 | 4.88 | 4.3 | 4.88 | 21.29 | 9.77 | 25.1 | 9.86 | 24.71 | 9.86 | 24.51
59.04 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.34 | 4.79 | 13.48 | 4.49 | 4.88 | 9.77 | 9.77 | 10.25 | 9.86 | 21.49 | 9.77 | 18.75 | 0.1 | 15.43 | 0.1 | 14.55 | 0 | 6.05 | 4.69 | 4.88 | 5.47 | 4.88 | 21.29 | 9.77 | 24.71 | 9.86 | 24.51 | 9.86 | 23.24
59.072 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.82 | 4.69 | 14.06 | 4.49 | 5.18 | 9.77 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 19.34 | 0.1 | 16.21 | 0 | 15.23 | 0 | 7.03 | 4.59 | 4.88 | 6.15 | 4.88 | 21.19 | 9.77 | 24.61 | 9.86 | 24.51 | 9.86 | 23.93
59.104 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.22 | 4.79 | 14.55 | 4.49 | 6.05 | 9.67 | 9.77 | 9.96 | 9.86 | 21.58 | 9.77 | 20.12 | 0 | 16.9 | 0.1 | 16.02 | 0 | 9.08 | 4.59 | 4.88 | 5.76 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.32 | 9.86 | 24.02
59.136 | 0.1 | 0.68 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 20.61 | 4.79 | 15.04 | 4.49 | 7.03 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.61 | 0 | 17.58 | 0.1 | 16.7 | 0 | 10.55 | 4.59 | 4.79 | 5.47 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.32 | 9.86 | 24.12
59.168 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 21 | 4.79 | 15.53 | 4.59 | 8.4 | 9.67 | 10.55 | 9.96 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 18.16 | 0.1 | 17.29 | 0.1 | 11.91 | 4.59 | 4.88 | 5.18 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.32 | 9.86 | 24.12
59.2 | 0.1 | 0.59 | 0.1 | 8.59 | 0.1 | 14.36 | 4.88 | 19.14 | 4.88 | 16.11 | 4.59 | 9.28 | 9.67 | 11.23 | 9.86 | 9.86 | 21.49 | 9.77 | 21.68 | 0.1 | 18.85 | 0.1 | 17.97 | 0.1 | 12.89 | 4.59 | 4.88 | 4.98 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.12
59.232 | 0.1 | 0.59 | 0.1 | 2.44 | 0.1 | 10.16 | 4.88 | 18.95 | 4.79 | 16.5 | 4.59 | 9.96 | 9.67 | 12.31 | 9.86 | 9.86 | 21.49 | 9.77 | 22.36 | 0.1 | 19.53 | 0 | 18.56 | 0 | 13.77 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
59.264 | 0.1 | 0.49 | 0 | 1.95 | 0 | 2.54 | 4.88 | 18.46 | 4.79 | 16.99 | 4.59 | 10.55 | 9.67 | 13.57 | 9.67 | 9.86 | 21.49 | 9.77 | 18.85 | 0 | 20.12 | 0 | 19.24 | 0 | 14.55 | 4.49 | 4.88 | 5.08 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
59.296 | 0.1 | 0.39 | 0.1 | 1.76 | 0.1 | 1.95 | 4.88 | 18.36 | 4.88 | 17.38 | 4.69 | 11.23 | 9.57 | 14.26 | 9.77 | 9.86 | 21.39 | 9.77 | 17.97 | 0 | 20.8 | 0.1 | 19.82 | 0 | 15.23 | 4.49 | 4.88 | 4.98 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.12
59.328 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.76 | 4.88 | 18.46 | 4.88 | 17.97 | 4.69 | 11.82 | 9.67 | 14.84 | 9.77 | 9.86 | 21.19 | 9.77 | 17.68 | 0.1 | 21.29 | 0.1 | 20.41 | 0.1 | 15.92 | 4.59 | 4.88 | 4.88 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
59.36 | 0.1 | 0.29 | 0.1 | 1.46 | 0.1 | 1.37 | 4.88 | 18.26 | 4.88 | 18.46 | 4.69 | 12.31 | 9.67 | 15.43 | 9.67 | 9.86 | 9.38 | 9.77 | 17.58 | 0.1 | 21.88 | 0.1 | 21 | 0 | 16.7 | 4.59 | 4.88 | 4.88 | 4.88 | 7.32 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.12
59.392 | 0.49 | 0.29 | 0.1 | 1.07 | 0.1 | 1.56 | 4.88 | 18.16 | 4.88 | 18.85 | 4.69 | 12.89 | 9.67 | 15.92 | 9.67 | 9.86 | 7.03 | 9.77 | 17.68 | 0.1 | 22.46 | 0 | 21.68 | 0.1 | 17.38 | 4.59 | 4.88 | 4.79 | 4.88 | 2.44 | 9.77 | 13.57 | 9.86 | 24.22 | 9.86 | 24.12
59.424 | 1.27 | 0.2 | 0 | 1.07 | 0 | 1.07 | 4.88 | 18.16 | 4.88 | 19.34 | 4.69 | 13.48 | 9.57 | 16.5 | 9.67 | 9.86 | 9.96 | 9.77 | 17.58 | 0 | 17.87 | 0.1 | 22.27 | 0 | 17.97 | 4.59 | 5.08 | 4.88 | 4.88 | 3.32 | 9.77 | 5.86 | 9.86 | 11.62 | 9.86 | 24.02
59.456 | 2.25 | 0.1 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.73 | 4.79 | 13.97 | 9.67 | 16.9 | 9.67 | 9.86 | 11.62 | 9.77 | 17.58 | 0 | 16.7 | 0.1 | 21.58 | 0 | 18.65 | 4.59 | 5.96 | 4.69 | 4.88 | 4.2 | 9.77 | 7.52 | 9.86 | 5.76 | 9.86 | 10.35
59.488 | 3.71 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 14.45 | 9.67 | 17.38 | 9.67 | 9.86 | 10.84 | 9.86 | 17.58 | 0.1 | 17.38 | 0.1 | 18.65 | 0.1 | 19.24 | 4.59 | 7.42 | 4.69 | 4.88 | 5.27 | 9.77 | 11.52 | 9.86 | 8.01 | 9.86 | 5.96
59.52 | 4.49 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 14.94 | 9.67 | 17.87 | 9.67 | 9.86 | 10.16 | 9.86 | 17.48 | 0 | 17.58 | 0.1 | 18.26 | 0 | 19.92 | 4.69 | 9.67 | 4.59 | 4.88 | 6.15 | 9.77 | 11.62 | 9.86 | 11.82 | 9.86 | 9.08
59.552 | 5.08 | 0.1 | 0.1 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.79 | 20.9 | 4.79 | 15.53 | 9.67 | 18.36 | 9.67 | 9.86 | 9.86 | 9.86 | 17.48 | 0.1 | 17.58 | 0 | 18.07 | 0 | 20.61 | 4.69 | 11.23 | 4.59 | 4.88 | 5.96 | 9.77 | 10.45 | 9.86 | 11.72 | 9.86 | 12.31
59.584 | 5.47 | 0 | 0.1 | 0.59 | 0 | 0.68 | 4.88 | 18.07 | 4.88 | 18.26 | 4.79 | 16.02 | 9.67 | 18.85 | 9.57 | 9.86 | 10.16 | 9.86 | 17.38 | 0 | 17.58 | 0 | 17.87 | 0 | 21.19 | 4.69 | 12.31 | 4.59 | 4.88 | 5.47 | 9.77 | 9.77 | 9.86 | 10.25 | 9.86 | 11.13
59.616 | 6.05 | 0 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 10.74 | 4.88 | 18.07 | 4.88 | 16.5 | 9.67 | 19.24 | 9.57 | 9.86 | 10.25 | 9.86 | 17.38 | 0 | 17.68 | 0.1 | 17.87 | 0 | 21.78 | 4.69 | 13.09 | 4.59 | 4.88 | 5.27 | 9.77 | 10.16 | 9.86 | 9.86 | 9.86 | 10.35
59.648 | 6.74 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 3.32 | 4.88 | 18.07 | 4.88 | 16.99 | 9.77 | 19.73 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 17.87 | 0 | 22.27 | 4.69 | 13.97 | 4.49 | 4.88 | 5.27 | 9.77 | 10.45 | 9.86 | 10.16 | 9.86 | 9.86
59.68 | 7.32 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 18.16 | 4.88 | 17.38 | 9.77 | 20.12 | 9.67 | 9.86 | 9.96 | 9.86 | 14.06 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 22.95 | 4.79 | 14.65 | 4.49 | 4.88 | 5.18 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86 | 10.35
59.712 | 7.91 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 3.71 | 4.88 | 18.16 | 4.79 | 17.77 | 9.77 | 20.51 | 9.67 | 9.86 | 9.96 | 9.86 | 3.71 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 23.44 | 4.79 | 15.43 | 4.49 | 5.08 | 5.08 | 9.77 | 10.25 | 9.86 | 10.45 | 9.86 | 10.45
59.744 | 8.5 | 0.1 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 4.88 | 4.88 | 18.16 | 4.79 | 18.26 | 9.77 | 21 | 9.57 | 9.86 | 9.96 | 9.86 | 2.15 | 0 | 17.58 | 0 | 17.68 | 0 | 24.02 | 4.79 | 16.11 | 4.49 | 6.25 | 5.08 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86 | 10.25
59.776 | 9.18 | 0.1 | 0.1 | 0.2 | 0.1 | 0.29 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 18.75 | 9.77 | 21.19 | 9.67 | 9.86 | 9.86 | 9.86 | 2.05 | 0 | 17.68 | 0.1 | 17.68 | 0 | 24.51 | 4.79 | 16.8 | 4.49 | 7.72 | 4.98 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86 | 9.96
59.808 | 9.77 | 0.1 | 1.07 | 0.2 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 19.24 | 9.77 | 21.68 | 9.67 | 9.86 | 9.77 | 9.86 | 1.86 | 0.1 | 17.68 | 0.1 | 17.68 | 0 | 25 | 4.79 | 17.58 | 4.49 | 9.77 | 4.88 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.16
59.84 | 10.35 | 0.1 | 1.86 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.63 | 9.77 | 22.17 | 9.67 | 10.55 | 9.77 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 25.39 | 4.79 | 18.16 | 4.49 | 11.33 | 4.88 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86 | 10.16
59.872 | 10.94 | 0.1 | 3.13 | 0.1 | 0.1 | 0.2 | 4.88 | 4.88 | 4.88 | 18.16 | 4.88 | 20.02 | 9.77 | 22.56 | 9.67 | 11.23 | 9.77 | 9.86 | 1.56 | 0.1 | 4.98 | 0 | 17.68 | 0 | 26.08 | 4.79 | 18.75 | 4.59 | 12.4 | 4.88 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
59.904 | 11.43 | 0.1 | 4.1 | 0.1 | 0 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.51 | 9.77 | 22.95 | 9.67 | 12.5 | 9.77 | 9.86 | 1.27 | 0 | 1.76 | 0.1 | 8.01 | 0 | 25.39 | 4.79 | 19.43 | 4.59 | 13.18 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86
59.936 | 12.01 | 0 | 4.69 | 0 | 0.2 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.8 | 9.77 | 23.24 | 9.67 | 13.57 | 9.67 | 9.86 | 1.46 | 0 | 2.25 | 0.1 | 1.95 | 0 | 22.85 | 4.79 | 20.12 | 4.59 | 14.06 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
59.968 | 12.6 | 0.1 | 5.37 | 0 | 1.07 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.68 | 9.77 | 23.63 | 9.67 | 14.36 | 9.67 | 9.86 | 1.37 | 0.1 | 1.95 | 0.1 | 2.44 | 0.1 | 22.17 | 4.79 | 20.8 | 4.59 | 14.75 | 4.69 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
60 | 13.18 | 0.1 | 5.86 | 0 | 2.15 | 0.1 | 5.08 | 4.79 | 4.88 | 17.19 | 4.88 | 18.07 | 9.86 | 24.12 | 9.77 | 14.94 | 9.67 | 9.86 | 1.07 | 0.1 | 1.76 | 0.1 | 2.05 | 0 | 21.78 | 4.79 | 21.19 | 4.59 | 15.53 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
60.032 | 13.67 | 0.1 | 6.54 | 0 | 3.71 | 0 | 5.86 | 4.69 | 4.88 | 4.59 | 4.88 | 18.07 | 9.86 | 24.41 | 9.77 | 15.43 | 9.67 | 9.86 | 1.07 | 0.1 | 1.66 | 0 | 1.86 | 0 | 21.68 | 4.79 | 21.88 | 4.69 | 16.21 | 4.59 | 10.16 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
60.064 | 14.26 | 0 | 7.13 | 0 | 4.39 | 0 | 6.93 | 4.69 | 4.88 | 2.93 | 4.88 | 18.07 | 9.77 | 22.36 | 9.77 | 15.92 | 9.67 | 9.86 | 0.88 | 0 | 1.56 | 0.1 | 1.66 | 0 | 21.58 | 4.79 | 22.46 | 4.69 | 16.9 | 4.59 | 11.33 | 9.67 | 9.86 | 9.77 | 9.96 | 9.77
60.096 | 14.84 | 0.1 | 7.62 | 0 | 4.98 | 0 | 8.4 | 4.69 | 4.88 | 3.71 | 4.88 | 18.07 | 9.86 | 22.27 | 9.77 | 16.5 | 9.57 | 9.96 | 0.78 | 0 | 1.37 | 0.1 | 1.56 | 0 | 21.49 | 4.79 | 23.05 | 4.69 | 17.58 | 4.59 | 12.7 | 9.67 | 9.96 | 9.67 | 10.94 | 9.77
60.128 | 15.23 | 0.1 | 8.3 | 0.1 | 5.47 | 0.1 | 9.28 | 4.59 | 4.88 | 4.79 | 4.88 | 18.07 | 9.86 | 21.88 | 9.77 | 16.99 | 9.67 | 9.86 | 0.88 | 0.1 | 1.37 | 0.1 | 1.56 | 0 | 21.49 | 4.88 | 23.44 | 4.69 | 18.26 | 4.59 | 14.84 | 9.67 | 10.94 | 9.77 | 12.21 | 9.77
60.16 | 15.72 | 0.1 | 8.89 | 0.1 | 6.05 | 0.1 | 9.86 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.58 | 9.67 | 9.96 | 0.68 | 0.1 | 1.07 | 0.1 | 1.27 | 0 | 21.49 | 4.88 | 24.12 | 4.69 | 18.85 | 4.59 | 16.31 | 9.67 | 12.21 | 9.67 | 14.16 | 9.67
60.192 | 16.21 | 0.1 | 9.57 | 0.1 | 6.64 | 0.1 | 10.55 | 4.59 | 4.88 | 5.76 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 17.87 | 9.67 | 9.86 | 0.68 | 0.1 | 1.07 | 0 | 1.27 | 0 | 21.39 | 4.88 | 24.61 | 4.79 | 19.53 | 4.49 | 16.99 | 9.67 | 14.26 | 9.67 | 15.82 | 9.67
60.224 | 16.8 | 0.1 | 10.16 | 0 | 7.23 | 0 | 11.23 | 4.49 | 4.88 | 5.37 | 4.88 | 18.07 | 9.77 | 21.58 | 9.77 | 18.46 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 0 | 21.39 | 4.79 | 25.1 | 4.79 | 20.12 | 4.59 | 17.87 | 9.57 | 15.92 | 9.67 | 16.8 | 9.67
60.256 | 17.19 | 0 | 10.64 | 0 | 7.81 | 0 | 11.82 | 4.49 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.85 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 0 | 21.39 | 4.79 | 25.49 | 4.79 | 20.8 | 4.49 | 18.65 | 9.67 | 16.9 | 9.67 | 17.58 | 9.67
60.288 | 14.06 | 0.1 | 11.33 | 0.1 | 8.5 | 0.1 | 12.31 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.24 | 9.67 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 21.39 | 4.88 | 26.08 | 4.79 | 21.19 | 4.59 | 19.24 | 9.67 | 17.77 | 9.67 | 18.46 | 9.67
60.32 | 14.45 | 0.1 | 11.82 | 0 | 8.98 | 0.1 | 12.89 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.73 | 9.67 | 9.96 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 21.78 | 4.79 | 21.88 | 4.59 | 19.92 | 9.67 | 18.46 | 9.67 | 19.14 | 9.67
60.352 | 14.45 | 0.1 | 12.4 | 0.1 | 9.67 | 0.1 | 13.48 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.49 | 9.86 | 20.12 | 9.67 | 10.16 | 0.29 | 0.1 | 0.68 | 0 | 0.78 | 0 | 9.86 | 4.79 | 20.22 | 4.79 | 22.46 | 4.59 | 20.61 | 9.67 | 19.24 | 9.67 | 19.82 | 9.57
60.384 | 14.55 | 0 | 12.99 | 0 | 10.25 | 0 | 13.97 | 4.49 | 4.88 | 4.88 | 4.88 | 14.06 | 9.77 | 21.39 | 9.77 | 20.51 | 9.67 | 10.94 | 0.29 | 0 | 0.59 | 0 | 0.59 | 0 | 2.54 | 4.79 | 20.9 | 4.79 | 23.15 | 4.59 | 21.19 | 9.57 | 19.82 | 9.67 | 20.31 | 9.57
60.416 | 14.45 | 0.1 | 13.38 | 0 | 10.94 | 0 | 14.45 | 4.49 | 4.88 | 4.79 | 4.88 | 4.59 | 9.77 | 21.39 | 9.77 | 20.9 | 9.67 | 11.91 | 0.2 | 0 | 0.59 | 0.1 | 0.59 | 0 | 3.32 | 4.88 | 21.19 | 4.79 | 23.54 | 4.69 | 21.68 | 9.67 | 20.41 | 9.67 | 21 | 9.67
60.448 | 14.45 | 0.1 | 13.97 | 0.1 | 11.43 | 0.1 | 15.04 | 4.49 | 4.88 | 4.69 | 4.88 | 2.93 | 9.86 | 21.39 | 9.86 | 21.39 | 9.67 | 13.28 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0 | 3.61 | 4.88 | 21.19 | 4.79 | 24.12 | 4.69 | 22.36 | 9.67 | 20.9 | 9.67 | 21.58 | 9.67
60.48 | 14.55 | 0.1 | 14.45 | 0.1 | 12.01 | 0.1 | 15.53 | 4.59 | 4.88 | 4.69 | 4.88 | 3.71 | 9.86 | 14.65 | 9.86 | 21.78 | 9.67 | 14.06 | 0.2 | 0 | 0.49 | 0.1 | 0.49 | 0.1 | 4.59 | 4.88 | 21.19 | 4.79 | 24.61 | 4.69 | 22.85 | 9.67 | 21.58 | 9.67 | 22.17 | 9.67
60.512 | 14.45 | 0.1 | 15.04 | 0 | 12.5 | 0.1 | 16.02 | 4.59 | 4.88 | 4.69 | 4.88 | 4.69 | 9.86 | 7.32 | 9.86 | 22.17 | 9.67 | 14.75 | 0.2 | 0 | 0.29 | 0.1 | 0.39 | 0 | 6.05 | 4.88 | 21.29 | 4.79 | 25.1 | 4.69 | 23.34 | 9.67 | 22.17 | 9.67 | 22.75 | 9.67
60.544 | 14.45 | 0 | 15.53 | 0 | 13.09 | 0 | 16.5 | 4.59 | 4.88 | 4.59 | 4.88 | 5.86 | 9.77 | 8.2 | 9.86 | 22.46 | 9.67 | 15.33 | 0.1 | 0 | 0.29 | 0.2 | 0.29 | 0 | 6.05 | 4.79 | 21.29 | 4.79 | 25.59 | 4.69 | 24.02 | 9.67 | 22.85 | 9.67 | 23.24 | 9.57
60.576 | 14.45 | 0 | 16.02 | 0 | 13.67 | 0 | 16.9 | 4.59 | 4.88 | 4.59 | 4.88 | 5.76 | 9.77 | 10.94 | 9.86 | 22.95 | 9.67 | 15.82 | 0.1 | 0 | 0.29 | 1.17 | 0.29 | 0 | 5.66 | 4.88 | 21.19 | 4.79 | 26.17 | 4.79 | 24.41 | 9.67 | 23.24 | 9.67 | 23.73 | 9.67
60.608 | 14.45 | 0.1 | 16.5 | 0 | 14.16 | 0.1 | 17.38 | 4.69 | 4.88 | 4.59 | 4.88 | 5.37 | 9.86 | 11.23 | 9.86 | 23.34 | 9.77 | 16.31 | 0.1 | 0.29 | 0.29 | 2.64 | 0.29 | 0 | 5.37 | 4.88 | 21.29 | 4.69 | 21.49 | 4.79 | 24.9 | 9.67 | 23.83 | 9.67 | 24.32 | 9.67
60.64 | 14.45 | 0.1 | 16.99 | 0 | 14.65 | 0.1 | 17.87 | 4.69 | 5.66 | 4.59 | 4.88 | 5.18 | 9.86 | 10.45 | 9.86 | 23.63 | 9.77 | 16.8 | 0.1 | 1.56 | 0.2 | 4.39 | 0.29 | 0 | 5.18 | 4.88 | 21.19 | 4.79 | 20.22 | 4.79 | 25.49 | 9.67 | 24.32 | 9.67 | 24.81 | 9.67
60.672 | 14.45 | 0.1 | 16.02 | 0.1 | 15.23 | 0.1 | 18.26 | 4.69 | 6.15 | 4.59 | 4.88 | 4.88 | 9.86 | 9.86 | 9.86 | 24.02 | 9.77 | 17.29 | 0.1 | 2.83 | 0.2 | 6.05 | 0.2 | 0 | 5.08 | 4.88 | 21.19 | 4.79 | 21 | 4.79 | 25.98 | 9.67 | 24.81 | 9.67 | 25.29 | 9.67
60.704 | 12.7 | 0 | 14.06 | 0 | 15.63 | 0.1 | 18.75 | 4.69 | 7.13 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 24.41 | 9.77 | 17.68 | 0.1 | 4.98 | 0.1 | 6.93 | 0.2 | 0.59 | 5.08 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 26.47 | 9.67 | 25.29 | 9.67 | 25.88 | 9.67
60.736 | 3.91 | 0 | 14.55 | 0 | 16.11 | 0 | 19.24 | 4.69 | 8.5 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 22.46 | 9.77 | 18.16 | 0 | 6.15 | 0.2 | 7.72 | 0.1 | 1.56 | 5.08 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 26.95 | 9.67 | 25.88 | 9.67 | 26.37 | 9.67
60.768 | 1.86 | 0.1 | 14.65 | 0.1 | 16.7 | 0.1 | 19.73 | 4.79 | 9.47 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 22.17 | 9.77 | 18.65 | 0.1 | 7.03 | 0.1 | 8.5 | 0.1 | 3.03 | 5.08 | 4.88 | 6.64 | 4.79 | 21.19 | 4.79 | 27.44 | 9.77 | 26.27 | 9.77 | 26.76 | 9.67
60.8 | 1.86 | 0.1 | 14.45 | 0.1 | 17.09 | 0.1 | 20.12 | 4.79 | 10.16 | 4.59 | 4.88 | 4.79 | 9.86 | 10.16 | 9.86 | 21.88 | 9.86 | 19.14 | 0.1 | 7.81 | 0.1 | 9.38 | 0.1 | 5.08 | 4.98 | 4.88 | 2.34 | 4.79 | 21.29 | 4.79 | 27.83 | 9.77 | 26.86 | 9.77 | 27.25 | 9.77
60.832 | 1.56 | 0.1 | 14.55 | 0 | 13.97 | 0.1 | 20.51 | 4.79 | 10.74 | 4.49 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 19.53 | 0.1 | 8.59 | 0.1 | 10.16 | 0.1 | 6.45 | 4.88 | 4.88 | 3.32 | 4.79 | 21.29 | 4.79 | 28.32 | 9.77 | 27.25 | 9.77 | 27.83 | 9.77
60.864 | 1.37 | 0 | 14.45 | 0 | 13.97 | 0 | 20.9 | 4.79 | 11.33 | 4.49 | 4.88 | 4.59 | 9.86 | 9.96 | 9.86 | 21.68 | 9.77 | 20.02 | 0.1 | 9.38 | 0.1 | 10.94 | 0.1 | 7.23 | 4.79 | 4.88 | 4.2 | 4.79 | 21.29 | 4.79 | 28.91 | 9.77 | 27.83 | 9.77 | 28.32 | 9.67
60.896 | 1.37 | 0 | 14.45 | 0 | 14.16 | 0 | 17.58 | 4.79 | 11.91 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.31 | 0 | 10.16 | 0.1 | 11.72 | 0.1 | 8.01 | 4.79 | 4.88 | 5.57 | 4.79 | 21.29 | 4.79 | 29.4 | 9.77 | 28.22 | 9.77 | 28.81 | 9.77
60.928 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 17.68 | 4.79 | 12.5 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.8 | 0.1 | 10.94 | 0.1 | 12.5 | 0.1 | 8.89 | 4.79 | 4.88 | 6.15 | 4.79 | 21.29 | 4.79 | 24.71 | 9.77 | 28.91 | 9.77 | 29.3 | 9.77
60.96 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 17.87 | 4.79 | 12.99 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.19 | 0.1 | 11.62 | 0.1 | 13.28 | 0.1 | 9.57 | 4.69 | 4.88 | 6.05 | 4.79 | 21.19 | 4.88 | 23.63 | 9.77 | 29.3 | 9.77 | 27.05 | 9.77
60.992 | 0.98 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 13.57 | 4.69 | 4.88 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 21.58 | 0.1 | 12.4 | 0.1 | 14.06 | 0.1 | 10.35 | 4.69 | 4.88 | 5.66 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 26.37 | 9.77 | 23.44 | 9.77
61.024 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 0 | 18.16 | 4.79 | 14.06 | 4.69 | 4.98 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 21.97 | 0.1 | 13.18 | 0.1 | 14.65 | 0.1 | 11.13 | 4.59 | 4.79 | 5.18 | 4.79 | 21 | 4.79 | 24.32 | 9.77 | 23.44 | 9.77 | 24.22 | 9.77
61.056 | 0.78 | 0 | 14.36 | 0 | 14.45 | 0 | 18.16 | 4.79 | 14.65 | 4.69 | 5.96 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.36 | 0.1 | 13.97 | 0.1 | 15.43 | 0.1 | 11.91 | 4.59 | 4.79 | 4.98 | 4.79 | 6.05 | 4.88 | 24.32 | 9.77 | 24.22 | 9.77 | 24.32 | 9.77
61.088 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.14 | 4.69 | 6.93 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.75 | 0.1 | 14.75 | 0.1 | 16.11 | 0.1 | 12.79 | 4.59 | 4.88 | 5.18 | 4.79 | 2.73 | 4.88 | 24.41 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77
61.12 | 0.68 | 0.1 | 6.15 | 0.1 | 14.45 | 0.1 | 18.16 | 4.79 | 15.72 | 4.69 | 8.4 | 4.59 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 0.1 | 15.33 | 0.1 | 16.9 | 0.1 | 13.48 | 4.59 | 5.66 | 5.08 | 4.79 | 3.42 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
61.152 | 0.59 | 0.1 | 1.86 | 0 | 14.45 | 0.1 | 18.16 | 4.88 | 16.11 | 4.79 | 9.28 | 4.49 | 9.86 | 9.67 | 9.86 | 13.28 | 9.86 | 23.54 | 0.1 | 16.11 | 0.1 | 17.48 | 0.1 | 14.26 | 4.59 | 6.64 | 5.08 | 4.79 | 4.3 | 4.79 | 24.41 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
61.184 | 0.49 | 0 | 1.95 | 0 | 14.45 | 0 | 18.16 | 4.79 | 16.6 | 4.79 | 9.86 | 4.49 | 9.86 | 9.67 | 9.86 | 7.13 | 9.86 | 23.93 | 0.1 | 16.7 | 0 | 18.07 | 0.1 | 14.94 | 4.49 | 8.5 | 4.88 | 4.79 | 5.57 | 4.88 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
61.216 | 0.49 | 0 | 1.66 | 0 | 14.45 | 0 | 18.07 | 4.88 | 17.09 | 4.79 | 10.55 | 4.49 | 10.55 | 9.67 | 9.86 | 7.91 | 9.86 | 24.22 | 0.1 | 17.48 | 0.1 | 18.75 | 0.1 | 15.63 | 4.59 | 10.25 | 4.88 | 4.79 | 6.25 | 4.88 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
61.248 | 0.39 | 0.1 | 1.56 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 11.23 | 9.67 | 9.86 | 10.94 | 9.96 | 20.8 | 0.1 | 18.07 | 0.1 | 19.34 | 0.1 | 16.41 | 4.59 | 11.62 | 4.88 | 4.79 | 5.96 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
61.28 | 0.29 | 0.1 | 1.46 | 0 | 5.08 | 0.1 | 18.07 | 4.88 | 17.97 | 4.79 | 11.82 | 4.59 | 12.5 | 9.57 | 9.86 | 11.33 | 9.96 | 20.9 | 0.1 | 18.85 | 0.1 | 20.02 | 0.1 | 16.99 | 4.59 | 12.6 | 4.79 | 4.79 | 5.47 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
61.312 | 0.29 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 18.07 | 4.88 | 18.46 | 4.79 | 12.4 | 4.59 | 13.48 | 9.57 | 9.86 | 10.35 | 9.96 | 21.29 | 0.1 | 19.43 | 0.1 | 20.61 | 0.1 | 17.68 | 4.59 | 13.48 | 4.79 | 4.79 | 5.27 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
61.344 | 0.2 | 0 | 1.07 | 0 | 1.86 | 0 | 4.98 | 4.79 | 18.95 | 4.79 | 12.79 | 4.59 | 14.36 | 9.67 | 9.86 | 9.86 | 9.96 | 21.39 | 0.1 | 19.92 | 0.1 | 21 | 0.1 | 18.46 | 4.59 | 14.26 | 4.69 | 4.79 | 5.18 | 4.79 | 21.68 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
61.376 | 0.2 | 0 | 0.98 | 0 | 1.66 | 0 | 2.83 | 4.79 | 19.34 | 4.88 | 13.38 | 4.69 | 14.84 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0 | 20.61 | 0.1 | 21.78 | 0.1 | 18.95 | 4.69 | 15.04 | 4.69 | 4.79 | 5.18 | 4.88 | 8.89 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
61.408 | 0.2 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.52 | 4.88 | 19.73 | 4.88 | 13.97 | 4.69 | 15.43 | 9.67 | 9.86 | 10.16 | 9.96 | 21.39 | 0.1 | 21.19 | 0.1 | 22.36 | 0.1 | 19.63 | 4.69 | 15.82 | 4.69 | 4.79 | 5.08 | 4.88 | 6.15 | 9.77 | 24.22 | 9.86 | 11.62 | 9.86
61.44 | 0.1 | 0.1 | 0.88 | 0 | 1.37 | 0.1 | 4.49 | 4.88 | 20.22 | 4.88 | 14.45 | 4.69 | 16.02 | 9.67 | 9.86 | 10.16 | 9.96 | 21.49 | 0.1 | 21.68 | 0.1 | 19.63 | 0.1 | 20.22 | 4.69 | 16.5 | 4.59 | 4.79 | 5.08 | 4.88 | 9.86 | 9.77 | 11.52 | 9.86 | 5.86 | 9.86
61.472 | 0.1 | 0.29 | 0.68 | 0.1 | 1.17 | 0.1 | 5.57 | 4.88 | 20.61 | 4.88 | 15.04 | 4.69 | 16.6 | 9.67 | 10.55 | 9.86 | 9.96 | 21.49 | 0.1 | 22.36 | 0.1 | 16.8 | 0.1 | 20.9 | 4.69 | 17.19 | 4.59 | 4.79 | 5.08 | 4.88 | 12.4 | 9.77 | 5.66 | 9.86 | 8.79 | 9.86
61.504 | 0.1 | 1.17 | 0.59 | 0 | 1.07 | 0.1 | 5.76 | 4.79 | 21 | 4.79 | 15.43 | 4.69 | 16.99 | 9.67 | 11.23 | 9.86 | 9.96 | 21.39 | 0.1 | 17.77 | 0.1 | 17.58 | 0.1 | 21.39 | 4.79 | 17.87 | 4.59 | 4.79 | 4.88 | 4.79 | 10.94 | 9.77 | 9.08 | 9.86 | 12.21 | 9.86
61.536 | 0.1 | 2.05 | 0.59 | 0 | 1.07 | 0 | 5.47 | 4.79 | 18.26 | 4.79 | 15.92 | 4.79 | 17.48 | 9.67 | 12.31 | 9.86 | 9.86 | 21.49 | 0 | 16.6 | 0.1 | 17.68 | 0.1 | 21.97 | 4.69 | 18.46 | 4.49 | 4.79 | 4.88 | 4.88 | 9.96 | 9.77 | 12.21 | 9.86 | 11.52 | 9.77
61.568 | 0.1 | 3.61 | 0.59 | 0.1 | 0.88 | 0.1 | 5.37 | 4.88 | 18.85 | 4.88 | 16.41 | 4.79 | 17.97 | 9.67 | 13.48 | 9.86 | 9.96 | 21.39 | 0.1 | 17.29 | 0.1 | 17.68 | 0.1 | 22.56 | 4.79 | 19.14 | 4.59 | 4.88 | 4.79 | 4.79 | 9.96 | 9.77 | 11.43 | 9.86 | 10.25 | 9.86
61.6 | 0.1 | 4.39 | 0.49 | 0 | 0.78 | 0.1 | 4.88 | 4.88 | 18.36 | 4.88 | 16.99 | 4.79 | 18.46 | 9.67 | 14.36 | 9.77 | 9.96 | 21.49 | 0.1 | 17.48 | 0.1 | 17.68 | 0.1 | 18.16 | 4.79 | 19.73 | 4.59 | 5.86 | 4.79 | 4.88 | 10.45 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
61.632 | 0.1 | 4.98 | 0.39 | 0 | 0.78 | 0 | 4.88 | 4.88 | 18.36 | 4.88 | 17.38 | 4.79 | 18.85 | 9.67 | 14.94 | 9.67 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.19 | 4.79 | 20.31 | 4.49 | 7.32 | 4.69 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 10.25 | 9.86
61.664 | 0.1 | 5.47 | 0.29 | 0 | 0.68 | 0 | 4.88 | 4.79 | 18.36 | 4.79 | 17.77 | 4.79 | 19.34 | 9.67 | 15.43 | 9.67 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.79 | 20.9 | 4.59 | 9.47 | 4.69 | 4.79 | 10.16 | 9.77 | 10.35 | 9.86 | 10.35 | 9.86
61.696 | 0.1 | 5.96 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.79 | 18.26 | 4.79 | 18.26 | 4.79 | 19.73 | 9.67 | 16.02 | 9.67 | 9.96 | 9.86 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 4.79 | 21.49 | 4.59 | 11.04 | 4.69 | 4.88 | 10.16 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86
61.728 | 0.1 | 6.64 | 0.2 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 18.75 | 4.79 | 20.22 | 9.77 | 16.5 | 9.67 | 9.96 | 6.93 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 22.07 | 4.59 | 12.21 | 4.59 | 4.88 | 10.16 | 9.86 | 10.45 | 9.86 | 10.16 | 9.86
61.76 | 0.1 | 7.23 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 19.24 | 4.79 | 20.61 | 9.77 | 16.99 | 9.67 | 9.96 | 9.67 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 22.66 | 4.59 | 13.09 | 4.59 | 4.88 | 10.16 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
61.792 | 0.1 | 7.81 | 0.1 | 0.1 | 0.29 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 19.63 | 4.79 | 20.9 | 9.77 | 17.38 | 9.67 | 9.96 | 11.43 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.24 | 4.59 | 13.87 | 4.59 | 4.79 | 9.96 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86
61.824 | 0.1 | 8.4 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.79 | 18.16 | 4.79 | 20.12 | 4.79 | 21.39 | 9.77 | 17.87 | 9.57 | 9.96 | 10.94 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.73 | 4.69 | 14.65 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
61.856 | 0.1 | 8.98 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.88 | 18.16 | 4.88 | 20.51 | 4.79 | 21.78 | 9.77 | 18.36 | 9.57 | 9.86 | 10.16 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 4.88 | 24.22 | 4.69 | 15.43 | 4.59 | 4.88 | 9.96 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86
61.888 | 0.1 | 9.67 | 0.1 | 0.68 | 0.29 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 20.9 | 4.88 | 22.27 | 9.77 | 18.85 | 9.57 | 9.96 | 9.86 | 0.1 | 17.58 | 0.1 | 6.84 | 0.1 | 17.77 | 4.88 | 24.71 | 4.69 | 16.02 | 4.59 | 4.88 | 9.86 | 9.77 | 10.06 | 9.86 | 9.86 | 9.86
61.92 | 0.1 | 10.25 | 0.1 | 1.37 | 0.2 | 0.68 | 4.59 | 4.88 | 18.07 | 4.88 | 17.38 | 4.88 | 22.66 | 9.86 | 19.24 | 9.57 | 9.96 | 9.96 | 0.1 | 17.48 | 0.1 | 1.66 | 0.1 | 17.68 | 4.88 | 25.29 | 4.69 | 16.8 | 4.59 | 4.88 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
61.952 | 0.1 | 10.74 | 0.1 | 2.64 | 0.1 | 1.37 | 4.59 | 4.88 | 12.99 | 4.88 | 17.68 | 4.88 | 22.95 | 9.86 | 19.73 | 9.57 | 9.96 | 10.25 | 0.1 | 5.18 | 0.1 | 2.25 | 0.1 | 17.68 | 4.88 | 25.88 | 4.79 | 17.48 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.77 | 9.86
61.984 | 0.1 | 11.43 | 0.1 | 3.81 | 0.1 | 2.64 | 4.49 | 4.79 | 4 | 4.88 | 17.77 | 4.79 | 23.34 | 9.77 | 20.12 | 9.57 | 9.96 | 10.16 | 0.1 | 1.95 | 0.1 | 2.05 | 0 | 17.68 | 4.88 | 26.27 | 4.79 | 18.07 | 4.59 | 4.88 | 9.77 | 9.77 | 9.96 | 9.77 | 9.77 | 9.86
62.016 | 0.1 | 11.91 | 0.1 | 4.49 | 0.1 | 3.71 | 4.49 | 4.79 | 3.13 | 4.88 | 18.07 | 4.88 | 23.63 | 9.77 | 20.51 | 9.57 | 9.96 | 9.86 | 0 | 2.15 | 0.1 | 1.76 | 0.1 | 17.19 | 4.88 | 22.75 | 4.69 | 18.65 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
62.048 | 0.1 | 12.5 | 0.1 | 5.08 | 0.1 | 4.59 | 4.59 | 4.88 | 3.71 | 4.88 | 18.07 | 4.88 | 24.12 | 9.86 | 20.9 | 9.57 | 10.84 | 9.86 | 0.1 | 2.05 | 0.1 | 1.66 | 0.1 | 4.1 | 4.98 | 22.07 | 4.79 | 19.34 | 4.69 | 5.76 | 9.67 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
62.08 | 0.1 | 12.99 | 0.1 | 5.66 | 0.1 | 5.08 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 4.88 | 24.41 | 9.86 | 21.39 | 9.57 | 11.82 | 9.86 | 0.1 | 1.76 | 0.1 | 1.56 | 0.1 | 1.86 | 4.88 | 21.68 | 4.79 | 20.02 | 4.69 | 7.13 | 9.67 | 10.84 | 9.86 | 9.86 | 9.67 | 9.86
62.112 | 0.1 | 13.57 | 0.1 | 6.25 | 0.1 | 5.66 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 22.46 | 9.86 | 21.68 | 9.57 | 13.18 | 9.77 | 0.1 | 1.66 | 0.1 | 1.56 | 0.1 | 2.15 | 4.88 | 21.58 | 4.79 | 20.7 | 4.69 | 9.38 | 9.67 | 11.91 | 9.77 | 10.94 | 9.77 | 9.96
62.144 | 0.1 | 14.06 | 0.1 | 6.84 | 0.1 | 6.15 | 4.49 | 4.79 | 5.66 | 4.88 | 18.07 | 4.88 | 22.17 | 9.86 | 22.17 | 9.67 | 14.06 | 9.77 | 0 | 1.37 | 0.1 | 1.37 | 0.1 | 1.76 | 4.88 | 21.58 | 4.79 | 21.19 | 4.69 | 10.74 | 9.67 | 13.67 | 9.77 | 12.01 | 9.67 | 11.04
62.176 | 0.1 | 14.65 | 0.1 | 7.42 | 0.1 | 6.74 | 4.59 | 4.88 | 5.47 | 4.79 | 18.07 | 4.88 | 21.88 | 9.77 | 22.46 | 9.67 | 14.65 | 9.77 | 0 | 1.37 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.49 | 4.79 | 21.68 | 4.69 | 12.01 | 9.67 | 15.23 | 9.67 | 13.67 | 9.67 | 12.21
62.208 | 0.1 | 15.14 | 0.1 | 8.01 | 0.1 | 7.42 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.68 | 9.86 | 22.95 | 9.67 | 15.33 | 9.77 | 0.1 | 1.27 | 0.1 | 0.98 | 0.1 | 1.56 | 4.88 | 21.39 | 4.79 | 22.36 | 4.79 | 13.09 | 9.67 | 16.5 | 9.67 | 15.43 | 9.67 | 14.26
62.24 | 0.1 | 15.72 | 0.1 | 8.59 | 0.1 | 8.01 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.78 | 9.86 | 23.34 | 9.67 | 15.82 | 9.67 | 0.1 | 1.17 | 0.1 | 1.07 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 22.85 | 4.79 | 13.77 | 9.67 | 17.38 | 9.67 | 16.6 | 9.67 | 15.82
62.272 | 0.1 | 16.11 | 0.1 | 9.28 | 0.1 | 8.59 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 23.63 | 9.67 | 16.31 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 23.44 | 4.79 | 14.65 | 9.67 | 18.26 | 9.67 | 17.38 | 9.67 | 16.9
62.304 | 0.1 | 16.6 | 0 | 9.77 | 0.1 | 9.28 | 4.69 | 5.47 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 24.02 | 9.67 | 16.9 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.17 | 4.98 | 21.29 | 4.79 | 23.93 | 4.79 | 15.33 | 9.67 | 18.95 | 9.67 | 18.16 | 9.67 | 17.68
62.336 | 0.1 | 17.09 | 0.1 | 10.45 | 0.1 | 9.67 | 4.69 | 6.25 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 24.41 | 9.67 | 17.19 | 9.67 | 0 | 0.88 | 0.1 | 0.78 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.32 | 4.79 | 16.11 | 9.67 | 19.53 | 9.67 | 18.85 | 9.67 | 18.36
62.368 | 0.1 | 13.77 | 0.1 | 11.04 | 0.1 | 10.45 | 4.69 | 7.32 | 4.79 | 4.88 | 5.96 | 4.88 | 21.49 | 9.86 | 22.46 | 9.67 | 17.77 | 9.67 | 0.68 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.29 | 4.88 | 25 | 4.79 | 16.8 | 9.67 | 20.12 | 9.67 | 19.53 | 9.67 | 19.14
62.4 | 0.1 | 13.97 | 0.1 | 11.62 | 0.1 | 11.04 | 4.69 | 8.59 | 4.79 | 4.88 | 2.73 | 4.88 | 21.49 | 9.86 | 22.27 | 9.67 | 18.26 | 9.67 | 1.66 | 0.68 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 21.29 | 4.88 | 25.49 | 4.79 | 17.58 | 9.67 | 20.8 | 9.67 | 20.22 | 9.67 | 19.82
62.432 | 0.1 | 14.16 | 0.1 | 12.11 | 0.1 | 11.62 | 4.79 | 9.47 | 4.69 | 4.88 | 3.52 | 4.88 | 21.49 | 9.86 | 21.88 | 9.67 | 18.65 | 9.57 | 3.13 | 0.68 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 14.26 | 4.88 | 25.98 | 4.79 | 18.07 | 9.67 | 21.39 | 9.67 | 20.8 | 9.67 | 20.31
62.464 | 0.1 | 14.45 | 0.1 | 12.7 | 0.1 | 12.11 | 4.79 | 10.06 | 4.69 | 4.88 | 4.59 | 4.79 | 21.39 | 9.86 | 21.68 | 9.67 | 19.24 | 9.57 | 5.18 | 0.49 | 0.1 | 0.49 | 0.1 | 0.68 | 4.88 | 3.22 | 4.79 | 21.78 | 4.79 | 18.56 | 9.67 | 21.97 | 9.67 | 21.39 | 9.67 | 20.9
62.496 | 0.1 | 14.36 | 0.1 | 13.18 | 0.1 | 12.6 | 4.79 | 10.64 | 4.69 | 4.88 | 5.57 | 4.88 | 21.49 | 9.77 | 21.68 | 9.77 | 19.43 | 9.57 | 6.25 | 0.49 | 0.1 | 0.49 | 0.1 | 0.68 | 4.88 | 3.13 | 4.88 | 20.31 | 4.79 | 19.24 | 9.67 | 22.46 | 9.67 | 21.88 | 9.67 | 21.58
62.528 | 0.1 | 14.45 | 0.1 | 13.87 | 0.1 | 13.28 | 4.79 | 11.33 | 4.59 | 4.88 | 5.86 | 4.88 | 14.06 | 9.86 | 21.58 | 9.77 | 19.92 | 9.57 | 7.13 | 0.49 | 0.2 | 0.39 | 0.1 | 0.59 | 4.98 | 3.52 | 4.88 | 20.9 | 4.88 | 19.92 | 9.67 | 22.95 | 9.67 | 22.56 | 9.67 | 22.17
62.56 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 13.77 | 4.88 | 11.91 | 4.59 | 4.88 | 5.47 | 4.88 | 7.23 | 9.86 | 21.49 | 9.77 | 20.41 | 9.57 | 7.81 | 0.29 | 1.46 | 0.29 | 0.1 | 0.59 | 4.88 | 4.3 | 4.88 | 21.19 | 4.88 | 20.61 | 9.67 | 23.63 | 9.67 | 23.05 | 9.67 | 22.75
62.592 | 0.1 | 14.45 | 0.1 | 14.84 | 0.1 | 14.36 | 4.88 | 12.5 | 4.59 | 4.88 | 5.18 | 4.88 | 8.01 | 9.86 | 21.49 | 9.77 | 20.8 | 9.57 | 8.59 | 0.29 | 2.73 | 0.2 | 1.07 | 0.49 | 4.88 | 5.86 | 4.88 | 21.19 | 4.88 | 21.19 | 9.67 | 24.22 | 9.67 | 23.63 | 9.67 | 23.15
62.624 | 0.1 | 14.36 | 0.1 | 15.33 | 0.1 | 14.84 | 4.79 | 13.09 | 4.49 | 4.88 | 4.98 | 4.79 | 11.13 | 9.86 | 21.39 | 9.77 | 21.29 | 9.57 | 9.47 | 0.2 | 4.49 | 0.2 | 2.05 | 0.39 | 4.88 | 6.05 | 4.88 | 21.19 | 4.88 | 21.58 | 9.67 | 24.61 | 9.67 | 24.22 | 9.67 | 23.73
62.656 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 15.33 | 4.79 | 13.48 | 4.49 | 4.88 | 4.98 | 4.88 | 11.23 | 9.86 | 21.39 | 9.77 | 21.58 | 9.67 | 10.25 | 0.2 | 6.05 | 0.2 | 3.91 | 0.29 | 4.88 | 5.76 | 4.79 | 21.19 | 4.79 | 22.27 | 9.67 | 25.1 | 9.67 | 24.71 | 9.67 | 24.22
62.688 | 0.1 | 14.36 | 0.1 | 16.31 | 0.1 | 15.82 | 4.88 | 14.06 | 4.59 | 5.08 | 4.98 | 4.88 | 10.35 | 9.86 | 21.39 | 9.77 | 21.97 | 9.67 | 11.13 | 0.2 | 7.03 | 0.2 | 5.47 | 0.29 | 4.98 | 5.47 | 4.88 | 21.19 | 4.88 | 22.95 | 9.77 | 25.59 | 9.77 | 25.1 | 9.67 | 24.81
62.72 | 0.1 | 14.36 | 0.1 | 16.8 | 0.1 | 16.31 | 4.88 | 14.65 | 4.49 | 5.86 | 4.88 | 4.88 | 9.96 | 9.86 | 21.39 | 9.77 | 22.36 | 9.67 | 11.72 | 0.2 | 7.81 | 0.1 | 6.74 | 0.29 | 4.88 | 5.18 | 4.88 | 21.19 | 4.88 | 23.34 | 9.77 | 26.17 | 9.77 | 25.68 | 9.67 | 25.29
62.752 | 0.1 | 14.36 | 0.1 | 17.29 | 0.1 | 16.8 | 4.88 | 15.14 | 4.59 | 6.93 | 4.79 | 4.88 | 9.96 | 9.86 | 21.39 | 9.77 | 22.75 | 9.67 | 12.6 | 0.2 | 8.69 | 0.1 | 7.52 | 0.2 | 4.88 | 5.18 | 4.88 | 21.19 | 4.88 | 24.02 | 9.77 | 26.66 | 9.77 | 26.17 | 9.67 | 25.88
62.784 | 0.1 | 14.16 | 0.1 | 15.43 | 0.1 | 17.19 | 4.88 | 15.63 | 4.59 | 8.4 | 4.69 | 4.79 | 10.25 | 9.86 | 13.77 | 9.77 | 23.15 | 9.67 | 13.28 | 0.1 | 9.38 | 0.1 | 8.4 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.88 | 24.51 | 9.77 | 27.15 | 9.77 | 26.76 | 9.67 | 26.27
62.816 | 0.1 | 5.08 | 0.1 | 15.33 | 0.1 | 14.94 | 4.88 | 16.11 | 4.59 | 9.28 | 4.69 | 4.88 | 10.16 | 9.86 | 7.23 | 9.77 | 23.54 | 9.67 | 14.06 | 0.1 | 10.25 | 0.1 | 9.08 | 0.1 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 25 | 9.77 | 27.54 | 9.77 | 27.25 | 9.77 | 26.76
62.848 | 0.1 | 1.66 | 0.1 | 14.84 | 0.1 | 15.23 | 4.88 | 16.6 | 4.59 | 9.96 | 4.69 | 4.88 | 9.96 | 9.86 | 8.01 | 9.86 | 23.93 | 9.67 | 14.75 | 0.1 | 11.04 | 0.1 | 9.86 | 0.1 | 4.98 | 5.08 | 4.88 | 21.19 | 4.88 | 25.39 | 9.77 | 28.13 | 9.77 | 27.74 | 9.77 | 27.25
62.88 | 0.1 | 1.95 | 0.1 | 14.75 | 0.1 | 14.75 | 4.88 | 16.99 | 4.69 | 10.64 | 4.59 | 4.88 | 9.86 | 9.86 | 11.04 | 9.86 | 24.32 | 9.67 | 15.43 | 0.1 | 11.82 | 0.1 | 10.64 | 0.1 | 4.88 | 4.98 | 4.88 | 21.09 | 4.88 | 25.88 | 9.77 | 28.42 | 9.77 | 28.13 | 9.77 | 27.74
62.912 | 0.1 | 1.66 | 0.1 | 14.75 | 0.1 | 14.65 | 4.88 | 17.58 | 4.59 | 11.23 | 4.59 | 4.88 | 9.96 | 9.86 | 11.33 | 9.77 | 20.9 | 9.67 | 16.11 | 0.1 | 12.6 | 0.1 | 11.43 | 0.1 | 4.88 | 4.88 | 4.88 | 7.03 | 4.88 | 23.15 | 9.77 | 29.01 | 9.77 | 28.52 | 9.77 | 28.22
62.944 | 0.1 | 1.46 | 0.1 | 14.55 | 0.1 | 14.65 | 4.88 | 17.97 | 4.69 | 11.82 | 4.49 | 4.79 | 9.96 | 9.86 | 10.35 | 9.86 | 20.9 | 9.67 | 16.8 | 0 | 13.18 | 0 | 12.31 | 0 | 4.88 | 4.79 | 4.88 | 2.34 | 4.88 | 20.8 | 9.77 | 29.49 | 9.86 | 29.01 | 9.77 | 28.71
62.976 | 0.1 | 1.46 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.36 | 4.69 | 12.31 | 4.59 | 4.88 | 9.86 | 9.86 | 9.86 | 9.86 | 21.19 | 9.67 | 17.58 | 0 | 14.06 | 0.1 | 12.99 | 0 | 4.88 | 4.79 | 4.79 | 3.32 | 4.88 | 21.39 | 9.77 | 26.17 | 9.77 | 29.4 | 9.77 | 29.1
63.008 | 0.1 | 1.07 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.75 | 4.79 | 12.89 | 4.59 | 4.88 | 9.77 | 9.86 | 10.06 | 9.86 | 21.49 | 9.77 | 18.26 | 0.1 | 14.75 | 0.1 | 13.77 | 0 | 4.88 | 4.69 | 4.88 | 4.1 | 4.88 | 21.39 | 9.77 | 25.29 | 9.86 | 25.78 | 9.77 | 26.37
63.04 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.34 | 4.79 | 13.38 | 4.59 | 4.88 | 9.77 | 9.86 | 10.25 | 9.86 | 21.39 | 9.77 | 18.65 | 0 | 15.43 | 0.1 | 14.45 | 0.1 | 5.86 | 4.69 | 4.88 | 5.27 | 4.88 | 21.29 | 9.77 | 24.81 | 9.86 | 25.1 | 9.77 | 23.54
63.072 | 0.1 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.82 | 4.69 | 13.87 | 4.49 | 4.98 | 9.77 | 9.86 | 10.25 | 9.86 | 21.49 | 9.77 | 19.43 | 0.1 | 16.21 | 0.1 | 15.33 | 0 | 6.93 | 4.59 | 4.88 | 6.25 | 4.88 | 21.19 | 9.77 | 24.61 | 9.86 | 24.71 | 9.77 | 24.12
63.104 | 0.1 | 0.88 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 20.12 | 4.79 | 14.45 | 4.49 | 5.96 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.02 | 0 | 16.8 | 0.1 | 16.02 | 0 | 8.89 | 4.59 | 4.88 | 5.96 | 4.88 | 21.29 | 9.77 | 24.51 | 9.77 | 24.51 | 9.86 | 24.12
63.136 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.61 | 4.79 | 15.04 | 4.59 | 6.84 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.51 | 0 | 17.48 | 0.1 | 16.6 | 0 | 10.45 | 4.59 | 4.79 | 5.47 | 4.88 | 21.29 | 9.77 | 24.51 | 9.77 | 24.41 | 9.86 | 24.22
63.168 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 21 | 4.88 | 15.53 | 4.59 | 8.2 | 9.67 | 10.55 | 9.96 | 9.86 | 21.49 | 9.77 | 21 | 0.1 | 18.07 | 0.1 | 17.29 | 0 | 11.82 | 4.59 | 4.88 | 5.18 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.41 | 9.86 | 24.12
63.2 | 0.1 | 0.59 | 0.1 | 9.96 | 0.1 | 14.36 | 4.88 | 18.85 | 4.88 | 16.02 | 4.59 | 9.18 | 9.67 | 11.13 | 9.86 | 9.86 | 21.49 | 9.77 | 21.68 | 0.1 | 18.85 | 0.1 | 17.97 | 0 | 12.7 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
63.232 | 0.1 | 0.59 | 0.1 | 2.93 | 0.1 | 11.04 | 4.88 | 18.95 | 4.88 | 16.5 | 4.59 | 9.86 | 9.67 | 12.21 | 9.77 | 9.86 | 21.49 | 9.77 | 22.17 | 0.1 | 19.34 | 0.1 | 18.65 | 0.1 | 13.67 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22
63.264 | 0.1 | 0.49 | 0 | 1.95 | 0.1 | 2.73 | 4.88 | 18.46 | 4.79 | 16.99 | 4.59 | 10.55 | 9.57 | 13.48 | 9.77 | 9.86 | 21.49 | 9.77 | 18.36 | 0 | 20.02 | 0 | 19.24 | 0 | 14.45 | 4.49 | 4.79 | 5.08 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77 | 24.12
63.296 | 0.1 | 0.39 | 0.1 | 1.66 | 0.1 | 2.05 | 4.88 | 18.36 | 4.88 | 17.38 | 4.69 | 11.13 | 9.67 | 14.26 | 9.77 | 9.86 | 21.39 | 9.86 | 17.48 | 0 | 20.7 | 0.1 | 19.82 | 0 | 15.14 | 4.49 | 4.88 | 4.88 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
63.328 | 0.1 | 0.39 | 0.1 | 1.56 | 0.1 | 1.76 | 4.88 | 18.36 | 4.88 | 17.87 | 4.69 | 11.72 | 9.67 | 14.84 | 9.67 | 9.86 | 21.39 | 9.77 | 17.58 | 0 | 21.19 | 0.1 | 20.51 | 0.1 | 15.92 | 4.59 | 4.79 | 4.88 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
63.36 | 0.1 | 0.29 | 0.1 | 1.56 | 0.1 | 1.66 | 4.88 | 18.26 | 4.88 | 18.26 | 4.69 | 12.31 | 9.67 | 15.33 | 9.67 | 9.86 | 10.64 | 9.86 | 17.58 | 0.1 | 21.78 | 0.1 | 21 | 0 | 16.7 | 4.59 | 4.79 | 4.88 | 4.88 | 7.81 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
63.392 | 0.2 | 0.29 | 0.1 | 1.17 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.75 | 4.69 | 12.79 | 9.67 | 15.92 | 9.67 | 9.86 | 6.74 | 9.77 | 17.58 | 0.1 | 22.36 | 0.1 | 21.68 | 0 | 17.29 | 4.59 | 4.79 | 4.79 | 4.88 | 2.34 | 9.77 | 14.55 | 9.86 | 24.22 | 9.86 | 24.12
63.424 | 1.27 | 0.2 | 0 | 0.98 | 0.1 | 1.27 | 4.88 | 18.16 | 4.88 | 19.24 | 4.69 | 13.38 | 9.67 | 16.41 | 9.67 | 9.86 | 9.67 | 9.86 | 17.58 | 0 | 18.26 | 0 | 22.27 | 0 | 17.87 | 4.59 | 5.08 | 4.79 | 4.88 | 3.22 | 9.77 | 6.15 | 9.86 | 12.89 | 9.86 | 24.12
63.456 | 2.15 | 0.1 | 0 | 0.88 | 0.1 | 1.07 | 4.88 | 18.07 | 4.88 | 19.73 | 4.79 | 13.87 | 9.67 | 16.99 | 9.67 | 9.86 | 11.62 | 9.86 | 17.58 | 0 | 16.9 | 0.1 | 22.75 | 0 | 18.56 | 4.59 | 5.86 | 4.69 | 4.88 | 4 | 9.77 | 7.32 | 9.86 | 6.35 | 9.86 | 11.52
63.488 | 3.61 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 14.45 | 9.67 | 17.38 | 9.67 | 9.86 | 10.94 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 19.53 | 0.1 | 19.14 | 4.59 | 7.42 | 4.59 | 4.88 | 5.18 | 9.77 | 11.33 | 9.86 | 7.52 | 9.86 | 5.76
63.52 | 4.39 | 0.1 | 0.1 | 0.88 | 0.1 | 0.98 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 14.84 | 9.67 | 17.87 | 9.67 | 9.86 | 10.16 | 9.86 | 17.48 | 0.1 | 17.68 | 0.1 | 18.75 | 0 | 19.92 | 4.69 | 9.47 | 4.59 | 4.88 | 6.25 | 9.77 | 11.82 | 9.86 | 11.33 | 9.86 | 8.5
63.552 | 4.98 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 9.67 | 18.36 | 9.57 | 9.86 | 9.86 | 9.86 | 17.48 | 0 | 17.58 | 0 | 18.26 | 0 | 20.41 | 4.69 | 11.04 | 4.59 | 4.88 | 5.96 | 9.77 | 10.45 | 9.86 | 11.82 | 9.86 | 12.11
63.584 | 5.57 | 0 | 0 | 0.59 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 17.68 | 4.79 | 15.82 | 9.67 | 18.75 | 9.57 | 9.86 | 10.16 | 9.86 | 17.38 | 0 | 17.58 | 0.1 | 18.07 | 0 | 21.19 | 4.69 | 12.31 | 4.59 | 4.88 | 5.47 | 9.77 | 10.16 | 9.86 | 10.45 | 9.86 | 11.33
63.616 | 6.05 | 0 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 12.21 | 4.88 | 17.68 | 4.79 | 16.41 | 9.67 | 19.24 | 9.57 | 9.86 | 10.16 | 9.86 | 17.38 | 0 | 17.68 | 0.1 | 18.07 | 0 | 21.68 | 4.69 | 13.09 | 4.49 | 4.88 | 5.27 | 9.77 | 10.25 | 9.77 | 9.96 | 9.86 | 10.25
63.648 | 6.64 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 3.61 | 4.88 | 17.97 | 4.79 | 16.9 | 9.77 | 19.63 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 17.97 | 0 | 22.27 | 4.79 | 13.97 | 4.49 | 4.88 | 5.18 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86 | 10.06
63.68 | 7.23 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 18.07 | 4.88 | 17.38 | 9.77 | 20.12 | 9.67 | 9.86 | 9.96 | 9.86 | 15.33 | 0.1 | 17.68 | 0.1 | 17.87 | 0 | 22.75 | 4.79 | 14.75 | 4.49 | 4.88 | 5.18 | 9.77 | 10.45 | 9.86 | 10.35 | 9.86 | 10.25
63.712 | 7.81 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 3.61 | 4.88 | 18.07 | 4.79 | 17.77 | 9.77 | 20.51 | 9.67 | 9.86 | 9.96 | 9.86 | 3.71 | 0.1 | 17.68 | 0 | 17.87 | 0 | 23.44 | 4.79 | 15.43 | 4.49 | 4.98 | 5.18 | 9.77 | 10.25 | 9.86 | 10.45 | 9.86 | 10.35
63.744 | 8.4 | 0 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 4.79 | 4.88 | 18.07 | 4.79 | 18.26 | 9.77 | 20.9 | 9.67 | 9.86 | 9.96 | 9.96 | 2.15 | 0 | 17.68 | 0.1 | 17.77 | 0 | 23.93 | 4.79 | 16.21 | 4.49 | 6.05 | 4.98 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86 | 10.25
63.776 | 9.08 | 0.1 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 18.65 | 9.77 | 21.29 | 9.67 | 9.86 | 9.86 | 9.86 | 2.34 | 0 | 17.68 | 0.1 | 17.77 | 0 | 24.51 | 4.79 | 16.8 | 4.49 | 7.62 | 4.98 | 9.77 | 10.06 | 9.77 | 10.06 | 9.86 | 10.06
63.808 | 9.67 | 0.1 | 0.88 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 19.14 | 9.77 | 21.68 | 9.67 | 9.86 | 9.77 | 9.86 | 1.86 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 25 | 4.79 | 17.48 | 4.49 | 9.67 | 4.88 | 9.77 | 10.06 | 9.86 | 10.06 | 9.86 | 10.06
63.84 | 10.25 | 0.1 | 1.76 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 22.07 | 9.67 | 10.55 | 9.77 | 9.96 | 1.86 | 0 | 17.58 | 0.1 | 17.68 | 0 | 25.39 | 4.88 | 18.16 | 4.49 | 11.13 | 4.88 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86 | 10.16
63.872 | 10.84 | 0.1 | 3.03 | 0.1 | 0.1 | 0.2 | 4.88 | 5.08 | 4.88 | 18.16 | 4.88 | 20.02 | 9.77 | 22.56 | 9.67 | 11.33 | 9.77 | 9.86 | 1.56 | 0.1 | 5.96 | 0.1 | 17.68 | 0 | 25.98 | 4.88 | 18.75 | 4.59 | 12.21 | 4.79 | 9.77 | 9.96 | 9.86 | 10.06 | 9.86 | 10.16
63.904 | 11.43 | 0.1 | 4.1 | 0.1 | 0.1 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.41 | 9.77 | 22.85 | 9.67 | 12.4 | 9.77 | 9.86 | 1.37 | 0 | 1.76 | 0.1 | 9.28 | 0 | 26.47 | 4.79 | 19.34 | 4.59 | 13.09 | 4.79 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86
63.936 | 11.91 | 0 | 4.69 | 0 | 0.2 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.8 | 9.77 | 23.15 | 9.67 | 13.57 | 9.67 | 9.96 | 1.37 | 0 | 2.34 | 0.1 | 2.54 | 0 | 23.34 | 4.79 | 20.02 | 4.59 | 13.87 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
63.968 | 12.5 | 0.1 | 5.27 | 0.1 | 1.17 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.38 | 9.86 | 23.63 | 9.77 | 14.36 | 9.67 | 9.96 | 1.17 | 0.1 | 2.15 | 0.1 | 2.44 | 0 | 22.36 | 4.79 | 20.7 | 4.59 | 14.75 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
64 | 13.09 | 0.1 | 5.86 | 0 | 2.15 | 0.1 | 4.98 | 4.79 | 4.98 | 17.87 | 4.88 | 17.58 | 9.86 | 24.02 | 9.77 | 14.94 | 9.67 | 9.96 | 1.17 | 0.1 | 1.66 | 0.1 | 2.05 | 0 | 21.97 | 4.88 | 21.19 | 4.59 | 15.43 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
64.032 | 13.57 | 0.1 | 6.45 | 0.1 | 3.61 | 0 | 5.66 | 4.79 | 4.88 | 5.18 | 4.88 | 17.77 | 9.86 | 24.41 | 9.77 | 15.43 | 9.67 | 9.86 | 0.88 | 0.1 | 1.66 | 0 | 1.76 | 0 | 21.78 | 4.88 | 21.78 | 4.69 | 16.11 | 4.59 | 9.96 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
64.064 | 14.16 | 0.1 | 7.13 | 0 | 4.3 | 0 | 6.74 | 4.69 | 4.88 | 2.93 | 4.88 | 17.97 | 9.77 | 22.27 | 9.77 | 15.92 | 9.67 | 9.86 | 0.88 | 0 | 1.46 | 0 | 1.76 | 0 | 21.68 | 4.88 | 22.46 | 4.69 | 16.8 | 4.59 | 11.04 | 9.67 | 9.77 | 9.77 | 9.96 | 9.77
64.096 | 14.65 | 0 | 7.62 | 0.1 | 4.98 | 0 | 8.11 | 4.69 | 4.88 | 3.61 | 4.88 | 17.97 | 9.77 | 22.27 | 9.77 | 16.5 | 9.57 | 9.96 | 0.88 | 0 | 1.46 | 0.1 | 1.46 | 0 | 21.58 | 4.88 | 23.05 | 4.69 | 17.58 | 4.59 | 12.4 | 9.67 | 9.86 | 9.67 | 10.84 | 9.77
64.128 | 15.14 | 0.1 | 8.3 | 0.1 | 5.47 | 0.1 | 9.08 | 4.59 | 4.88 | 4.69 | 4.88 | 18.07 | 9.86 | 21.78 | 9.77 | 16.9 | 9.67 | 9.96 | 0.78 | 0.1 | 1.27 | 0.1 | 1.56 | 0 | 21.49 | 4.88 | 23.54 | 4.69 | 18.16 | 4.59 | 14.45 | 9.67 | 10.74 | 9.77 | 12.21 | 9.77
64.16 | 15.63 | 0.1 | 8.89 | 0.1 | 6.05 | 0.1 | 9.86 | 4.59 | 4.88 | 5.66 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 17.38 | 9.57 | 9.96 | 0.78 | 0.1 | 1.27 | 0.1 | 1.27 | 0 | 21.49 | 4.88 | 24.02 | 4.69 | 18.75 | 4.59 | 15.92 | 9.67 | 12.11 | 9.67 | 14.16 | 9.67
64.192 | 16.11 | 0.1 | 9.57 | 0.1 | 6.64 | 0.1 | 10.45 | 4.49 | 4.88 | 5.86 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 17.87 | 9.67 | 9.86 | 0.59 | 0.1 | 1.07 | 0.1 | 1.17 | 0 | 21.39 | 4.88 | 24.51 | 4.79 | 19.43 | 4.59 | 16.99 | 9.67 | 13.97 | 9.67 | 15.82 | 9.67
64.224 | 16.7 | 0 | 10.16 | 0 | 7.13 | 0 | 11.13 | 4.49 | 4.88 | 5.37 | 4.88 | 18.07 | 9.77 | 21.58 | 9.77 | 18.46 | 9.67 | 9.86 | 0.49 | 0 | 1.07 | 0 | 1.17 | 0 | 21.39 | 4.88 | 25 | 4.79 | 20.12 | 4.49 | 17.87 | 9.67 | 15.63 | 9.67 | 16.8 | 9.67
64.256 | 17.09 | 0 | 10.55 | 0 | 7.81 | 0 | 11.72 | 4.49 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.75 | 9.57 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 0 | 21.39 | 4.88 | 25.49 | 4.79 | 20.7 | 4.59 | 18.46 | 9.57 | 16.7 | 9.67 | 17.58 | 9.67
64.288 | 13.77 | 0.1 | 11.23 | 0.1 | 8.4 | 0.1 | 12.21 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.24 | 9.67 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 26.08 | 4.69 | 21.29 | 4.59 | 19.24 | 9.67 | 17.58 | 9.67 | 18.46 | 9.67
64.32 | 13.97 | 0.1 | 11.82 | 0.1 | 8.98 | 0.1 | 12.89 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.63 | 9.67 | 9.96 | 0.39 | 0.1 | 0.68 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 21.97 | 4.79 | 21.88 | 4.59 | 19.82 | 9.67 | 18.36 | 9.67 | 18.95 | 9.67
64.352 | 14.16 | 0.1 | 12.4 | 0.1 | 9.67 | 0.1 | 13.48 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 20.12 | 9.67 | 10.06 | 0.29 | 0 | 0.68 | 0 | 0.78 | 0 | 11.72 | 4.88 | 20.31 | 4.79 | 22.46 | 4.59 | 20.41 | 9.67 | 18.95 | 9.67 | 19.73 | 9.67
64.384 | 14.45 | 0 | 12.89 | 0 | 10.16 | 0 | 13.97 | 4.49 | 4.88 | 4.88 | 4.88 | 17.77 | 9.86 | 21.39 | 9.86 | 20.51 | 9.67 | 10.94 | 0.29 | 0 | 0.49 | 0 | 0.68 | 0 | 2.93 | 4.88 | 20.9 | 4.79 | 22.95 | 4.59 | 21.09 | 9.57 | 19.53 | 9.67 | 20.31 | 9.67
64.416 | 14.36 | 0 | 13.48 | 0 | 10.74 | 0 | 14.45 | 4.49 | 4.88 | 4.69 | 4.88 | 5.76 | 9.77 | 21.39 | 9.77 | 20.9 | 9.67 | 11.72 | 0.2 | 0 | 0.59 | 0.1 | 0.59 | 0 | 3.22 | 4.88 | 21.19 | 4.79 | 23.44 | 4.69 | 21.78 | 9.57 | 20.12 | 9.67 | 20.9 | 9.57
64.448 | 14.45 | 0.1 | 13.97 | 0.1 | 11.43 | 0.1 | 14.94 | 4.59 | 4.88 | 4.69 | 4.88 | 2.73 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 13.09 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 3.61 | 4.88 | 21.19 | 4.79 | 24.02 | 4.69 | 22.27 | 9.67 | 20.9 | 9.67 | 21.49 | 9.67
64.48 | 14.55 | 0.1 | 14.55 | 0 | 12.01 | 0.1 | 15.43 | 4.59 | 4.88 | 4.69 | 4.88 | 3.52 | 9.86 | 16.9 | 9.86 | 21.78 | 9.67 | 13.97 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0 | 4.49 | 4.88 | 21.19 | 4.79 | 24.61 | 4.69 | 22.85 | 9.67 | 21.49 | 9.67 | 22.17 | 9.67
64.512 | 14.45 | 0.1 | 15.04 | 0.1 | 12.5 | 0.1 | 16.02 | 4.59 | 4.88 | 4.69 | 4.88 | 4.59 | 9.86 | 7.72 | 9.86 | 22.07 | 9.77 | 14.65 | 0.2 | 0 | 0.39 | 0.1 | 0.49 | 0.1 | 5.96 | 4.88 | 21.19 | 4.79 | 25.1 | 4.69 | 23.34 | 9.67 | 22.07 | 9.67 | 22.66 | 9.67
64.544 | 14.45 | 0 | 15.53 | 0 | 12.99 | 0.1 | 16.5 | 4.59 | 4.88 | 4.59 | 4.88 | 5.66 | 9.86 | 7.81 | 9.86 | 22.46 | 9.67 | 15.33 | 0.1 | 0 | 0.29 | 0.1 | 0.29 | 0 | 6.05 | 4.88 | 21.29 | 4.79 | 25.59 | 4.69 | 23.93 | 9.67 | 22.66 | 9.67 | 23.15 | 9.67
64.576 | 14.45 | 0 | 15.92 | 0 | 13.57 | 0 | 16.9 | 4.59 | 4.88 | 4.59 | 4.88 | 5.86 | 9.77 | 10.55 | 9.86 | 22.95 | 9.77 | 15.72 | 0.1 | 0 | 0.29 | 1.07 | 0.29 | 0 | 5.76 | 4.88 | 21.19 | 4.79 | 26.17 | 4.79 | 24.41 | 9.67 | 23.24 | 9.67 | 23.83 | 9.67
64.608 | 14.45 | 0.1 | 16.5 | 0.1 | 14.06 | 0.1 | 17.38 | 4.69 | 4.88 | 4.59 | 4.88 | 5.47 | 9.86 | 11.33 | 9.86 | 23.34 | 9.77 | 16.21 | 0.1 | 0.2 | 0.29 | 2.34 | 0.29 | 0.1 | 5.37 | 4.88 | 21.29 | 4.79 | 21.78 | 4.79 | 24.9 | 9.67 | 23.73 | 9.67 | 24.22 | 9.67
64.64 | 14.45 | 0.1 | 16.9 | 0.1 | 14.65 | 0.1 | 17.87 | 4.69 | 5.27 | 4.59 | 4.88 | 5.18 | 9.86 | 10.45 | 9.86 | 23.63 | 9.77 | 16.8 | 0.1 | 1.46 | 0.2 | 4.2 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.79 | 20.22 | 4.79 | 25.49 | 9.67 | 24.22 | 9.67 | 24.81 | 9.67
64.672 | 14.45 | 0.1 | 16.6 | 0.1 | 15.14 | 0.1 | 18.36 | 4.69 | 6.05 | 4.59 | 4.88 | 4.98 | 9.86 | 9.86 | 9.86 | 24.02 | 9.77 | 17.29 | 0.1 | 2.64 | 0.2 | 5.66 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.79 | 20.9 | 4.79 | 25.88 | 9.67 | 24.81 | 9.67 | 25.29 | 9.67
64.704 | 14.16 | 0.1 | 14.16 | 0 | 15.63 | 0.1 | 18.75 | 4.69 | 7.03 | 4.49 | 4.88 | 4.88 | 9.86 | 9.96 | 9.86 | 24.41 | 9.77 | 17.68 | 0.1 | 4.59 | 0.1 | 6.84 | 0.1 | 0.49 | 5.18 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 26.47 | 9.67 | 25.2 | 9.67 | 25.78 | 9.67
64.736 | 4.88 | 0.1 | 14.65 | 0 | 16.21 | 0 | 19.24 | 4.69 | 8.5 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 22.46 | 9.77 | 18.16 | 0 | 5.86 | 0.2 | 7.52 | 0.1 | 1.46 | 5.08 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 26.95 | 9.67 | 25.78 | 9.67 | 26.37 | 9.67
64.768 | 1.86 | 0.1 | 14.75 | 0.1 | 16.6 | 0.1 | 19.63 | 4.79 | 9.38 | 4.49 | 4.88 | 4.88 | 9.86 | 10.25 | 9.86 | 22.27 | 9.86 | 18.56 | 0.1 | 6.84 | 0.2 | 8.4 | 0.2 | 3.03 | 5.08 | 4.88 | 6.84 | 4.79 | 21.19 | 4.79 | 27.34 | 9.77 | 26.27 | 9.77 | 26.86 | 9.77
64.8 | 1.95 | 0.1 | 14.55 | 0.1 | 17.19 | 0.1 | 20.12 | 4.79 | 10.06 | 4.59 | 4.88 | 4.79 | 9.86 | 10.06 | 9.86 | 21.88 | 9.86 | 19.04 | 0.1 | 7.62 | 0.1 | 9.28 | 0.1 | 4.98 | 4.98 | 4.88 | 2.44 | 4.79 | 21.29 | 4.79 | 27.83 | 9.77 | 26.76 | 9.77 | 27.34 | 9.77
64.832 | 1.66 | 0.1 | 14.55 | 0.1 | 13.87 | 0.1 | 20.51 | 4.79 | 10.64 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.68 | 9.86 | 19.43 | 0.1 | 8.4 | 0.1 | 9.96 | 0.1 | 6.35 | 4.98 | 4.88 | 3.42 | 4.79 | 21.29 | 4.79 | 28.32 | 9.77 | 27.25 | 9.77 | 27.74 | 9.77
64.864 | 1.46 | 0 | 14.45 | 0 | 13.97 | 0 | 20.9 | 4.79 | 11.23 | 4.49 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.68 | 9.86 | 19.92 | 0.1 | 9.28 | 0.1 | 10.74 | 0.1 | 7.13 | 4.79 | 4.88 | 4.2 | 4.79 | 21.29 | 4.79 | 28.81 | 9.77 | 27.83 | 9.77 | 28.13 | 9.67
64.896 | 1.27 | 0 | 14.45 | 0 | 14.16 | 0 | 17.58 | 4.79 | 11.91 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.31 | 0 | 10.16 | 0.1 | 11.62 | 0.1 | 8.01 | 4.79 | 4.79 | 5.47 | 4.79 | 21.29 | 4.79 | 29.3 | 9.77 | 28.13 | 9.77 | 28.71 | 9.77
64.928 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 17.68 | 4.79 | 12.4 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 20.8 | 0.1 | 10.84 | 0.1 | 12.31 | 0.1 | 8.69 | 4.79 | 4.88 | 6.35 | 4.79 | 21.19 | 4.79 | 24.71 | 9.77 | 28.61 | 9.77 | 29.1 | 9.77
64.96 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 17.97 | 4.79 | 12.99 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.19 | 0.1 | 11.52 | 0.1 | 13.09 | 0.1 | 9.47 | 4.79 | 4.88 | 5.96 | 4.79 | 21.29 | 4.79 | 23.44 | 9.77 | 29.1 | 9.77 | 27.34 | 9.77
64.992 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.16 | 4.79 | 13.48 | 4.69 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.39 | 9.86 | 21.58 | 0.1 | 12.4 | 0.1 | 13.87 | 0.1 | 10.25 | 4.69 | 4.88 | 5.57 | 4.79 | 21.19 | 4.79 | 24.02 | 9.77 | 28.42 | 9.77 | 23.63 | 9.77
65.024 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 0 | 18.07 | 4.79 | 14.06 | 4.69 | 4.98 | 4.49 | 9.86 | 9.77 | 9.86 | 21.39 | 9.86 | 21.97 | 0.1 | 13.18 | 0.1 | 14.45 | 0.1 | 11.13 | 4.59 | 4.88 | 5.27 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 23.93 | 9.77 | 24.32 | 9.77
65.056 | 0.78 | 0 | 14.36 | 0 | 14.45 | 0 | 18.16 | 4.79 | 14.55 | 4.69 | 5.76 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.36 | 0.1 | 13.87 | 0.1 | 15.33 | 0.1 | 11.82 | 4.59 | 4.88 | 5.18 | 4.79 | 6.54 | 4.88 | 24.32 | 9.77 | 24.51 | 9.77 | 24.41 | 9.77
65.088 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.04 | 4.79 | 6.74 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.96 | 22.75 | 0.1 | 14.55 | 0.1 | 16.02 | 0.1 | 12.7 | 4.59 | 4.88 | 5.08 | 4.79 | 2.54 | 4.88 | 24.22 | 9.77 | 24.51 | 9.86 | 24.32 | 9.86
65.12 | 0.68 | 0.1 | 7.13 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.53 | 4.69 | 8.2 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.96 | 23.15 | 0.1 | 15.33 | 0.1 | 16.6 | 0.1 | 13.48 | 4.59 | 5.76 | 4.98 | 4.79 | 3.42 | 4.88 | 24.32 | 9.77 | 24.41 | 9.86 | 24.22 | 9.86
65.152 | 0.59 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.02 | 4.79 | 9.18 | 4.49 | 9.86 | 9.67 | 9.86 | 13.87 | 9.96 | 23.54 | 0.1 | 16.02 | 0.1 | 17.29 | 0.1 | 14.16 | 4.59 | 6.54 | 5.08 | 4.79 | 4.2 | 4.79 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
65.184 | 0.49 | 0 | 2.25 | 0 | 14.45 | 0 | 18.16 | 4.79 | 16.5 | 4.79 | 9.86 | 4.49 | 9.86 | 9.67 | 9.86 | 7.23 | 9.86 | 23.83 | 0.1 | 16.8 | 0.1 | 17.97 | 0 | 14.94 | 4.49 | 8.4 | 4.98 | 4.79 | 5.57 | 4.88 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
65.216 | 0.49 | 0 | 1.66 | 0 | 14.45 | 0 | 18.16 | 4.79 | 16.99 | 4.79 | 10.45 | 4.49 | 10.55 | 9.67 | 9.86 | 8.01 | 9.86 | 24.22 | 0 | 17.38 | 0.1 | 18.65 | 0.1 | 15.63 | 4.59 | 10.25 | 4.88 | 4.79 | 6.35 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
65.248 | 0.49 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 17.48 | 4.79 | 11.13 | 4.49 | 11.33 | 9.67 | 9.86 | 10.74 | 9.96 | 21 | 0.1 | 17.97 | 0.1 | 19.24 | 0.1 | 16.31 | 4.59 | 11.52 | 4.79 | 4.79 | 5.96 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
65.28 | 0.39 | 0.1 | 1.37 | 0.1 | 5.37 | 0.1 | 18.16 | 4.88 | 17.97 | 4.79 | 11.72 | 4.59 | 12.4 | 9.57 | 9.86 | 11.23 | 9.96 | 20.9 | 0.1 | 18.65 | 0.1 | 19.92 | 0.1 | 16.9 | 4.59 | 12.6 | 4.88 | 4.79 | 5.47 | 4.88 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
65.312 | 0.29 | 0.1 | 1.27 | 0.1 | 1.86 | 0.1 | 18.07 | 4.88 | 18.36 | 4.79 | 12.4 | 4.59 | 13.48 | 9.67 | 9.86 | 10.45 | 9.96 | 21.09 | 0.1 | 19.24 | 0.1 | 20.61 | 0.1 | 17.58 | 4.59 | 13.48 | 4.79 | 4.79 | 5.18 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
65.344 | 0.2 | 0 | 1.07 | 0 | 1.95 | 0 | 6.35 | 4.79 | 18.75 | 4.79 | 12.89 | 4.59 | 14.36 | 9.67 | 9.86 | 9.96 | 9.96 | 21.39 | 0.1 | 19.92 | 0.1 | 21.09 | 0.1 | 18.26 | 4.59 | 14.26 | 4.69 | 4.79 | 5.18 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12 | 9.77
65.376 | 0.2 | 0 | 0.98 | 0 | 1.76 | 0 | 2.64 | 4.88 | 19.24 | 4.79 | 13.38 | 4.59 | 14.94 | 9.57 | 9.86 | 9.86 | 9.86 | 21.39 | 0 | 20.41 | 0.1 | 21.68 | 0.1 | 18.85 | 4.59 | 14.84 | 4.69 | 4.79 | 5.18 | 4.88 | 10.25 | 9.77 | 24.22 | 9.77 | 24.12 | 9.77
65.408 | 0.2 | 0.1 | 0.98 | 0.1 | 1.56 | 0.1 | 3.42 | 4.88 | 19.73 | 4.88 | 13.97 | 4.69 | 15.53 | 9.67 | 9.86 | 9.96 | 9.96 | 21.39 | 0.1 | 20.9 | 0.1 | 22.27 | 0.1 | 19.43 | 4.69 | 15.63 | 4.59 | 4.79 | 5.08 | 4.79 | 5.86 | 9.77 | 24.22 | 9.86 | 11.91 | 9.86
65.44 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 4.39 | 4.88 | 20.12 | 4.88 | 14.45 | 4.69 | 16.02 | 9.67 | 9.86 | 10.16 | 9.96 | 21.49 | 0.1 | 21.58 | 0.1 | 21.68 | 0.1 | 20.12 | 4.69 | 16.41 | 4.59 | 4.79 | 5.08 | 4.88 | 9.47 | 9.77 | 12.5 | 9.86 | 5.76 | 9.86
65.472 | 0.1 | 0.2 | 0.78 | 0.1 | 1.17 | 0.1 | 5.37 | 4.88 | 20.61 | 4.88 | 14.94 | 4.69 | 16.5 | 9.67 | 10.45 | 9.86 | 9.96 | 21.49 | 0.1 | 22.17 | 0.1 | 17.29 | 0.1 | 20.8 | 4.69 | 16.99 | 4.59 | 4.79 | 4.98 | 4.79 | 12.31 | 9.86 | 5.96 | 9.86 | 8.59 | 9.86
65.504 | 0.1 | 1.07 | 0.59 | 0 | 1.17 | 0.1 | 6.05 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 16.99 | 9.67 | 11.23 | 9.96 | 9.96 | 21.39 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 21.29 | 4.69 | 17.68 | 4.59 | 4.79 | 4.88 | 4.79 | 11.13 | 9.77 | 8.4 | 9.86 | 12.11 | 9.77
65.536 | 0.1 | 5.47 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.26 | 4.88 | 17.87 | 4.79 | 19.34 | 9.67 | 15.43 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 20.9 | 4.59 | 9.18 | 4.69 | 4.88 | 10.35 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86
65.568 | 0.1 | 6.05 | 0.29 | 0 | 0.59 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 18.26 | 4.79 | 19.82 | 9.67 | 15.92 | 9.67 | 9.96 | 10.25 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 21.58 | 4.59 | 10.74 | 4.69 | 4.88 | 10.16 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86
65.6 | 0.1 | 6.64 | 0.29 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 18.75 | 4.79 | 20.22 | 9.77 | 16.5 | 9.67 | 9.96 | 6.74 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 22.17 | 4.59 | 12.01 | 4.59 | 4.79 | 9.96 | 9.77 | 10.45 | 9.86 | 10.16 | 9.86
65.632 | 0.1 | 7.32 | 0.2 | 0 | 0.49 | 0 | 4.79 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 20.61 | 9.67 | 16.99 | 9.67 | 9.96 | 9.77 | 0 | 17.58 | 0.1 | 17.58 | 0.1 | 17.87 | 4.88 | 22.66 | 4.59 | 12.89 | 4.59 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
65.664 | 0.1 | 7.81 | 0.1 | 0 | 0.49 | 0 | 4.69 | 4.88 | 18.16 | 4.79 | 19.53 | 4.79 | 21 | 9.77 | 17.48 | 9.67 | 9.96 | 11.52 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.15 | 4.59 | 13.67 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.77 | 10.06 | 9.77
65.696 | 0.1 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 20.02 | 4.88 | 21.58 | 9.77 | 17.97 | 9.67 | 9.96 | 10.94 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.73 | 4.69 | 14.55 | 4.59 | 4.88 | 10.06 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
65.728 | 0.1 | 8.98 | 0.1 | 0 | 0.29 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 20.51 | 4.88 | 21.88 | 9.77 | 18.46 | 9.57 | 9.96 | 10.16 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 17.68 | 4.88 | 24.32 | 4.69 | 15.23 | 4.59 | 4.79 | 9.86 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
65.76 | 0.1 | 9.67 | 0.1 | 0.68 | 0.2 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 20.9 | 4.88 | 22.27 | 9.77 | 18.75 | 9.57 | 9.96 | 9.86 | 0.1 | 17.58 | 0.1 | 7.32 | 0.1 | 17.77 | 4.88 | 24.81 | 4.69 | 15.92 | 4.59 | 4.88 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
65.792 | 0.1 | 10.25 | 0 | 1.37 | 0.2 | 0.68 | 4.49 | 4.88 | 18.07 | 4.88 | 17.48 | 4.79 | 22.66 | 9.77 | 19.24 | 9.57 | 9.96 | 10.06 | 0 | 17.29 | 0.1 | 1.66 | 0.1 | 17.68 | 4.88 | 25.29 | 4.69 | 16.6 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.77 | 9.86 | 9.77
65.824 | 0.1 | 10.74 | 0 | 2.54 | 0.2 | 1.27 | 4.59 | 4.88 | 15.33 | 4.79 | 17.58 | 4.79 | 22.95 | 9.77 | 19.73 | 9.57 | 9.86 | 10.16 | 0 | 4.98 | 0.1 | 2.44 | 0.1 | 17.68 | 4.88 | 25.78 | 4.69 | 17.38 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
65.856 | 0.1 | 11.43 | 0.1 | 3.81 | 0.1 | 2.54 | 4.59 | 4.88 | 4.49 | 4.88 | 17.77 | 4.88 | 23.44 | 9.77 | 20.12 | 9.57 | 9.96 | 10.06 | 0.1 | 1.76 | 0.1 | 2.05 | 0.1 | 17.68 | 4.88 | 26.37 | 4.79 | 17.97 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
65.888 | 0.1 | 11.91 | 0.1 | 4.59 | 0.1 | 3.81 | 4.49 | 4.88 | 3.13 | 4.88 | 18.07 | 4.88 | 23.73 | 9.77 | 20.61 | 9.57 | 10.16 | 9.96 | 0.1 | 2.15 | 0.1 | 1.66 | 0.1 | 15.14 | 4.88 | 23.15 | 4.79 | 18.65 | 4.59 | 4.79 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
65.92 | 0.1 | 12.6 | 0.1 | 5.08 | 0.1 | 4.49 | 4.59 | 4.88 | 3.81 | 4.88 | 18.07 | 4.88 | 24.12 | 9.77 | 20.9 | 9.67 | 10.84 | 9.86 | 0.1 | 1.86 | 0.1 | 1.66 | 0.1 | 3.71 | 4.88 | 22.27 | 4.79 | 19.34 | 4.59 | 5.76 | 9.67 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
65.952 | 0.1 | 13.09 | 0.1 | 5.66 | 0.1 | 5.08 | 4.49 | 4.88 | 4.69 | 4.88 | 18.07 | 4.79 | 23.73 | 9.77 | 21.29 | 9.57 | 11.91 | 9.86 | 0 | 1.66 | 0.1 | 1.46 | 0.1 | 2.25 | 4.88 | 21.78 | 4.79 | 19.92 | 4.59 | 7.23 | 9.67 | 10.74 | 9.77 | 9.86 | 9.67 | 9.77
65.984 | 0.1 | 13.57 | 0.1 | 6.25 | 0.1 | 5.57 | 4.49 | 4.88 | 5.76 | 4.79 | 18.16 | 4.79 | 22.27 | 9.77 | 21.78 | 9.57 | 13.18 | 9.77 | 0 | 1.56 | 0.1 | 1.37 | 0.1 | 2.15 | 4.88 | 21.58 | 4.79 | 20.51 | 4.59 | 8.98 | 9.67 | 11.91 | 9.77 | 10.94 | 9.67 | 9.96
66.016 | 0.1 | 14.06 | 0.1 | 6.84 | 0.1 | 6.15 | 4.59 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 22.17 | 9.77 | 22.17 | 9.67 | 14.06 | 9.77 | 0.1 | 1.46 | 0.1 | 1.27 | 0.1 | 1.86 | 4.88 | 21.58 | 4.79 | 21.19 | 4.69 | 10.74 | 9.67 | 13.67 | 9.77 | 12.01 | 9.67 | 10.84
66.048 | 0.1 | 14.65 | 0.1 | 7.42 | 0.1 | 6.74 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 21.88 | 9.86 | 22.66 | 9.67 | 14.75 | 9.77 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.76 | 4.88 | 21.49 | 4.79 | 21.68 | 4.69 | 12.01 | 9.67 | 15.33 | 9.67 | 13.87 | 9.67 | 12.21
66.08 | 0.1 | 15.14 | 0.1 | 8.01 | 0.1 | 7.42 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.78 | 9.86 | 22.95 | 9.67 | 15.33 | 9.77 | 0.1 | 1.27 | 0.1 | 1.17 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 22.27 | 4.69 | 12.89 | 9.67 | 16.6 | 9.67 | 15.43 | 9.67 | 14.16
66.112 | 0.1 | 15.63 | 0.1 | 8.59 | 0.1 | 8.01 | 4.59 | 4.79 | 4.98 | 4.88 | 18.07 | 4.88 | 21.68 | 9.86 | 23.34 | 9.67 | 15.82 | 9.67 | 0 | 1.07 | 0.1 | 0.88 | 0 | 1.37 | 4.88 | 21.39 | 4.79 | 22.75 | 4.79 | 13.77 | 9.67 | 17.38 | 9.67 | 16.7 | 9.67 | 16.02
66.144 | 0.1 | 16.21 | 0 | 9.28 | 0.1 | 8.59 | 4.59 | 4.79 | 4.98 | 4.88 | 18.07 | 4.79 | 21.58 | 9.77 | 23.73 | 9.67 | 16.31 | 9.67 | 0 | 0.98 | 0.1 | 0.88 | 0 | 1.37 | 4.98 | 21.29 | 4.79 | 23.44 | 4.79 | 14.55 | 9.57 | 18.16 | 9.67 | 17.58 | 9.67 | 16.8
66.176 | 0.1 | 16.6 | 0.1 | 9.86 | 0.1 | 9.28 | 4.69 | 5.57 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 24.02 | 9.67 | 16.8 | 9.67 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 1.27 | 4.98 | 21.29 | 4.79 | 23.93 | 4.79 | 15.23 | 9.67 | 18.75 | 9.67 | 18.36 | 9.67 | 17.68
66.208 | 0.1 | 17.19 | 0.1 | 10.45 | 0.1 | 9.77 | 4.69 | 6.25 | 4.98 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 24.41 | 9.67 | 17.38 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.41 | 4.79 | 15.92 | 9.67 | 19.53 | 9.67 | 18.95 | 9.67 | 18.46
66.24 | 0.1 | 13.67 | 0.1 | 11.04 | 0.1 | 10.45 | 4.69 | 7.42 | 4.88 | 4.88 | 5.86 | 4.88 | 21.49 | 9.86 | 22.46 | 9.67 | 17.77 | 9.67 | 0.78 | 0.78 | 0.1 | 0.68 | 0.1 | 0.88 | 4.88 | 21.19 | 4.79 | 24.9 | 4.79 | 16.6 | 9.67 | 20.22 | 9.57 | 19.73 | 9.67 | 19.14
66.272 | 0.1 | 14.36 | 0.1 | 11.62 | 0.1 | 11.04 | 4.69 | 8.69 | 4.69 | 4.88 | 2.93 | 4.88 | 21.49 | 9.77 | 22.07 | 9.67 | 18.16 | 9.57 | 1.66 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 21.19 | 4.79 | 25.49 | 4.79 | 17.29 | 9.67 | 20.9 | 9.57 | 20.22 | 9.67 | 19.82
66.304 | 0.1 | 14.26 | 0 | 12.21 | 0.1 | 11.52 | 4.69 | 9.47 | 4.69 | 4.79 | 3.42 | 4.79 | 21.49 | 9.77 | 21.88 | 9.67 | 18.65 | 9.57 | 3.22 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 13.38 | 4.79 | 25.88 | 4.79 | 17.97 | 9.67 | 21.39 | 9.67 | 20.9 | 9.67 | 20.31
66.336 | 0.1 | 14.45 | 0.1 | 12.7 | 0.1 | 12.21 | 4.79 | 10.16 | 4.69 | 4.88 | 4.59 | 4.88 | 21.49 | 9.86 | 21.68 | 9.77 | 19.14 | 9.67 | 5.08 | 0.49 | 0.1 | 0.49 | 0.1 | 0.68 | 4.98 | 3.13 | 4.88 | 24.12 | 4.79 | 18.65 | 9.67 | 22.07 | 9.67 | 21.49 | 9.67 | 20.9
66.368 | 0.1 | 14.36 | 0.1 | 13.18 | 0.1 | 12.7 | 4.79 | 10.74 | 4.59 | 4.88 | 5.66 | 4.88 | 21.39 | 9.86 | 21.78 | 9.77 | 19.53 | 9.57 | 6.35 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 4.88 | 3.22 | 4.79 | 20.9 | 4.79 | 19.24 | 9.67 | 22.56 | 9.67 | 21.97 | 9.67 | 21.49
66.4 | 0.1 | 14.45 | 0.1 | 13.77 | 0.1 | 13.28 | 4.79 | 11.43 | 4.59 | 4.88 | 5.96 | 4.88 | 13.28 | 9.86 | 21.58 | 9.77 | 20.02 | 9.57 | 7.23 | 0.39 | 0.2 | 0.29 | 0.1 | 0.59 | 4.88 | 3.42 | 4.79 | 21.09 | 4.79 | 20.02 | 9.67 | 23.15 | 9.67 | 22.56 | 9.67 | 22.07
66.432 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 13.77 | 4.79 | 12.01 | 4.59 | 4.88 | 5.37 | 4.79 | 6.93 | 9.77 | 21.49 | 9.77 | 20.41 | 9.57 | 8.01 | 0.29 | 1.46 | 0.29 | 0.1 | 0.49 | 4.88 | 4.59 | 4.79 | 21.29 | 4.79 | 20.51 | 9.67 | 23.63 | 9.67 | 23.15 | 9.67 | 22.66
66.464 | 0.1 | 14.45 | 0 | 14.84 | 0.1 | 14.26 | 4.79 | 12.5 | 4.59 | 4.79 | 5.18 | 4.88 | 8.11 | 9.77 | 21.49 | 9.77 | 20.8 | 9.57 | 8.69 | 0.29 | 2.54 | 0.29 | 1.07 | 0.49 | 4.88 | 5.96 | 4.79 | 21.29 | 4.79 | 21.09 | 9.67 | 24.12 | 9.67 | 23.54 | 9.67 | 23.15
66.496 | 0.1 | 14.45 | 0.1 | 15.33 | 0.1 | 14.84 | 4.88 | 13.09 | 4.59 | 4.88 | 4.98 | 4.88 | 11.13 | 9.86 | 21.49 | 9.77 | 21.19 | 9.67 | 9.47 | 0.29 | 4.39 | 0.2 | 2.05 | 0.49 | 4.88 | 6.15 | 4.88 | 21.29 | 4.79 | 21.68 | 9.77 | 24.71 | 9.67 | 24.12 | 9.67 | 23.73
66.528 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 15.33 | 4.88 | 13.67 | 4.59 | 4.88 | 4.98 | 4.88 | 11.23 | 9.86 | 21.49 | 9.77 | 21.68 | 9.67 | 10.35 | 0.2 | 5.86 | 0.2 | 3.91 | 0.29 | 4.88 | 5.86 | 4.79 | 21.29 | 4.88 | 22.27 | 9.77 | 25.1 | 9.67 | 24.71 | 9.67 | 24.32
66.56 | 0.1 | 14.45 | 0.1 | 16.31 | 0.1 | 15.82 | 4.88 | 14.16 | 4.49 | 5.18 | 4.88 | 4.88 | 10.45 | 9.86 | 21.49 | 9.77 | 22.07 | 9.67 | 11.13 | 0.2 | 6.84 | 0.2 | 5.47 | 0.29 | 4.88 | 5.37 | 4.88 | 21.19 | 4.88 | 22.85 | 9.77 | 25.68 | 9.67 | 25.2 | 9.67 | 24.71
66.592 | 0.1 | 14.45 | 0.1 | 16.8 | 0.1 | 16.31 | 4.88 | 14.75 | 4.59 | 6.05 | 4.79 | 4.79 | 9.86 | 9.77 | 21.39 | 9.77 | 22.36 | 9.67 | 11.82 | 0.1 | 7.72 | 0.1 | 6.64 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 23.34 | 9.77 | 26.27 | 9.67 | 25.68 | 9.67 | 25.29
66.624 | 0.1 | 14.36 | 0 | 17.29 | 0.1 | 16.8 | 4.88 | 15.14 | 4.59 | 6.93 | 4.88 | 4.88 | 9.96 | 9.77 | 21.39 | 9.77 | 22.75 | 9.67 | 12.6 | 0.1 | 8.4 | 0.1 | 7.42 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 23.83 | 9.77 | 26.76 | 9.77 | 26.08 | 9.67 | 25.78
66.656 | 0.1 | 14.26 | 0.1 | 15.33 | 0.1 | 17.29 | 4.88 | 15.72 | 4.59 | 8.4 | 4.79 | 4.88 | 10.25 | 9.77 | 14.06 | 9.77 | 23.24 | 9.67 | 13.28 | 0.1 | 9.38 | 0.1 | 8.2 | 0.2 | 4.88 | 5.18 | 4.79 | 21.19 | 4.88 | 24.51 | 9.77 | 27.25 | 9.77 | 26.66 | 9.77 | 26.37
66.688 | 0.1 | 4.39 | 0.1 | 15.23 | 0.1 | 14.94 | 4.88 | 16.21 | 4.59 | 9.38 | 4.69 | 4.88 | 10.16 | 9.77 | 7.23 | 9.77 | 23.63 | 9.67 | 14.16 | 0.1 | 10.06 | 0.1 | 9.08 | 0.1 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 24.9 | 9.77 | 27.74 | 9.77 | 27.15 | 9.77 | 26.76
66.72 | 0.1 | 1.76 | 0.1 | 14.84 | 0.1 | 15.23 | 4.88 | 16.7 | 4.59 | 9.96 | 4.69 | 4.88 | 9.96 | 9.77 | 8.01 | 9.77 | 23.93 | 9.67 | 14.84 | 0.1 | 10.94 | 0.1 | 9.96 | 0.1 | 4.88 | 4.98 | 4.79 | 21.09 | 4.88 | 25.39 | 9.77 | 28.22 | 9.77 | 27.64 | 9.77 | 27.34
66.752 | 0.1 | 1.86 | 0.1 | 14.65 | 0.1 | 14.84 | 4.88 | 17.19 | 4.59 | 10.64 | 4.59 | 4.79 | 9.96 | 9.77 | 11.04 | 9.77 | 24.32 | 9.67 | 15.63 | 0 | 11.72 | 0.1 | 10.74 | 0.1 | 4.88 | 4.98 | 4.79 | 21.09 | 4.88 | 25.98 | 9.77 | 28.61 | 9.77 | 28.13 | 9.77 | 27.74
66.784 | 0.1 | 1.66 | 0 | 14.65 | 0.1 | 14.65 | 4.88 | 17.58 | 4.59 | 11.13 | 4.59 | 4.79 | 9.96 | 9.77 | 11.23 | 9.77 | 20.9 | 9.67 | 16.31 | 0.1 | 12.5 | 0 | 11.52 | 0.1 | 4.88 | 4.88 | 4.79 | 7.81 | 4.88 | 25.1 | 9.77 | 29.1 | 9.77 | 28.61 | 9.77 | 28.13
66.816 | 0.1 | 1.56 | 0.1 | 14.55 | 0.1 | 14.65 | 4.88 | 18.07 | 4.69 | 11.82 | 4.59 | 4.88 | 9.96 | 9.77 | 10.55 | 9.86 | 21 | 9.67 | 16.8 | 0.1 | 13.18 | 0.1 | 12.4 | 0.1 | 4.88 | 4.79 | 4.88 | 2.54 | 4.88 | 21.68 | 9.77 | 29.49 | 9.77 | 29.01 | 9.77 | 28.71
66.848 | 0.1 | 1.27 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.46 | 4.69 | 12.4 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 21.19 | 9.67 | 17.48 | 0.1 | 13.97 | 0.1 | 13.09 | 0.1 | 4.88 | 4.79 | 4.88 | 3.22 | 4.88 | 21.58 | 9.77 | 25.88 | 9.77 | 29.01 | 9.77 | 29.1
66.88 | 0.1 | 1.27 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.95 | 4.79 | 12.99 | 4.59 | 4.88 | 9.77 | 9.86 | 9.96 | 9.77 | 21.49 | 9.67 | 18.16 | 0.1 | 14.65 | 0.1 | 13.77 | 0 | 4.88 | 4.79 | 4.79 | 4 | 4.88 | 21.49 | 9.77 | 25.1 | 9.77 | 25.49 | 9.86 | 26.56
66.912 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.43 | 4.79 | 13.48 | 4.49 | 4.88 | 9.77 | 9.77 | 10.25 | 9.77 | 21.39 | 9.67 | 18.75 | 0 | 15.33 | 0.1 | 14.45 | 0 | 5.96 | 4.69 | 4.79 | 5.08 | 4.88 | 21.29 | 9.77 | 24.61 | 9.77 | 24.61 | 9.77 | 23.24
66.944 | 0.1 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.82 | 4.79 | 13.97 | 4.49 | 5.08 | 9.67 | 9.77 | 10.16 | 9.77 | 21.49 | 9.67 | 19.43 | 0 | 16.02 | 0.1 | 15.14 | 0 | 6.93 | 4.59 | 4.79 | 6.25 | 4.79 | 21.39 | 9.77 | 24.61 | 9.77 | 24.61 | 9.86 | 24.12
66.976 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.22 | 4.79 | 14.55 | 4.49 | 5.86 | 9.77 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 20.12 | 0.1 | 16.7 | 0.1 | 15.92 | 0 | 8.59 | 4.59 | 4.79 | 5.96 | 4.88 | 21.29 | 9.77 | 24.51 | 9.77 | 24.51 | 9.86 | 24.22
67.008 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.7 | 4.79 | 15.04 | 4.49 | 6.93 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.51 | 0.1 | 17.58 | 0.1 | 16.7 | 0 | 10.45 | 4.59 | 4.79 | 5.57 | 4.88 | 21.29 | 9.77 | 24.41 | 9.77 | 24.41 | 9.86 | 24.22
67.04 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 21.09 | 4.88 | 15.53 | 4.49 | 8.2 | 9.67 | 10.45 | 9.86 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 18.07 | 0.1 | 17.29 | 0.1 | 11.91 | 4.59 | 4.79 | 5.18 | 4.88 | 21.29 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
67.072 | 0.1 | 0.59 | 0.1 | 9.67 | 0.1 | 14.36 | 4.88 | 19.04 | 4.79 | 16.11 | 4.59 | 9.28 | 9.67 | 11.23 | 9.86 | 9.86 | 21.49 | 9.77 | 21.78 | 0 | 18.75 | 0.1 | 17.97 | 0 | 12.89 | 4.59 | 4.79 | 5.08 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
67.104 | 0.1 | 0.49 | 0 | 2.93 | 0.1 | 11.04 | 4.88 | 18.95 | 4.88 | 16.5 | 4.59 | 9.86 | 9.67 | 12.21 | 9.77 | 9.86 | 21.49 | 9.77 | 22.27 | 0 | 19.34 | 0.1 | 18.56 | 0 | 13.67 | 4.49 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
67.136 | 0.1 | 0.49 | 0.1 | 2.05 | 0.1 | 3.03 | 4.88 | 18.56 | 4.88 | 16.99 | 4.59 | 10.45 | 9.67 | 13.48 | 9.77 | 9.86 | 21.49 | 9.77 | 18.46 | 0.1 | 20.02 | 0.1 | 19.34 | 0.1 | 14.45 | 4.49 | 4.79 | 5.08 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
67.168 | 0.1 | 0.49 | 0.1 | 1.66 | 0.1 | 1.95 | 4.88 | 18.46 | 4.88 | 17.48 | 4.59 | 11.13 | 9.67 | 14.26 | 9.77 | 9.86 | 21.49 | 9.77 | 17.97 | 0.1 | 20.61 | 0.1 | 19.82 | 0.1 | 15.14 | 4.49 | 4.79 | 5.08 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
67.2 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.86 | 4.88 | 18.46 | 4.88 | 17.87 | 4.69 | 11.72 | 9.67 | 14.84 | 9.77 | 9.86 | 21.49 | 9.77 | 17.68 | 0.1 | 21.19 | 0.1 | 20.51 | 0 | 15.92 | 4.59 | 4.79 | 4.88 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
67.232 | 0.1 | 0.29 | 0.1 | 1.37 | 0.1 | 1.56 | 4.88 | 18.26 | 4.88 | 18.46 | 4.69 | 12.31 | 9.67 | 15.33 | 9.67 | 9.86 | 10.25 | 9.77 | 17.68 | 0 | 21.68 | 0 | 21 | 0 | 16.5 | 4.59 | 4.79 | 4.88 | 4.88 | 7.72 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.12
67.264 | 0.2 | 0.2 | 0.1 | 1.17 | 0.1 | 1.56 | 4.88 | 18.16 | 4.88 | 18.75 | 4.69 | 12.79 | 9.67 | 15.82 | 9.67 | 9.86 | 6.64 | 9.77 | 17.68 | 0.1 | 22.27 | 0 | 21.68 | 0 | 17.19 | 4.59 | 4.79 | 4.88 | 4.79 | 2.44 | 9.77 | 13.57 | 9.86 | 24.12 | 9.86 | 24.12
67.296 | 1.17 | 0.2 | 0.1 | 1.17 | 0.1 | 1.17 | 4.88 | 18.26 | 4.88 | 19.24 | 4.69 | 13.48 | 9.67 | 16.41 | 9.67 | 9.86 | 9.38 | 9.86 | 17.58 | 0.1 | 20.12 | 0.1 | 22.27 | 0.1 | 17.77 | 4.59 | 4.88 | 4.79 | 4.88 | 3.22 | 9.77 | 6.05 | 9.86 | 13.18 | 9.86 | 24.12
67.328 | 2.15 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.73 | 4.79 | 13.97 | 9.67 | 16.9 | 9.67 | 9.86 | 11.72 | 9.86 | 17.58 | 0.1 | 17.09 | 0.1 | 22.85 | 0 | 18.46 | 4.59 | 5.96 | 4.69 | 4.88 | 3.81 | 9.77 | 7.81 | 9.86 | 5.96 | 9.86 | 11.82
67.36 | 3.52 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 14.36 | 9.67 | 17.38 | 9.67 | 9.86 | 11.04 | 9.86 | 17.58 | 0.1 | 17.58 | 0.1 | 19.34 | 0 | 19.14 | 4.59 | 7.42 | 4.69 | 4.88 | 5.08 | 9.77 | 11.52 | 9.86 | 7.72 | 9.86 | 5.66
67.392 | 4.39 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 14.94 | 9.67 | 17.77 | 9.67 | 9.86 | 10.06 | 9.86 | 17.48 | 0 | 17.58 | 0 | 18.36 | 0 | 19.82 | 4.69 | 9.67 | 4.59 | 4.79 | 6.15 | 9.77 | 11.62 | 9.77 | 11.72 | 9.86 | 8.79
67.424 | 4.98 | 0 | 0.1 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 9.67 | 18.16 | 9.67 | 9.86 | 9.86 | 9.86 | 17.48 | 0 | 17.68 | 0.1 | 18.16 | 0 | 20.41 | 4.69 | 11.04 | 4.59 | 4.79 | 6.05 | 9.77 | 10.45 | 9.86 | 11.82 | 9.86 | 12.01
67.456 | 5.57 | 0.1 | 0.1 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.98 | 17.77 | 4.79 | 15.92 | 9.77 | 18.75 | 9.57 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 21.19 | 4.69 | 12.21 | 4.59 | 4.88 | 5.57 | 9.77 | 9.96 | 9.86 | 10.45 | 9.86 | 11.62
67.488 | 6.05 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 10.94 | 4.88 | 18.07 | 4.79 | 16.41 | 9.67 | 19.24 | 9.57 | 9.86 | 10.25 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 21.68 | 4.69 | 13.09 | 4.59 | 4.88 | 5.47 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.35
67.52 | 6.64 | 0.1 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 3.42 | 4.88 | 18.07 | 4.88 | 16.9 | 9.77 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 17.87 | 0 | 22.17 | 4.69 | 13.87 | 4.49 | 4.88 | 5.08 | 9.77 | 10.35 | 9.86 | 10.25 | 9.86 | 9.86
67.552 | 7.32 | 0.1 | 0.1 | 0.39 | 0.1 | 0.59 | 4.88 | 3.32 | 4.88 | 18.16 | 4.88 | 17.38 | 9.77 | 20.02 | 9.67 | 9.86 | 9.96 | 9.86 | 12.31 | 0 | 17.68 | 0.1 | 17.77 | 0 | 22.85 | 4.79 | 14.75 | 4.49 | 4.79 | 5.08 | 9.77 | 10.45 | 9.77 | 10.45 | 9.86 | 10.45
67.584 | 7.81 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 4.88 | 3.81 | 4.88 | 18.16 | 4.88 | 17.77 | 9.77 | 20.41 | 9.67 | 9.86 | 9.86 | 9.86 | 3.42 | 0 | 17.68 | 0 | 17.77 | 0 | 23.24 | 4.79 | 15.43 | 4.49 | 5.08 | 5.08 | 9.77 | 10.16 | 9.86 | 10.45 | 9.86 | 10.45
67.616 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 4.88 | 4.88 | 18.16 | 4.88 | 18.26 | 9.77 | 20.9 | 9.67 | 9.86 | 9.96 | 9.86 | 2.25 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 23.83 | 4.79 | 16.11 | 4.59 | 6.05 | 5.08 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.45
67.648 | 8.98 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 18.75 | 9.77 | 21.39 | 9.67 | 9.86 | 9.86 | 9.86 | 2.15 | 0.1 | 17.58 | 0.1 | 17.77 | 0 | 24.41 | 4.79 | 16.8 | 4.49 | 7.62 | 4.98 | 9.77 | 10.16 | 9.77 | 10.06 | 9.86 | 10.06
67.68 | 9.67 | 0.1 | 0.78 | 0.2 | 0.1 | 0.29 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 19.14 | 9.77 | 21.78 | 9.67 | 9.86 | 9.77 | 9.86 | 1.76 | 0 | 17.58 | 0.1 | 17.77 | 0 | 24.81 | 4.79 | 17.58 | 4.49 | 9.77 | 4.88 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 10.16
67.712 | 10.25 | 0.1 | 1.86 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 22.07 | 9.67 | 10.55 | 9.77 | 9.86 | 1.86 | 0 | 17.58 | 0 | 17.68 | 0 | 25.39 | 4.79 | 18.16 | 4.49 | 11.33 | 4.79 | 9.77 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16
67.744 | 10.74 | 0.1 | 3.03 | 0.1 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 20.02 | 9.77 | 22.46 | 9.67 | 11.23 | 9.77 | 9.86 | 1.46 | 0 | 6.84 | 0 | 17.68 | 0 | 25.88 | 4.79 | 18.85 | 4.59 | 12.31 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
67.776 | 11.43 | 0.1 | 4.1 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 9.86 | 22.85 | 9.67 | 12.31 | 9.77 | 9.86 | 1.56 | 0.1 | 1.66 | 0.1 | 9.28 | 0 | 26.47 | 4.88 | 19.43 | 4.59 | 13.18 | 4.79 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
67.808 | 12.01 | 0.1 | 4.79 | 0.1 | 0.2 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.8 | 9.86 | 23.24 | 9.67 | 13.48 | 9.67 | 9.96 | 1.27 | 0.1 | 2.34 | 0.1 | 2.25 | 0.1 | 23.34 | 4.88 | 20.02 | 4.59 | 14.06 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
67.84 | 12.5 | 0.1 | 5.27 | 0.1 | 1.17 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.38 | 9.86 | 23.63 | 9.67 | 14.36 | 9.67 | 9.86 | 1.37 | 0 | 2.05 | 0.1 | 2.34 | 0 | 22.36 | 4.79 | 20.7 | 4.59 | 14.84 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
67.872 | 13.09 | 0.1 | 5.86 | 0 | 2.15 | 0 | 4.98 | 4.79 | 4.88 | 15.53 | 4.88 | 17.68 | 9.77 | 24.02 | 9.67 | 14.84 | 9.67 | 9.86 | 0.98 | 0 | 1.66 | 0 | 2.05 | 0 | 21.88 | 4.79 | 21.29 | 4.59 | 15.53 | 4.59 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
67.904 | 13.57 | 0.1 | 6.35 | 0 | 3.52 | 0 | 5.66 | 4.69 | 4.88 | 4.59 | 4.88 | 17.77 | 9.77 | 24.32 | 9.77 | 15.43 | 9.67 | 9.86 | 0.88 | 0.1 | 1.56 | 0 | 1.66 | 0 | 21.68 | 4.79 | 21.88 | 4.69 | 16.21 | 4.59 | 9.96 | 9.67 | 9.77 | 9.77 | 9.86 | 9.77
67.936 | 14.16 | 0.1 | 7.03 | 0.1 | 4.39 | 0.1 | 6.74 | 4.69 | 4.88 | 3.03 | 4.88 | 18.07 | 9.86 | 21.97 | 9.77 | 15.92 | 9.67 | 9.96 | 0.88 | 0.1 | 1.37 | 0.1 | 1.76 | 0.1 | 21.68 | 4.88 | 22.46 | 4.69 | 16.9 | 4.59 | 11.04 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
67.968 | 14.65 | 0.1 | 7.62 | 0.1 | 4.98 | 0.1 | 8.2 | 4.69 | 4.88 | 3.71 | 4.88 | 17.97 | 9.86 | 22.27 | 9.77 | 16.5 | 9.67 | 9.96 | 0.78 | 0.1 | 1.66 | 0.1 | 1.56 | 0 | 21.58 | 4.88 | 22.95 | 4.69 | 17.58 | 4.59 | 12.5 | 9.77 | 9.86 | 9.77 | 10.74 | 9.77
68 | 15.14 | 0.1 | 8.3 | 0.1 | 5.57 | 0.1 | 9.08 | 4.59 | 4.88 | 4.69 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 16.9 | 9.67 | 9.96 | 0.78 | 0 | 1.37 | 0.1 | 1.46 | 0 | 21.39 | 4.88 | 23.44 | 4.69 | 18.16 | 4.59 | 14.55 | 9.67 | 10.74 | 9.67 | 12.11 | 9.77
68.032 | 15.63 | 0.1 | 8.89 | 0 | 6.05 | 0 | 9.96 | 4.49 | 4.88 | 5.76 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 17.38 | 9.57 | 9.86 | 0.68 | 0 | 1.17 | 0 | 1.37 | 0 | 21.39 | 4.88 | 24.02 | 4.69 | 18.85 | 4.59 | 16.21 | 9.67 | 12.11 | 9.67 | 14.06 | 9.77
68.064 | 16.11 | 0 | 9.47 | 0 | 6.64 | 0.1 | 10.45 | 4.49 | 4.88 | 5.76 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 17.77 | 9.67 | 9.86 | 0.59 | 0.1 | 1.07 | 0 | 1.17 | 0 | 21.39 | 4.79 | 24.51 | 4.79 | 19.43 | 4.49 | 16.99 | 9.67 | 13.97 | 9.67 | 15.53 | 9.67
68.096 | 16.6 | 0.1 | 10.06 | 0.1 | 7.23 | 0.1 | 11.13 | 4.49 | 4.88 | 5.37 | 4.88 | 18.07 | 9.86 | 21.58 | 9.77 | 18.36 | 9.67 | 9.96 | 0.59 | 0.1 | 1.07 | 0.1 | 1.17 | 0 | 21.39 | 4.88 | 25.1 | 4.79 | 20.12 | 4.59 | 17.68 | 9.67 | 15.53 | 9.67 | 16.8 | 9.67
68.128 | 17.19 | 0.1 | 10.64 | 0.1 | 7.91 | 0.1 | 11.72 | 4.49 | 4.88 | 5.18 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.65 | 9.67 | 9.96 | 0.59 | 0 | 0.88 | 0.1 | 0.98 | 0 | 21.29 | 4.88 | 25.59 | 4.79 | 20.7 | 4.59 | 18.56 | 9.67 | 16.7 | 9.67 | 17.58 | 9.67
68.16 | 13.67 | 0.1 | 11.23 | 0 | 8.5 | 0.1 | 12.31 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 19.14 | 9.67 | 9.96 | 0.49 | 0 | 0.88 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 26.08 | 4.69 | 21.29 | 4.59 | 19.24 | 9.67 | 17.58 | 9.67 | 18.26 | 9.67
68.192 | 14.16 | 0.1 | 11.82 | 0 | 8.98 | 0 | 12.89 | 4.39 | 4.88 | 4.98 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 19.63 | 9.67 | 9.86 | 0.29 | 0 | 0.68 | 0 | 0.78 | 0 | 21.29 | 4.88 | 21.49 | 4.79 | 21.97 | 4.59 | 19.92 | 9.67 | 18.46 | 9.57 | 19.04 | 9.67
68.224 | 14.16 | 0 | 12.4 | 0 | 9.67 | 0 | 13.38 | 4.49 | 4.88 | 4.79 | 4.88 | 17.97 | 9.86 | 21.49 | 9.86 | 20.12 | 9.67 | 9.96 | 0.29 | 0 | 0.68 | 0 | 0.78 | 0 | 11.72 | 4.88 | 20.31 | 4.79 | 22.36 | 4.59 | 20.41 | 9.57 | 18.95 | 9.57 | 19.73 | 9.57
68.256 | 14.45 | 0.1 | 12.89 | 0.1 | 10.25 | 0.1 | 13.97 | 4.49 | 4.88 | 4.79 | 4.88 | 17.97 | 9.86 | 21.39 | 9.86 | 20.51 | 9.67 | 10.84 | 0.29 | 0 | 0.59 | 0.1 | 0.68 | 0 | 3.13 | 4.88 | 21 | 4.79 | 22.95 | 4.59 | 21.09 | 9.67 | 19.73 | 9.67 | 20.22 | 9.67
68.288 | 14.36 | 0.1 | 13.48 | 0.1 | 10.84 | 0.1 | 14.45 | 4.49 | 4.88 | 4.79 | 4.88 | 5.66 | 9.86 | 21.39 | 9.86 | 20.9 | 9.67 | 11.82 | 0.29 | 0.1 | 0.59 | 0.1 | 0.59 | 0.1 | 3.32 | 4.88 | 21.19 | 4.79 | 23.54 | 4.69 | 21.68 | 9.67 | 20.22 | 9.67 | 20.8 | 9.67
68.32 | 14.45 | 0.1 | 13.97 | 0.1 | 11.43 | 0.1 | 15.04 | 4.59 | 4.88 | 4.79 | 4.88 | 2.83 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 13.18 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0 | 3.81 | 4.88 | 21.19 | 4.79 | 24.02 | 4.69 | 22.27 | 9.67 | 20.8 | 9.67 | 21.49 | 9.67
68.352 | 14.45 | 0.1 | 14.55 | 0 | 12.01 | 0 | 15.53 | 4.59 | 4.88 | 4.69 | 4.88 | 3.42 | 9.86 | 15.53 | 9.86 | 21.78 | 9.67 | 14.06 | 0.2 | 0 | 0.39 | 0.1 | 0.49 | 0.1 | 4.49 | 4.88 | 21.19 | 4.79 | 24.61 | 4.69 | 22.85 | 9.67 | 21.49 | 9.57 | 22.17 | 9.67
68.384 | 14.45 | 0 | 14.94 | 0 | 12.5 | 0.1 | 16.02 | 4.59 | 4.88 | 4.59 | 4.88 | 4.49 | 9.86 | 7.72 | 9.86 | 22.17 | 9.67 | 14.75 | 0.1 | 0 | 0.29 | 0 | 0.39 | 0 | 6.05 | 4.79 | 21.29 | 4.79 | 25.1 | 4.69 | 23.34 | 9.67 | 21.97 | 9.67 | 22.66 | 9.67
68.416 | 14.45 | 0.1 | 15.53 | 0.1 | 13.09 | 0.1 | 16.5 | 4.59 | 4.88 | 4.59 | 4.88 | 5.86 | 9.86 | 7.72 | 9.86 | 22.46 | 9.77 | 15.23 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0 | 6.15 | 4.88 | 21.29 | 4.79 | 25.59 | 4.79 | 23.93 | 9.67 | 22.66 | 9.67 | 23.24 | 9.67
68.448 | 14.45 | 0.1 | 16.02 | 0.1 | 13.67 | 0.1 | 16.99 | 4.69 | 4.88 | 4.59 | 4.88 | 5.76 | 9.86 | 10.64 | 9.86 | 22.85 | 9.77 | 15.82 | 0.1 | 0.1 | 0.29 | 1.17 | 0.29 | 0 | 5.86 | 4.88 | 21.29 | 4.79 | 26.08 | 4.79 | 24.51 | 9.67 | 23.15 | 9.67 | 23.73 | 9.67
68.48 | 14.45 | 0.1 | 16.41 | 0.1 | 14.16 | 0.1 | 17.48 | 4.69 | 4.88 | 4.59 | 4.88 | 5.37 | 9.86 | 11.33 | 9.86 | 23.24 | 9.77 | 16.31 | 0.1 | 0.2 | 0.29 | 2.25 | 0.29 | 0 | 5.47 | 4.88 | 21.29 | 4.79 | 21.49 | 4.79 | 24.9 | 9.67 | 23.63 | 9.67 | 24.32 | 9.67
68.512 | 14.45 | 0.1 | 16.9 | 0 | 14.65 | 0 | 17.87 | 4.69 | 5.57 | 4.49 | 4.88 | 4.98 | 9.86 | 10.45 | 9.86 | 23.63 | 9.77 | 16.8 | 0.1 | 1.46 | 0.2 | 4.1 | 0.2 | 0 | 5.18 | 4.88 | 21.19 | 4.79 | 20.51 | 4.79 | 25.39 | 9.67 | 24.22 | 9.67 | 24.81 | 9.67
68.544 | 14.36 | 0.1 | 16.5 | 0.1 | 15.14 | 0.1 | 18.36 | 4.69 | 6.05 | 4.59 | 4.88 | 4.98 | 9.77 | 9.86 | 9.86 | 24.02 | 9.77 | 17.29 | 0.1 | 2.64 | 0.2 | 5.57 | 0.2 | 0 | 5.08 | 4.79 | 21.19 | 4.79 | 21 | 4.79 | 25.98 | 9.67 | 24.81 | 9.67 | 25.2 | 9.67
68.576 | 14.45 | 0.1 | 14.65 | 0.1 | 15.63 | 0.1 | 18.75 | 4.69 | 7.03 | 4.49 | 4.88 | 4.88 | 9.86 | 10.06 | 9.86 | 24.41 | 9.77 | 17.77 | 0.1 | 4.59 | 0.2 | 6.74 | 0.2 | 0.29 | 5.27 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 26.47 | 9.77 | 25.29 | 9.67 | 25.78 | 9.67
68.608 | 5.08 | 0.1 | 14.65 | 0.1 | 16.21 | 0.1 | 19.24 | 4.79 | 8.4 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 22.27 | 9.77 | 18.16 | 0.1 | 6.05 | 0.2 | 7.62 | 0.2 | 1.46 | 5.08 | 4.88 | 21.19 | 4.79 | 21.19 | 4.79 | 26.95 | 9.77 | 25.68 | 9.77 | 26.27 | 9.77
68.64 | 1.76 | 0.1 | 14.65 | 0.1 | 16.7 | 0.1 | 19.63 | 4.79 | 9.38 | 4.49 | 4.88 | 4.79 | 9.86 | 10.25 | 9.86 | 22.27 | 9.77 | 18.56 | 0.1 | 6.93 | 0.1 | 8.4 | 0.1 | 3.03 | 4.98 | 4.88 | 6.93 | 4.79 | 21.19 | 4.79 | 27.34 | 9.77 | 26.17 | 9.77 | 26.86 | 9.77
68.672 | 2.05 | 0.1 | 14.45 | 0 | 17.09 | 0 | 20.12 | 4.79 | 10.06 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.88 | 9.77 | 19.14 | 0 | 7.72 | 0.1 | 9.08 | 0.1 | 5.08 | 4.88 | 4.88 | 2.54 | 4.79 | 21.29 | 4.79 | 28.03 | 9.77 | 26.76 | 9.77 | 27.34 | 9.67
68.704 | 1.76 | 0 | 14.55 | 0 | 13.77 | 0.1 | 20.51 | 4.79 | 10.55 | 4.49 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.68 | 9.77 | 19.53 | 0 | 8.5 | 0.1 | 9.86 | 0.1 | 6.25 | 4.88 | 4.88 | 3.22 | 4.79 | 21.29 | 4.79 | 28.32 | 9.77 | 27.25 | 9.77 | 27.74 | 9.77
68.736 | 1.37 | 0.1 | 14.55 | 0.1 | 13.97 | 0.1 | 20.9 | 4.79 | 11.23 | 4.59 | 4.88 | 4.69 | 9.86 | 9.96 | 9.86 | 21.68 | 9.77 | 19.92 | 0.1 | 9.28 | 0.1 | 10.74 | 0.1 | 7.13 | 4.88 | 4.88 | 4.2 | 4.79 | 21.19 | 4.79 | 28.91 | 9.77 | 27.74 | 9.77 | 28.03 | 9.77
68.768 | 1.37 | 0.1 | 14.45 | 0.1 | 14.16 | 0.1 | 17.77 | 4.79 | 11.82 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.31 | 0.1 | 9.96 | 0.1 | 11.62 | 0.1 | 8.01 | 4.79 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 29.4 | 9.77 | 28.22 | 9.77 | 28.61 | 9.77
68.8 | 1.17 | 0.1 | 14.45 | 0 | 14.45 | 0.1 | 17.68 | 4.79 | 12.4 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.8 | 0.1 | 10.84 | 0.1 | 12.31 | 0.1 | 8.69 | 4.69 | 4.88 | 6.25 | 4.79 | 21.19 | 4.79 | 24.51 | 9.77 | 28.61 | 9.77 | 29.1 | 9.77
68.832 | 1.07 | 0 | 14.45 | 0 | 14.45 | 0 | 17.87 | 4.79 | 12.89 | 4.59 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 11.72 | 0.1 | 13.18 | 0 | 9.57 | 4.59 | 4.88 | 5.96 | 4.79 | 21.19 | 4.79 | 23.54 | 9.77 | 29.1 | 9.77 | 27.25 | 9.77
68.864 | 1.07 | 0 | 14.36 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 13.48 | 4.69 | 4.88 | 4.49 | 9.77 | 9.77 | 9.86 | 21.39 | 9.86 | 21.58 | 0.1 | 12.5 | 0.1 | 13.67 | 0.1 | 10.35 | 4.69 | 4.79 | 5.47 | 4.79 | 21.19 | 4.79 | 24.02 | 9.77 | 28.42 | 9.77 | 24.61 | 9.77
68.896 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.07 | 4.88 | 14.06 | 4.69 | 4.98 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.97 | 0.1 | 13.18 | 0.1 | 14.45 | 0.1 | 11.23 | 4.59 | 4.88 | 5.18 | 4.79 | 21.09 | 4.79 | 24.22 | 9.77 | 24.81 | 9.77 | 24.51 | 9.77
68.928 | 0.88 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 18.16 | 4.88 | 14.55 | 4.69 | 5.66 | 4.59 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.36 | 0.1 | 13.97 | 0.1 | 15.33 | 0.1 | 11.91 | 4.59 | 4.88 | 5.18 | 4.79 | 5.47 | 4.79 | 24.32 | 9.77 | 24.51 | 9.77 | 24.41 | 9.86
68.96 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.04 | 4.69 | 6.64 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.75 | 0.1 | 14.65 | 0.1 | 15.92 | 0.1 | 12.6 | 4.59 | 4.88 | 5.18 | 4.79 | 2.73 | 4.79 | 24.22 | 9.77 | 24.51 | 9.77 | 24.41 | 9.86
68.992 | 0.59 | 0 | 6.45 | 0 | 14.45 | 0 | 18.16 | 4.88 | 15.63 | 4.79 | 8.2 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 0 | 15.43 | 0.1 | 16.7 | 0.1 | 13.38 | 4.59 | 5.76 | 5.08 | 4.79 | 3.42 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
69.024 | 0.59 | 0 | 1.95 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 16.02 | 4.69 | 8.98 | 4.49 | 9.77 | 9.67 | 9.86 | 15.72 | 9.86 | 23.54 | 0.1 | 16.11 | 0.1 | 17.29 | 0 | 14.06 | 4.49 | 6.64 | 4.98 | 4.79 | 4.2 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86
69.056 | 0.59 | 0.1 | 1.86 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.6 | 4.79 | 9.77 | 4.49 | 9.86 | 9.67 | 9.86 | 7.42 | 9.86 | 23.83 | 0.1 | 16.7 | 0.1 | 17.97 | 0.1 | 14.84 | 4.49 | 8.3 | 4.98 | 4.79 | 5.66 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
69.088 | 0.49 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.99 | 4.79 | 10.45 | 4.49 | 10.55 | 9.67 | 9.86 | 7.81 | 9.86 | 24.22 | 0.1 | 17.38 | 0.1 | 18.65 | 0.1 | 15.53 | 4.59 | 10.35 | 4.88 | 4.79 | 6.25 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
69.12 | 0.39 | 0.1 | 1.37 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 17.58 | 4.79 | 11.04 | 4.59 | 11.23 | 9.57 | 9.86 | 10.64 | 9.86 | 20.8 | 0.1 | 17.97 | 0.1 | 19.24 | 0.1 | 16.31 | 4.59 | 11.62 | 4.79 | 4.79 | 5.96 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
69.152 | 0.29 | 0.1 | 1.37 | 0 | 5.47 | 0 | 18.07 | 4.79 | 17.87 | 4.79 | 11.72 | 4.49 | 12.6 | 9.57 | 9.86 | 11.33 | 9.86 | 20.9 | 0 | 18.65 | 0.1 | 20.02 | 0.1 | 17.19 | 4.59 | 12.79 | 4.79 | 4.79 | 5.37 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
69.184 | 0.29 | 0 | 1.17 | 0 | 1.76 | 0.1 | 18.07 | 4.79 | 18.36 | 4.79 | 12.31 | 4.59 | 13.48 | 9.57 | 9.86 | 10.45 | 9.86 | 21.09 | 0 | 19.24 | 0.1 | 20.51 | 0.1 | 17.58 | 4.59 | 13.48 | 4.69 | 4.79 | 5.27 | 4.79 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
69.216 | 0.29 | 0.1 | 1.27 | 0.1 | 2.15 | 0.1 | 6.54 | 4.88 | 18.85 | 4.79 | 12.79 | 4.59 | 14.26 | 9.67 | 9.86 | 9.86 | 9.96 | 21.39 | 0.1 | 19.92 | 0.1 | 21.09 | 0.1 | 18.16 | 4.59 | 14.36 | 4.69 | 4.79 | 5.08 | 4.79 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
69.248 | 0.2 | 0.1 | 0.88 | 0.1 | 1.76 | 0.1 | 2.83 | 4.88 | 19.24 | 4.88 | 13.28 | 4.69 | 14.84 | 9.67 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 20.51 | 0.1 | 21.68 | 0.1 | 18.85 | 4.59 | 15.04 | 4.69 | 4.79 | 5.18 | 4.79 | 9.67 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
69.28 | 0.2 | 0.1 | 0.98 | 0.1 | 1.46 | 0.1 | 3.42 | 4.88 | 19.73 | 4.88 | 13.87 | 4.69 | 15.43 | 9.67 | 9.86 | 10.16 | 9.96 | 21.39 | 0.1 | 21 | 0.1 | 22.27 | 0.1 | 19.43 | 4.69 | 15.72 | 4.59 | 4.79 | 5.08 | 4.88 | 5.86 | 9.86 | 24.22 | 9.86 | 12.7 | 9.86
69.312 | 0.1 | 0.1 | 0.78 | 0 | 1.46 | 0 | 4.3 | 4.79 | 20.12 | 4.88 | 14.36 | 4.69 | 16.02 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0 | 21.58 | 0.1 | 21.78 | 0.1 | 20.12 | 4.59 | 16.41 | 4.59 | 4.79 | 4.98 | 4.79 | 9.86 | 9.77 | 11.52 | 9.86 | 5.76 | 9.86
69.344 | 0.1 | 0.29 | 0.78 | 0.1 | 1.27 | 0.1 | 5.66 | 4.88 | 20.51 | 4.79 | 14.94 | 4.69 | 16.5 | 9.67 | 10.55 | 9.86 | 9.86 | 21.39 | 0 | 22.17 | 0.1 | 18.65 | 0 | 20.8 | 4.69 | 16.99 | 4.59 | 4.79 | 4.98 | 4.79 | 12.31 | 9.77 | 5.76 | 9.86 | 8.01 | 9.86
69.376 | 0.1 | 1.07 | 0.68 | 0.1 | 1.17 | 0.1 | 5.76 | 4.88 | 21 | 4.88 | 15.43 | 4.69 | 16.99 | 9.67 | 11.23 | 9.96 | 9.96 | 21.49 | 0.1 | 18.07 | 0.1 | 18.26 | 0.1 | 21.39 | 4.69 | 17.77 | 4.59 | 4.79 | 4.88 | 4.79 | 11.04 | 9.86 | 8.4 | 9.86 | 11.72 | 9.86
69.408 | 0.1 | 2.15 | 0.59 | 0 | 1.17 | 0.1 | 5.57 | 4.88 | 17.68 | 4.88 | 15.92 | 4.79 | 17.38 | 9.67 | 12.4 | 9.96 | 9.96 | 21.49 | 0.1 | 16.6 | 0.1 | 18.07 | 0.1 | 21.88 | 4.69 | 18.46 | 4.59 | 4.79 | 4.88 | 4.79 | 10.16 | 9.86 | 12.01 | 9.86 | 11.62 | 9.86
69.44 | 0.1 | 3.61 | 0.49 | 0 | 0.88 | 0.1 | 5.27 | 4.88 | 17.97 | 4.88 | 16.41 | 4.79 | 17.97 | 9.67 | 13.57 | 9.86 | 9.96 | 21.39 | 0.1 | 17.29 | 0.1 | 17.97 | 0.1 | 22.46 | 4.79 | 19.04 | 4.59 | 4.88 | 4.79 | 4.79 | 10.06 | 9.86 | 11.52 | 9.86 | 10.35 | 9.86
69.472 | 0.1 | 4.39 | 0.49 | 0 | 0.88 | 0 | 4.98 | 4.88 | 18.16 | 4.88 | 16.9 | 4.79 | 18.46 | 9.67 | 14.36 | 9.77 | 9.96 | 21.39 | 0 | 17.48 | 0.1 | 17.77 | 0 | 18.07 | 4.79 | 19.82 | 4.59 | 5.76 | 4.69 | 4.79 | 10.35 | 9.77 | 10.25 | 9.86 | 9.86 | 9.86
69.504 | 0.1 | 4.98 | 0.39 | 0 | 0.78 | 0.1 | 4.88 | 4.88 | 18.16 | 4.79 | 17.29 | 4.79 | 18.75 | 9.67 | 14.94 | 9.67 | 9.96 | 21.39 | 0 | 17.48 | 0.1 | 17.77 | 0 | 16.9 | 4.79 | 20.31 | 4.49 | 7.03 | 4.69 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
69.536 | 0.1 | 5.47 | 0.29 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 17.77 | 4.79 | 19.24 | 9.67 | 15.43 | 9.67 | 9.96 | 21.39 | 0.1 | 17.48 | 0.1 | 17.77 | 0.1 | 17.58 | 4.79 | 21 | 4.59 | 8.98 | 4.69 | 4.79 | 10.25 | 9.86 | 10.25 | 9.86 | 10.45 | 9.86
69.568 | 0.1 | 6.05 | 0.29 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.26 | 4.79 | 19.73 | 9.67 | 15.92 | 9.67 | 9.96 | 10.64 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.68 | 4.79 | 21.49 | 4.59 | 10.64 | 4.69 | 4.79 | 10.16 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86
69.6 | 0.1 | 6.64 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 18.65 | 4.79 | 20.12 | 9.77 | 16.41 | 9.67 | 9.96 | 6.64 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.68 | 4.88 | 22.07 | 4.59 | 12.01 | 4.59 | 4.79 | 10.16 | 9.77 | 10.45 | 9.86 | 10.16 | 9.86
69.632 | 0.1 | 7.23 | 0.2 | 0 | 0.49 | 0 | 4.69 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 20.61 | 9.67 | 16.99 | 9.67 | 9.96 | 9.47 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 22.66 | 4.59 | 12.79 | 4.59 | 4.79 | 10.06 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
69.664 | 0.1 | 7.81 | 0.1 | 0 | 0.49 | 0.1 | 4.69 | 4.79 | 18.16 | 4.79 | 19.53 | 4.79 | 21 | 9.77 | 17.58 | 9.67 | 9.96 | 11.62 | 0 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 4.88 | 23.24 | 4.59 | 13.67 | 4.59 | 4.79 | 10.06 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
69.696 | 0.1 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 20.12 | 4.88 | 21.49 | 9.77 | 17.87 | 9.67 | 9.96 | 10.94 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.73 | 4.69 | 14.55 | 4.59 | 4.88 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16 | 9.86
69.728 | 0.1 | 8.98 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 20.51 | 4.88 | 21.78 | 9.77 | 18.36 | 9.67 | 9.96 | 10.16 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 4.88 | 24.22 | 4.69 | 15.23 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
69.76 | 0.1 | 9.67 | 0.1 | 0.68 | 0.2 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 20.9 | 4.88 | 22.17 | 9.77 | 18.75 | 9.57 | 9.96 | 9.77 | 0.1 | 17.58 | 0.1 | 8.59 | 0.1 | 17.68 | 4.88 | 24.81 | 4.69 | 16.02 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
69.792 | 0.1 | 10.25 | 0 | 1.37 | 0.2 | 0.59 | 4.59 | 4.79 | 18.07 | 4.88 | 17.48 | 4.79 | 22.66 | 9.77 | 19.24 | 9.57 | 9.96 | 9.96 | 0 | 17.48 | 0.1 | 1.95 | 0 | 17.68 | 4.88 | 25.29 | 4.69 | 16.6 | 4.49 | 4.79 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
69.824 | 0.1 | 10.74 | 0 | 2.44 | 0.2 | 1.27 | 4.49 | 4.88 | 17.58 | 4.88 | 17.58 | 4.79 | 22.95 | 9.77 | 19.73 | 9.57 | 9.96 | 10.16 | 0 | 5.86 | 0.1 | 2.34 | 0.1 | 17.68 | 4.88 | 25.78 | 4.69 | 17.29 | 4.49 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
69.856 | 0.1 | 11.43 | 0.1 | 3.81 | 0.2 | 2.34 | 4.59 | 4.88 | 4.79 | 4.88 | 17.87 | 4.88 | 23.34 | 9.77 | 20.12 | 9.57 | 9.96 | 9.96 | 0.1 | 1.66 | 0.1 | 2.15 | 0.1 | 17.68 | 4.88 | 26.27 | 4.79 | 17.97 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
69.888 | 0.1 | 12.01 | 0.1 | 4.49 | 0.1 | 3.71 | 4.49 | 4.88 | 2.93 | 4.88 | 18.07 | 4.88 | 23.73 | 9.77 | 20.51 | 9.57 | 9.96 | 9.86 | 0.1 | 2.34 | 0.1 | 1.76 | 0.1 | 17.48 | 4.88 | 22.85 | 4.79 | 18.56 | 4.59 | 4.79 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
69.92 | 0.1 | 12.5 | 0.1 | 5.08 | 0.1 | 4.49 | 4.59 | 4.88 | 3.61 | 4.88 | 18.07 | 4.88 | 24.12 | 9.77 | 20.9 | 9.57 | 10.64 | 9.96 | 0.1 | 1.95 | 0.1 | 1.86 | 0.1 | 4.49 | 4.88 | 22.07 | 4.79 | 19.24 | 4.59 | 5.96 | 9.67 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
69.952 | 0.1 | 13.09 | 0.1 | 5.57 | 0.1 | 5.08 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 4.79 | 24.41 | 9.77 | 21.29 | 9.57 | 11.72 | 9.96 | 0 | 1.66 | 0.1 | 1.37 | 0 | 2.44 | 4.88 | 21.88 | 4.79 | 19.92 | 4.59 | 7.03 | 9.67 | 10.94 | 9.67 | 9.86 | 9.77 | 9.86
69.984 | 0.1 | 13.57 | 0.1 | 6.15 | 0.1 | 5.57 | 4.59 | 4.79 | 5.66 | 4.88 | 18.16 | 4.79 | 22.46 | 9.77 | 21.78 | 9.57 | 12.99 | 9.77 | 0 | 1.56 | 0.1 | 1.56 | 0.1 | 2.05 | 4.88 | 21.58 | 4.79 | 20.51 | 4.59 | 8.79 | 9.67 | 11.82 | 9.67 | 11.04 | 9.67 | 9.96
70.016 | 0.1 | 14.06 | 0.1 | 6.74 | 0.1 | 6.15 | 4.59 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 22.36 | 9.86 | 22.17 | 9.67 | 13.97 | 9.77 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 1.86 | 4.88 | 21.58 | 4.79 | 21.09 | 4.69 | 10.45 | 9.67 | 13.48 | 9.77 | 12.01 | 9.67 | 10.94
70.048 | 0.1 | 14.65 | 0.1 | 7.42 | 0.1 | 6.74 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 21.97 | 9.86 | 22.46 | 9.67 | 14.65 | 9.67 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.49 | 4.79 | 21.68 | 4.69 | 11.82 | 9.67 | 15.23 | 9.67 | 13.77 | 9.67 | 12.21
70.08 | 0.1 | 15.14 | 0.1 | 8.01 | 0.1 | 7.32 | 4.59 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 21.78 | 9.86 | 22.95 | 9.67 | 15.23 | 9.77 | 0.1 | 1.27 | 0.1 | 1.07 | 0.1 | 1.56 | 4.88 | 21.39 | 4.79 | 22.27 | 4.69 | 12.79 | 9.67 | 16.41 | 9.67 | 15.43 | 9.67 | 14.06
70.112 | 0.1 | 15.63 | 0.1 | 8.59 | 0.1 | 8.01 | 4.59 | 4.79 | 4.88 | 4.79 | 18.07 | 4.79 | 21.78 | 9.77 | 23.34 | 9.67 | 15.72 | 9.67 | 0 | 1.17 | 0.1 | 0.98 | 0 | 1.66 | 4.88 | 21.39 | 4.79 | 22.75 | 4.69 | 13.67 | 9.67 | 17.38 | 9.67 | 16.7 | 9.67 | 15.82
70.144 | 0.1 | 16.11 | 0 | 9.28 | 0.1 | 8.5 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 21.58 | 9.77 | 23.63 | 9.67 | 16.21 | 9.67 | 0 | 1.07 | 0.1 | 0.88 | 0 | 1.37 | 4.88 | 21.29 | 4.79 | 23.34 | 4.69 | 14.45 | 9.67 | 18.07 | 9.67 | 17.48 | 9.67 | 16.7
70.176 | 0.1 | 16.5 | 0.1 | 9.77 | 0.1 | 9.18 | 4.69 | 5.57 | 4.98 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 24.02 | 9.67 | 16.8 | 9.67 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 23.83 | 4.79 | 15.14 | 9.67 | 18.85 | 9.67 | 18.26 | 9.67 | 17.68
70.208 | 0.1 | 17.09 | 0.1 | 10.35 | 0.1 | 9.67 | 4.69 | 6.15 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 24.41 | 9.67 | 17.29 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.32 | 4.79 | 15.82 | 9.57 | 19.53 | 9.67 | 18.85 | 9.67 | 18.46
70.24 | 0.1 | 13.77 | 0.1 | 11.04 | 0.1 | 10.35 | 4.69 | 7.32 | 4.88 | 4.88 | 6.93 | 4.88 | 21.49 | 9.86 | 22.46 | 9.67 | 17.68 | 9.67 | 0.49 | 0.78 | 0.1 | 0.68 | 0.1 | 0.88 | 4.88 | 21.19 | 4.79 | 24.9 | 4.79 | 16.6 | 9.67 | 20.12 | 9.67 | 19.53 | 9.57 | 19.04
70.272 | 0.1 | 14.16 | 0.1 | 11.52 | 0.1 | 10.94 | 4.69 | 8.5 | 4.69 | 4.88 | 2.64 | 4.79 | 21.49 | 9.77 | 22.17 | 9.67 | 18.07 | 9.57 | 1.66 | 0.68 | 0.1 | 0.59 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.39 | 4.79 | 17.29 | 9.67 | 20.8 | 9.57 | 20.22 | 9.57 | 19.82
70.304 | 0.1 | 14.16 | 0.1 | 12.11 | 0.1 | 11.52 | 4.69 | 9.47 | 4.79 | 4.88 | 3.32 | 4.79 | 21.49 | 9.77 | 21.88 | 9.67 | 18.56 | 9.57 | 3.03 | 0.59 | 0 | 0.49 | 0 | 0.78 | 4.88 | 14.45 | 4.79 | 25.88 | 4.79 | 17.87 | 9.67 | 21.39 | 9.57 | 20.9 | 9.67 | 20.22
70.336 | 0.1 | 14.45 | 0.1 | 12.6 | 0.1 | 12.11 | 4.79 | 10.16 | 4.69 | 4.88 | 4.49 | 4.88 | 21.49 | 9.86 | 21.68 | 9.67 | 18.95 | 9.57 | 4.98 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 3.42 | 4.79 | 25.49 | 4.79 | 18.46 | 9.67 | 21.88 | 9.67 | 21.49 | 9.67 | 20.9
70.368 | 0.1 | 14.36 | 0.1 | 13.18 | 0.1 | 12.6 | 4.79 | 10.74 | 4.59 | 4.88 | 5.66 | 4.88 | 21.39 | 9.86 | 21.68 | 9.67 | 19.43 | 9.57 | 6.25 | 0.49 | 0.1 | 0.49 | 0.1 | 0.68 | 4.88 | 3.22 | 4.79 | 21.88 | 4.79 | 19.14 | 9.67 | 22.46 | 9.67 | 21.97 | 9.67 | 21.58
70.4 | 0.1 | 14.45 | 0.1 | 13.77 | 0.1 | 13.18 | 4.79 | 11.33 | 4.59 | 4.88 | 5.86 | 4.88 | 14.55 | 9.86 | 21.58 | 9.67 | 19.92 | 9.57 | 7.13 | 0.39 | 0.1 | 0.39 | 0.1 | 0.59 | 4.88 | 3.42 | 4.79 | 21.39 | 4.79 | 19.73 | 9.67 | 23.05 | 9.67 | 22.56 | 9.67 | 21.97
70.432 | 0.1 | 14.45 | 0.1 | 14.26 | 0.1 | 13.77 | 4.79 | 12.01 | 4.59 | 4.88 | 5.47 | 4.79 | 7.13 | 9.77 | 21.49 | 9.67 | 20.31 | 9.57 | 7.81 | 0.29 | 1.37 | 0.29 | 0.1 | 0.49 | 4.88 | 4.39 | 4.79 | 21.49 | 4.79 | 20.41 | 9.67 | 23.54 | 9.67 | 23.05 | 9.67 | 22.66
70.464 | 0.1 | 14.36 | 0.1 | 14.75 | 0.1 | 14.16 | 4.79 | 12.5 | 4.59 | 4.79 | 5.27 | 4.79 | 7.91 | 9.77 | 21.49 | 9.67 | 20.7 | 9.57 | 8.59 | 0.29 | 2.54 | 0.29 | 1.07 | 0.39 | 4.88 | 5.86 | 4.79 | 21.39 | 4.79 | 21 | 9.67 | 24.12 | 9.67 | 23.54 | 9.67 | 23.15
70.496 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 14.75 | 4.88 | 13.09 | 4.59 | 4.88 | 4.98 | 4.88 | 11.04 | 9.86 | 21.49 | 9.77 | 21.19 | 9.67 | 9.47 | 0.29 | 4.39 | 0.2 | 2.05 | 0.49 | 4.88 | 6.25 | 4.79 | 21.39 | 4.79 | 21.49 | 9.67 | 24.61 | 9.67 | 24.22 | 9.67 | 23.73
70.528 | 0.1 | 14.45 | 0.1 | 15.72 | 0.1 | 15.23 | 4.88 | 13.57 | 4.49 | 4.88 | 4.88 | 4.88 | 11.33 | 9.86 | 21.49 | 9.77 | 21.49 | 9.67 | 10.16 | 0.2 | 5.76 | 0.2 | 3.81 | 0.29 | 4.88 | 5.76 | 4.79 | 21.29 | 4.79 | 22.17 | 9.77 | 25.1 | 9.67 | 24.61 | 9.67 | 24.22
70.56 | 0.1 | 14.36 | 0.1 | 16.31 | 0.1 | 15.72 | 4.88 | 14.06 | 4.59 | 4.98 | 4.88 | 4.88 | 10.45 | 9.86 | 21.39 | 9.77 | 21.88 | 9.67 | 11.04 | 0.2 | 6.84 | 0.2 | 5.47 | 0.29 | 4.88 | 5.47 | 4.79 | 21.19 | 4.79 | 22.66 | 9.77 | 25.68 | 9.67 | 25.2 | 9.67 | 24.71
70.592 | 0.1 | 14.45 | 0.1 | 16.8 | 0.1 | 16.21 | 4.79 | 14.65 | 4.59 | 5.96 | 4.88 | 4.79 | 9.86 | 9.77 | 21.39 | 9.77 | 22.36 | 9.67 | 11.72 | 0.1 | 7.62 | 0.1 | 6.64 | 0.2 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 23.24 | 9.67 | 26.27 | 9.67 | 25.68 | 9.67 | 25.2
70.624 | 0.1 | 14.36 | 0.1 | 17.19 | 0.1 | 16.7 | 4.88 | 15.04 | 4.59 | 6.84 | 4.79 | 4.79 | 9.96 | 9.86 | 21.39 | 9.77 | 22.66 | 9.67 | 12.5 | 0.1 | 8.4 | 0.1 | 7.52 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 23.73 | 9.67 | 26.56 | 9.67 | 26.08 | 9.67 | 25.78
70.656 | 0.1 | 14.36 | 0.1 | 14.75 | 0.1 | 17.19 | 4.88 | 15.63 | 4.59 | 8.3 | 4.69 | 4.88 | 10.25 | 9.86 | 14.65 | 9.77 | 23.15 | 9.67 | 13.18 | 0.2 | 9.28 | 0.1 | 8.3 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 27.15 | 9.77 | 26.66 | 9.67 | 26.27
70.688 | 0.1 | 4.88 | 0.1 | 15.23 | 0.1 | 14.26 | 4.88 | 16.11 | 4.59 | 9.28 | 4.69 | 4.88 | 10.16 | 9.86 | 7.23 | 9.77 | 23.54 | 9.67 | 14.06 | 0.1 | 9.96 | 0.1 | 9.08 | 0.2 | 4.88 | 5.18 | 4.79 | 21.19 | 4.88 | 24.9 | 9.77 | 27.54 | 9.77 | 27.25 | 9.77 | 26.66
70.72 | 0.1 | 1.76 | 0.1 | 14.65 | 0.1 | 14.75 | 4.88 | 16.6 | 4.69 | 9.86 | 4.69 | 4.88 | 9.96 | 9.86 | 8.01 | 9.77 | 23.83 | 9.67 | 14.84 | 0.1 | 10.74 | 0.1 | 9.77 | 0.1 | 4.88 | 5.08 | 4.79 | 21.19 | 4.88 | 25.29 | 9.77 | 28.13 | 9.77 | 27.54 | 9.77 | 27.25
70.752 | 0.1 | 1.95 | 0 | 14.65 | 0.1 | 14.65 | 4.88 | 17.19 | 4.59 | 10.55 | 4.59 | 4.79 | 9.96 | 9.77 | 10.94 | 9.77 | 24.22 | 9.67 | 15.33 | 0 | 11.52 | 0.1 | 10.84 | 0.1 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.88 | 9.77 | 28.61 | 9.77 | 28.13 | 9.77 | 27.74
70.784 | 0.1 | 1.76 | 0 | 14.65 | 0.1 | 14.55 | 4.88 | 17.58 | 4.69 | 11.13 | 4.59 | 4.79 | 9.96 | 9.77 | 11.33 | 9.77 | 22.46 | 9.67 | 16.11 | 0.1 | 12.31 | 0.1 | 11.43 | 0 | 4.88 | 4.88 | 4.79 | 8.79 | 4.79 | 26.27 | 9.77 | 29.01 | 9.77 | 28.52 | 9.77 | 28.22
70.816 | 0.1 | 1.46 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.07 | 4.69 | 11.72 | 4.59 | 4.88 | 9.96 | 9.86 | 10.45 | 9.77 | 21.09 | 9.67 | 16.8 | 0.1 | 13.18 | 0.1 | 12.21 | 0.1 | 4.88 | 4.79 | 4.88 | 2.34 | 4.88 | 22.75 | 9.77 | 29.49 | 9.77 | 29.1 | 9.77 | 28.61
70.848 | 0.1 | 1.37 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.46 | 4.69 | 12.31 | 4.49 | 4.88 | 9.86 | 9.86 | 9.86 | 9.77 | 21.39 | 9.67 | 17.38 | 0.1 | 13.87 | 0.1 | 12.99 | 0.1 | 4.88 | 4.69 | 4.79 | 3.32 | 4.88 | 22.07 | 9.77 | 26.37 | 9.77 | 29.4 | 9.77 | 29.1
70.88 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 18.85 | 4.79 | 12.79 | 4.59 | 4.88 | 9.86 | 9.86 | 9.96 | 9.77 | 21.58 | 9.77 | 18.07 | 0.1 | 14.55 | 0.1 | 13.67 | 0.1 | 4.88 | 4.69 | 4.79 | 3.81 | 4.88 | 21.68 | 9.77 | 25.29 | 9.77 | 25.78 | 9.77 | 27.54
70.912 | 0.1 | 0.98 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.34 | 4.79 | 13.38 | 4.49 | 4.79 | 9.77 | 9.86 | 10.16 | 9.77 | 21.39 | 9.67 | 18.75 | 0 | 15.33 | 0.1 | 14.45 | 0 | 5.76 | 4.69 | 4.79 | 4.88 | 4.79 | 21.49 | 9.77 | 24.81 | 9.77 | 24.81 | 9.77 | 23.44
70.944 | 0.1 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.73 | 4.79 | 13.87 | 4.49 | 4.98 | 9.77 | 9.77 | 10.16 | 9.77 | 21.49 | 9.67 | 19.43 | 0 | 16.02 | 0.1 | 15.23 | 0 | 6.84 | 4.59 | 4.79 | 6.25 | 4.79 | 21.49 | 9.77 | 24.61 | 9.77 | 24.61 | 9.77 | 24.41
70.976 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.22 | 4.79 | 14.45 | 4.49 | 5.76 | 9.77 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.02 | 0.1 | 16.7 | 0.1 | 15.92 | 0.1 | 8.5 | 4.59 | 4.88 | 6.15 | 4.88 | 21.39 | 9.77 | 24.61 | 9.77 | 24.61 | 9.77 | 24.41
71.008 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.61 | 4.79 | 14.94 | 4.49 | 6.74 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.61 | 0.1 | 17.38 | 0.1 | 16.5 | 0.1 | 10.35 | 4.59 | 4.88 | 5.57 | 4.88 | 21.29 | 9.77 | 24.51 | 9.77 | 24.51 | 9.77 | 24.32
71.04 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 21 | 4.88 | 15.43 | 4.49 | 8.11 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 18.07 | 0.1 | 17.29 | 0.1 | 11.72 | 4.59 | 4.79 | 5.18 | 4.88 | 21.29 | 9.77 | 24.41 | 9.77 | 24.41 | 9.86 | 24.22
71.072 | 0.1 | 0.59 | 0.1 | 10.35 | 0.1 | 14.36 | 4.88 | 18.95 | 4.79 | 16.02 | 4.59 | 9.08 | 9.67 | 11.13 | 9.86 | 9.77 | 21.49 | 9.77 | 21.78 | 0 | 18.75 | 0 | 17.77 | 0 | 12.7 | 4.59 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
71.104 | 0.1 | 0.59 | 0.1 | 2.93 | 0.1 | 12.01 | 4.88 | 18.85 | 4.88 | 16.41 | 4.59 | 9.86 | 9.67 | 12.01 | 9.77 | 9.86 | 21.39 | 9.77 | 22.27 | 0 | 19.34 | 0.1 | 18.56 | 0 | 13.57 | 4.49 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22
71.136 | 0.1 | 0.49 | 0.1 | 2.15 | 0.1 | 4 | 4.88 | 18.46 | 4.88 | 16.9 | 4.59 | 10.45 | 9.67 | 13.38 | 9.77 | 9.86 | 21.49 | 9.77 | 17.97 | 0.1 | 20.02 | 0.1 | 19.14 | 0.1 | 14.36 | 4.49 | 4.88 | 5.08 | 4.88 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
71.168 | 0.1 | 0.39 | 0.1 | 1.66 | 0.1 | 1.95 | 4.88 | 18.36 | 4.88 | 17.38 | 4.59 | 11.13 | 9.67 | 14.16 | 9.77 | 9.86 | 21.39 | 9.77 | 17.09 | 0.1 | 20.61 | 0.1 | 19.92 | 0.1 | 15.04 | 4.59 | 4.79 | 5.08 | 4.88 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
71.2 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.86 | 4.88 | 18.36 | 4.88 | 17.77 | 4.69 | 11.72 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 9.77 | 17.48 | 0.1 | 21.19 | 0.1 | 20.51 | 0 | 15.72 | 4.59 | 4.79 | 4.98 | 4.88 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
71.232 | 0.1 | 0.29 | 0.1 | 1.46 | 0.1 | 1.46 | 4.88 | 18.26 | 4.88 | 18.26 | 4.69 | 12.31 | 9.67 | 15.33 | 9.67 | 9.77 | 11.13 | 9.77 | 17.48 | 0 | 21.78 | 0.1 | 21 | 0 | 16.5 | 4.59 | 4.79 | 4.88 | 4.79 | 9.47 | 9.77 | 24.22 | 9.77 | 24.12 | 9.77 | 24.12
71.264 | 0.2 | 0.2 | 0.1 | 1.17 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.65 | 4.69 | 12.7 | 9.67 | 15.82 | 9.67 | 9.86 | 6.54 | 9.77 | 17.48 | 0.1 | 22.27 | 0.1 | 21.58 | 0 | 17.09 | 4.59 | 4.79 | 4.88 | 4.79 | 2.73 | 9.77 | 15.43 | 9.77 | 24.22 | 9.86 | 24.12
71.296 | 1.17 | 0.2 | 0.1 | 1.17 | 0.1 | 1.27 | 4.88 | 18.26 | 4.88 | 19.24 | 4.69 | 13.38 | 9.67 | 16.31 | 9.67 | 9.86 | 8.89 | 9.77 | 17.58 | 0.1 | 20.61 | 0.1 | 22.07 | 0.1 | 17.87 | 4.59 | 4.88 | 4.79 | 4.88 | 3.22 | 9.77 | 6.74 | 9.77 | 13.87 | 9.86 | 24.12
71.328 | 1.95 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.63 | 4.69 | 13.87 | 9.67 | 16.9 | 9.67 | 9.86 | 11.43 | 9.77 | 17.48 | 0.1 | 17.29 | 0.1 | 22.66 | 0.1 | 18.46 | 4.59 | 5.76 | 4.69 | 4.79 | 3.61 | 9.77 | 7.23 | 9.77 | 6.25 | 9.86 | 12.21
71.36 | 3.42 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 14.45 | 9.67 | 17.29 | 9.67 | 9.86 | 11.13 | 9.86 | 17.48 | 0.1 | 17.77 | 0.1 | 19.43 | 0.1 | 19.14 | 4.59 | 7.23 | 4.69 | 4.79 | 4.88 | 9.77 | 11.33 | 9.77 | 7.72 | 9.86 | 5.86
71.392 | 4.3 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 14.94 | 9.67 | 17.77 | 9.67 | 9.86 | 10.25 | 9.77 | 17.38 | 0 | 17.68 | 0 | 18.46 | 0 | 19.82 | 4.69 | 9.28 | 4.69 | 4.79 | 6.15 | 9.77 | 11.82 | 9.77 | 11.72 | 9.86 | 8.5
71.424 | 4.88 | 0 | 0 | 0.68 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 9.67 | 18.26 | 9.57 | 9.86 | 9.96 | 9.77 | 17.38 | 0 | 17.68 | 0 | 18.26 | 0 | 20.41 | 4.59 | 10.74 | 4.59 | 4.79 | 6.05 | 9.77 | 10.55 | 9.77 | 11.82 | 9.86 | 11.91
71.456 | 5.37 | 0.1 | 0.1 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 17.48 | 4.79 | 15.92 | 9.67 | 18.65 | 9.67 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 20.9 | 4.69 | 12.11 | 4.59 | 4.88 | 5.57 | 9.77 | 9.96 | 9.77 | 10.45 | 9.86 | 11.62
71.488 | 5.96 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 12.99 | 4.88 | 17.68 | 4.79 | 16.41 | 9.67 | 19.24 | 9.67 | 9.86 | 10.25 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 21.58 | 4.69 | 12.99 | 4.59 | 4.88 | 5.27 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.45
71.52 | 6.54 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 3.71 | 4.88 | 17.87 | 4.79 | 16.9 | 9.67 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.17 | 4.69 | 13.77 | 4.49 | 4.88 | 5.27 | 9.77 | 10.45 | 9.77 | 10.25 | 9.86 | 9.96
71.552 | 7.13 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 18.07 | 4.88 | 17.29 | 9.67 | 20.02 | 9.57 | 9.86 | 9.96 | 9.86 | 14.65 | 0 | 17.68 | 0 | 17.77 | 0.1 | 22.75 | 4.79 | 14.55 | 4.49 | 4.79 | 5.27 | 9.77 | 10.45 | 9.77 | 10.45 | 9.86 | 10.25
71.584 | 7.72 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 3.71 | 4.88 | 17.97 | 4.88 | 17.77 | 9.77 | 20.41 | 9.67 | 9.86 | 9.96 | 9.86 | 3.91 | 0 | 17.68 | 0 | 17.77 | 0 | 23.24 | 4.69 | 15.33 | 4.59 | 4.88 | 5.08 | 9.77 | 10.16 | 9.77 | 10.35 | 9.86 | 10.45
71.616 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 4.79 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 20.9 | 9.67 | 9.86 | 9.96 | 9.86 | 2.05 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 23.83 | 4.79 | 16.02 | 4.49 | 5.96 | 5.08 | 9.77 | 10.16 | 9.77 | 10.35 | 9.86 | 10.45
71.648 | 8.98 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 5.66 | 4.88 | 18.16 | 4.79 | 18.65 | 9.77 | 21.29 | 9.67 | 9.86 | 9.86 | 9.86 | 2.15 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 24.32 | 4.79 | 16.6 | 4.49 | 7.42 | 4.98 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.16
71.68 | 9.67 | 0.1 | 0.78 | 0.2 | 0.1 | 0.29 | 4.88 | 5.66 | 4.88 | 18.07 | 4.88 | 19.14 | 9.77 | 21.68 | 9.67 | 9.86 | 9.86 | 9.86 | 1.86 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.9 | 4.79 | 17.29 | 4.59 | 9.57 | 4.98 | 9.77 | 10.16 | 9.77 | 10.06 | 9.86 | 9.96
71.712 | 10.25 | 0.1 | 1.66 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 21.97 | 9.67 | 10.45 | 9.77 | 9.86 | 1.76 | 0 | 17.58 | 0 | 17.68 | 0 | 25.39 | 4.79 | 17.97 | 4.59 | 11.13 | 4.88 | 9.77 | 9.96 | 9.77 | 10.06 | 9.86 | 10.16
71.744 | 10.74 | 0.1 | 2.83 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 20.02 | 9.77 | 22.36 | 9.67 | 11.13 | 9.77 | 9.86 | 1.56 | 0.1 | 7.23 | 0.1 | 17.68 | 0 | 25.88 | 4.79 | 18.65 | 4.59 | 12.11 | 4.79 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86 | 10.16
71.776 | 11.33 | 0.1 | 3.91 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.41 | 9.77 | 22.85 | 9.67 | 12.21 | 9.77 | 9.86 | 1.37 | 0.1 | 1.76 | 0.1 | 9.67 | 0.1 | 26.37 | 4.79 | 19.34 | 4.59 | 13.09 | 4.79 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86 | 9.96
71.808 | 11.91 | 0.1 | 4.59 | 0.1 | 0.2 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.8 | 9.77 | 23.24 | 9.67 | 13.48 | 9.67 | 9.86 | 1.46 | 0.1 | 2.25 | 0.1 | 2.34 | 0.1 | 23.15 | 4.88 | 19.92 | 4.59 | 13.87 | 4.79 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
71.84 | 12.5 | 0.1 | 5.18 | 0.1 | 1.17 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.58 | 9.86 | 23.63 | 9.67 | 14.26 | 9.67 | 9.86 | 1.37 | 0.1 | 2.05 | 0.1 | 2.25 | 0 | 22.27 | 4.88 | 20.51 | 4.59 | 14.75 | 4.69 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
71.872 | 12.99 | 0.1 | 5.76 | 0 | 1.95 | 0.1 | 4.98 | 4.79 | 4.88 | 17.77 | 4.88 | 17.58 | 9.77 | 23.93 | 9.67 | 14.84 | 9.67 | 9.86 | 1.07 | 0 | 1.76 | 0 | 2.05 | 0 | 21.88 | 4.88 | 21 | 4.59 | 15.43 | 4.59 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
71.904 | 13.48 | 0 | 6.25 | 0 | 3.32 | 0 | 5.66 | 4.69 | 4.88 | 5.86 | 4.88 | 17.77 | 9.77 | 24.32 | 9.67 | 15.33 | 9.67 | 9.86 | 1.07 | 0 | 1.56 | 0.1 | 1.66 | 0 | 21.68 | 4.79 | 21.78 | 4.69 | 16.11 | 4.59 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
71.936 | 14.06 | 0.1 | 6.93 | 0.1 | 4.2 | 0 | 6.74 | 4.79 | 4.88 | 2.83 | 4.88 | 18.07 | 9.86 | 21.29 | 9.77 | 15.92 | 9.67 | 9.86 | 0.88 | 0.1 | 1.37 | 0.1 | 1.56 | 0.1 | 21.68 | 4.88 | 22.27 | 4.69 | 16.9 | 4.59 | 10.94 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77
71.968 | 14.55 | 0.1 | 7.52 | 0.1 | 4.88 | 0.1 | 8.11 | 4.69 | 4.88 | 3.61 | 4.88 | 17.97 | 9.86 | 21.78 | 9.77 | 16.41 | 9.67 | 9.86 | 0.88 | 0.1 | 1.46 | 0.1 | 1.56 | 0 | 21.49 | 4.88 | 22.85 | 4.69 | 17.58 | 4.59 | 12.4 | 9.77 | 9.77 | 9.77 | 11.04 | 9.77
72 | 15.14 | 0.1 | 8.11 | 0.1 | 5.57 | 0.1 | 9.08 | 4.59 | 4.88 | 4.59 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 16.9 | 9.67 | 9.86 | 0.78 | 0.1 | 1.27 | 0.1 | 1.37 | 0.1 | 21.39 | 4.88 | 23.44 | 4.69 | 18.16 | 4.59 | 14.45 | 9.67 | 10.94 | 9.77 | 11.91 | 9.77
72.032 | 15.63 | 0.1 | 8.79 | 0 | 5.96 | 0.1 | 9.77 | 4.49 | 4.88 | 5.66 | 4.88 | 18.07 | 9.86 | 21.58 | 9.77 | 17.38 | 9.57 | 9.86 | 0.68 | 0 | 1.17 | 0 | 1.27 | 0 | 21.39 | 4.88 | 23.83 | 4.69 | 18.85 | 4.59 | 16.02 | 9.67 | 12.01 | 9.67 | 13.87 | 9.77
72.064 | 16.11 | 0.1 | 9.38 | 0.1 | 6.54 | 0.1 | 10.45 | 4.49 | 4.88 | 5.86 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 17.77 | 9.57 | 9.86 | 0.59 | 0.1 | 1.07 | 0 | 1.27 | 0 | 21.39 | 4.88 | 24.41 | 4.79 | 19.43 | 4.49 | 16.9 | 9.67 | 13.77 | 9.67 | 15.43 | 9.67
72.096 | 16.6 | 0.1 | 9.96 | 0.1 | 7.23 | 0.1 | 11.13 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.26 | 9.67 | 9.96 | 0.59 | 0.1 | 0.88 | 0.1 | 1.17 | 0.1 | 21.39 | 4.88 | 24.81 | 4.79 | 20.02 | 4.59 | 17.87 | 9.67 | 15.43 | 9.67 | 16.7 | 9.67
72.128 | 17.09 | 0.1 | 10.55 | 0 | 7.72 | 0.1 | 11.72 | 4.49 | 4.88 | 5.18 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.65 | 9.67 | 9.96 | 0.49 | 0.1 | 0.98 | 0.1 | 0.98 | 0.1 | 21.29 | 4.88 | 25.49 | 4.79 | 20.7 | 4.59 | 18.46 | 9.67 | 16.5 | 9.67 | 17.38 | 9.67
72.16 | 13.87 | 0.1 | 11.13 | 0.1 | 8.4 | 0.1 | 12.31 | 4.49 | 4.88 | 4.98 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 19.24 | 9.67 | 9.96 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 21.19 | 4.59 | 19.24 | 9.67 | 17.58 | 9.67 | 18.26 | 9.67
72.192 | 13.97 | 0.1 | 11.72 | 0 | 8.98 | 0 | 12.89 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 19.63 | 9.67 | 9.86 | 0.29 | 0 | 0.68 | 0 | 0.78 | 0 | 21.29 | 4.88 | 23.34 | 4.79 | 21.78 | 4.59 | 19.82 | 9.67 | 18.26 | 9.67 | 18.95 | 9.67
72.224 | 14.16 | 0.1 | 12.21 | 0 | 9.67 | 0 | 13.38 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 20.02 | 9.67 | 9.96 | 0.29 | 0 | 0.59 | 0 | 0.78 | 0 | 13.48 | 4.79 | 20.9 | 4.79 | 22.36 | 4.59 | 20.31 | 9.67 | 18.85 | 9.57 | 19.53 | 9.57
72.256 | 14.45 | 0.1 | 12.89 | 0.1 | 10.16 | 0.1 | 13.87 | 4.49 | 4.88 | 4.79 | 4.88 | 17.97 | 9.86 | 21.49 | 9.86 | 20.51 | 9.67 | 10.64 | 0.29 | 0.1 | 0.59 | 0.1 | 0.68 | 0 | 3.32 | 4.88 | 21.39 | 4.79 | 22.95 | 4.69 | 21.09 | 9.67 | 19.53 | 9.67 | 20.12 | 9.67
72.288 | 14.36 | 0.1 | 13.38 | 0 | 10.74 | 0.1 | 14.45 | 4.49 | 4.88 | 4.79 | 4.88 | 6.54 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 11.62 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.22 | 4.88 | 21.39 | 4.79 | 23.44 | 4.59 | 21.68 | 9.67 | 20.22 | 9.67 | 20.8 | 9.67
72.32 | 14.36 | 0.1 | 13.87 | 0.1 | 11.33 | 0.1 | 14.94 | 4.49 | 4.88 | 4.79 | 4.88 | 2.64 | 9.86 | 21.39 | 9.86 | 21.19 | 9.67 | 12.99 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.52 | 4.88 | 21.29 | 4.79 | 23.93 | 4.69 | 22.17 | 9.67 | 20.8 | 9.67 | 21.39 | 9.67
72.352 | 14.45 | 0.1 | 14.55 | 0 | 11.91 | 0 | 15.43 | 4.59 | 4.88 | 4.69 | 4.88 | 3.61 | 9.86 | 16.8 | 9.77 | 21.68 | 9.67 | 14.06 | 0.2 | 0 | 0.39 | 0.1 | 0.49 | 0.1 | 4.49 | 4.88 | 21.19 | 4.79 | 24.41 | 4.69 | 22.85 | 9.67 | 21.39 | 9.67 | 21.97 | 9.67
72.384 | 14.36 | 0.1 | 14.94 | 0 | 12.5 | 0.1 | 15.92 | 4.49 | 4.88 | 4.59 | 4.88 | 4.39 | 9.86 | 8.3 | 9.77 | 22.07 | 9.67 | 14.55 | 0.1 | 0.1 | 0.29 | 0 | 0.39 | 0 | 5.86 | 4.88 | 21.29 | 4.79 | 25 | 4.69 | 23.34 | 9.67 | 21.88 | 9.67 | 22.36 | 9.67
72.416 | 14.45 | 0.1 | 15.43 | 0.1 | 12.99 | 0.1 | 16.41 | 4.59 | 4.88 | 4.59 | 4.88 | 5.57 | 9.86 | 7.32 | 9.86 | 22.46 | 9.77 | 15.14 | 0.2 | 0.1 | 0.29 | 0.1 | 0.39 | 0 | 6.05 | 4.88 | 21.29 | 4.88 | 25.59 | 4.69 | 23.83 | 9.67 | 22.46 | 9.67 | 23.15 | 9.67
72.448 | 14.45 | 0.1 | 15.92 | 0.1 | 13.48 | 0.1 | 16.9 | 4.59 | 4.88 | 4.59 | 4.88 | 5.86 | 9.86 | 10.35 | 9.86 | 22.85 | 9.77 | 15.72 | 0.1 | 0.1 | 0.29 | 1.07 | 0.29 | 0.1 | 5.76 | 4.88 | 21.29 | 4.88 | 25.98 | 4.69 | 24.51 | 9.67 | 23.05 | 9.67 | 23.63 | 9.67
72.48 | 14.36 | 0.1 | 16.41 | 0.1 | 14.06 | 0.1 | 17.38 | 4.69 | 4.88 | 4.59 | 4.88 | 5.47 | 9.86 | 11.52 | 9.86 | 23.24 | 9.77 | 16.21 | 0.1 | 0.2 | 0.2 | 2.05 | 0.29 | 0.1 | 5.37 | 4.88 | 21.19 | 4.79 | 21.78 | 4.79 | 24.9 | 9.67 | 23.63 | 9.67 | 24.22 | 9.67
72.512 | 14.36 | 0.1 | 16.9 | 0 | 14.55 | 0 | 17.87 | 4.59 | 5.27 | 4.49 | 4.88 | 5.08 | 9.86 | 10.55 | 9.77 | 23.54 | 9.77 | 16.7 | 0.1 | 1.46 | 0.2 | 4 | 0.2 | 0 | 5.18 | 4.88 | 21.19 | 4.79 | 20.22 | 4.69 | 25.39 | 9.67 | 24.22 | 9.67 | 24.71 | 9.67
72.544 | 14.36 | 0.1 | 17.29 | 0 | 15.14 | 0.1 | 18.26 | 4.69 | 5.96 | 4.59 | 4.79 | 4.88 | 9.86 | 10.16 | 9.86 | 24.02 | 9.77 | 17.19 | 0.1 | 2.64 | 0.2 | 5.47 | 0.2 | 0 | 5.08 | 4.79 | 21.19 | 4.88 | 21 | 4.79 | 25.88 | 9.67 | 24.61 | 9.67 | 25.29 | 9.67
72.576 | 14.36 | 0.1 | 15.33 | 0.1 | 15.63 | 0.1 | 18.75 | 4.69 | 6.93 | 4.59 | 4.88 | 4.98 | 9.86 | 9.96 | 9.86 | 24.32 | 9.77 | 17.68 | 0.1 | 4.59 | 0.2 | 6.64 | 0.2 | 0.2 | 5.08 | 4.88 | 21.19 | 4.88 | 21.19 | 4.79 | 26.47 | 9.77 | 25.2 | 9.67 | 25.78 | 9.77
72.608 | 5.47 | 0.1 | 15.04 | 0.1 | 16.11 | 0.1 | 19.24 | 4.69 | 8.4 | 4.59 | 4.88 | 4.98 | 9.86 | 10.16 | 9.86 | 21.58 | 9.77 | 18.07 | 0.1 | 5.96 | 0.1 | 7.52 | 0.2 | 1.46 | 5.08 | 4.88 | 21.19 | 4.88 | 21.19 | 4.79 | 26.86 | 9.77 | 25.68 | 9.67 | 26.17 | 9.77
72.64 | 1.56 | 0.1 | 14.84 | 0.1 | 16.6 | 0.1 | 19.63 | 4.69 | 9.38 | 4.49 | 4.88 | 4.79 | 9.86 | 10.25 | 9.86 | 22.07 | 9.77 | 18.56 | 0.1 | 6.93 | 0.1 | 8.2 | 0.1 | 2.83 | 4.98 | 4.88 | 7.72 | 4.79 | 21.19 | 4.79 | 27.34 | 9.77 | 26.17 | 9.77 | 26.56 | 9.77
72.672 | 1.95 | 0.1 | 14.55 | 0 | 16.99 | 0 | 20.02 | 4.69 | 10.06 | 4.49 | 4.88 | 4.79 | 9.86 | 10.16 | 9.77 | 21.68 | 9.77 | 19.04 | 0 | 7.72 | 0.1 | 9.08 | 0.1 | 4.88 | 4.88 | 4.88 | 2.34 | 4.79 | 21.19 | 4.79 | 28.03 | 9.77 | 26.86 | 9.67 | 27.25 | 9.77
72.704 | 1.66 | 0 | 14.65 | 0 | 15.23 | 0 | 20.51 | 4.79 | 10.64 | 4.49 | 4.79 | 4.79 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 19.43 | 0.1 | 8.5 | 0.1 | 9.77 | 0.1 | 6.15 | 4.79 | 4.79 | 3.22 | 4.88 | 21.19 | 4.79 | 28.32 | 9.77 | 27.25 | 9.77 | 27.74 | 9.67
72.736 | 1.46 | 0.1 | 14.55 | 0.1 | 14.06 | 0.1 | 20.9 | 4.79 | 11.23 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.58 | 9.86 | 19.92 | 0.1 | 9.28 | 0.1 | 10.55 | 0.1 | 7.03 | 4.79 | 4.88 | 3.91 | 4.88 | 21.19 | 4.79 | 28.91 | 9.77 | 27.64 | 9.77 | 28.13 | 9.77
72.768 | 1.37 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 17.87 | 4.79 | 11.82 | 4.59 | 4.88 | 4.59 | 9.86 | 9.96 | 9.86 | 21.49 | 9.86 | 20.31 | 0.1 | 10.16 | 0.1 | 11.52 | 0.1 | 7.81 | 4.79 | 4.88 | 5.08 | 4.88 | 21.29 | 4.79 | 29.3 | 9.77 | 28.13 | 9.77 | 28.61 | 9.77
72.8 | 1.17 | 0.1 | 14.55 | 0.1 | 14.45 | 0.1 | 17.68 | 4.79 | 12.4 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.7 | 0.1 | 10.94 | 0.1 | 12.31 | 0.1 | 8.69 | 4.69 | 4.88 | 6.15 | 4.79 | 21.19 | 4.79 | 24.61 | 9.86 | 28.61 | 9.77 | 29.01 | 9.77
72.832 | 1.07 | 0.1 | 14.45 | 0 | 14.45 | 0 | 17.87 | 4.79 | 12.99 | 4.59 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21 | 0.1 | 11.62 | 0 | 13.09 | 0 | 9.57 | 4.59 | 4.79 | 6.05 | 4.79 | 21.19 | 4.79 | 23.34 | 9.77 | 29.1 | 9.77 | 28.91 | 9.77
72.864 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.16 | 4.79 | 13.48 | 4.59 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.39 | 9.86 | 21.49 | 0.1 | 12.31 | 0 | 13.87 | 0.1 | 10.25 | 4.59 | 4.79 | 5.57 | 4.88 | 21.19 | 4.79 | 24.02 | 9.77 | 29.01 | 9.77 | 25.78 | 9.77
72.896 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.07 | 4.88 | 13.97 | 4.69 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.88 | 0.1 | 13.18 | 0.1 | 14.55 | 0.1 | 11.04 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 4.79 | 24.22 | 9.86 | 25.49 | 9.86 | 25 | 9.86
72.928 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 0.1 | 18.16 | 4.88 | 14.55 | 4.69 | 5.57 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.27 | 0.1 | 13.97 | 0.1 | 15.33 | 0.1 | 11.82 | 4.59 | 4.88 | 5.08 | 4.88 | 6.93 | 4.79 | 24.22 | 9.86 | 24.9 | 9.77 | 24.61 | 9.86
72.96 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.04 | 4.69 | 6.54 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.66 | 0.1 | 14.65 | 0.1 | 16.02 | 0.1 | 12.6 | 4.59 | 4.88 | 5.08 | 4.79 | 2.44 | 4.79 | 24.22 | 9.86 | 24.61 | 9.77 | 24.51 | 9.86
72.992 | 0.59 | 0.1 | 7.42 | 0 | 14.45 | 0 | 18.16 | 4.79 | 15.63 | 4.69 | 8.11 | 4.49 | 9.86 | 9.67 | 9.77 | 21.39 | 9.86 | 23.05 | 0 | 15.33 | 0.1 | 16.6 | 0 | 13.28 | 4.59 | 5.76 | 5.08 | 4.79 | 3.32 | 4.79 | 24.32 | 9.77 | 24.41 | 9.77 | 24.32 | 9.86
73.024 | 0.59 | 0 | 2.25 | 0 | 14.45 | 0 | 18.16 | 4.79 | 16.02 | 4.69 | 8.89 | 4.49 | 9.77 | 9.67 | 9.77 | 17.29 | 9.86 | 23.44 | 0.1 | 16.21 | 0 | 17.19 | 0 | 13.97 | 4.49 | 6.45 | 5.08 | 4.88 | 4.1 | 4.79 | 24.32 | 9.77 | 24.41 | 9.77 | 24.41 | 9.86
73.056 | 0.59 | 0.1 | 2.05 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.5 | 4.79 | 9.77 | 4.49 | 9.86 | 9.67 | 9.86 | 8.4 | 9.86 | 23.83 | 0.1 | 16.8 | 0.1 | 17.87 | 0.1 | 14.75 | 4.49 | 8.01 | 5.08 | 4.88 | 5.27 | 4.79 | 24.32 | 9.86 | 24.41 | 9.86 | 24.22 | 9.86
73.088 | 0.49 | 0.1 | 1.76 | 0.1 | 14.45 | 0.1 | 18.07 | 4.88 | 16.99 | 4.79 | 10.45 | 4.49 | 10.45 | 9.67 | 9.86 | 7.23 | 9.96 | 24.22 | 0.1 | 17.38 | 0.1 | 18.65 | 0.1 | 15.43 | 4.59 | 10.25 | 4.88 | 4.88 | 6.25 | 4.79 | 24.32 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86
73.12 | 0.39 | 0.1 | 1.56 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 17.48 | 4.79 | 11.04 | 4.49 | 11.13 | 9.57 | 9.86 | 10.45 | 9.86 | 22.27 | 0.1 | 17.97 | 0.1 | 19.24 | 0.1 | 16.21 | 4.59 | 11.52 | 4.79 | 4.79 | 5.96 | 4.79 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
73.152 | 0.29 | 0.1 | 1.27 | 0 | 6.15 | 0 | 18.07 | 4.79 | 17.97 | 4.79 | 11.62 | 4.49 | 12.4 | 9.57 | 9.77 | 11.43 | 9.86 | 20.9 | 0 | 18.75 | 0.1 | 19.82 | 0.1 | 16.9 | 4.59 | 12.6 | 4.79 | 4.79 | 5.27 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
73.184 | 0.2 | 0 | 1.37 | 0 | 1.76 | 0.1 | 18.07 | 4.79 | 18.46 | 4.79 | 12.21 | 4.59 | 13.38 | 9.57 | 9.77 | 10.45 | 9.86 | 21.29 | 0 | 19.34 | 0 | 20.31 | 0 | 17.48 | 4.59 | 13.38 | 4.79 | 4.79 | 5.18 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
73.216 | 0.2 | 0.1 | 1.07 | 0.1 | 1.95 | 0.1 | 6.93 | 4.88 | 18.75 | 4.79 | 12.79 | 4.59 | 14.26 | 9.67 | 9.86 | 9.96 | 9.96 | 21.49 | 0.1 | 19.92 | 0.1 | 21.09 | 0.1 | 18.16 | 4.59 | 14.26 | 4.69 | 4.79 | 5.27 | 4.79 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
73.248 | 0.2 | 0.1 | 1.07 | 0.1 | 2.15 | 0.1 | 2.64 | 4.88 | 19.24 | 4.88 | 13.38 | 4.69 | 14.84 | 9.67 | 9.86 | 9.96 | 9.96 | 21.49 | 0.1 | 20.51 | 0.1 | 21.68 | 0.1 | 18.85 | 4.59 | 14.94 | 4.69 | 4.79 | 5.18 | 4.79 | 10.35 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
73.28 | 0.2 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.42 | 4.88 | 19.73 | 4.88 | 13.87 | 4.69 | 15.43 | 9.67 | 9.86 | 10.16 | 9.96 | 21.49 | 0.1 | 21.19 | 0.1 | 22.27 | 0.1 | 19.34 | 4.59 | 15.72 | 4.59 | 4.79 | 5.08 | 4.79 | 5.76 | 9.86 | 24.12 | 9.86 | 13.57 | 9.86
73.312 | 0.1 | 0 | 0.78 | 0 | 1.37 | 0 | 4.69 | 4.79 | 20.12 | 4.79 | 14.36 | 4.69 | 15.92 | 9.67 | 9.77 | 10.16 | 9.96 | 21.39 | 0 | 21.68 | 0.1 | 22.17 | 0 | 20.02 | 4.59 | 16.41 | 4.59 | 4.79 | 4.98 | 4.79 | 9.67 | 9.77 | 12.21 | 9.86 | 5.86 | 9.86
73.344 | 0 | 0.2 | 0.68 | 0.1 | 1.17 | 0.1 | 5.47 | 4.79 | 20.61 | 4.88 | 14.84 | 4.69 | 16.41 | 9.67 | 10.45 | 9.96 | 9.96 | 21.49 | 0.1 | 22.17 | 0.1 | 19.24 | 0.1 | 20.7 | 4.69 | 17.09 | 4.59 | 4.88 | 4.98 | 4.79 | 12.31 | 9.77 | 5.96 | 9.86 | 7.72 | 9.86
73.376 | 0.1 | 1.17 | 0.68 | 0.1 | 1.17 | 0.1 | 5.66 | 4.88 | 20.9 | 4.88 | 15.33 | 4.69 | 16.9 | 9.67 | 11.13 | 9.86 | 9.96 | 21.49 | 0.1 | 18.16 | 0.1 | 18.56 | 0.1 | 21.39 | 4.69 | 17.68 | 4.59 | 4.88 | 4.98 | 4.79 | 11.23 | 9.86 | 7.72 | 9.86 | 11.52 | 9.86
73.408 | 0.1 | 1.95 | 0.59 | 0.1 | 0.88 | 0.1 | 5.47 | 4.88 | 17.58 | 4.88 | 15.82 | 4.79 | 17.38 | 9.67 | 12.21 | 9.86 | 9.96 | 21.49 | 0.1 | 16.7 | 0.1 | 18.16 | 0.1 | 21.97 | 4.69 | 18.36 | 4.59 | 4.79 | 4.88 | 4.79 | 9.96 | 9.86 | 11.62 | 9.86 | 11.91 | 9.86
73.44 | 0.1 | 3.32 | 0.59 | 0.1 | 0.88 | 0.1 | 5.18 | 4.88 | 17.87 | 4.88 | 16.31 | 4.79 | 17.87 | 9.67 | 13.48 | 9.86 | 9.96 | 21.39 | 0.1 | 17.38 | 0.1 | 18.07 | 0.1 | 22.46 | 4.79 | 19.14 | 4.59 | 4.88 | 4.88 | 4.79 | 10.16 | 9.86 | 11.91 | 9.86 | 10.45 | 9.86
73.472 | 0.1 | 4.3 | 0.39 | 0 | 0.88 | 0 | 4.88 | 4.79 | 18.07 | 4.88 | 16.8 | 4.69 | 18.36 | 9.67 | 14.26 | 9.77 | 9.96 | 21.39 | 0 | 17.48 | 0.1 | 17.87 | 0 | 17.87 | 4.79 | 19.63 | 4.49 | 5.86 | 4.69 | 4.79 | 10.45 | 9.77 | 10.45 | 9.86 | 9.86 | 9.86
73.504 | 0.1 | 4.88 | 0.29 | 0 | 0.68 | 0 | 4.98 | 4.79 | 18.16 | 4.79 | 17.29 | 4.79 | 18.75 | 9.67 | 14.84 | 9.67 | 9.96 | 21.39 | 0 | 17.58 | 0 | 17.87 | 0.1 | 16.8 | 4.79 | 20.22 | 4.59 | 7.03 | 4.69 | 4.79 | 10.45 | 9.77 | 9.86 | 9.86 | 10.16 | 9.86
73.536 | 0.1 | 5.37 | 0.29 | 0 | 0.68 | 0.1 | 4.98 | 4.88 | 18.16 | 4.88 | 17.77 | 4.79 | 19.24 | 9.77 | 15.43 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.48 | 4.79 | 20.9 | 4.59 | 8.98 | 4.69 | 4.79 | 10.35 | 9.86 | 10.25 | 9.86 | 10.45 | 9.86
73.568 | 0.1 | 5.96 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 18.16 | 4.79 | 19.73 | 9.77 | 15.82 | 9.77 | 9.96 | 11.52 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.68 | 4.79 | 21.49 | 4.59 | 10.55 | 4.59 | 4.79 | 10.16 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86
73.6 | 0.1 | 6.54 | 0.2 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.65 | 4.79 | 20.12 | 9.77 | 16.31 | 9.67 | 9.96 | 6.84 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.68 | 4.79 | 22.07 | 4.59 | 11.91 | 4.59 | 4.79 | 10.16 | 9.86 | 10.45 | 9.86 | 10.06 | 9.86
73.632 | 0.1 | 7.13 | 0.2 | 0 | 0.49 | 0 | 4.69 | 4.79 | 18.16 | 4.88 | 19.14 | 4.79 | 20.51 | 9.77 | 16.9 | 9.67 | 9.96 | 8.89 | 0 | 17.58 | 0 | 17.68 | 0 | 17.68 | 4.79 | 22.66 | 4.59 | 12.79 | 4.49 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
73.664 | 0.1 | 7.72 | 0.1 | 0 | 0.49 | 0 | 4.69 | 4.79 | 18.16 | 4.88 | 19.53 | 4.79 | 20.9 | 9.77 | 17.29 | 9.67 | 9.96 | 11.33 | 0 | 17.58 | 0 | 17.68 | 0 | 17.68 | 4.79 | 23.15 | 4.59 | 13.77 | 4.59 | 4.79 | 10.06 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86
73.696 | 0.1 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.92 | 4.79 | 21.29 | 9.77 | 17.87 | 9.67 | 9.96 | 11.13 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.63 | 4.59 | 14.55 | 4.59 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 10.16 | 9.96
73.728 | 0.1 | 8.98 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 20.31 | 4.79 | 21.78 | 9.86 | 18.26 | 9.57 | 9.96 | 10.25 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.68 | 4.88 | 24.22 | 4.69 | 15.23 | 4.59 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 9.96 | 9.96
73.76 | 0.1 | 9.67 | 0.1 | 0.68 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.8 | 4.88 | 22.07 | 9.77 | 18.65 | 9.57 | 9.96 | 9.86 | 0.1 | 17.58 | 0.1 | 9.08 | 0.1 | 17.68 | 4.88 | 24.71 | 4.69 | 15.92 | 4.49 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
73.792 | 0.1 | 10.25 | 0 | 1.27 | 0.2 | 0.49 | 4.49 | 4.79 | 18.07 | 4.88 | 17.77 | 4.79 | 22.46 | 9.77 | 19.14 | 9.57 | 9.96 | 9.86 | 0 | 17.58 | 0 | 2.05 | 0 | 17.68 | 4.88 | 25.29 | 4.69 | 16.5 | 4.49 | 4.79 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
73.824 | 0.1 | 10.64 | 0.1 | 2.44 | 0.2 | 1.27 | 4.49 | 4.79 | 17.87 | 4.79 | 17.68 | 4.79 | 22.85 | 9.77 | 19.53 | 9.57 | 9.96 | 10.16 | 0.1 | 6.25 | 0 | 2.25 | 0 | 17.68 | 4.88 | 25.68 | 4.69 | 17.29 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
73.856 | 0.1 | 11.33 | 0.1 | 3.81 | 0.1 | 2.34 | 4.59 | 4.88 | 5.08 | 4.88 | 17.97 | 4.88 | 23.24 | 9.86 | 20.02 | 9.57 | 9.96 | 10.16 | 0.1 | 1.56 | 0.1 | 2.05 | 0.1 | 17.68 | 4.88 | 26.27 | 4.69 | 17.97 | 4.59 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
73.888 | 0.1 | 11.91 | 0.1 | 4.49 | 0.1 | 3.61 | 4.59 | 4.88 | 2.93 | 4.88 | 18.16 | 4.79 | 23.63 | 9.86 | 20.51 | 9.57 | 9.96 | 9.86 | 0.1 | 2.15 | 0.1 | 1.86 | 0.1 | 17.68 | 4.88 | 22.07 | 4.79 | 18.56 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
73.92 | 0.1 | 12.4 | 0.1 | 5.08 | 0.1 | 4.49 | 4.49 | 4.88 | 3.71 | 4.88 | 18.07 | 4.88 | 24.02 | 9.86 | 20.8 | 9.57 | 10.64 | 9.86 | 0.1 | 1.95 | 0.1 | 1.86 | 0.1 | 5.76 | 4.88 | 21.39 | 4.79 | 19.24 | 4.59 | 5.86 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
73.952 | 0 | 12.99 | 0.1 | 5.57 | 0.1 | 4.98 | 4.49 | 4.88 | 4.59 | 4.88 | 18.07 | 4.79 | 24.41 | 9.77 | 21.29 | 9.57 | 11.62 | 9.86 | 0 | 1.76 | 0 | 1.46 | 0.1 | 1.95 | 4.88 | 21.68 | 4.79 | 19.82 | 4.59 | 6.84 | 9.67 | 10.64 | 9.77 | 9.77 | 9.77 | 9.86
73.984 | 0.1 | 13.48 | 0.1 | 6.25 | 0.1 | 5.47 | 4.49 | 4.79 | 5.66 | 4.79 | 18.07 | 4.79 | 21.97 | 9.86 | 21.68 | 9.57 | 12.89 | 9.86 | 0.1 | 1.56 | 0 | 1.46 | 0 | 2.25 | 4.88 | 21.39 | 4.79 | 20.31 | 4.59 | 8.59 | 9.67 | 11.62 | 9.77 | 10.74 | 9.67 | 9.96
74.016 | 0.1 | 14.06 | 0.1 | 6.74 | 0.1 | 6.05 | 4.59 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 22.27 | 9.86 | 21.97 | 9.67 | 13.87 | 9.77 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 2.05 | 4.88 | 21.49 | 4.79 | 21 | 4.69 | 10.45 | 9.67 | 13.28 | 9.77 | 11.82 | 9.67 | 10.84
74.048 | 0.1 | 14.55 | 0.1 | 7.42 | 0.1 | 6.74 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 4.88 | 21.78 | 9.86 | 22.46 | 9.67 | 14.55 | 9.77 | 0.1 | 1.46 | 0.1 | 1.27 | 0.1 | 1.66 | 4.88 | 21.39 | 4.79 | 21.58 | 4.69 | 11.72 | 9.67 | 15.14 | 9.67 | 13.48 | 9.67 | 12.21
74.08 | 0.1 | 15.04 | 0.1 | 8.01 | 0.1 | 7.32 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.79 | 21.68 | 9.86 | 22.85 | 9.67 | 15.14 | 9.77 | 0.1 | 1.27 | 0.1 | 1.07 | 0.1 | 1.76 | 4.88 | 21.39 | 4.79 | 22.27 | 4.69 | 12.7 | 9.57 | 16.31 | 9.67 | 15.23 | 9.67 | 14.06
74.112 | 0.1 | 15.63 | 0 | 8.59 | 0.1 | 7.91 | 4.59 | 4.88 | 4.98 | 4.79 | 18.07 | 4.79 | 21.78 | 9.86 | 23.24 | 9.67 | 15.72 | 9.67 | 0 | 1.17 | 0 | 1.07 | 0.1 | 1.37 | 4.88 | 21.39 | 4.79 | 22.66 | 4.69 | 13.57 | 9.67 | 17.29 | 9.67 | 16.5 | 9.57 | 15.82
74.144 | 0 | 16.11 | 0 | 9.28 | 0.1 | 8.5 | 4.59 | 4.79 | 4.88 | 4.79 | 18.07 | 4.79 | 21.58 | 9.86 | 23.63 | 9.67 | 16.21 | 9.67 | 0.1 | 1.07 | 0 | 0.88 | 0 | 1.37 | 4.88 | 21.29 | 4.79 | 23.34 | 4.69 | 14.36 | 9.57 | 18.07 | 9.67 | 17.29 | 9.67 | 16.8
74.176 | 0.1 | 16.5 | 0.1 | 9.77 | 0.1 | 9.08 | 4.69 | 5.66 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 23.93 | 9.67 | 16.6 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 23.73 | 4.79 | 15.23 | 9.57 | 18.75 | 9.67 | 18.16 | 9.67 | 17.68
74.208 | 0.1 | 17.09 | 0.1 | 10.45 | 0.1 | 9.67 | 4.69 | 6.05 | 4.88 | 4.88 | 18.07 | 4.79 | 21.49 | 9.86 | 24.32 | 9.67 | 17.19 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.41 | 4.79 | 15.82 | 9.67 | 19.43 | 9.67 | 18.85 | 9.67 | 18.36
74.24 | 0.1 | 13.97 | 0.1 | 10.94 | 0.1 | 10.25 | 4.69 | 7.03 | 4.79 | 4.79 | 7.32 | 4.79 | 21.49 | 9.86 | 21.97 | 9.67 | 17.68 | 9.67 | 0.29 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.9 | 4.79 | 16.5 | 9.67 | 20.12 | 9.57 | 19.43 | 9.67 | 19.04
74.272 | 0.1 | 13.97 | 0 | 11.52 | 0.1 | 10.94 | 4.69 | 8.5 | 4.79 | 4.88 | 2.73 | 4.79 | 21.49 | 9.86 | 22.27 | 9.67 | 18.16 | 9.57 | 1.56 | 0.68 | 0 | 0.59 | 0.1 | 0.88 | 4.88 | 21.19 | 4.79 | 25.39 | 4.79 | 17.19 | 9.67 | 20.7 | 9.57 | 20.12 | 9.57 | 19.73
74.304 | 0 | 14.16 | 0.1 | 12.01 | 0.1 | 11.43 | 4.69 | 9.28 | 4.69 | 4.79 | 3.32 | 4.79 | 21.49 | 9.86 | 21.68 | 9.67 | 18.56 | 9.57 | 2.93 | 0.68 | 0.1 | 0.49 | 0 | 0.88 | 4.88 | 17.77 | 4.79 | 25.88 | 4.79 | 17.77 | 9.57 | 21.29 | 9.67 | 20.7 | 9.57 | 20.31
74.336 | 0.1 | 14.45 | 0.1 | 12.7 | 0.1 | 12.01 | 4.79 | 10.06 | 4.69 | 4.88 | 4.3 | 4.88 | 21.49 | 9.86 | 21.68 | 9.67 | 18.95 | 9.57 | 4.98 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 4.1 | 4.79 | 25.59 | 4.79 | 18.46 | 9.67 | 21.78 | 9.67 | 21.29 | 9.67 | 20.9
74.368 | 0.1 | 14.36 | 0.1 | 13.18 | 0.1 | 12.6 | 4.79 | 10.55 | 4.69 | 4.88 | 5.47 | 4.88 | 21.39 | 9.86 | 21.68 | 9.77 | 19.43 | 9.57 | 6.15 | 0.49 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 3.03 | 4.79 | 22.17 | 4.79 | 19.14 | 9.67 | 22.36 | 9.67 | 21.88 | 9.67 | 21.49
74.4 | 0.1 | 14.45 | 0.1 | 13.77 | 0.1 | 13.09 | 4.79 | 11.23 | 4.59 | 4.88 | 5.86 | 4.88 | 16.7 | 9.86 | 21.58 | 9.77 | 19.92 | 9.57 | 7.13 | 0.49 | 0.1 | 0.39 | 0.1 | 0.59 | 4.88 | 3.52 | 4.79 | 21.49 | 4.79 | 19.82 | 9.67 | 22.95 | 9.67 | 22.46 | 9.67 | 21.97
74.432 | 0.1 | 14.45 | 0.1 | 14.26 | 0.1 | 13.67 | 4.79 | 11.82 | 4.59 | 4.88 | 5.57 | 4.79 | 7.32 | 9.86 | 21.49 | 9.67 | 20.22 | 9.57 | 7.81 | 0.29 | 1.07 | 0.29 | 0.1 | 0.49 | 4.88 | 4.39 | 4.79 | 21.49 | 4.79 | 20.51 | 9.67 | 23.63 | 9.67 | 23.05 | 9.67 | 22.66
74.464 | 0.1 | 14.36 | 0 | 14.75 | 0.1 | 14.16 | 4.79 | 12.4 | 4.49 | 4.79 | 5.18 | 4.79 | 7.62 | 9.86 | 21.49 | 9.77 | 20.7 | 9.57 | 8.5 | 0.29 | 2.15 | 0.2 | 0.88 | 0.49 | 4.88 | 5.57 | 4.79 | 21.49 | 4.79 | 21 | 9.67 | 24.02 | 9.67 | 23.54 | 9.67 | 23.05
74.496 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 14.75 | 4.88 | 12.99 | 4.59 | 4.88 | 4.98 | 4.88 | 10.64 | 9.86 | 21.39 | 9.77 | 21.19 | 9.67 | 9.38 | 0.29 | 4.1 | 0.2 | 1.86 | 0.49 | 4.88 | 6.25 | 4.88 | 21.39 | 4.79 | 21.58 | 9.67 | 24.61 | 9.67 | 24.22 | 9.67 | 23.63
74.528 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 15.23 | 4.88 | 13.48 | 4.49 | 4.88 | 4.98 | 4.88 | 11.43 | 9.86 | 21.39 | 9.77 | 21.49 | 9.67 | 10.06 | 0.29 | 5.57 | 0.2 | 3.61 | 0.39 | 4.88 | 5.76 | 4.79 | 21.39 | 4.79 | 22.07 | 9.77 | 25.1 | 9.67 | 24.61 | 9.67 | 24.12
74.56 | 0.1 | 14.36 | 0.1 | 16.21 | 0.1 | 15.63 | 4.88 | 14.06 | 4.59 | 4.88 | 4.88 | 4.88 | 10.45 | 9.86 | 21.39 | 9.77 | 21.88 | 9.67 | 10.94 | 0.2 | 6.54 | 0.2 | 5.37 | 0.29 | 4.88 | 5.47 | 4.79 | 21.29 | 4.79 | 22.66 | 9.77 | 25.68 | 9.77 | 25.2 | 9.67 | 24.71
74.592 | 0 | 14.45 | 0 | 16.7 | 0.1 | 16.21 | 4.79 | 14.55 | 4.59 | 5.66 | 4.79 | 4.79 | 9.86 | 9.86 | 21.39 | 9.77 | 22.36 | 9.67 | 11.72 | 0.1 | 7.52 | 0.1 | 6.54 | 0.2 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 23.24 | 9.67 | 26.17 | 9.67 | 25.59 | 9.67 | 25.2
74.624 | 0.1 | 14.36 | 0 | 17.29 | 0.1 | 16.7 | 4.79 | 14.94 | 4.49 | 6.64 | 4.79 | 4.79 | 10.06 | 9.86 | 21.39 | 9.77 | 22.66 | 9.67 | 12.5 | 0.1 | 8.4 | 0.1 | 7.32 | 0.2 | 4.88 | 4.98 | 4.79 | 21.29 | 4.79 | 23.73 | 9.77 | 26.56 | 9.77 | 26.08 | 9.67 | 25.78
74.656 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 17.19 | 4.88 | 15.53 | 4.59 | 8.01 | 4.79 | 4.79 | 10.25 | 9.86 | 16.99 | 9.77 | 23.15 | 9.67 | 13.18 | 0.2 | 9.08 | 0.1 | 8.2 | 0.2 | 4.88 | 5.18 | 4.79 | 21.19 | 4.79 | 24.32 | 9.77 | 27.15 | 9.77 | 26.56 | 9.67 | 26.27
74.688 | 0.1 | 5.66 | 0.1 | 15.14 | 0.1 | 13.97 | 4.88 | 16.11 | 4.59 | 9.08 | 4.69 | 4.88 | 10.16 | 9.86 | 8.01 | 9.86 | 23.44 | 9.67 | 13.97 | 0.1 | 9.96 | 0.1 | 8.98 | 0.2 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 24.81 | 9.77 | 27.54 | 9.77 | 27.15 | 9.77 | 26.76
74.72 | 0.1 | 1.66 | 0.1 | 14.65 | 0.1 | 14.36 | 4.88 | 16.5 | 4.59 | 9.77 | 4.69 | 4.88 | 9.96 | 9.86 | 7.52 | 9.86 | 23.83 | 9.67 | 14.65 | 0.1 | 10.55 | 0.1 | 9.67 | 0.2 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.29 | 9.77 | 28.03 | 9.77 | 27.64 | 9.77 | 27.25
74.752 | 0 | 1.95 | 0.1 | 14.65 | 0.1 | 14.45 | 4.88 | 17.09 | 4.59 | 10.45 | 4.59 | 4.79 | 9.86 | 9.77 | 10.64 | 9.77 | 24.22 | 9.67 | 15.43 | 0 | 11.52 | 0.1 | 10.64 | 0.1 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.88 | 9.77 | 28.52 | 9.77 | 28.13 | 9.77 | 27.74
74.784 | 0.1 | 1.66 | 0 | 14.65 | 0.1 | 14.45 | 4.88 | 17.38 | 4.69 | 11.04 | 4.59 | 4.79 | 9.86 | 9.86 | 11.52 | 9.77 | 23.05 | 9.67 | 15.92 | 0.1 | 12.21 | 0.1 | 11.33 | 0 | 4.88 | 4.88 | 4.79 | 9.28 | 4.79 | 26.37 | 9.77 | 29.01 | 9.77 | 28.52 | 9.77 | 28.13
74.816 | 0.1 | 1.56 | 0.1 | 14.55 | 0.1 | 14.45 | 4.88 | 17.97 | 4.69 | 11.72 | 4.59 | 4.88 | 9.86 | 9.86 | 10.45 | 9.86 | 21.19 | 9.67 | 16.7 | 0.1 | 13.09 | 0.1 | 12.21 | 0.1 | 4.88 | 4.79 | 4.79 | 2.54 | 4.79 | 23.15 | 9.77 | 29.49 | 9.77 | 29.01 | 9.77 | 28.61
74.848 | 0.1 | 1.37 | 0.1 | 14.55 | 0.1 | 14.45 | 4.88 | 18.36 | 4.69 | 12.21 | 4.59 | 4.88 | 9.86 | 9.86 | 9.86 | 9.86 | 21.58 | 9.67 | 17.29 | 0.1 | 13.77 | 0.1 | 12.89 | 0.1 | 4.88 | 4.69 | 4.79 | 3.22 | 4.88 | 22.27 | 9.77 | 25.88 | 9.77 | 29.4 | 9.77 | 29.1
74.88 | 0.1 | 1.17 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.85 | 4.69 | 12.79 | 4.59 | 4.88 | 9.77 | 9.86 | 9.96 | 9.86 | 21.68 | 9.67 | 17.97 | 0.1 | 14.55 | 0.1 | 13.67 | 0.1 | 4.88 | 4.69 | 4.79 | 3.91 | 4.79 | 21.78 | 9.77 | 25.1 | 9.77 | 26.17 | 9.77 | 28.32
74.912 | 0 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.69 | 13.28 | 4.59 | 4.79 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 9.67 | 18.65 | 0 | 15.33 | 0 | 14.36 | 0 | 5.86 | 4.69 | 4.79 | 4.79 | 4.79 | 21.58 | 9.77 | 24.81 | 9.77 | 25.2 | 9.77 | 23.63
74.944 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.63 | 4.79 | 13.87 | 4.49 | 4.88 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 9.67 | 19.24 | 0 | 15.92 | 0.1 | 15.04 | 0 | 6.64 | 4.69 | 4.79 | 6.15 | 4.79 | 21.58 | 9.77 | 24.61 | 9.77 | 24.81 | 9.77 | 24.51
74.976 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.12 | 4.79 | 14.36 | 4.49 | 5.57 | 9.77 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 19.92 | 0.1 | 16.7 | 0.1 | 15.82 | 0.1 | 8.4 | 4.59 | 4.79 | 6.15 | 4.79 | 21.49 | 9.77 | 24.61 | 9.77 | 24.61 | 9.77 | 24.51
75.008 | 0.1 | 0.78 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.51 | 4.79 | 14.94 | 4.49 | 6.64 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.51 | 0.1 | 17.38 | 0.1 | 16.6 | 0.1 | 10.35 | 4.59 | 4.79 | 5.57 | 4.88 | 21.29 | 9.77 | 24.51 | 9.77 | 24.51 | 9.77 | 24.41
75.04 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.9 | 4.79 | 15.43 | 4.49 | 8.01 | 9.67 | 10.16 | 9.86 | 9.86 | 21.49 | 9.77 | 21 | 0.1 | 17.97 | 0.1 | 17.29 | 0.1 | 11.72 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 9.77 | 24.41 | 9.77 | 24.41 | 9.86 | 24.32
75.072 | 0 | 0.59 | 0 | 11.04 | 0.1 | 14.36 | 4.88 | 18.07 | 4.79 | 15.92 | 4.59 | 9.08 | 9.67 | 11.04 | 9.86 | 9.86 | 21.49 | 9.77 | 21.68 | 0 | 18.65 | 0.1 | 17.87 | 0 | 12.89 | 4.59 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.41 | 9.77 | 24.22 | 9.86 | 24.22
75.104 | 0.1 | 0.49 | 0 | 3.32 | 0.1 | 14.06 | 4.88 | 18.46 | 4.79 | 16.41 | 4.49 | 9.77 | 9.67 | 11.91 | 9.77 | 9.86 | 21.39 | 9.77 | 22.17 | 0 | 19.24 | 0 | 18.46 | 0 | 13.57 | 4.49 | 4.79 | 5.08 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22
75.136 | 0.1 | 0.59 | 0.1 | 1.86 | 0.1 | 3.91 | 4.88 | 18.16 | 4.88 | 16.8 | 4.59 | 10.45 | 9.67 | 13.38 | 9.77 | 9.86 | 21.49 | 9.77 | 17.68 | 0.1 | 19.92 | 0.1 | 19.14 | 0.1 | 14.36 | 4.49 | 4.79 | 5.18 | 4.88 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22
75.168 | 0.1 | 0.39 | 0.1 | 1.86 | 0.1 | 1.86 | 4.88 | 18.16 | 4.88 | 17.38 | 4.59 | 11.04 | 9.67 | 14.16 | 9.77 | 9.86 | 21.39 | 9.77 | 16.6 | 0.1 | 20.51 | 0.1 | 19.73 | 0.1 | 15.14 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22
75.2 | 0.1 | 0.39 | 0.1 | 1.46 | 0 | 1.95 | 4.88 | 18.26 | 4.88 | 17.77 | 4.69 | 11.62 | 9.57 | 14.75 | 9.67 | 9.86 | 21.39 | 9.77 | 17.29 | 0.1 | 21.09 | 0.1 | 20.31 | 0.1 | 15.72 | 4.59 | 4.79 | 4.88 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22
75.232 | 0.1 | 0.29 | 0 | 1.46 | 0.1 | 1.56 | 4.88 | 18.16 | 4.88 | 18.26 | 4.69 | 12.21 | 9.67 | 15.33 | 9.67 | 9.86 | 11.62 | 9.77 | 17.38 | 0 | 21.68 | 0.1 | 21 | 0 | 16.41 | 4.49 | 4.79 | 4.88 | 4.79 | 9.96 | 9.77 | 24.22 | 9.77 | 24.12 | 9.77 | 24.22
75.264 | 0.2 | 0.2 | 0 | 1.37 | 0.1 | 1.56 | 4.88 | 18.16 | 4.88 | 18.65 | 4.69 | 12.7 | 9.67 | 15.82 | 9.67 | 9.86 | 6.64 | 9.77 | 17.38 | 0.1 | 22.17 | 0 | 21.58 | 0 | 17.19 | 4.59 | 4.79 | 4.88 | 4.79 | 3.03 | 9.77 | 17.77 | 9.77 | 24.22 | 9.77 | 24.12
75.296 | 1.17 | 0.2 | 0.1 | 1.27 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 19.14 | 4.69 | 13.28 | 9.67 | 16.31 | 9.67 | 9.86 | 8.59 | 9.86 | 17.38 | 0.1 | 22.07 | 0.1 | 22.17 | 0.1 | 17.87 | 4.59 | 4.79 | 4.79 | 4.88 | 3.22 | 9.77 | 7.03 | 9.86 | 15.23 | 9.86 | 24.22
75.328 | 1.95 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.07 | 4.88 | 19.53 | 4.69 | 13.77 | 9.67 | 16.8 | 9.67 | 9.86 | 11.33 | 9.77 | 17.38 | 0.1 | 18.56 | 0.1 | 22.66 | 0.1 | 18.46 | 4.59 | 5.76 | 4.79 | 4.79 | 3.61 | 9.77 | 7.03 | 9.86 | 6.54 | 9.86 | 12.79
75.36 | 3.32 | 0.1 | 0.1 | 0.88 | 0.1 | 1.17 | 4.88 | 18.07 | 4.88 | 20.02 | 4.79 | 14.36 | 9.67 | 17.29 | 9.67 | 9.86 | 11.13 | 9.77 | 17.48 | 0.1 | 18.26 | 0.1 | 19.53 | 0.1 | 19.24 | 4.59 | 7.03 | 4.69 | 4.88 | 4.69 | 9.77 | 11.04 | 9.77 | 7.32 | 9.86 | 5.86
75.392 | 4.2 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.41 | 4.79 | 14.84 | 9.67 | 17.77 | 9.67 | 9.86 | 10.16 | 9.77 | 17.48 | 0 | 17.87 | 0 | 18.75 | 0 | 19.82 | 4.69 | 9.08 | 4.69 | 4.79 | 6.15 | 9.77 | 11.91 | 9.77 | 11.43 | 9.77 | 8.3
75.424 | 4.88 | 0.1 | 0 | 0.68 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.8 | 4.79 | 15.43 | 9.67 | 18.16 | 9.57 | 9.86 | 9.86 | 9.77 | 17.48 | 0 | 17.87 | 0.1 | 18.26 | 0 | 20.41 | 4.69 | 10.64 | 4.69 | 4.79 | 5.96 | 9.77 | 10.55 | 9.77 | 11.82 | 9.86 | 11.82
75.456 | 5.37 | 0.1 | 0.1 | 0.68 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 17.77 | 4.79 | 15.82 | 9.67 | 18.65 | 9.57 | 9.86 | 10.06 | 9.86 | 17.38 | 0.1 | 17.87 | 0.1 | 18.07 | 0.1 | 20.9 | 4.69 | 11.91 | 4.59 | 4.88 | 5.57 | 9.77 | 10.16 | 9.86 | 10.55 | 9.86 | 11.62
75.488 | 5.96 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 16.41 | 4.88 | 17.58 | 4.79 | 16.31 | 9.67 | 19.14 | 9.67 | 9.86 | 10.25 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 18.07 | 0.1 | 21.58 | 4.69 | 12.89 | 4.59 | 4.88 | 5.27 | 9.77 | 10.25 | 9.86 | 9.86 | 9.86 | 10.45
75.52 | 6.54 | 0.1 | 0.1 | 0.49 | 0.1 | 0.68 | 4.88 | 4.69 | 4.88 | 17.87 | 4.79 | 16.8 | 9.77 | 19.53 | 9.67 | 9.86 | 10.06 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.97 | 0.1 | 22.17 | 4.79 | 13.77 | 4.59 | 4.88 | 5.27 | 9.77 | 10.45 | 9.86 | 10.25 | 9.86 | 9.96
75.552 | 7.13 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 18.07 | 4.79 | 17.29 | 9.77 | 20.02 | 9.67 | 9.86 | 9.86 | 9.77 | 16.9 | 0 | 17.68 | 0 | 17.77 | 0.1 | 22.85 | 4.79 | 14.55 | 4.59 | 4.79 | 5.08 | 9.77 | 10.55 | 9.77 | 10.45 | 9.86 | 10.35
75.584 | 7.72 | 0.1 | 0 | 0.39 | 0.1 | 0.49 | 4.88 | 3.52 | 4.88 | 17.97 | 4.79 | 17.68 | 9.67 | 20.41 | 9.67 | 9.86 | 9.96 | 9.86 | 4.49 | 0 | 17.68 | 0 | 17.77 | 0 | 23.44 | 4.79 | 15.23 | 4.59 | 4.88 | 5.08 | 9.77 | 10.35 | 9.77 | 10.45 | 9.86 | 10.45
75.616 | 8.4 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 4.88 | 4.69 | 4.88 | 18.07 | 4.88 | 18.16 | 9.77 | 20.9 | 9.67 | 9.86 | 9.96 | 9.86 | 1.86 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 23.83 | 4.79 | 16.02 | 4.59 | 5.86 | 5.18 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86 | 10.45
75.648 | 8.98 | 0.1 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 18.65 | 9.77 | 21.19 | 9.67 | 9.86 | 9.86 | 9.86 | 2.15 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 24.32 | 4.79 | 16.6 | 4.59 | 7.23 | 4.98 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 10.16
75.68 | 9.57 | 0.1 | 0.68 | 0.2 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 19.14 | 9.77 | 21.58 | 9.67 | 9.86 | 9.77 | 9.86 | 1.95 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.81 | 4.79 | 17.19 | 4.49 | 9.28 | 4.98 | 9.77 | 10.25 | 9.86 | 10.16 | 9.86 | 10.06
75.712 | 10.16 | 0 | 1.56 | 0.2 | 0.1 | 0.2 | 4.88 | 5.37 | 4.88 | 18.07 | 4.79 | 19.53 | 9.77 | 22.07 | 9.67 | 10.55 | 9.77 | 9.86 | 1.76 | 0 | 17.58 | 0.1 | 17.68 | 0 | 25.39 | 4.79 | 17.97 | 4.59 | 10.94 | 4.88 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86 | 10.25
75.744 | 10.74 | 0.1 | 2.73 | 0.1 | 0.1 | 0.2 | 4.88 | 5.08 | 4.88 | 18.07 | 4.79 | 19.92 | 9.77 | 22.36 | 9.67 | 11.04 | 9.77 | 9.86 | 1.56 | 0 | 8.59 | 0.1 | 17.68 | 0 | 25.88 | 4.79 | 18.56 | 4.59 | 12.01 | 4.79 | 9.77 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16
75.776 | 11.33 | 0.1 | 3.91 | 0.1 | 0.1 | 0.2 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.31 | 9.77 | 22.85 | 9.67 | 12.01 | 9.67 | 9.86 | 1.46 | 0.1 | 2.05 | 0.1 | 11.82 | 0.1 | 26.37 | 4.88 | 19.24 | 4.59 | 12.89 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
75.808 | 11.91 | 0.1 | 4.59 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.8 | 9.77 | 23.24 | 9.67 | 13.38 | 9.67 | 9.86 | 1.46 | 0.1 | 2.25 | 0.1 | 2.83 | 0.1 | 22.95 | 4.88 | 19.92 | 4.59 | 13.87 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
75.84 | 12.5 | 0.1 | 5.18 | 0.1 | 0.98 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.77 | 9.77 | 23.54 | 9.67 | 14.16 | 9.67 | 9.86 | 1.17 | 0.1 | 1.95 | 0.1 | 2.34 | 0.1 | 22.17 | 4.79 | 20.51 | 4.59 | 14.55 | 4.69 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86 | 9.96
75.872 | 12.99 | 0.1 | 5.76 | 0 | 1.76 | 0.1 | 4.98 | 4.79 | 4.88 | 18.07 | 4.88 | 17.58 | 9.77 | 23.93 | 9.67 | 14.84 | 9.67 | 9.86 | 1.17 | 0 | 1.56 | 0 | 1.95 | 0 | 21.88 | 4.79 | 21 | 4.69 | 15.33 | 4.59 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.96
75.904 | 13.57 | 0.1 | 6.25 | 0 | 3.03 | 0.1 | 5.47 | 4.69 | 4.88 | 6.84 | 4.88 | 17.77 | 9.77 | 24.32 | 9.77 | 15.33 | 9.67 | 9.86 | 1.07 | 0 | 1.66 | 0.1 | 1.66 | 0 | 21.68 | 4.79 | 21.68 | 4.69 | 16.02 | 4.59 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
75.936 | 13.97 | 0.1 | 6.93 | 0.1 | 4.1 | 0.1 | 6.54 | 4.69 | 4.88 | 2.64 | 4.88 | 18.07 | 9.86 | 21 | 9.77 | 15.82 | 9.67 | 9.86 | 0.98 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 21.68 | 4.88 | 22.27 | 4.69 | 16.7 | 4.59 | 10.74 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
75.968 | 14.55 | 0.1 | 7.52 | 0.1 | 4.79 | 0.1 | 7.81 | 4.69 | 4.88 | 3.42 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 16.41 | 9.67 | 9.86 | 0.88 | 0.1 | 1.37 | 0.1 | 1.56 | 0.1 | 21.58 | 4.88 | 22.85 | 4.69 | 17.38 | 4.59 | 12.21 | 9.77 | 9.77 | 9.77 | 10.94 | 9.86
76 | 15.04 | 0.1 | 8.11 | 0.1 | 5.47 | 0.1 | 8.89 | 4.59 | 4.88 | 4.39 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 16.9 | 9.67 | 9.86 | 0.78 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 21.39 | 4.88 | 23.34 | 4.69 | 18.07 | 4.49 | 14.16 | 9.67 | 10.94 | 9.77 | 11.91 | 9.77
76.032 | 15.63 | 0.1 | 8.69 | 0 | 5.86 | 0.1 | 9.77 | 4.49 | 4.88 | 5.57 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 17.29 | 9.67 | 9.86 | 0.68 | 0 | 1.17 | 0 | 1.27 | 0 | 21.49 | 4.88 | 23.93 | 4.69 | 18.65 | 4.59 | 15.82 | 9.67 | 12.01 | 9.67 | 13.67 | 9.67
76.064 | 16.02 | 0.1 | 9.38 | 0.1 | 6.45 | 0.1 | 10.35 | 4.49 | 4.88 | 5.86 | 4.79 | 17.97 | 9.77 | 21.49 | 9.77 | 17.77 | 9.57 | 9.86 | 0.59 | 0.1 | 1.07 | 0.1 | 1.17 | 0 | 21.39 | 4.79 | 24.41 | 4.79 | 19.34 | 4.49 | 16.8 | 9.67 | 13.67 | 9.67 | 15.33 | 9.67
76.096 | 16.5 | 0.1 | 9.86 | 0.1 | 7.03 | 0.1 | 11.04 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.26 | 9.67 | 9.86 | 0.59 | 0.1 | 1.07 | 0.1 | 1.27 | 0.1 | 21.39 | 4.88 | 24.81 | 4.79 | 19.92 | 4.59 | 17.58 | 9.67 | 15.43 | 9.67 | 16.5 | 9.67
76.128 | 16.99 | 0.1 | 10.45 | 0.1 | 7.72 | 0.1 | 11.62 | 4.49 | 4.88 | 5.27 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.75 | 9.67 | 9.86 | 0.49 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 25.39 | 4.79 | 20.51 | 4.59 | 18.46 | 9.67 | 16.6 | 9.67 | 17.38 | 9.67
76.16 | 14.06 | 0.1 | 11.13 | 0.1 | 8.3 | 0.1 | 12.21 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 19.14 | 9.67 | 9.86 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 21.19 | 4.49 | 19.14 | 9.67 | 17.58 | 9.67 | 18.07 | 9.67
76.192 | 13.97 | 0.1 | 11.72 | 0 | 8.89 | 0 | 12.7 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 19.53 | 9.67 | 9.86 | 0.29 | 0 | 0.68 | 0 | 0.88 | 0 | 21.29 | 4.88 | 24.41 | 4.79 | 21.68 | 4.59 | 19.82 | 9.67 | 18.16 | 9.67 | 18.95 | 9.67
76.224 | 14.16 | 0.1 | 12.21 | 0 | 9.47 | 0.1 | 13.28 | 4.49 | 4.88 | 4.79 | 4.79 | 17.97 | 9.86 | 21.49 | 9.77 | 19.92 | 9.67 | 9.96 | 0.29 | 0.1 | 0.59 | 0.1 | 0.68 | 0 | 15.04 | 4.79 | 21.78 | 4.79 | 22.27 | 4.59 | 20.31 | 9.67 | 18.95 | 9.57 | 19.43 | 9.67
76.256 | 14.45 | 0.1 | 12.79 | 0.1 | 10.06 | 0.1 | 13.87 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 20.41 | 9.67 | 10.55 | 0.29 | 0.1 | 0.68 | 0.1 | 0.59 | 0.1 | 3.52 | 4.88 | 21.68 | 4.79 | 22.85 | 4.59 | 21 | 9.67 | 19.53 | 9.67 | 20.12 | 9.67
76.288 | 14.36 | 0.1 | 13.28 | 0.1 | 10.55 | 0.1 | 14.36 | 4.49 | 4.88 | 4.79 | 4.88 | 6.84 | 9.86 | 21.39 | 9.86 | 20.8 | 9.67 | 11.52 | 0.2 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 3.32 | 4.88 | 21.49 | 4.88 | 23.34 | 4.69 | 21.49 | 9.67 | 20.12 | 9.67 | 20.7 | 9.67
76.32 | 14.36 | 0.1 | 13.87 | 0.1 | 11.23 | 0.1 | 14.94 | 4.49 | 4.88 | 4.79 | 4.79 | 2.64 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 12.79 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.42 | 4.88 | 21.39 | 4.79 | 23.83 | 4.69 | 22.17 | 9.67 | 20.8 | 9.67 | 21.29 | 9.67
76.352 | 14.45 | 0.1 | 14.45 | 0 | 11.82 | 0 | 15.43 | 4.49 | 4.88 | 4.79 | 4.88 | 3.32 | 9.86 | 17.97 | 9.86 | 21.68 | 9.67 | 13.77 | 0.2 | 0 | 0.39 | 0 | 0.49 | 0.1 | 4.2 | 4.88 | 21.29 | 4.79 | 24.41 | 4.69 | 22.66 | 9.67 | 21.39 | 9.67 | 21.88 | 9.67
76.384 | 14.36 | 0 | 14.84 | 0 | 12.4 | 0.1 | 15.92 | 4.49 | 4.88 | 4.59 | 4.88 | 4.39 | 9.77 | 8.69 | 9.86 | 22.07 | 9.67 | 14.45 | 0.1 | 0 | 0.29 | 0 | 0.39 | 0 | 5.76 | 4.79 | 21.29 | 4.88 | 24.9 | 4.69 | 23.24 | 9.67 | 21.97 | 9.67 | 22.46 | 9.67
76.416 | 14.45 | 0.1 | 15.43 | 0.1 | 12.89 | 0.1 | 16.41 | 4.59 | 4.88 | 4.59 | 4.88 | 5.57 | 9.86 | 7.23 | 9.86 | 22.46 | 9.67 | 15.04 | 0.2 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 6.25 | 4.88 | 21.29 | 4.88 | 25.39 | 4.69 | 23.73 | 9.67 | 22.46 | 9.67 | 23.05 | 9.67
76.448 | 14.45 | 0.1 | 15.92 | 0.1 | 13.48 | 0.1 | 16.9 | 4.59 | 4.88 | 4.59 | 4.88 | 5.86 | 9.86 | 10.25 | 9.86 | 22.75 | 9.67 | 15.63 | 0.1 | 0.1 | 0.29 | 1.07 | 0.29 | 0.1 | 5.86 | 4.88 | 21.29 | 4.88 | 25.98 | 4.79 | 24.32 | 9.67 | 23.05 | 9.67 | 23.54 | 9.67
76.48 | 14.36 | 0.1 | 16.41 | 0.1 | 13.97 | 0.1 | 17.29 | 4.69 | 4.88 | 4.59 | 4.88 | 5.47 | 9.86 | 11.52 | 9.86 | 23.15 | 9.77 | 16.11 | 0.1 | 0.1 | 0.2 | 2.05 | 0.29 | 0.1 | 5.37 | 4.88 | 21.19 | 4.79 | 24.22 | 4.79 | 24.81 | 9.77 | 23.63 | 9.67 | 24.02 | 9.67
76.512 | 14.36 | 0.1 | 16.9 | 0 | 14.45 | 0 | 17.77 | 4.59 | 5.08 | 4.49 | 4.88 | 5.08 | 9.86 | 10.55 | 9.86 | 23.54 | 9.67 | 16.6 | 0.1 | 1.17 | 0.2 | 3.91 | 0.2 | 0 | 5.18 | 4.88 | 21.19 | 4.79 | 20.61 | 4.79 | 25.39 | 9.67 | 24.22 | 9.67 | 24.61 | 9.67
76.544 | 14.36 | 0.1 | 17.29 | 0.1 | 15.04 | 0.1 | 18.26 | 4.59 | 5.66 | 4.49 | 4.79 | 4.79 | 9.86 | 9.96 | 9.86 | 23.93 | 9.67 | 17.09 | 0.1 | 2.44 | 0.1 | 5.47 | 0.2 | 0 | 4.98 | 4.79 | 21.19 | 4.88 | 21.39 | 4.79 | 25.88 | 9.77 | 24.61 | 9.67 | 25.1 | 9.67
76.576 | 14.36 | 0.1 | 15.43 | 0.1 | 15.53 | 0.1 | 18.65 | 4.69 | 6.74 | 4.59 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 24.32 | 9.77 | 17.58 | 0.1 | 4.3 | 0.2 | 6.64 | 0.2 | 0.2 | 5.08 | 4.88 | 21.19 | 4.88 | 21.39 | 4.79 | 26.27 | 9.77 | 25.1 | 9.67 | 25.59 | 9.77
76.608 | 5.66 | 0.1 | 15.23 | 0.1 | 16.02 | 0.1 | 19.14 | 4.69 | 8.11 | 4.59 | 4.88 | 4.88 | 9.86 | 10.25 | 9.86 | 21.29 | 9.77 | 18.07 | 0.1 | 5.76 | 0.1 | 7.52 | 0.1 | 1.46 | 5.08 | 4.88 | 21.19 | 4.88 | 21.29 | 4.79 | 26.76 | 9.77 | 25.68 | 9.77 | 26.27 | 9.77
76.64 | 1.76 | 0.1 | 14.84 | 0.1 | 16.5 | 0.1 | 19.53 | 4.69 | 9.08 | 4.49 | 4.88 | 4.88 | 9.86 | 10.35 | 9.86 | 21.68 | 9.77 | 18.46 | 0.1 | 6.84 | 0.1 | 8.2 | 0.1 | 2.73 | 4.98 | 4.88 | 8.69 | 4.79 | 21.29 | 4.79 | 27.25 | 9.77 | 26.27 | 9.77 | 26.66 | 9.77
76.672 | 1.86 | 0.1 | 14.65 | 0 | 16.99 | 0 | 20.02 | 4.69 | 9.86 | 4.59 | 4.79 | 4.69 | 9.77 | 9.96 | 9.86 | 21.58 | 9.77 | 18.95 | 0 | 7.62 | 0.1 | 8.98 | 0.1 | 4.69 | 4.88 | 4.79 | 2.44 | 4.79 | 21.19 | 4.79 | 27.83 | 9.77 | 26.76 | 9.77 | 27.05 | 9.77
76.704 | 1.76 | 0 | 14.75 | 0 | 16.7 | 0 | 20.41 | 4.69 | 10.45 | 4.49 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 19.34 | 0.1 | 8.4 | 0 | 9.77 | 0.1 | 5.96 | 4.88 | 4.79 | 3.13 | 4.88 | 21.19 | 4.79 | 28.32 | 9.77 | 27.15 | 9.77 | 27.64 | 9.77
76.736 | 1.56 | 0.1 | 14.55 | 0.1 | 14.65 | 0.1 | 20.8 | 4.79 | 11.13 | 4.59 | 4.88 | 4.69 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 19.82 | 0.1 | 9.28 | 0.1 | 10.64 | 0.1 | 7.03 | 4.79 | 4.88 | 3.81 | 4.88 | 21.29 | 4.79 | 28.91 | 9.77 | 27.74 | 9.77 | 28.03 | 9.77
76.768 | 1.37 | 0.1 | 14.55 | 0.1 | 14.65 | 0.1 | 19.63 | 4.79 | 11.72 | 4.59 | 4.88 | 4.59 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.22 | 0.1 | 9.96 | 0.1 | 11.43 | 0.1 | 7.81 | 4.69 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 29.2 | 9.86 | 28.13 | 9.77 | 28.52 | 9.77
76.8 | 1.17 | 0.1 | 14.55 | 0.1 | 14.65 | 0.1 | 17.68 | 4.79 | 12.31 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.61 | 0.1 | 10.74 | 0.1 | 12.31 | 0.1 | 8.59 | 4.79 | 4.79 | 6.05 | 4.79 | 21.19 | 4.79 | 26.56 | 9.86 | 28.61 | 9.77 | 28.91 | 9.77
76.832 | 1.07 | 0.1 | 14.45 | 0 | 14.55 | 0 | 17.97 | 4.79 | 12.89 | 4.59 | 4.79 | 4.49 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21 | 0 | 11.62 | 0 | 13.09 | 0 | 9.47 | 4.69 | 4.79 | 5.86 | 4.88 | 21.19 | 4.79 | 23.44 | 9.77 | 29.1 | 9.77 | 29.4 | 9.77
76.864 | 0.88 | 0.1 | 14.45 | 0 | 14.55 | 0.1 | 18.26 | 4.79 | 13.48 | 4.69 | 4.79 | 4.49 | 9.77 | 9.77 | 9.86 | 21.39 | 9.77 | 21.49 | 0.1 | 12.4 | 0 | 13.87 | 0 | 10.16 | 4.69 | 4.79 | 5.47 | 4.88 | 21.19 | 4.79 | 24.22 | 9.77 | 29.4 | 9.77 | 26.08 | 9.77
76.896 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.16 | 4.79 | 13.87 | 4.69 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.88 | 0.1 | 13.09 | 0.1 | 14.55 | 0.1 | 11.04 | 4.59 | 4.88 | 5.47 | 4.88 | 21.19 | 4.88 | 24.32 | 9.86 | 25.78 | 9.77 | 25.2 | 9.86
76.928 | 0.78 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 14.45 | 4.69 | 5.47 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.27 | 0.1 | 13.87 | 0.1 | 15.14 | 0.1 | 11.72 | 4.59 | 4.79 | 5.08 | 4.88 | 8.11 | 4.79 | 24.32 | 9.86 | 25 | 9.86 | 24.81 | 9.86
76.96 | 0.68 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.16 | 4.79 | 14.94 | 4.79 | 6.45 | 4.59 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.66 | 0.1 | 14.55 | 0.1 | 16.02 | 0.1 | 12.6 | 4.59 | 4.79 | 5.18 | 4.79 | 2.34 | 4.88 | 24.32 | 9.86 | 24.71 | 9.86 | 24.51 | 9.86
76.992 | 0.59 | 0.1 | 8.01 | 0 | 14.45 | 0 | 18.07 | 4.79 | 15.53 | 4.79 | 7.91 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 23.05 | 0.1 | 15.33 | 0 | 16.8 | 0 | 13.28 | 4.49 | 5.76 | 5.08 | 4.79 | 3.42 | 4.79 | 24.32 | 9.77 | 24.51 | 9.77 | 24.41 | 9.86
77.024 | 0.49 | 0 | 2.34 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 15.92 | 4.79 | 8.79 | 4.49 | 9.77 | 9.67 | 9.86 | 19.24 | 9.86 | 23.44 | 0.1 | 15.92 | 0 | 17.29 | 0 | 13.97 | 4.49 | 6.25 | 4.98 | 4.88 | 3.91 | 4.79 | 24.32 | 9.77 | 24.51 | 9.77 | 24.41 | 9.86
77.056 | 0.49 | 0.1 | 2.15 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.41 | 4.79 | 9.67 | 4.49 | 9.77 | 9.67 | 9.86 | 8.59 | 9.86 | 23.73 | 0.1 | 16.6 | 0.1 | 17.97 | 0.1 | 14.84 | 4.49 | 7.81 | 4.98 | 4.88 | 5.18 | 4.79 | 24.32 | 9.86 | 24.41 | 9.86 | 24.22 | 9.86
77.088 | 0.49 | 0.1 | 1.66 | 0.1 | 14.45 | 0.1 | 18.07 | 4.88 | 16.9 | 4.79 | 10.25 | 4.59 | 10.55 | 9.67 | 9.86 | 7.23 | 9.86 | 24.12 | 0.1 | 17.38 | 0.1 | 18.56 | 0.1 | 15.53 | 4.49 | 9.96 | 4.88 | 4.88 | 6.25 | 4.88 | 24.32 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86
77.12 | 0.29 | 0.1 | 1.37 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 17.38 | 4.79 | 10.94 | 4.59 | 11.04 | 9.57 | 9.86 | 10.16 | 9.86 | 23.73 | 0.1 | 18.07 | 0.1 | 19.24 | 0.1 | 16.21 | 4.49 | 11.43 | 4.88 | 4.79 | 6.05 | 4.79 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
77.152 | 0.29 | 0.1 | 1.46 | 0 | 7.13 | 0 | 18.07 | 4.88 | 17.87 | 4.79 | 11.52 | 4.59 | 12.31 | 9.57 | 9.86 | 11.43 | 9.86 | 21.09 | 0 | 18.65 | 0 | 19.92 | 0 | 16.9 | 4.49 | 12.5 | 4.69 | 4.88 | 5.47 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
77.184 | 0.2 | 0 | 1.27 | 0 | 2.05 | 0 | 18.07 | 4.79 | 18.36 | 4.79 | 12.21 | 4.59 | 13.38 | 9.67 | 9.86 | 10.64 | 9.86 | 21.68 | 0.1 | 19.24 | 0 | 20.51 | 0 | 17.38 | 4.49 | 13.18 | 4.69 | 4.88 | 5.18 | 4.79 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
77.216 | 0.2 | 0.1 | 1.17 | 0.1 | 2.05 | 0.1 | 7.72 | 4.88 | 18.75 | 4.88 | 12.7 | 4.59 | 14.16 | 9.67 | 9.86 | 9.96 | 9.86 | 21.68 | 0.1 | 19.82 | 0.1 | 21 | 0.1 | 18.16 | 4.59 | 14.06 | 4.69 | 4.88 | 5.18 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
77.248 | 0.2 | 0.1 | 1.07 | 0.1 | 1.76 | 0.1 | 2.64 | 4.88 | 19.24 | 4.88 | 13.28 | 4.69 | 14.75 | 9.67 | 9.86 | 9.77 | 9.86 | 21.58 | 0.1 | 20.51 | 0.1 | 21.58 | 0.1 | 18.85 | 4.59 | 14.84 | 4.69 | 4.88 | 5.08 | 4.79 | 11.33 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
77.28 | 0.1 | 0.1 | 0.88 | 0.1 | 1.56 | 0.1 | 3.22 | 4.88 | 19.63 | 4.88 | 13.77 | 4.69 | 15.33 | 9.67 | 9.86 | 10.06 | 9.86 | 21.58 | 0.1 | 21 | 0.1 | 22.17 | 0.1 | 19.53 | 4.59 | 15.53 | 4.59 | 4.79 | 5.08 | 4.79 | 5.57 | 9.86 | 24.22 | 9.86 | 14.55 | 9.86
77.312 | 0.1 | 0.1 | 0.78 | 0 | 1.37 | 0 | 4.1 | 4.79 | 20.12 | 4.88 | 14.36 | 4.59 | 15.92 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 0 | 21.68 | 0.1 | 22.56 | 0 | 20.12 | 4.59 | 16.31 | 4.59 | 4.79 | 5.08 | 4.79 | 8.98 | 9.86 | 12.7 | 9.86 | 6.15 | 9.86
77.344 | 0 | 0.2 | 0.78 | 0.1 | 1.27 | 0 | 5.27 | 4.79 | 20.51 | 4.79 | 14.75 | 4.69 | 16.21 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 0 | 22.17 | 0 | 19.43 | 0 | 20.7 | 4.69 | 16.9 | 4.49 | 4.88 | 4.98 | 4.79 | 12.21 | 9.77 | 6.15 | 9.86 | 7.42 | 9.86
77.376 | 0.1 | 0.88 | 0.59 | 0.1 | 1.07 | 0.1 | 5.86 | 4.88 | 20.9 | 4.88 | 15.33 | 4.69 | 16.8 | 9.67 | 11.13 | 9.86 | 9.96 | 21.49 | 0.1 | 20.51 | 0.1 | 18.65 | 0.1 | 21.39 | 4.69 | 17.58 | 4.59 | 4.88 | 5.08 | 4.88 | 11.33 | 9.86 | 7.62 | 9.86 | 11.43 | 9.96
77.408 | 0.1 | 1.86 | 0.59 | 0.1 | 1.07 | 0.1 | 5.66 | 4.88 | 17.58 | 4.88 | 15.82 | 4.69 | 17.29 | 9.67 | 12.11 | 9.86 | 9.96 | 21.49 | 0.1 | 17.19 | 0.1 | 18.26 | 0.1 | 21.97 | 4.69 | 18.16 | 4.59 | 4.88 | 4.88 | 4.88 | 10.25 | 9.86 | 11.62 | 9.86 | 11.82 | 9.86
77.44 | 0.1 | 3.13 | 0.49 | 0.1 | 0.88 | 0.1 | 5.27 | 4.88 | 17.58 | 4.88 | 16.31 | 4.79 | 17.77 | 9.67 | 13.48 | 9.86 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 22.36 | 4.69 | 18.85 | 4.59 | 4.88 | 4.79 | 4.88 | 10.16 | 9.86 | 11.72 | 9.86 | 10.55 | 9.86
77.472 | 0 | 4.1 | 0.49 | 0 | 0.78 | 0 | 4.88 | 4.79 | 17.87 | 4.88 | 16.8 | 4.69 | 18.26 | 9.67 | 14.26 | 9.77 | 9.86 | 21.39 | 0 | 17.58 | 0 | 17.97 | 0.1 | 17.77 | 4.69 | 19.53 | 4.49 | 5.86 | 4.79 | 4.79 | 10.25 | 9.86 | 10.45 | 9.86 | 9.86 | 9.86
77.504 | 0 | 4.79 | 0.39 | 0 | 0.78 | 0.1 | 4.88 | 4.79 | 18.07 | 4.88 | 17.29 | 4.79 | 18.65 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 0.1 | 17.68 | 0 | 17.87 | 0 | 16.9 | 4.79 | 20.12 | 4.49 | 6.93 | 4.69 | 4.79 | 10.35 | 9.77 | 9.96 | 9.86 | 10.16 | 9.86
77.536 | 0.1 | 5.37 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 17.77 | 4.79 | 19.24 | 9.67 | 15.33 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.58 | 4.79 | 20.8 | 4.49 | 8.69 | 4.69 | 4.79 | 10.35 | 9.86 | 10.16 | 9.86 | 10.35 | 9.96
77.568 | 0.1 | 5.86 | 0.29 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 18.16 | 4.79 | 19.53 | 9.77 | 15.82 | 9.77 | 9.86 | 12.5 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.77 | 4.79 | 21.39 | 4.59 | 10.45 | 4.59 | 4.79 | 10.16 | 9.86 | 10.35 | 9.86 | 10.55 | 9.86
77.6 | 0.1 | 6.45 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.56 | 4.79 | 20.02 | 9.77 | 16.31 | 9.67 | 9.86 | 7.03 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.68 | 4.79 | 21.88 | 4.59 | 11.82 | 4.59 | 4.79 | 10.06 | 9.86 | 10.35 | 9.86 | 10.25 | 9.86
77.632 | 0.1 | 7.03 | 0.2 | 0 | 0.49 | 0 | 4.79 | 4.79 | 18.07 | 4.88 | 19.14 | 4.79 | 20.51 | 9.77 | 16.9 | 9.67 | 9.86 | 8.5 | 0.1 | 17.58 | 0 | 17.68 | 0 | 17.68 | 4.79 | 22.46 | 4.59 | 12.7 | 4.59 | 4.79 | 10.16 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
77.664 | 0 | 7.62 | 0.1 | 0.1 | 0.39 | 0 | 4.69 | 4.79 | 18.07 | 4.88 | 19.43 | 4.79 | 20.8 | 9.77 | 17.38 | 9.67 | 9.86 | 11.23 | 0.1 | 17.58 | 0 | 17.68 | 0 | 17.77 | 4.79 | 22.95 | 4.59 | 13.67 | 4.49 | 4.79 | 10.06 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
77.696 | 0.1 | 8.3 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.92 | 4.79 | 21.29 | 9.77 | 17.77 | 9.67 | 9.96 | 11.33 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.54 | 4.59 | 14.45 | 4.59 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
77.728 | 0.1 | 8.89 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.31 | 4.79 | 21.68 | 9.77 | 18.26 | 9.67 | 9.86 | 10.35 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 17.77 | 4.88 | 24.12 | 4.69 | 15.14 | 4.59 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
77.76 | 0.1 | 9.47 | 0.1 | 0.39 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.8 | 4.79 | 22.07 | 9.77 | 18.65 | 9.57 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 9.67 | 0.1 | 17.77 | 4.88 | 24.61 | 4.69 | 16.02 | 4.59 | 4.79 | 9.86 | 9.86 | 10.06 | 9.86 | 9.86 | 9.86
77.792 | 0 | 10.16 | 0 | 1.27 | 0.2 | 0.59 | 4.49 | 4.79 | 18.07 | 4.88 | 18.65 | 4.79 | 22.56 | 9.77 | 19.14 | 9.57 | 9.86 | 9.86 | 0 | 17.58 | 0 | 2.05 | 0 | 17.68 | 4.88 | 25.1 | 4.69 | 16.5 | 4.49 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
77.824 | 0 | 10.64 | 0 | 2.25 | 0.2 | 1.17 | 4.49 | 4.79 | 18.07 | 4.88 | 17.68 | 4.79 | 22.85 | 9.77 | 19.53 | 9.57 | 9.86 | 10.25 | 0.1 | 7.23 | 0 | 2.34 | 0 | 17.68 | 4.88 | 25.59 | 4.69 | 17.29 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
77.856 | 0.1 | 11.23 | 0.1 | 3.71 | 0.1 | 2.25 | 4.59 | 4.88 | 6.54 | 4.88 | 18.07 | 4.88 | 23.24 | 9.86 | 20.02 | 9.57 | 9.86 | 10.16 | 0.1 | 1.95 | 0.1 | 2.05 | 0.1 | 17.68 | 4.88 | 26.08 | 4.69 | 17.87 | 4.59 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
77.888 | 0.1 | 11.82 | 0.1 | 4.49 | 0.1 | 3.61 | 4.49 | 4.88 | 2.64 | 4.88 | 18.26 | 4.79 | 23.63 | 9.77 | 20.51 | 9.57 | 9.96 | 9.96 | 0.1 | 2.25 | 0.1 | 1.66 | 0.1 | 17.68 | 4.88 | 21.58 | 4.69 | 18.56 | 4.59 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
77.92 | 0.1 | 12.4 | 0.1 | 5.08 | 0.1 | 4.49 | 4.49 | 4.88 | 3.32 | 4.88 | 18.07 | 4.88 | 23.93 | 9.86 | 20.9 | 9.57 | 10.45 | 9.86 | 0.1 | 2.15 | 0.1 | 1.66 | 0.1 | 6.45 | 4.88 | 20.51 | 4.69 | 19.24 | 4.59 | 5.76 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
77.952 | 0 | 12.99 | 0.1 | 5.57 | 0.1 | 5.08 | 4.49 | 4.88 | 4.59 | 4.88 | 18.07 | 4.79 | 24.41 | 9.77 | 21.29 | 9.57 | 11.52 | 9.86 | 0 | 1.66 | 0 | 1.46 | 0 | 1.95 | 4.88 | 21.49 | 4.69 | 19.92 | 4.59 | 6.84 | 9.67 | 10.64 | 9.77 | 9.86 | 9.67 | 9.86
77.984 | 0 | 13.38 | 0.1 | 6.05 | 0.1 | 5.47 | 4.49 | 4.79 | 5.57 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 21.58 | 9.57 | 12.79 | 9.86 | 0 | 1.66 | 0 | 1.37 | 0 | 2.25 | 4.88 | 21.19 | 4.69 | 20.51 | 4.59 | 8.5 | 9.67 | 11.52 | 9.77 | 10.74 | 9.67 | 9.86
78.016 | 0.1 | 13.97 | 0.1 | 6.74 | 0.1 | 6.05 | 4.49 | 4.79 | 5.76 | 4.88 | 18.07 | 4.79 | 21.97 | 9.86 | 22.07 | 9.67 | 13.77 | 9.77 | 0.1 | 1.46 | 0 | 1.37 | 0.1 | 2.05 | 4.88 | 21.29 | 4.79 | 21 | 4.69 | 10.35 | 9.67 | 13.18 | 9.77 | 11.72 | 9.67 | 11.04
78.048 | 0.1 | 14.45 | 0.1 | 7.32 | 0.1 | 6.64 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 22.36 | 9.67 | 14.45 | 9.77 | 0.1 | 1.46 | 0.1 | 1.27 | 0.1 | 1.76 | 4.88 | 21.29 | 4.79 | 21.58 | 4.69 | 11.72 | 9.67 | 15.14 | 9.67 | 13.28 | 9.67 | 12.01
78.08 | 0.1 | 15.04 | 0.1 | 7.91 | 0.1 | 7.32 | 4.59 | 4.79 | 5.27 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 22.75 | 9.67 | 15.04 | 9.77 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.29 | 4.79 | 22.17 | 4.69 | 12.6 | 9.67 | 16.31 | 9.67 | 15.14 | 9.67 | 13.77
78.112 | 0 | 15.53 | 0.1 | 8.5 | 0.1 | 7.91 | 4.59 | 4.79 | 4.98 | 4.88 | 17.97 | 4.79 | 21.68 | 9.77 | 23.24 | 9.67 | 15.63 | 9.67 | 0 | 1.17 | 0 | 0.98 | 0 | 1.56 | 4.88 | 21.29 | 4.79 | 22.66 | 4.69 | 13.57 | 9.67 | 17.29 | 9.67 | 16.41 | 9.57 | 15.63
78.144 | 0 | 16.02 | 0.1 | 9.08 | 0.1 | 8.4 | 4.59 | 4.79 | 4.88 | 4.88 | 18.07 | 4.79 | 21.49 | 9.86 | 23.63 | 9.67 | 16.11 | 9.67 | 0 | 0.98 | 0 | 0.88 | 0 | 1.27 | 4.88 | 21.19 | 4.79 | 23.24 | 4.69 | 14.36 | 9.57 | 18.07 | 9.67 | 17.29 | 9.67 | 16.6
78.176 | 0.1 | 16.5 | 0.1 | 9.77 | 0.1 | 9.08 | 4.69 | 5.18 | 4.88 | 4.88 | 18.07 | 4.79 | 21.49 | 9.86 | 23.93 | 9.67 | 16.7 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 23.73 | 4.79 | 15.14 | 9.67 | 18.75 | 9.67 | 18.07 | 9.57 | 17.58
78.208 | 0.1 | 16.99 | 0.1 | 10.35 | 0.1 | 9.67 | 4.69 | 6.05 | 4.88 | 4.88 | 17.97 | 4.79 | 21.49 | 9.86 | 24.32 | 9.67 | 17.09 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.17 | 4.88 | 21.19 | 4.79 | 24.32 | 4.79 | 15.82 | 9.67 | 19.34 | 9.67 | 18.85 | 9.67 | 18.26
78.24 | 0.1 | 15.92 | 0.1 | 10.94 | 0.1 | 10.25 | 4.69 | 6.93 | 4.79 | 4.88 | 7.72 | 4.79 | 21.39 | 9.86 | 21.29 | 9.67 | 17.58 | 9.67 | 0.29 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.19 | 4.79 | 24.9 | 4.79 | 16.5 | 9.67 | 20.02 | 9.67 | 19.43 | 9.57 | 18.95
78.272 | 0 | 14.06 | 0.1 | 11.52 | 0.1 | 10.94 | 4.69 | 8.4 | 4.69 | 4.88 | 2.54 | 4.79 | 21.39 | 9.86 | 22.07 | 9.67 | 18.07 | 9.57 | 1.46 | 0.78 | 0 | 0.49 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.39 | 4.79 | 17.19 | 9.67 | 20.7 | 9.67 | 20.02 | 9.57 | 19.53
78.304 | 0 | 14.45 | 0 | 12.01 | 0.1 | 11.43 | 4.69 | 9.28 | 4.69 | 4.79 | 3.22 | 4.79 | 21.39 | 9.77 | 21.49 | 9.67 | 18.46 | 9.57 | 2.83 | 0.59 | 0 | 0.59 | 0 | 0.88 | 4.88 | 20.9 | 4.79 | 25.78 | 4.79 | 17.77 | 9.67 | 21.19 | 9.67 | 20.61 | 9.57 | 20.22
78.336 | 0.1 | 14.65 | 0.1 | 12.6 | 0.1 | 12.01 | 4.79 | 9.96 | 4.59 | 4.88 | 4.3 | 4.79 | 21.39 | 9.86 | 21.49 | 9.77 | 18.95 | 9.57 | 4.79 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 5.18 | 4.79 | 26.27 | 4.79 | 18.46 | 9.67 | 21.78 | 9.67 | 21.19 | 9.67 | 20.8
78.368 | 0.1 | 14.45 | 0.1 | 13.09 | 0.1 | 12.6 | 4.79 | 10.55 | 4.59 | 4.88 | 5.37 | 4.79 | 21.39 | 9.86 | 21.58 | 9.77 | 19.43 | 9.67 | 6.15 | 0.59 | 0.1 | 0.49 | 0.1 | 0.68 | 4.88 | 2.83 | 4.79 | 23.05 | 4.79 | 19.14 | 9.67 | 22.36 | 9.67 | 21.97 | 9.67 | 21.39
78.4 | 0.1 | 14.45 | 0.1 | 13.67 | 0.1 | 13.09 | 4.79 | 11.23 | 4.59 | 4.88 | 5.86 | 4.79 | 17.87 | 9.86 | 21.49 | 9.77 | 19.82 | 9.67 | 7.03 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 4.88 | 3.32 | 4.79 | 22.07 | 4.79 | 19.73 | 9.67 | 22.95 | 9.67 | 22.36 | 9.67 | 21.97
78.432 | 0 | 14.45 | 0 | 14.26 | 0.1 | 13.67 | 4.79 | 11.82 | 4.59 | 4.88 | 5.57 | 4.79 | 8.01 | 9.86 | 21.49 | 9.67 | 20.22 | 9.57 | 7.81 | 0.39 | 1.07 | 0.29 | 0 | 0.49 | 4.88 | 4.3 | 4.79 | 21.78 | 4.79 | 20.31 | 9.67 | 23.44 | 9.67 | 22.95 | 9.67 | 22.66
78.464 | 0 | 14.45 | 0 | 14.75 | 0.1 | 14.16 | 4.79 | 12.31 | 4.49 | 4.88 | 5.18 | 4.79 | 7.32 | 9.86 | 21.49 | 9.77 | 20.61 | 9.57 | 8.5 | 0.29 | 2.05 | 0.29 | 0.78 | 0.49 | 4.88 | 5.57 | 4.79 | 21.58 | 4.79 | 20.9 | 9.67 | 24.02 | 9.67 | 23.54 | 9.67 | 23.05
78.496 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 14.75 | 4.88 | 12.89 | 4.59 | 4.88 | 4.98 | 4.79 | 10.45 | 9.86 | 21.39 | 9.77 | 21.09 | 9.67 | 9.38 | 0.29 | 3.81 | 0.2 | 1.76 | 0.49 | 4.88 | 6.15 | 4.79 | 21.58 | 4.79 | 21.58 | 9.67 | 24.61 | 9.67 | 24.02 | 9.67 | 23.63
78.528 | 0.1 | 14.45 | 0.1 | 15.72 | 0.1 | 15.23 | 4.88 | 13.48 | 4.49 | 4.88 | 4.88 | 4.79 | 11.43 | 9.86 | 21.49 | 9.77 | 21.49 | 9.67 | 10.06 | 0.29 | 5.47 | 0.2 | 3.32 | 0.49 | 4.88 | 5.86 | 4.79 | 21.49 | 4.79 | 22.07 | 9.67 | 25 | 9.67 | 24.61 | 9.67 | 24.12
78.56 | 0.1 | 14.36 | 0.1 | 16.21 | 0.1 | 15.72 | 4.88 | 13.97 | 4.49 | 4.88 | 4.98 | 4.79 | 10.55 | 9.86 | 21.39 | 9.77 | 21.88 | 9.67 | 10.94 | 0.2 | 6.64 | 0.2 | 5.27 | 0.39 | 4.88 | 5.47 | 4.79 | 21.29 | 4.79 | 22.66 | 9.67 | 25.59 | 9.77 | 25.1 | 9.67 | 24.71
78.592 | 0 | 14.45 | 0 | 16.7 | 0.1 | 16.11 | 4.79 | 14.45 | 4.49 | 5.57 | 4.88 | 4.79 | 9.86 | 9.86 | 21.39 | 9.77 | 22.27 | 9.67 | 11.62 | 0.2 | 7.52 | 0.1 | 6.45 | 0.29 | 4.88 | 5.18 | 4.79 | 21.39 | 4.79 | 23.24 | 9.67 | 26.08 | 9.67 | 25.59 | 9.67 | 25.1
78.624 | 0 | 14.36 | 0 | 17.19 | 0.1 | 16.6 | 4.88 | 14.94 | 4.59 | 6.45 | 4.79 | 4.79 | 9.96 | 9.86 | 21.39 | 9.77 | 22.66 | 9.67 | 12.31 | 0.2 | 8.11 | 0.1 | 7.23 | 0.2 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 23.73 | 9.67 | 26.56 | 9.77 | 26.08 | 9.67 | 25.78
78.656 | 0.1 | 14.36 | 0.1 | 13.97 | 0.1 | 17.19 | 4.88 | 15.53 | 4.59 | 7.81 | 4.79 | 4.79 | 10.16 | 9.86 | 19.73 | 9.86 | 23.05 | 9.67 | 13.18 | 0.2 | 9.08 | 0.1 | 8.11 | 0.2 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 24.22 | 9.77 | 27.05 | 9.77 | 26.56 | 9.77 | 26.27
78.688 | 0.1 | 6.74 | 0.1 | 14.45 | 0.1 | 13.87 | 4.88 | 15.92 | 4.59 | 8.89 | 4.69 | 4.79 | 10.25 | 9.86 | 8.59 | 9.86 | 23.44 | 9.67 | 13.87 | 0.1 | 9.86 | 0.1 | 8.89 | 0.2 | 4.88 | 4.98 | 4.79 | 21.29 | 4.79 | 24.81 | 9.77 | 27.54 | 9.77 | 27.15 | 9.77 | 26.66
78.72 | 0.1 | 1.86 | 0.1 | 14.45 | 0.1 | 14.16 | 4.88 | 16.5 | 4.59 | 9.67 | 4.69 | 4.79 | 9.96 | 9.86 | 7.32 | 9.86 | 23.83 | 9.67 | 14.55 | 0.1 | 10.64 | 0.1 | 9.77 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 25.29 | 9.77 | 28.03 | 9.77 | 27.54 | 9.77 | 27.25
78.752 | 0 | 1.95 | 0.1 | 14.55 | 0.1 | 14.36 | 4.88 | 16.99 | 4.59 | 10.35 | 4.59 | 4.79 | 9.86 | 9.86 | 10.45 | 9.77 | 24.22 | 9.67 | 15.33 | 0.1 | 11.43 | 0.1 | 10.45 | 0.1 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.78 | 9.77 | 28.52 | 9.77 | 28.03 | 9.67 | 27.64
78.784 | 0 | 1.95 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 17.38 | 4.69 | 10.94 | 4.59 | 4.79 | 9.96 | 9.86 | 11.43 | 9.77 | 23.73 | 9.67 | 16.02 | 0.1 | 12.21 | 0.1 | 11.23 | 0.1 | 4.88 | 4.88 | 4.79 | 10.45 | 4.79 | 26.27 | 9.77 | 29.01 | 9.77 | 28.42 | 9.77 | 28.13
78.816 | 0.1 | 1.37 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 17.87 | 4.69 | 11.62 | 4.59 | 4.79 | 9.86 | 9.86 | 10.64 | 9.86 | 21.58 | 9.67 | 16.7 | 0.1 | 12.99 | 0.1 | 12.11 | 0.1 | 4.88 | 4.88 | 4.79 | 2.93 | 4.79 | 23.24 | 9.77 | 29.49 | 9.77 | 29.01 | 9.77 | 28.61
78.848 | 0.1 | 1.37 | 0.1 | 14.55 | 0.1 | 14.45 | 4.88 | 18.36 | 4.69 | 12.21 | 4.49 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 21.68 | 9.67 | 17.29 | 0.1 | 13.87 | 0.1 | 12.79 | 0.1 | 4.88 | 4.69 | 4.79 | 3.13 | 4.79 | 22.27 | 9.77 | 25.29 | 9.77 | 29.4 | 9.77 | 29.1
78.88 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 18.75 | 4.69 | 12.79 | 4.49 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 18.07 | 0.1 | 14.45 | 0.1 | 13.57 | 0.1 | 4.88 | 4.88 | 4.79 | 3.71 | 4.88 | 21.78 | 9.77 | 24.71 | 9.77 | 25.88 | 9.77 | 29.49
78.912 | 0 | 1.07 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.69 | 13.38 | 4.49 | 4.79 | 9.77 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 18.65 | 0 | 15.23 | 0.1 | 14.26 | 0 | 5.76 | 4.69 | 4.79 | 4.69 | 4.79 | 21.58 | 9.77 | 24.71 | 9.77 | 25.2 | 9.77 | 24.61
78.944 | 0 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.63 | 4.79 | 13.87 | 4.49 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.58 | 9.67 | 19.24 | 0.1 | 15.92 | 0.1 | 15.14 | 0 | 6.45 | 4.69 | 4.79 | 5.96 | 4.79 | 21.58 | 9.77 | 24.51 | 9.77 | 24.71 | 9.77 | 25
78.976 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.12 | 4.79 | 14.36 | 4.49 | 5.47 | 9.67 | 9.86 | 10.06 | 9.86 | 21.49 | 9.77 | 19.92 | 0.1 | 16.6 | 0.1 | 15.72 | 0.1 | 8.11 | 4.59 | 4.79 | 6.15 | 4.79 | 21.49 | 9.77 | 24.51 | 9.77 | 24.51 | 9.77 | 24.71
79.008 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.51 | 4.79 | 14.84 | 4.49 | 6.54 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.61 | 0.1 | 17.19 | 0.1 | 16.41 | 0.1 | 10.25 | 4.59 | 4.79 | 5.66 | 4.79 | 21.29 | 9.77 | 24.41 | 9.86 | 24.51 | 9.77 | 24.61
79.04 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.9 | 4.79 | 15.33 | 4.49 | 7.81 | 9.67 | 10.06 | 9.86 | 9.86 | 21.49 | 9.77 | 21 | 0.1 | 17.87 | 0.1 | 17.09 | 0.1 | 11.52 | 4.59 | 4.79 | 5.27 | 4.88 | 21.39 | 9.77 | 24.41 | 9.86 | 24.41 | 9.77 | 24.41
79.072 | 0 | 0.78 | 0 | 13.09 | 0.1 | 14.36 | 4.88 | 18.07 | 4.79 | 15.82 | 4.49 | 8.89 | 9.67 | 11.04 | 9.86 | 9.86 | 21.49 | 9.77 | 21.58 | 0.1 | 18.46 | 0 | 17.77 | 0.1 | 12.7 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 9.77 | 24.41 | 9.77 | 24.32 | 9.77 | 24.32
79.104 | 0 | 0.49 | 0 | 4 | 0.1 | 14.16 | 4.88 | 17.87 | 4.79 | 16.31 | 4.59 | 9.57 | 9.67 | 11.91 | 9.86 | 9.86 | 21.39 | 9.77 | 22.27 | 0.1 | 19.24 | 0.1 | 18.46 | 0 | 13.48 | 4.49 | 4.79 | 5.18 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22
79.136 | 0.1 | 0.59 | 0.1 | 1.95 | 0.1 | 4.39 | 4.88 | 18.07 | 4.79 | 16.8 | 4.59 | 10.35 | 9.67 | 13.28 | 9.77 | 9.86 | 21.39 | 9.77 | 17.68 | 0.1 | 19.92 | 0.1 | 19.14 | 0.1 | 14.26 | 4.59 | 4.79 | 5.08 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
79.168 | 0.1 | 0.49 | 0.1 | 1.86 | 0.1 | 1.86 | 4.88 | 18.07 | 4.79 | 17.29 | 4.59 | 10.94 | 9.67 | 14.06 | 9.77 | 9.86 | 21.39 | 9.77 | 16.6 | 0.1 | 20.41 | 0.1 | 19.73 | 0.1 | 15.04 | 4.59 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22
79.2 | 0.1 | 0.39 | 0.1 | 1.56 | 0.1 | 1.95 | 4.88 | 18.07 | 4.79 | 17.68 | 4.69 | 11.52 | 9.57 | 14.75 | 9.77 | 9.86 | 21.39 | 9.77 | 17.19 | 0.1 | 20.9 | 0.1 | 20.31 | 0.1 | 15.72 | 4.59 | 4.79 | 4.98 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22
79.232 | 0 | 0.29 | 0 | 1.56 | 0.1 | 1.66 | 4.88 | 18.16 | 4.79 | 18.16 | 4.69 | 12.11 | 9.57 | 15.23 | 9.67 | 9.86 | 11.82 | 9.77 | 17.38 | 0 | 21.58 | 0.1 | 20.8 | 0 | 16.5 | 4.59 | 4.79 | 4.88 | 4.79 | 10.55 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22
79.264 | 0.1 | 0.29 | 0 | 1.37 | 0.1 | 1.46 | 4.88 | 18.16 | 4.79 | 18.56 | 4.69 | 12.7 | 9.67 | 15.72 | 9.67 | 9.86 | 6.74 | 9.77 | 17.38 | 0.1 | 22.17 | 0 | 21.39 | 0 | 17.09 | 4.49 | 4.79 | 4.98 | 4.79 | 3.03 | 9.77 | 18.46 | 9.77 | 24.22 | 9.86 | 24.12
79.296 | 0.98 | 0.2 | 0.1 | 1.17 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 19.14 | 4.69 | 13.18 | 9.67 | 16.31 | 9.67 | 9.86 | 8.5 | 9.86 | 17.38 | 0.1 | 22.66 | 0.1 | 22.07 | 0.1 | 17.77 | 4.59 | 4.79 | 4.79 | 4.79 | 3.22 | 9.77 | 7.62 | 9.86 | 17.77 | 9.86 | 24.22
79.328 | 1.86 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.07 | 4.88 | 19.43 | 4.79 | 13.77 | 9.67 | 16.8 | 9.67 | 9.86 | 11.33 | 9.86 | 17.38 | 0.1 | 19.24 | 0.1 | 22.56 | 0.1 | 18.46 | 4.59 | 5.86 | 4.79 | 4.79 | 3.61 | 9.77 | 6.64 | 9.86 | 7.13 | 9.86 | 13.57
79.36 | 3.13 | 0.2 | 0.1 | 0.98 | 0.1 | 1.27 | 4.88 | 18.07 | 4.88 | 19.92 | 4.79 | 14.16 | 9.67 | 17.29 | 9.67 | 9.86 | 11.23 | 9.86 | 17.48 | 0.1 | 18.56 | 0.1 | 19.14 | 0.1 | 19.14 | 4.69 | 7.03 | 4.69 | 4.79 | 4.69 | 9.77 | 10.64 | 9.86 | 6.93 | 9.86 | 6.15
79.392 | 4.1 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.79 | 20.41 | 4.79 | 14.75 | 9.67 | 17.68 | 9.67 | 9.86 | 10.16 | 9.77 | 17.38 | 0 | 17.97 | 0.1 | 18.65 | 0 | 19.73 | 4.59 | 8.98 | 4.69 | 4.79 | 6.05 | 9.77 | 12.01 | 9.77 | 10.94 | 9.77 | 7.91
79.424 | 4.69 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.79 | 20.8 | 4.79 | 15.33 | 9.67 | 18.16 | 9.57 | 9.86 | 9.77 | 9.86 | 17.48 | 0.1 | 17.97 | 0 | 18.07 | 0 | 20.31 | 4.69 | 10.55 | 4.69 | 4.79 | 6.15 | 9.77 | 10.64 | 9.86 | 11.91 | 9.77 | 11.52
79.456 | 5.27 | 0.1 | 0.1 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 19.63 | 4.79 | 15.82 | 9.67 | 18.65 | 9.57 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.97 | 0.1 | 17.97 | 0.1 | 20.9 | 4.69 | 11.91 | 4.59 | 4.79 | 5.76 | 9.77 | 10.25 | 9.86 | 10.74 | 9.86 | 11.91
79.488 | 5.86 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 17.68 | 4.88 | 17.77 | 4.79 | 16.31 | 9.67 | 19.14 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.87 | 0.1 | 17.97 | 0.1 | 21.58 | 4.69 | 12.79 | 4.59 | 4.79 | 5.37 | 9.77 | 10.25 | 9.86 | 10.06 | 9.86 | 10.45
79.52 | 6.54 | 0.1 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 4.69 | 4.79 | 18.16 | 4.79 | 16.8 | 9.67 | 19.53 | 9.67 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.87 | 0.1 | 22.07 | 4.69 | 13.67 | 4.59 | 4.79 | 5.18 | 9.77 | 10.35 | 9.86 | 10.16 | 9.86 | 9.96
79.552 | 7.03 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.79 | 18.16 | 4.88 | 17.19 | 9.67 | 20.02 | 9.57 | 9.86 | 9.86 | 9.86 | 17.19 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.75 | 4.79 | 14.45 | 4.59 | 4.79 | 5.18 | 9.77 | 10.45 | 9.77 | 10.35 | 9.77 | 10.25
79.584 | 7.62 | 0 | 0 | 0.39 | 0.1 | 0.49 | 4.88 | 3.52 | 4.79 | 18.07 | 4.79 | 17.58 | 9.67 | 20.31 | 9.57 | 9.86 | 9.77 | 9.86 | 5.08 | 0.1 | 17.68 | 0.1 | 17.77 | 0 | 23.24 | 4.69 | 15.14 | 4.59 | 4.88 | 5.18 | 9.77 | 10.45 | 9.77 | 10.45 | 9.86 | 10.45
79.616 | 8.3 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 4.88 | 4.79 | 4.79 | 18.16 | 4.88 | 18.07 | 9.77 | 20.8 | 9.67 | 9.86 | 9.96 | 9.86 | 1.86 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.73 | 4.79 | 16.02 | 4.59 | 5.76 | 5.08 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86 | 10.45
79.648 | 8.89 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 18.56 | 9.77 | 21.19 | 9.67 | 9.96 | 9.86 | 9.86 | 2.34 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 24.32 | 4.79 | 16.6 | 4.59 | 7.03 | 5.08 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86 | 10.25
79.68 | 9.47 | 0.1 | 0.68 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 18.95 | 9.77 | 21.68 | 9.67 | 9.86 | 9.77 | 9.86 | 2.15 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.71 | 4.79 | 17.29 | 4.59 | 8.98 | 4.98 | 9.77 | 10.25 | 9.77 | 10.16 | 9.86 | 10.16
79.712 | 10.16 | 0.1 | 1.46 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.79 | 18.07 | 4.88 | 19.43 | 9.77 | 21.97 | 9.67 | 10.64 | 9.77 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 25.39 | 4.79 | 17.97 | 4.59 | 10.74 | 4.88 | 9.77 | 10.16 | 9.77 | 10.16 | 9.77 | 10.16
79.744 | 10.55 | 0.1 | 2.64 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.79 | 18.07 | 4.88 | 19.92 | 9.77 | 22.36 | 9.67 | 11.04 | 9.77 | 9.86 | 1.46 | 0.1 | 8.98 | 0 | 17.68 | 0 | 25.78 | 4.79 | 18.56 | 4.59 | 11.82 | 4.79 | 9.77 | 9.96 | 9.77 | 10.06 | 9.77 | 10.06
79.776 | 11.23 | 0.1 | 3.81 | 0.1 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.31 | 9.77 | 22.75 | 9.67 | 12.11 | 9.77 | 9.86 | 1.37 | 0.1 | 2.05 | 0.1 | 12.99 | 0.1 | 26.27 | 4.79 | 19.14 | 4.59 | 12.89 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
79.808 | 11.82 | 0.1 | 4.59 | 0.1 | 0.1 | 0.1 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 20.7 | 9.77 | 23.15 | 9.67 | 13.38 | 9.67 | 9.86 | 1.37 | 0.1 | 2.34 | 0.1 | 2.93 | 0.1 | 22.17 | 4.88 | 19.92 | 4.69 | 13.67 | 4.79 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86 | 9.86
79.84 | 12.4 | 0.1 | 5.18 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 19.43 | 9.77 | 23.54 | 9.67 | 14.16 | 9.67 | 9.96 | 1.27 | 0.1 | 1.95 | 0.1 | 2.15 | 0.1 | 21.49 | 4.88 | 20.31 | 4.69 | 14.55 | 4.79 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
79.872 | 12.99 | 0.1 | 5.76 | 0 | 1.66 | 0.1 | 4.88 | 4.79 | 4.79 | 18.07 | 4.88 | 17.58 | 9.77 | 23.93 | 9.67 | 14.84 | 9.67 | 9.86 | 0.98 | 0.1 | 1.66 | 0.1 | 2.05 | 0.1 | 21.68 | 4.79 | 21 | 4.69 | 15.23 | 4.59 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
79.904 | 13.48 | 0.1 | 6.25 | 0 | 3.03 | 0.1 | 5.57 | 4.69 | 4.79 | 7.72 | 4.88 | 18.07 | 9.77 | 24.32 | 9.67 | 15.33 | 9.67 | 9.86 | 0.88 | 0.1 | 1.76 | 0.1 | 1.66 | 0 | 21.39 | 4.79 | 21.58 | 4.69 | 16.02 | 4.59 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
79.936 | 13.97 | 0.1 | 6.93 | 0.1 | 3.91 | 0.1 | 6.35 | 4.69 | 4.88 | 2.73 | 4.88 | 18.16 | 9.86 | 20.9 | 9.77 | 15.82 | 9.67 | 9.96 | 0.88 | 0.1 | 1.56 | 0.1 | 1.86 | 0.1 | 21.49 | 4.88 | 22.27 | 4.69 | 16.7 | 4.59 | 10.74 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
79.968 | 14.45 | 0.1 | 7.52 | 0.1 | 4.69 | 0.1 | 7.72 | 4.69 | 4.88 | 3.22 | 4.88 | 18.07 | 9.86 | 21.09 | 9.77 | 16.31 | 9.67 | 9.96 | 0.88 | 0.1 | 1.37 | 0.1 | 1.56 | 0.1 | 21.49 | 4.88 | 22.66 | 4.79 | 17.38 | 4.59 | 12.11 | 9.77 | 9.77 | 9.77 | 10.94 | 9.86
80 | 15.04 | 0.1 | 8.11 | 0.1 | 5.27 | 0.1 | 8.79 | 4.69 | 4.88 | 4.3 | 4.88 | 18.07 | 9.77 | 21.29 | 9.77 | 16.9 | 9.57 | 9.96 | 0.78 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 21.39 | 4.88 | 23.24 | 4.69 | 17.97 | 4.59 | 13.97 | 9.77 | 10.84 | 9.77 | 11.91 | 9.77
80.032 | 15.53 | 0.1 | 8.69 | 0.1 | 5.76 | 0.1 | 9.67 | 4.59 | 4.79 | 5.47 | 4.88 | 17.97 | 9.77 | 21.39 | 9.77 | 17.38 | 9.47 | 9.86 | 0.78 | 0 | 1.17 | 0.1 | 1.37 | 0 | 21.39 | 4.88 | 23.83 | 4.79 | 18.56 | 4.59 | 15.72 | 9.67 | 11.91 | 9.67 | 13.67 | 9.77
80.064 | 15.92 | 0.1 | 9.28 | 0.1 | 6.35 | 0.1 | 10.25 | 4.49 | 4.79 | 5.86 | 4.88 | 17.97 | 9.77 | 21.39 | 9.77 | 17.77 | 9.57 | 9.86 | 0.59 | 0.1 | 1.17 | 0 | 1.27 | 0 | 21.39 | 4.88 | 24.22 | 4.79 | 19.24 | 4.49 | 16.7 | 9.67 | 13.48 | 9.67 | 15.23 | 9.67
80.096 | 16.5 | 0.1 | 9.86 | 0.1 | 7.03 | 0.1 | 10.94 | 4.59 | 4.88 | 5.66 | 4.88 | 18.07 | 9.86 | 21.39 | 9.77 | 18.26 | 9.67 | 9.96 | 0.59 | 0.1 | 0.88 | 0.1 | 1.27 | 0.1 | 21.39 | 4.88 | 24.81 | 4.79 | 19.92 | 4.59 | 17.58 | 9.67 | 15.23 | 9.67 | 16.5 | 9.67
80.128 | 16.99 | 0.1 | 10.45 | 0.1 | 7.52 | 0.1 | 11.52 | 4.59 | 4.79 | 5.18 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.65 | 9.67 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 0.98 | 0.1 | 21.29 | 4.88 | 25.29 | 4.79 | 20.51 | 4.59 | 18.26 | 9.67 | 16.41 | 9.67 | 17.29 | 9.67
80.16 | 15.82 | 0.1 | 11.13 | 0.1 | 8.2 | 0.1 | 12.11 | 4.59 | 4.79 | 4.98 | 4.88 | 17.97 | 9.86 | 21.49 | 9.77 | 19.24 | 9.57 | 9.96 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 21 | 4.59 | 18.95 | 9.67 | 17.38 | 9.67 | 18.07 | 9.67
80.192 | 14.06 | 0.1 | 11.62 | 0 | 8.79 | 0 | 12.7 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.77 | 21.39 | 9.77 | 19.63 | 9.57 | 9.86 | 0.29 | 0.1 | 0.68 | 0.1 | 0.88 | 0 | 21.29 | 4.88 | 24.9 | 4.79 | 21.78 | 4.59 | 19.73 | 9.67 | 18.16 | 9.67 | 18.75 | 9.67
80.224 | 14.36 | 0 | 12.21 | 0.1 | 9.47 | 0.1 | 13.18 | 4.49 | 4.79 | 4.88 | 4.88 | 17.97 | 9.77 | 21.49 | 9.77 | 20.02 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0.1 | 0.78 | 0 | 17.87 | 4.79 | 22.07 | 4.79 | 22.17 | 4.59 | 20.31 | 9.67 | 18.75 | 9.57 | 19.43 | 9.67
80.256 | 14.65 | 0.1 | 12.79 | 0.1 | 9.96 | 0.1 | 13.77 | 4.49 | 4.79 | 4.88 | 4.88 | 17.97 | 9.86 | 21.39 | 9.86 | 20.41 | 9.67 | 10.64 | 0.29 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 4.1 | 4.88 | 21.78 | 4.88 | 22.85 | 4.59 | 20.9 | 9.67 | 19.43 | 9.57 | 20.12 | 9.67
80.288 | 14.45 | 0.1 | 13.28 | 0.1 | 10.55 | 0.1 | 14.26 | 4.59 | 4.88 | 4.79 | 4.88 | 7.72 | 9.86 | 21.39 | 9.86 | 20.9 | 9.67 | 11.43 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.13 | 4.88 | 21.58 | 4.88 | 23.24 | 4.59 | 21.49 | 9.67 | 20.12 | 9.67 | 20.61 | 9.67
80.32 | 14.45 | 0.1 | 13.77 | 0.1 | 11.23 | 0.1 | 14.84 | 4.59 | 4.79 | 4.79 | 4.88 | 2.73 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 12.7 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.42 | 4.88 | 21.49 | 4.79 | 23.83 | 4.69 | 22.07 | 9.67 | 20.7 | 9.67 | 21.19 | 9.67
80.352 | 14.45 | 0.1 | 14.45 | 0 | 11.72 | 0 | 15.33 | 4.59 | 4.79 | 4.69 | 4.88 | 3.32 | 9.86 | 20.9 | 9.77 | 21.58 | 9.67 | 13.77 | 0.2 | 0 | 0.29 | 0.1 | 0.49 | 0.1 | 4.2 | 4.88 | 21.29 | 4.79 | 24.32 | 4.69 | 22.66 | 9.67 | 21.29 | 9.67 | 21.88 | 9.67
80.384 | 14.36 | 0.1 | 14.84 | 0.1 | 12.4 | 0.1 | 15.82 | 4.59 | 4.79 | 4.59 | 4.88 | 4.2 | 9.77 | 9.38 | 9.86 | 21.97 | 9.67 | 14.45 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0 | 5.66 | 4.79 | 21.29 | 4.88 | 24.9 | 4.69 | 23.24 | 9.67 | 21.88 | 9.67 | 22.36 | 9.67
80.416 | 14.45 | 0.1 | 15.43 | 0.1 | 12.89 | 0.1 | 16.31 | 4.59 | 4.88 | 4.59 | 4.88 | 5.47 | 9.86 | 6.93 | 9.86 | 22.46 | 9.67 | 15.04 | 0.2 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 6.05 | 4.88 | 21.29 | 4.88 | 25.39 | 4.69 | 23.73 | 9.67 | 22.46 | 9.67 | 22.95 | 9.67
80.448 | 14.36 | 0.1 | 15.92 | 0.1 | 13.38 | 0.1 | 16.8 | 4.69 | 4.88 | 4.59 | 4.88 | 5.86 | 9.86 | 9.86 | 9.86 | 22.85 | 9.67 | 15.53 | 0.1 | 0.1 | 0.29 | 0.88 | 0.29 | 0.1 | 5.96 | 4.88 | 21.29 | 4.88 | 25.88 | 4.69 | 24.22 | 9.67 | 23.05 | 9.67 | 23.44 | 9.67
80.48 | 14.36 | 0.1 | 16.31 | 0.1 | 13.97 | 0.1 | 17.29 | 4.69 | 4.79 | 4.59 | 4.88 | 5.57 | 9.86 | 11.62 | 9.86 | 23.15 | 9.67 | 16.11 | 0.1 | 0.1 | 0.2 | 1.86 | 0.29 | 0.1 | 5.47 | 4.88 | 21.19 | 4.79 | 25.59 | 4.79 | 24.71 | 9.77 | 23.54 | 9.67 | 24.02 | 9.67
80.512 | 14.36 | 0.1 | 16.9 | 0 | 14.45 | 0 | 17.68 | 4.69 | 4.88 | 4.49 | 4.88 | 5.08 | 9.77 | 10.64 | 9.86 | 23.54 | 9.67 | 16.6 | 0.1 | 1.07 | 0.2 | 3.61 | 0.2 | 0.1 | 5.18 | 4.79 | 21.19 | 4.79 | 21 | 4.79 | 25.29 | 9.77 | 24.12 | 9.67 | 24.41 | 9.67
80.544 | 14.36 | 0.1 | 17.29 | 0.1 | 15.04 | 0.1 | 18.16 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 9.77 | 9.96 | 9.86 | 23.93 | 9.67 | 17.09 | 0.1 | 2.25 | 0.1 | 5.37 | 0.2 | 0 | 5.08 | 4.79 | 21.19 | 4.88 | 21.58 | 4.79 | 25.78 | 9.67 | 24.61 | 9.67 | 25 | 9.67
80.576 | 14.36 | 0.1 | 15.43 | 0.1 | 15.43 | 0.1 | 18.65 | 4.79 | 6.54 | 4.59 | 4.88 | 4.88 | 9.86 | 9.96 | 9.86 | 24.32 | 9.77 | 17.58 | 0.1 | 4.1 | 0.2 | 6.54 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.88 | 21.49 | 4.79 | 26.17 | 9.77 | 25.1 | 9.77 | 25.49 | 9.67
80.608 | 6.64 | 0.1 | 15.23 | 0.1 | 15.92 | 0.1 | 19.14 | 4.79 | 8.01 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 21.09 | 9.77 | 18.07 | 0.1 | 5.57 | 0.1 | 7.32 | 0.1 | 1.37 | 5.08 | 4.88 | 21.19 | 4.88 | 21.49 | 4.79 | 26.86 | 9.77 | 25.59 | 9.77 | 25.98 | 9.67
80.64 | 1.76 | 0.1 | 14.94 | 0.1 | 16.41 | 0.1 | 19.53 | 4.79 | 8.98 | 4.49 | 4.88 | 4.98 | 9.86 | 10.25 | 9.86 | 21.49 | 9.77 | 18.46 | 0.1 | 6.64 | 0.1 | 8.2 | 0.1 | 2.64 | 5.08 | 4.88 | 9.28 | 4.79 | 21.39 | 4.79 | 27.25 | 9.77 | 26.08 | 9.77 | 26.56 | 9.77
80.672 | 1.95 | 0.1 | 14.65 | 0 | 16.9 | 0.1 | 19.92 | 4.79 | 9.77 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 18.85 | 0.1 | 7.52 | 0 | 8.89 | 0.1 | 4.59 | 4.88 | 4.79 | 2.44 | 4.79 | 21.29 | 4.79 | 27.74 | 9.77 | 26.56 | 9.77 | 27.05 | 9.67
80.704 | 1.66 | 0.1 | 14.75 | 0.1 | 16.9 | 0.1 | 20.31 | 4.79 | 10.35 | 4.49 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 19.24 | 0 | 8.3 | 0.1 | 9.67 | 0.1 | 5.86 | 4.88 | 4.79 | 3.32 | 4.88 | 21.29 | 4.79 | 28.03 | 9.77 | 27.05 | 9.77 | 27.54 | 9.77
80.736 | 1.46 | 0.1 | 14.55 | 0.1 | 14.94 | 0.1 | 20.8 | 4.79 | 11.04 | 4.59 | 4.88 | 4.69 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 19.82 | 0.1 | 9.08 | 0.1 | 10.55 | 0.1 | 7.03 | 4.79 | 4.88 | 3.52 | 4.88 | 21.29 | 4.88 | 28.61 | 9.77 | 27.64 | 9.77 | 28.03 | 9.77
80.768 | 1.37 | 0.1 | 14.55 | 0.1 | 14.94 | 0.1 | 20.7 | 4.79 | 11.62 | 4.59 | 4.88 | 4.69 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.22 | 0.1 | 9.77 | 0.1 | 11.33 | 0.1 | 7.91 | 4.79 | 4.88 | 4.79 | 4.88 | 21.29 | 4.79 | 29.1 | 9.77 | 28.13 | 9.77 | 28.52 | 9.77
80.8 | 1.27 | 0.1 | 14.55 | 0.1 | 14.75 | 0.1 | 18.65 | 4.88 | 12.21 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 20.61 | 0.1 | 10.64 | 0.1 | 12.11 | 0.1 | 8.69 | 4.79 | 4.88 | 6.15 | 4.79 | 21.19 | 4.79 | 27.54 | 9.86 | 28.52 | 9.77 | 28.91 | 9.77
80.832 | 1.07 | 0 | 14.45 | 0.1 | 14.65 | 0 | 18.36 | 4.79 | 12.79 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21 | 0.1 | 11.43 | 0 | 12.89 | 0 | 9.47 | 4.69 | 4.79 | 6.15 | 4.79 | 21.19 | 4.79 | 23.73 | 9.86 | 28.91 | 9.77 | 29.4 | 9.77
80.864 | 1.07 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 18.46 | 4.79 | 13.18 | 4.59 | 4.88 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 21.39 | 0 | 12.21 | 0 | 13.57 | 0 | 10.16 | 4.59 | 4.79 | 5.66 | 4.88 | 21.19 | 4.79 | 24.41 | 9.77 | 29.49 | 9.77 | 26.17 | 9.77
80.896 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.36 | 4.88 | 13.87 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 12.89 | 0.1 | 14.36 | 0.1 | 11.04 | 4.59 | 4.88 | 5.37 | 4.88 | 21.19 | 4.79 | 24.41 | 9.86 | 26.27 | 9.86 | 25.1 | 9.77
80.928 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.36 | 4.88 | 14.36 | 4.69 | 5.57 | 4.49 | 9.86 | 9.67 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 13.67 | 0.1 | 15.14 | 0.1 | 11.82 | 4.59 | 4.79 | 5.18 | 4.88 | 8.59 | 4.88 | 24.41 | 9.86 | 25.2 | 9.86 | 24.71 | 9.77
80.96 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 14.94 | 4.69 | 6.45 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 22.66 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 12.5 | 4.59 | 4.88 | 4.98 | 4.79 | 2.44 | 4.88 | 24.32 | 9.86 | 24.81 | 9.86 | 24.51 | 9.77
80.992 | 0.49 | 0.1 | 8.59 | 0 | 14.45 | 0 | 18.16 | 4.88 | 15.43 | 4.69 | 7.81 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 0 | 15.14 | 0 | 16.5 | 0 | 13.28 | 4.59 | 5.08 | 4.98 | 4.79 | 3.22 | 4.79 | 24.32 | 9.86 | 24.61 | 9.77 | 24.41 | 9.77
81.024 | 0.49 | 0 | 2.54 | 0 | 14.45 | 0.1 | 18.16 | 4.79 | 15.82 | 4.79 | 8.79 | 4.49 | 9.77 | 9.67 | 9.86 | 20.41 | 9.77 | 23.34 | 0 | 15.82 | 0 | 17.09 | 0 | 13.97 | 4.59 | 6.05 | 5.08 | 4.88 | 4.1 | 4.88 | 24.32 | 9.77 | 24.61 | 9.86 | 24.41 | 9.77
81.056 | 0.49 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.31 | 4.79 | 9.67 | 4.49 | 9.86 | 9.67 | 9.86 | 8.79 | 9.86 | 23.73 | 0.1 | 16.5 | 0.1 | 17.87 | 0.1 | 14.75 | 4.49 | 7.62 | 5.08 | 4.88 | 5.18 | 4.88 | 24.32 | 9.86 | 24.51 | 9.86 | 24.22 | 9.86
81.088 | 0.49 | 0.1 | 1.76 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 16.8 | 4.79 | 10.25 | 4.59 | 10.64 | 9.67 | 9.86 | 7.23 | 9.86 | 24.12 | 0.1 | 17.29 | 0.1 | 18.46 | 0.1 | 15.53 | 4.49 | 9.67 | 4.88 | 4.88 | 6.25 | 4.88 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
81.12 | 0.39 | 0.1 | 1.46 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 17.29 | 4.79 | 10.94 | 4.59 | 11.04 | 9.67 | 9.86 | 10.25 | 9.86 | 24.12 | 0.1 | 17.97 | 0.1 | 19.04 | 0.1 | 16.21 | 4.59 | 11.33 | 4.98 | 4.79 | 5.96 | 4.88 | 24.22 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
81.152 | 0.29 | 0.1 | 1.37 | 0 | 7.62 | 0 | 18.07 | 4.88 | 17.77 | 4.79 | 11.52 | 4.59 | 12.21 | 9.57 | 9.86 | 11.52 | 9.86 | 21.68 | 0 | 18.65 | 0 | 19.73 | 0 | 16.9 | 4.49 | 12.5 | 4.79 | 4.79 | 5.37 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
81.184 | 0.29 | 0 | 1.17 | 0.1 | 2.15 | 0.1 | 18.07 | 4.88 | 18.26 | 4.79 | 12.21 | 4.59 | 13.38 | 9.67 | 9.86 | 10.64 | 9.86 | 21.97 | 0 | 19.24 | 0 | 20.31 | 0 | 17.38 | 4.59 | 13.18 | 4.69 | 4.88 | 5.18 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
81.216 | 0.2 | 0.1 | 1.17 | 0.1 | 1.95 | 0.1 | 8.79 | 4.88 | 18.65 | 4.88 | 12.7 | 4.59 | 14.16 | 9.67 | 9.86 | 10.16 | 9.86 | 21.78 | 0.1 | 19.82 | 0.1 | 20.9 | 0.1 | 18.07 | 4.59 | 14.06 | 4.69 | 4.88 | 5.08 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
81.248 | 0.2 | 0.1 | 0.88 | 0.1 | 1.76 | 0.1 | 2.83 | 4.88 | 19.14 | 4.88 | 13.18 | 4.69 | 14.75 | 9.67 | 9.86 | 9.86 | 9.86 | 21.68 | 0.1 | 20.31 | 0.1 | 21.58 | 0.1 | 18.75 | 4.59 | 14.75 | 4.69 | 4.88 | 5.18 | 4.88 | 12.11 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
81.28 | 0.1 | 0.1 | 0.88 | 0.1 | 1.37 | 0.1 | 3.13 | 4.88 | 19.53 | 4.88 | 13.67 | 4.69 | 15.23 | 9.67 | 9.86 | 10.16 | 9.86 | 21.68 | 0 | 20.9 | 0.1 | 22.17 | 0.1 | 19.43 | 4.59 | 15.53 | 4.59 | 4.79 | 5.18 | 4.88 | 5.96 | 9.86 | 24.22 | 9.86 | 15.33 | 9.86
81.312 | 0.1 | 0.1 | 0.68 | 0 | 1.37 | 0 | 4 | 4.88 | 20.02 | 4.88 | 14.26 | 4.69 | 15.82 | 9.67 | 9.86 | 10.25 | 9.86 | 21.49 | 0 | 21.58 | 0 | 22.56 | 0 | 20.12 | 4.59 | 16.21 | 4.59 | 4.88 | 5.08 | 4.79 | 8.5 | 9.86 | 13.67 | 9.86 | 6.45 | 9.86
81.344 | 0 | 0.1 | 0.68 | 0.1 | 1.27 | 0 | 5.08 | 4.88 | 20.31 | 4.79 | 14.75 | 4.69 | 16.31 | 9.67 | 10.16 | 9.96 | 9.86 | 21.49 | 0 | 21.97 | 0 | 19.43 | 0 | 20.61 | 4.69 | 16.8 | 4.59 | 4.88 | 4.98 | 4.79 | 12.01 | 9.77 | 6.74 | 9.86 | 7.13 | 9.86
81.376 | 0.1 | 0.78 | 0.59 | 0.1 | 1.27 | 0.1 | 5.86 | 4.88 | 20.8 | 4.88 | 15.33 | 4.79 | 16.8 | 9.67 | 11.04 | 9.86 | 9.86 | 21.49 | 0.1 | 21.97 | 0.1 | 18.65 | 0.1 | 21.29 | 4.69 | 17.58 | 4.59 | 4.88 | 4.88 | 4.88 | 11.62 | 9.86 | 7.23 | 9.86 | 11.13 | 9.86
81.408 | 0.1 | 1.66 | 0.59 | 0.1 | 1.07 | 0.1 | 5.66 | 4.88 | 18.95 | 4.88 | 15.72 | 4.79 | 17.29 | 9.67 | 11.91 | 9.86 | 9.86 | 21.49 | 0.1 | 18.56 | 0.1 | 18.26 | 0.1 | 21.78 | 4.69 | 18.16 | 4.59 | 4.88 | 4.88 | 4.79 | 10.35 | 9.86 | 11.13 | 9.86 | 11.91 | 9.86
81.44 | 0.1 | 3.03 | 0.59 | 0.1 | 0.88 | 0.1 | 5.37 | 4.88 | 17.68 | 4.88 | 16.21 | 4.79 | 17.87 | 9.67 | 13.28 | 9.96 | 9.86 | 21.49 | 0.1 | 18.16 | 0.1 | 18.07 | 0.1 | 22.46 | 4.69 | 18.75 | 4.59 | 4.79 | 4.79 | 4.79 | 10.06 | 9.86 | 11.91 | 9.86 | 10.55 | 9.86
81.472 | 0 | 4 | 0.49 | 0 | 0.78 | 0.1 | 4.88 | 4.88 | 17.87 | 4.88 | 16.8 | 4.79 | 18.26 | 9.67 | 14.16 | 9.77 | 9.86 | 21.39 | 0 | 17.87 | 0 | 17.97 | 0 | 18.26 | 4.69 | 19.43 | 4.49 | 5.66 | 4.69 | 4.79 | 10.35 | 9.86 | 10.55 | 9.86 | 9.86 | 9.86
81.504 | 0 | 4.69 | 0.29 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 17.29 | 4.79 | 18.65 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 0 | 17.87 | 0 | 17.87 | 0 | 16.99 | 4.69 | 20.12 | 4.49 | 6.84 | 4.69 | 4.79 | 10.45 | 9.77 | 9.86 | 9.86 | 10.16 | 9.86
81.536 | 0.1 | 5.27 | 0.29 | 0.1 | 0.68 | 0.1 | 4.98 | 4.88 | 18.07 | 4.88 | 17.68 | 4.79 | 19.14 | 9.67 | 15.33 | 9.77 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.77 | 0.1 | 17.68 | 4.79 | 20.61 | 4.59 | 8.5 | 4.69 | 4.88 | 10.35 | 9.86 | 10.06 | 9.86 | 10.45 | 9.86
81.568 | 0.1 | 5.86 | 0.29 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 18.07 | 4.88 | 19.53 | 9.67 | 15.82 | 9.77 | 9.86 | 12.89 | 0.1 | 17.77 | 0.1 | 17.77 | 0.1 | 17.77 | 4.79 | 21.29 | 4.59 | 10.35 | 4.59 | 4.79 | 10.16 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86
81.6 | 0.1 | 6.45 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 18.56 | 4.88 | 20.02 | 9.77 | 16.31 | 9.67 | 9.86 | 6.93 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 17.77 | 4.79 | 21.88 | 4.59 | 11.72 | 4.59 | 4.79 | 10.16 | 9.86 | 10.55 | 9.86 | 10.16 | 9.86
81.632 | 0 | 7.03 | 0.2 | 0 | 0.49 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 18.95 | 4.79 | 20.41 | 9.67 | 16.8 | 9.67 | 9.86 | 8.3 | 0 | 17.58 | 0 | 17.68 | 0 | 17.68 | 4.79 | 22.46 | 4.59 | 12.7 | 4.49 | 4.79 | 10.16 | 9.77 | 10.25 | 9.86 | 10.06 | 9.86
81.664 | 0 | 7.62 | 0.1 | 0.1 | 0.49 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 19.34 | 4.79 | 20.9 | 9.77 | 17.19 | 9.67 | 9.86 | 11.04 | 0 | 17.58 | 0 | 17.68 | 0 | 17.77 | 4.79 | 22.95 | 4.59 | 13.57 | 4.59 | 4.79 | 10.06 | 9.77 | 9.96 | 9.77 | 10.06 | 9.86
81.696 | 0.1 | 8.2 | 0.1 | 0.1 | 0.39 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.92 | 4.88 | 21.39 | 9.77 | 17.77 | 9.67 | 9.86 | 11.33 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.54 | 4.59 | 14.26 | 4.59 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 10.16 | 9.86
81.728 | 0.1 | 8.79 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.22 | 4.88 | 21.78 | 9.77 | 18.16 | 9.67 | 9.86 | 10.45 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 24.02 | 4.59 | 15.04 | 4.59 | 4.79 | 9.86 | 9.86 | 10.16 | 9.86 | 10.06 | 9.86
81.76 | 0.1 | 9.47 | 0.1 | 0.2 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.7 | 4.88 | 22.07 | 9.77 | 18.56 | 9.67 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 11.52 | 0.1 | 17.77 | 4.88 | 24.51 | 4.69 | 15.82 | 4.59 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
81.792 | 0 | 9.96 | 0 | 1.27 | 0.2 | 0.2 | 4.59 | 4.88 | 18.07 | 4.88 | 20.22 | 4.79 | 22.46 | 9.77 | 19.14 | 9.57 | 9.86 | 9.96 | 0 | 17.58 | 0 | 2.54 | 0 | 17.68 | 4.79 | 25 | 4.69 | 16.5 | 4.49 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
81.824 | 0 | 10.55 | 0 | 2.15 | 0.2 | 1.07 | 4.49 | 4.88 | 17.97 | 4.88 | 18.56 | 4.88 | 22.85 | 9.77 | 19.43 | 9.57 | 9.86 | 10.16 | 0 | 8.4 | 0 | 2.25 | 0 | 17.68 | 4.88 | 25.49 | 4.69 | 17.09 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
81.856 | 0.1 | 11.23 | 0.1 | 3.61 | 0.1 | 2.05 | 4.49 | 4.88 | 7.32 | 4.88 | 18.46 | 4.88 | 23.24 | 9.77 | 20.02 | 9.57 | 9.86 | 10.16 | 0.1 | 2.25 | 0.1 | 2.15 | 0.1 | 17.68 | 4.88 | 26.08 | 4.69 | 17.77 | 4.59 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
81.888 | 0.1 | 11.82 | 0.1 | 4.39 | 0.1 | 3.52 | 4.49 | 4.88 | 2.54 | 4.88 | 18.46 | 4.88 | 23.63 | 9.77 | 20.31 | 9.57 | 9.86 | 9.96 | 0.1 | 2.25 | 0.1 | 1.76 | 0.1 | 17.68 | 4.88 | 21.39 | 4.69 | 18.56 | 4.59 | 4.79 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
81.92 | 0.1 | 12.31 | 0.1 | 4.98 | 0 | 4.3 | 4.49 | 4.88 | 3.32 | 4.88 | 18.26 | 4.88 | 24.02 | 9.77 | 20.8 | 9.57 | 10.55 | 9.86 | 0.1 | 2.15 | 0.1 | 1.86 | 0.1 | 6.74 | 4.88 | 20.31 | 4.79 | 19.14 | 4.59 | 5.66 | 9.67 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
81.952 | 0 | 12.89 | 0.1 | 5.47 | 0.1 | 4.98 | 4.49 | 4.88 | 4.39 | 4.88 | 18.16 | 4.79 | 24.41 | 9.77 | 21.19 | 9.57 | 11.52 | 9.86 | 0 | 1.56 | 0 | 1.46 | 0 | 1.66 | 4.88 | 21.29 | 4.69 | 19.82 | 4.59 | 6.74 | 9.67 | 10.64 | 9.77 | 9.86 | 9.67 | 9.86
81.984 | 0 | 13.38 | 0 | 6.05 | 0.1 | 5.37 | 4.49 | 4.79 | 5.47 | 4.88 | 18.16 | 4.88 | 21.68 | 9.77 | 21.58 | 9.67 | 12.7 | 9.86 | 0 | 1.66 | 0 | 1.46 | 0 | 2.25 | 4.88 | 21.19 | 4.79 | 20.31 | 4.59 | 8.3 | 9.67 | 11.43 | 9.77 | 10.74 | 9.67 | 9.86
82.016 | 0.1 | 13.97 | 0.1 | 6.64 | 0.1 | 5.96 | 4.49 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 21.78 | 9.86 | 21.97 | 9.67 | 13.67 | 9.77 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 2.05 | 4.88 | 21.19 | 4.79 | 20.9 | 4.69 | 10.35 | 9.67 | 12.99 | 9.77 | 11.52 | 9.67 | 10.84
82.048 | 0.1 | 14.45 | 0.1 | 7.23 | 0.1 | 6.64 | 4.59 | 4.88 | 5.57 | 4.88 | 18.16 | 4.88 | 21.49 | 9.86 | 22.36 | 9.67 | 14.45 | 9.77 | 0.1 | 1.56 | 0.1 | 1.27 | 0.1 | 1.76 | 4.88 | 21.19 | 4.79 | 21.58 | 4.69 | 11.62 | 9.67 | 15.04 | 9.77 | 13.09 | 9.67 | 11.91
82.08 | 0.1 | 14.94 | 0.1 | 7.81 | 0.1 | 7.13 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 22.75 | 9.67 | 15.04 | 9.67 | 0.1 | 1.27 | 0 | 1.17 | 0.1 | 1.66 | 4.88 | 21.29 | 4.79 | 22.17 | 4.69 | 12.6 | 9.67 | 16.31 | 9.67 | 15.04 | 9.67 | 13.48
82.112 | 0 | 15.43 | 0 | 8.4 | 0.1 | 7.81 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.58 | 9.77 | 23.24 | 9.67 | 15.63 | 9.67 | 0 | 1.17 | 0 | 0.98 | 0 | 1.37 | 4.88 | 21.29 | 4.79 | 22.66 | 4.69 | 13.38 | 9.67 | 17.19 | 9.67 | 16.31 | 9.67 | 15.43
82.144 | 0 | 15.92 | 0 | 9.08 | 0.1 | 8.4 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 23.54 | 9.67 | 16.02 | 9.67 | 0 | 1.07 | 0 | 0.88 | 0 | 1.37 | 4.88 | 21.19 | 4.79 | 23.15 | 4.69 | 14.26 | 9.57 | 17.97 | 9.67 | 17.19 | 9.67 | 16.5
82.176 | 0.1 | 16.41 | 0.1 | 9.67 | 0.1 | 8.98 | 4.69 | 5.08 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.93 | 9.67 | 16.6 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 23.83 | 4.79 | 14.94 | 9.57 | 18.65 | 9.67 | 18.07 | 9.67 | 17.38
82.208 | 0.1 | 16.9 | 0.1 | 10.25 | 0.1 | 9.57 | 4.69 | 5.86 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 24.32 | 9.67 | 17.09 | 9.67 | 0.1 | 0.88 | 0.1 | 0.78 | 0.1 | 1.17 | 4.88 | 21.19 | 4.79 | 24.32 | 4.79 | 15.72 | 9.57 | 19.24 | 9.67 | 18.75 | 9.57 | 18.16
82.24 | 0.1 | 16.9 | 0.1 | 10.84 | 0.1 | 10.16 | 4.69 | 6.93 | 4.88 | 4.88 | 8.79 | 4.88 | 21.49 | 9.86 | 20.9 | 9.67 | 17.58 | 9.67 | 0.2 | 0.88 | 0.1 | 0.78 | 0.1 | 1.07 | 4.88 | 21.19 | 4.79 | 24.71 | 4.79 | 16.41 | 9.57 | 20.02 | 9.67 | 19.24 | 9.57 | 18.85
82.272 | 0 | 14.55 | 0 | 11.43 | 0.1 | 10.74 | 4.69 | 8.3 | 4.59 | 4.88 | 2.83 | 4.88 | 21.49 | 9.86 | 21.58 | 9.67 | 17.97 | 9.57 | 1.56 | 0.78 | 0 | 0.59 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.29 | 4.79 | 17.19 | 9.67 | 20.7 | 9.67 | 20.02 | 9.57 | 19.53
82.304 | 0 | 14.75 | 0 | 12.01 | 0.1 | 11.33 | 4.69 | 9.28 | 4.69 | 4.79 | 3.42 | 4.88 | 21.39 | 9.86 | 21.39 | 9.67 | 18.46 | 9.57 | 2.64 | 0.68 | 0 | 0.49 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.88 | 4.79 | 17.68 | 9.67 | 21.09 | 9.57 | 20.61 | 9.67 | 20.12
82.336 | 0.1 | 14.75 | 0.1 | 12.6 | 0.1 | 12.01 | 4.79 | 9.96 | 4.69 | 4.88 | 4.1 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 18.85 | 9.67 | 4.49 | 0.59 | 0.1 | 0.59 | 0.1 | 0.78 | 4.88 | 5.96 | 4.79 | 26.27 | 4.79 | 18.46 | 9.67 | 21.68 | 9.67 | 21.19 | 9.67 | 20.8
82.368 | 0.1 | 14.65 | 0.1 | 13.09 | 0.1 | 12.5 | 4.79 | 10.55 | 4.59 | 4.88 | 5.27 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 19.24 | 9.57 | 5.96 | 0.49 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 2.54 | 4.79 | 23.05 | 4.79 | 19.14 | 9.67 | 22.36 | 9.67 | 21.78 | 9.67 | 21.39
82.4 | 0.1 | 14.55 | 0.1 | 13.67 | 0.1 | 13.09 | 4.79 | 11.23 | 4.59 | 4.88 | 5.86 | 4.88 | 18.95 | 9.86 | 21.49 | 9.77 | 19.82 | 9.57 | 6.84 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 4.88 | 3.32 | 4.79 | 22.27 | 4.79 | 19.73 | 9.67 | 22.85 | 9.67 | 22.36 | 9.67 | 21.78
82.432 | 0 | 14.45 | 0 | 14.16 | 0.1 | 13.57 | 4.79 | 11.82 | 4.49 | 4.88 | 5.57 | 4.79 | 8.11 | 9.86 | 21.49 | 9.77 | 20.22 | 9.57 | 7.62 | 0.49 | 1.07 | 0.29 | 0 | 0.59 | 4.88 | 4.3 | 4.79 | 21.78 | 4.79 | 20.31 | 9.67 | 23.54 | 9.67 | 22.95 | 9.67 | 22.36
82.464 | 0 | 14.45 | 0 | 14.65 | 0.1 | 14.06 | 4.79 | 12.31 | 4.49 | 4.88 | 5.18 | 4.88 | 7.32 | 9.86 | 21.49 | 9.77 | 20.61 | 9.57 | 8.5 | 0.49 | 2.05 | 0.29 | 0.88 | 0.49 | 4.88 | 5.47 | 4.79 | 21.68 | 4.79 | 20.9 | 9.67 | 23.93 | 9.67 | 23.34 | 9.67 | 22.95
82.496 | 0.1 | 14.45 | 0.1 | 15.14 | 0.1 | 14.65 | 4.88 | 12.89 | 4.59 | 4.88 | 5.18 | 4.88 | 10.45 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 9.18 | 0.29 | 3.71 | 0.29 | 1.76 | 0.49 | 4.88 | 6.15 | 4.79 | 21.58 | 4.79 | 21.49 | 9.67 | 24.51 | 9.67 | 24.02 | 9.67 | 23.54
82.528 | 0.1 | 14.45 | 0.1 | 15.63 | 0.1 | 15.14 | 4.79 | 13.48 | 4.49 | 4.88 | 4.88 | 4.88 | 11.43 | 9.86 | 21.39 | 9.77 | 21.39 | 9.67 | 9.96 | 0.29 | 5.47 | 0.2 | 3.32 | 0.39 | 4.88 | 6.05 | 4.79 | 21.49 | 4.79 | 21.97 | 9.67 | 25.1 | 9.67 | 24.51 | 9.67 | 24.02
82.56 | 0.1 | 14.45 | 0.1 | 16.11 | 0.1 | 15.63 | 4.88 | 13.97 | 4.49 | 4.88 | 4.98 | 4.88 | 10.45 | 9.86 | 21.39 | 9.77 | 21.78 | 9.67 | 10.64 | 0.2 | 6.64 | 0.2 | 5.27 | 0.39 | 4.88 | 5.57 | 4.79 | 21.39 | 4.79 | 22.66 | 9.67 | 25.49 | 9.67 | 25 | 9.67 | 24.51
82.592 | 0 | 14.45 | 0 | 16.7 | 0.1 | 16.11 | 4.79 | 14.45 | 4.49 | 5.47 | 4.88 | 4.88 | 9.96 | 9.86 | 21.39 | 9.77 | 22.27 | 9.67 | 11.52 | 0.2 | 7.42 | 0.1 | 6.45 | 0.29 | 4.88 | 5.27 | 4.79 | 21.39 | 4.79 | 23.15 | 9.67 | 25.98 | 9.67 | 25.59 | 9.67 | 25.2
82.624 | 0 | 14.36 | 0 | 17.09 | 0.1 | 16.6 | 4.88 | 14.94 | 4.49 | 6.35 | 4.79 | 4.79 | 9.96 | 9.86 | 21.39 | 9.77 | 22.56 | 9.67 | 12.31 | 0.2 | 8.2 | 0.1 | 7.32 | 0.29 | 4.88 | 4.98 | 4.79 | 21.29 | 4.79 | 23.63 | 9.67 | 26.47 | 9.67 | 25.98 | 9.67 | 25.59
82.656 | 0.1 | 14.36 | 0.1 | 13.77 | 0.1 | 17.09 | 4.88 | 15.53 | 4.49 | 7.81 | 4.79 | 4.88 | 10.16 | 9.86 | 21 | 9.86 | 23.05 | 9.67 | 12.99 | 0.2 | 8.98 | 0.1 | 8.01 | 0.29 | 4.88 | 5.18 | 4.79 | 21.39 | 4.79 | 24.22 | 9.77 | 27.05 | 9.77 | 26.47 | 9.77 | 26.08
82.688 | 0.1 | 7.52 | 0.1 | 13.97 | 0.1 | 13.77 | 4.88 | 16.02 | 4.59 | 8.79 | 4.69 | 4.88 | 10.16 | 9.86 | 8.79 | 9.86 | 23.34 | 9.67 | 13.77 | 0.2 | 9.86 | 0.1 | 8.89 | 0.2 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 24.71 | 9.77 | 27.54 | 9.77 | 26.95 | 9.77 | 26.66
82.72 | 0.1 | 2.05 | 0.1 | 14.16 | 0.1 | 13.97 | 4.88 | 16.5 | 4.59 | 9.67 | 4.69 | 4.88 | 10.06 | 9.86 | 7.13 | 9.86 | 23.73 | 9.67 | 14.45 | 0.1 | 10.55 | 0.1 | 9.67 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 25.2 | 9.77 | 28.03 | 9.77 | 27.44 | 9.77 | 27.05
82.752 | 0 | 1.95 | 0 | 14.45 | 0.1 | 14.16 | 4.88 | 17.09 | 4.59 | 10.35 | 4.59 | 4.79 | 9.96 | 9.77 | 10.25 | 9.77 | 24.12 | 9.67 | 15.14 | 0.1 | 11.43 | 0.1 | 10.45 | 0.2 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.68 | 9.77 | 28.52 | 9.77 | 27.93 | 9.77 | 27.54
82.784 | 0 | 1.66 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.38 | 4.59 | 10.94 | 4.59 | 4.79 | 9.96 | 9.86 | 11.43 | 9.86 | 24.12 | 9.67 | 15.82 | 0.1 | 12.11 | 0.1 | 11.23 | 0.1 | 4.88 | 4.98 | 4.79 | 13.67 | 4.79 | 26.37 | 9.77 | 28.91 | 9.77 | 28.42 | 9.77 | 28.03
82.816 | 0.1 | 1.46 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 17.87 | 4.69 | 11.52 | 4.49 | 4.88 | 9.86 | 9.86 | 10.74 | 9.86 | 22.07 | 9.77 | 16.5 | 0.1 | 12.99 | 0.1 | 12.01 | 0.1 | 4.88 | 4.79 | 4.79 | 3.32 | 4.79 | 22.85 | 9.77 | 29.49 | 9.77 | 28.91 | 9.77 | 28.52
82.848 | 0.1 | 1.37 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.26 | 4.69 | 12.11 | 4.59 | 4.88 | 9.86 | 9.86 | 10.06 | 9.86 | 21.97 | 9.77 | 17.19 | 0.1 | 13.57 | 0.1 | 12.7 | 0.1 | 4.88 | 4.79 | 4.79 | 3.13 | 4.88 | 22.07 | 9.77 | 25.29 | 9.77 | 29.4 | 9.77 | 29.01
82.88 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.75 | 4.69 | 12.7 | 4.49 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 21.78 | 9.77 | 17.87 | 0.1 | 14.36 | 0.1 | 13.57 | 0.1 | 4.88 | 4.69 | 4.79 | 3.71 | 4.79 | 21.68 | 9.77 | 24.41 | 9.77 | 25.2 | 9.77 | 29.49
82.912 | 0 | 1.17 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.69 | 13.28 | 4.49 | 4.79 | 9.67 | 9.77 | 10.16 | 9.86 | 21.58 | 9.67 | 18.65 | 0.1 | 15.14 | 0.1 | 14.26 | 0.1 | 5.37 | 4.69 | 4.79 | 4.69 | 4.79 | 21.58 | 9.77 | 24.61 | 9.77 | 24.9 | 9.77 | 25.78
82.944 | 0 | 0.98 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.63 | 4.69 | 13.77 | 4.49 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 19.14 | 0.1 | 15.92 | 0 | 14.94 | 0.1 | 6.35 | 4.59 | 4.79 | 6.15 | 4.79 | 21.49 | 9.77 | 24.41 | 9.77 | 24.51 | 9.77 | 25.2
82.976 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.12 | 4.79 | 14.26 | 4.49 | 5.47 | 9.67 | 9.86 | 10.06 | 9.86 | 21.58 | 9.77 | 19.82 | 0.1 | 16.5 | 0.1 | 15.72 | 0.1 | 7.81 | 4.69 | 4.79 | 6.25 | 4.79 | 21.49 | 9.77 | 24.51 | 9.86 | 24.41 | 9.77 | 24.81
83.008 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.51 | 4.79 | 14.75 | 4.49 | 6.45 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.31 | 0.1 | 17.29 | 0.1 | 16.31 | 0.1 | 9.86 | 4.59 | 4.79 | 5.86 | 4.88 | 21.29 | 9.77 | 24.41 | 9.86 | 24.41 | 9.77 | 24.61
83.04 | 0.1 | 0.78 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 21 | 4.79 | 15.23 | 4.49 | 7.81 | 9.67 | 9.96 | 9.86 | 9.96 | 21.49 | 9.77 | 21.09 | 0.1 | 17.87 | 0.1 | 16.99 | 0.1 | 11.43 | 4.59 | 4.79 | 5.37 | 4.88 | 21.39 | 9.77 | 24.32 | 9.86 | 24.41 | 9.77 | 24.51
83.072 | 0 | 0.68 | 0 | 14.16 | 0.1 | 14.36 | 4.88 | 17.68 | 4.79 | 15.82 | 4.59 | 8.89 | 9.67 | 10.74 | 9.86 | 9.86 | 21.39 | 9.77 | 21.68 | 0.1 | 18.65 | 0 | 17.77 | 0.1 | 12.5 | 4.59 | 4.79 | 5.27 | 4.79 | 21.19 | 9.77 | 24.41 | 9.77 | 24.32 | 9.77 | 24.41
83.104 | 0 | 0.59 | 0 | 4.88 | 0.1 | 14.36 | 4.88 | 17.77 | 4.79 | 16.31 | 4.59 | 9.57 | 9.57 | 11.72 | 9.77 | 9.86 | 21.39 | 9.77 | 22.17 | 0.1 | 19.14 | 0.1 | 18.46 | 0.1 | 13.38 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.32 | 9.77 | 24.32
83.136 | 0.1 | 0.59 | 0.1 | 1.86 | 0.1 | 5.57 | 4.88 | 17.97 | 4.79 | 16.7 | 4.59 | 10.35 | 9.67 | 13.09 | 9.77 | 9.86 | 21.39 | 9.86 | 17.97 | 0.1 | 19.73 | 0.1 | 18.95 | 0.1 | 14.16 | 4.59 | 4.79 | 5.18 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
83.168 | 0.1 | 0.49 | 0.1 | 1.95 | 0.1 | 1.56 | 4.88 | 18.07 | 4.79 | 17.19 | 4.59 | 10.94 | 9.67 | 13.97 | 9.77 | 9.96 | 21.39 | 9.77 | 16.6 | 0.1 | 20.31 | 0.1 | 19.73 | 0.1 | 14.94 | 4.49 | 4.79 | 5.18 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22
83.2 | 0.1 | 0.39 | 0.1 | 1.76 | 0.1 | 1.86 | 4.88 | 18.16 | 4.88 | 17.68 | 4.59 | 11.52 | 9.57 | 14.65 | 9.77 | 9.86 | 21.39 | 9.86 | 17.29 | 0.1 | 20.9 | 0.1 | 20.22 | 0.1 | 15.72 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22
83.232 | 0 | 0.29 | 0 | 1.46 | 0.1 | 1.76 | 4.88 | 18.07 | 4.79 | 18.16 | 4.69 | 12.11 | 9.57 | 15.23 | 9.67 | 9.86 | 12.6 | 9.77 | 17.38 | 0.1 | 21.58 | 0.1 | 20.9 | 0.1 | 16.41 | 4.59 | 4.79 | 4.88 | 4.79 | 13.18 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77 | 24.22
83.264 | 0.1 | 0.29 | 0 | 1.27 | 0.1 | 1.46 | 4.88 | 18.16 | 4.79 | 18.65 | 4.69 | 12.6 | 9.57 | 15.72 | 9.67 | 9.86 | 6.93 | 9.77 | 17.38 | 0.1 | 22.17 | 0.1 | 21.39 | 0 | 17.09 | 4.49 | 4.79 | 4.79 | 4.79 | 3.71 | 9.77 | 19.53 | 9.86 | 24.22 | 9.77 | 24.12
83.296 | 0.78 | 0.2 | 0.1 | 1.27 | 0.1 | 1.27 | 4.88 | 18.16 | 4.88 | 19.14 | 4.69 | 13.18 | 9.67 | 16.21 | 9.67 | 9.86 | 8.3 | 9.86 | 17.38 | 0.1 | 22.66 | 0.1 | 21.97 | 0.1 | 17.68 | 4.59 | 4.79 | 4.88 | 4.79 | 3.13 | 9.77 | 7.72 | 9.86 | 19.73 | 9.86 | 24.22
83.328 | 1.66 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.34 | 4.69 | 13.77 | 9.67 | 16.7 | 9.67 | 9.96 | 11.13 | 9.86 | 17.38 | 0.1 | 19.34 | 0.1 | 22.56 | 0.1 | 18.46 | 4.59 | 5.96 | 4.79 | 4.79 | 3.42 | 9.77 | 6.64 | 9.86 | 7.62 | 9.86 | 15.82
83.36 | 2.93 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.92 | 4.79 | 14.26 | 9.67 | 17.19 | 9.67 | 9.86 | 11.33 | 9.86 | 17.48 | 0.1 | 18.56 | 0.1 | 18.65 | 0.1 | 19.14 | 4.59 | 6.84 | 4.69 | 4.79 | 4.49 | 9.77 | 10.55 | 9.86 | 6.84 | 9.86 | 6.74
83.392 | 4 | 0.1 | 0 | 0.78 | 0 | 0.88 | 4.88 | 18.07 | 4.88 | 20.31 | 4.79 | 14.75 | 9.67 | 17.68 | 9.57 | 9.86 | 10.25 | 9.86 | 17.38 | 0.1 | 18.07 | 0.1 | 18.46 | 0.1 | 19.73 | 4.69 | 8.79 | 4.69 | 4.79 | 5.96 | 9.77 | 11.91 | 9.77 | 10.84 | 9.77 | 7.32
83.424 | 4.69 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.79 | 20.7 | 4.79 | 15.33 | 9.67 | 18.07 | 9.57 | 9.86 | 9.86 | 9.86 | 17.48 | 0.1 | 17.97 | 0 | 17.97 | 0 | 20.31 | 4.59 | 10.45 | 4.59 | 4.79 | 6.25 | 9.77 | 10.74 | 9.86 | 12.11 | 9.77 | 11.23
83.456 | 5.27 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.31 | 4.79 | 15.82 | 9.67 | 18.56 | 9.57 | 9.96 | 9.86 | 9.86 | 17.38 | 0.1 | 17.97 | 0.1 | 17.87 | 0.1 | 20.9 | 4.69 | 11.72 | 4.69 | 4.79 | 5.76 | 9.77 | 10.25 | 9.86 | 10.84 | 9.86 | 12.01
83.488 | 5.76 | 0.1 | 0.1 | 0.68 | 0.1 | 0.68 | 4.88 | 17.87 | 4.88 | 18.26 | 4.79 | 16.21 | 9.67 | 19.04 | 9.67 | 9.96 | 10.16 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.87 | 0.1 | 21.39 | 4.69 | 12.7 | 4.59 | 4.79 | 5.27 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86 | 10.55
83.52 | 6.35 | 0.1 | 0.1 | 0.59 | 0 | 0.59 | 4.88 | 5.27 | 4.88 | 18.26 | 4.79 | 16.8 | 9.67 | 19.43 | 9.57 | 9.96 | 10.16 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.87 | 0.1 | 21.97 | 4.79 | 13.67 | 4.59 | 4.79 | 5.27 | 9.77 | 10.35 | 9.86 | 9.96 | 9.86 | 10.16
83.552 | 6.93 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.03 | 4.88 | 18.26 | 4.79 | 17.19 | 9.67 | 19.92 | 9.57 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.66 | 4.79 | 14.45 | 4.59 | 4.79 | 5.08 | 9.77 | 10.55 | 9.86 | 10.35 | 9.77 | 10.16
83.584 | 7.52 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.42 | 4.88 | 18.16 | 4.88 | 17.68 | 9.67 | 20.22 | 9.57 | 9.86 | 9.86 | 9.86 | 6.25 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.15 | 4.79 | 15.23 | 4.59 | 4.79 | 5.27 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86 | 10.35
83.616 | 8.2 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 4.88 | 4.49 | 4.88 | 18.26 | 4.79 | 18.07 | 9.77 | 20.8 | 9.67 | 9.96 | 9.96 | 9.86 | 1.66 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.63 | 4.79 | 15.82 | 4.49 | 5.86 | 5.08 | 9.77 | 10.16 | 9.86 | 10.25 | 9.86 | 10.45
83.648 | 8.79 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 4.88 | 5.66 | 4.88 | 18.16 | 4.79 | 18.56 | 9.77 | 21.19 | 9.67 | 9.96 | 9.86 | 9.86 | 2.25 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.12 | 4.79 | 16.5 | 4.59 | 7.03 | 5.08 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 10.16
83.68 | 9.38 | 0.1 | 0.68 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 18.95 | 9.77 | 21.58 | 9.67 | 9.96 | 9.77 | 9.86 | 1.95 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.61 | 4.79 | 17.19 | 4.49 | 8.98 | 4.88 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.16
83.712 | 9.96 | 0.1 | 1.37 | 0.2 | 0.1 | 0.2 | 4.88 | 5.37 | 4.88 | 18.07 | 4.79 | 19.43 | 9.77 | 21.97 | 9.67 | 10.16 | 9.67 | 9.86 | 1.66 | 0 | 17.58 | 0 | 17.68 | 0.1 | 25.29 | 4.79 | 17.87 | 4.59 | 10.74 | 4.88 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.25
83.744 | 10.55 | 0.1 | 2.44 | 0.2 | 0.1 | 0.2 | 4.88 | 5.08 | 4.79 | 18.07 | 4.79 | 19.82 | 9.77 | 22.36 | 9.67 | 10.94 | 9.77 | 9.86 | 1.56 | 0.1 | 9.67 | 0.1 | 17.58 | 0.1 | 25.78 | 4.79 | 18.56 | 4.59 | 11.91 | 4.88 | 9.77 | 10.06 | 9.77 | 10.06 | 9.86 | 10.16
83.776 | 11.13 | 0.1 | 3.81 | 0.1 | 0.1 | 0.2 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.31 | 9.77 | 22.75 | 9.67 | 11.91 | 9.77 | 9.96 | 1.46 | 0.1 | 2.44 | 0.1 | 13.87 | 0.1 | 26.17 | 4.79 | 19.14 | 4.69 | 12.89 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
83.808 | 11.72 | 0.1 | 4.49 | 0.1 | 0.1 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.7 | 9.77 | 23.15 | 9.67 | 13.28 | 9.67 | 9.96 | 1.37 | 0.1 | 2.25 | 0.1 | 3.42 | 0.1 | 21.68 | 4.79 | 19.73 | 4.69 | 13.77 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
83.84 | 12.31 | 0.1 | 5.08 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.51 | 9.77 | 23.54 | 9.67 | 14.16 | 9.67 | 9.96 | 1.27 | 0.1 | 1.95 | 0.1 | 2.25 | 0.1 | 20.61 | 4.88 | 20.41 | 4.59 | 14.55 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
83.872 | 12.79 | 0.1 | 5.66 | 0 | 1.56 | 0.1 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 17.87 | 9.77 | 23.83 | 9.67 | 14.75 | 9.67 | 9.86 | 1.17 | 0.1 | 1.66 | 0.1 | 1.95 | 0.1 | 21.39 | 4.79 | 20.9 | 4.69 | 15.23 | 4.69 | 9.77 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86
83.904 | 13.38 | 0.1 | 6.15 | 0 | 2.93 | 0 | 5.57 | 4.69 | 4.79 | 8.4 | 4.79 | 18.26 | 9.77 | 24.22 | 9.67 | 15.23 | 9.67 | 9.86 | 1.07 | 0.1 | 1.56 | 0.1 | 1.86 | 0 | 21.29 | 4.79 | 21.49 | 4.79 | 15.92 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
83.936 | 13.87 | 0.1 | 6.74 | 0.1 | 3.91 | 0.1 | 6.25 | 4.79 | 4.88 | 2.83 | 4.88 | 18.36 | 9.86 | 20.8 | 9.77 | 15.82 | 9.67 | 9.96 | 0.88 | 0.1 | 1.56 | 0.1 | 1.66 | 0.1 | 21.39 | 4.88 | 22.17 | 4.79 | 16.7 | 4.59 | 10.64 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
83.968 | 14.45 | 0.1 | 7.42 | 0.1 | 4.69 | 0.1 | 7.62 | 4.79 | 4.88 | 3.22 | 4.88 | 18.26 | 9.77 | 20.9 | 9.77 | 16.31 | 9.67 | 9.96 | 0.88 | 0.1 | 1.46 | 0.1 | 1.56 | 0.1 | 21.29 | 4.88 | 22.66 | 4.79 | 17.29 | 4.59 | 12.01 | 9.77 | 9.86 | 9.77 | 10.64 | 9.77
84 | 14.94 | 0.1 | 8.01 | 0.1 | 5.27 | 0.1 | 8.69 | 4.69 | 4.88 | 4.2 | 4.88 | 18.16 | 9.77 | 21.09 | 9.77 | 16.8 | 9.57 | 9.96 | 0.88 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 21.39 | 4.88 | 23.24 | 4.79 | 17.97 | 4.59 | 13.87 | 9.77 | 10.64 | 9.77 | 11.62 | 9.77
84.032 | 15.43 | 0.1 | 8.59 | 0.1 | 5.76 | 0.1 | 9.57 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 17.29 | 9.57 | 9.86 | 0.68 | 0.1 | 1.37 | 0.1 | 1.27 | 0.1 | 21.39 | 4.88 | 23.73 | 4.79 | 18.75 | 4.59 | 15.53 | 9.77 | 11.82 | 9.67 | 13.38 | 9.77
84.064 | 15.92 | 0.1 | 9.18 | 0.1 | 6.35 | 0.1 | 10.25 | 4.49 | 4.88 | 5.96 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 17.77 | 9.57 | 9.86 | 0.59 | 0.1 | 1.07 | 0 | 1.17 | 0 | 21.29 | 4.79 | 24.22 | 4.79 | 19.24 | 4.49 | 16.6 | 9.67 | 13.28 | 9.67 | 15.14 | 9.67
84.096 | 16.41 | 0.1 | 9.77 | 0.1 | 6.93 | 0.1 | 10.94 | 4.59 | 4.88 | 5.57 | 4.88 | 18.16 | 9.86 | 21.39 | 9.77 | 18.26 | 9.57 | 9.96 | 0.59 | 0.1 | 1.07 | 0.1 | 1.17 | 0.1 | 21.39 | 4.88 | 24.81 | 4.88 | 19.92 | 4.59 | 17.48 | 9.67 | 15.14 | 9.67 | 16.31 | 9.67
84.128 | 16.9 | 0.1 | 10.45 | 0.1 | 7.62 | 0.1 | 11.43 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.65 | 9.57 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 25.29 | 4.88 | 20.51 | 4.59 | 18.26 | 9.67 | 16.31 | 9.67 | 17.19 | 9.67
84.16 | 16.9 | 0.1 | 11.04 | 0.1 | 8.2 | 0.1 | 12.01 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 19.04 | 9.67 | 9.96 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 21 | 4.59 | 18.95 | 9.67 | 17.29 | 9.67 | 18.07 | 9.67
84.192 | 14.55 | 0.1 | 11.62 | 0.1 | 8.79 | 0 | 12.6 | 4.49 | 4.88 | 4.88 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 19.53 | 9.57 | 9.86 | 0.39 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 21.29 | 4.88 | 26.17 | 4.79 | 21.58 | 4.59 | 19.73 | 9.67 | 18.07 | 9.67 | 18.75 | 9.67
84.224 | 14.84 | 0.1 | 12.11 | 0.1 | 9.38 | 0.1 | 13.18 | 4.49 | 4.88 | 4.88 | 4.79 | 17.97 | 9.77 | 21.39 | 9.77 | 19.92 | 9.67 | 9.96 | 0.29 | 0.1 | 0.59 | 0 | 0.78 | 0.1 | 21 | 4.79 | 23.05 | 4.88 | 22.17 | 4.59 | 20.31 | 9.67 | 18.75 | 9.67 | 19.34 | 9.67
84.256 | 14.75 | 0.1 | 12.7 | 0.1 | 9.96 | 0.1 | 13.67 | 4.59 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.39 | 9.77 | 20.31 | 9.67 | 10.74 | 0.29 | 0.1 | 0.59 | 0.1 | 0.78 | 0.1 | 5.27 | 4.88 | 22.27 | 4.88 | 22.75 | 4.59 | 20.9 | 9.67 | 19.34 | 9.57 | 20.12 | 9.67
84.288 | 14.65 | 0.1 | 13.18 | 0.1 | 10.55 | 0.1 | 14.26 | 4.49 | 4.88 | 4.88 | 4.88 | 8.5 | 9.77 | 21.39 | 9.77 | 20.9 | 9.67 | 11.43 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.73 | 4.88 | 21.78 | 4.88 | 23.24 | 4.59 | 21.49 | 9.67 | 20.02 | 9.67 | 20.7 | 9.67
84.32 | 14.55 | 0.1 | 13.77 | 0.1 | 11.13 | 0.1 | 14.75 | 4.59 | 4.88 | 4.79 | 4.88 | 2.83 | 9.86 | 21.39 | 9.77 | 21.19 | 9.67 | 12.7 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.42 | 4.88 | 21.58 | 4.88 | 23.93 | 4.69 | 22.07 | 9.67 | 20.61 | 9.67 | 21.29 | 9.67
84.352 | 14.45 | 0.1 | 14.36 | 0.1 | 11.72 | 0.1 | 15.33 | 4.59 | 4.88 | 4.69 | 4.88 | 3.32 | 9.86 | 21.39 | 9.77 | 21.68 | 9.67 | 13.77 | 0.2 | 0.1 | 0.39 | 0 | 0.49 | 0.1 | 4 | 4.88 | 21.49 | 4.88 | 24.41 | 4.69 | 22.75 | 9.67 | 21.19 | 9.67 | 21.78 | 9.67
84.384 | 14.45 | 0 | 14.75 | 0.1 | 12.31 | 0.1 | 15.82 | 4.59 | 4.79 | 4.59 | 4.88 | 4.1 | 9.77 | 10.64 | 9.86 | 21.97 | 9.67 | 14.45 | 0.1 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 5.47 | 4.88 | 21.39 | 4.88 | 24.81 | 4.69 | 23.24 | 9.67 | 21.88 | 9.67 | 22.36 | 9.67
84.416 | 14.45 | 0.1 | 15.33 | 0.1 | 12.89 | 0.1 | 16.31 | 4.59 | 4.88 | 4.59 | 4.88 | 5.27 | 9.86 | 6.64 | 9.86 | 22.46 | 9.67 | 15.04 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 6.15 | 4.88 | 21.29 | 4.88 | 25.39 | 4.69 | 23.63 | 9.67 | 22.46 | 9.67 | 22.95 | 9.67
84.448 | 14.45 | 0.1 | 15.72 | 0.1 | 13.38 | 0.1 | 16.8 | 4.69 | 4.88 | 4.59 | 4.88 | 6.05 | 9.77 | 9.57 | 9.86 | 22.85 | 9.67 | 15.63 | 0.1 | 0.1 | 0.29 | 0.78 | 0.29 | 0.1 | 6.05 | 4.88 | 21.29 | 4.88 | 25.88 | 4.79 | 24.22 | 9.67 | 22.95 | 9.67 | 23.44 | 9.67
84.48 | 14.36 | 0.1 | 16.31 | 0.1 | 13.97 | 0.1 | 17.29 | 4.69 | 4.88 | 4.59 | 4.88 | 5.66 | 9.86 | 11.52 | 9.86 | 23.15 | 9.67 | 16.02 | 0.1 | 0.1 | 0.2 | 1.76 | 0.29 | 0.1 | 5.57 | 4.88 | 21.19 | 4.88 | 25.68 | 4.79 | 24.81 | 9.77 | 23.54 | 9.67 | 24.02 | 9.67
84.512 | 14.36 | 0.1 | 16.7 | 0.1 | 14.45 | 0.1 | 17.68 | 4.69 | 4.88 | 4.59 | 4.88 | 5.18 | 9.77 | 10.74 | 9.86 | 23.63 | 9.67 | 16.6 | 0.1 | 1.17 | 0.2 | 3.42 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.88 | 21.58 | 4.79 | 25.29 | 9.77 | 24.02 | 9.67 | 24.51 | 9.67
84.544 | 14.36 | 0.1 | 17.19 | 0.1 | 15.04 | 0.1 | 18.16 | 4.69 | 5.47 | 4.59 | 4.79 | 4.88 | 9.77 | 10.16 | 9.86 | 23.93 | 9.67 | 17.09 | 0.1 | 2.05 | 0.1 | 5.27 | 0.2 | 0 | 5.08 | 4.79 | 21.19 | 4.88 | 21.78 | 4.79 | 25.68 | 9.77 | 24.51 | 9.67 | 25.1 | 9.67
84.576 | 14.36 | 0.1 | 14.94 | 0.1 | 15.53 | 0.1 | 18.56 | 4.79 | 6.45 | 4.59 | 4.88 | 4.98 | 9.86 | 9.86 | 9.86 | 24.32 | 9.77 | 17.58 | 0.1 | 3.81 | 0.1 | 6.54 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.88 | 21.58 | 4.79 | 26.27 | 9.77 | 25 | 9.77 | 25.49 | 9.67
84.608 | 7.52 | 0.1 | 15.23 | 0.1 | 15.92 | 0.1 | 19.04 | 4.79 | 7.81 | 4.49 | 4.88 | 4.98 | 9.86 | 10.06 | 9.86 | 20.9 | 9.77 | 17.97 | 0.1 | 5.47 | 0.1 | 7.32 | 0.1 | 1.17 | 5.08 | 4.88 | 21.19 | 4.88 | 21.49 | 4.79 | 26.76 | 9.77 | 25.68 | 9.77 | 25.98 | 9.67
84.64 | 2.15 | 0.1 | 14.75 | 0.1 | 16.41 | 0.1 | 19.43 | 4.79 | 8.79 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 20.9 | 9.77 | 18.46 | 0.1 | 6.54 | 0.1 | 8.2 | 0.1 | 2.44 | 4.98 | 4.88 | 10.64 | 4.88 | 21.39 | 4.79 | 27.25 | 9.77 | 26.08 | 9.77 | 26.47 | 9.67
84.672 | 2.05 | 0.1 | 14.65 | 0.1 | 16.9 | 0 | 19.92 | 4.79 | 9.67 | 4.49 | 4.88 | 4.79 | 9.77 | 10.16 | 9.86 | 21.29 | 9.67 | 18.85 | 0.1 | 7.42 | 0.1 | 8.89 | 0.1 | 4.3 | 4.98 | 4.88 | 2.64 | 4.88 | 21.29 | 4.79 | 27.83 | 9.77 | 26.56 | 9.77 | 26.95 | 9.77
84.704 | 1.66 | 0.1 | 14.65 | 0.1 | 17.38 | 0.1 | 20.22 | 4.79 | 10.25 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 21.39 | 9.77 | 19.34 | 0 | 8.11 | 0.1 | 9.67 | 0.1 | 5.66 | 4.88 | 4.79 | 3.22 | 4.88 | 21.29 | 4.79 | 28.22 | 9.77 | 26.95 | 9.77 | 27.44 | 9.67
84.736 | 1.46 | 0.1 | 14.55 | 0.1 | 15.23 | 0.1 | 20.7 | 4.79 | 10.94 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.39 | 9.77 | 19.82 | 0.1 | 8.98 | 0.1 | 10.55 | 0.1 | 6.84 | 4.79 | 4.88 | 3.52 | 4.88 | 21.29 | 4.79 | 28.71 | 9.86 | 27.54 | 9.77 | 28.03 | 9.77
84.768 | 1.37 | 0.1 | 14.55 | 0.1 | 15.04 | 0.1 | 21.19 | 4.88 | 11.52 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.22 | 0.1 | 9.86 | 0.1 | 11.33 | 0.1 | 7.62 | 4.79 | 4.88 | 4.69 | 4.88 | 21.29 | 4.79 | 29.2 | 9.86 | 28.03 | 9.77 | 28.42 | 9.77
84.8 | 1.17 | 0.1 | 14.55 | 0.1 | 14.75 | 0.1 | 19.14 | 4.88 | 12.11 | 4.59 | 4.79 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.61 | 0.1 | 10.74 | 0.1 | 11.91 | 0.1 | 8.5 | 4.69 | 4.88 | 6.15 | 4.88 | 21.19 | 4.79 | 28.61 | 9.86 | 28.52 | 9.86 | 28.91 | 9.77
84.832 | 1.07 | 0.1 | 14.45 | 0.1 | 14.65 | 0 | 18.65 | 4.79 | 12.7 | 4.59 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21 | 0 | 11.43 | 0 | 12.79 | 0 | 9.28 | 4.69 | 4.88 | 5.96 | 4.88 | 21.19 | 4.79 | 24.02 | 9.86 | 29.01 | 9.77 | 29.4 | 9.77
84.864 | 0.98 | 0.1 | 14.45 | 0.1 | 14.65 | 0 | 18.56 | 4.79 | 13.18 | 4.59 | 4.79 | 4.49 | 9.77 | 9.67 | 9.86 | 21.49 | 9.77 | 21.39 | 0 | 12.21 | 0 | 13.57 | 0 | 9.96 | 4.59 | 4.88 | 5.66 | 4.88 | 21.19 | 4.79 | 24.61 | 9.77 | 29.4 | 9.77 | 25.88 | 9.77
84.896 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.77 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 12.89 | 0.1 | 14.36 | 0.1 | 10.84 | 4.59 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 24.51 | 9.86 | 25.88 | 9.86 | 25 | 9.77
84.928 | 0.78 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 14.36 | 4.69 | 5.57 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 22.27 | 0.1 | 13.77 | 0.1 | 15.04 | 0.1 | 11.62 | 4.59 | 4.88 | 5.08 | 4.88 | 9.28 | 4.88 | 24.51 | 9.86 | 25.1 | 9.86 | 24.61 | 9.77
84.96 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 14.84 | 4.79 | 6.25 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 22.56 | 0.1 | 14.45 | 0.1 | 15.72 | 0.1 | 12.4 | 4.59 | 4.88 | 5.08 | 4.88 | 2.64 | 4.79 | 24.41 | 9.86 | 24.71 | 9.86 | 24.41 | 9.77
84.992 | 0.59 | 0.1 | 9.96 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.43 | 4.79 | 7.62 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 0 | 15.23 | 0 | 16.5 | 0 | 13.28 | 4.49 | 5.08 | 5.08 | 4.88 | 3.22 | 4.79 | 24.32 | 9.86 | 24.61 | 9.86 | 24.51 | 9.77
85.024 | 0.59 | 0.1 | 2.93 | 0.1 | 14.45 | 0.1 | 18.16 | 4.79 | 15.82 | 4.79 | 8.59 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 23.34 | 0 | 15.92 | 0 | 17.09 | 0 | 13.87 | 4.59 | 5.96 | 4.98 | 4.88 | 3.81 | 4.79 | 24.32 | 9.77 | 24.51 | 9.86 | 24.41 | 9.77
85.056 | 0.59 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.31 | 4.79 | 9.57 | 4.59 | 9.86 | 9.67 | 9.86 | 10.06 | 9.86 | 23.73 | 0.1 | 16.6 | 0.1 | 17.68 | 0.1 | 14.65 | 4.59 | 7.52 | 4.88 | 4.88 | 4.98 | 4.88 | 24.32 | 9.86 | 24.41 | 9.86 | 24.32 | 9.86
85.088 | 0.49 | 0.1 | 1.86 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 16.8 | 4.79 | 10.16 | 4.59 | 10.06 | 9.67 | 9.86 | 6.84 | 9.86 | 24.12 | 0.1 | 17.29 | 0.1 | 18.46 | 0.1 | 15.33 | 4.49 | 9.67 | 4.88 | 4.88 | 6.15 | 4.88 | 24.32 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86
85.12 | 0.39 | 0.1 | 1.46 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 17.29 | 4.79 | 10.74 | 4.59 | 10.94 | 9.67 | 9.86 | 9.77 | 9.86 | 24.41 | 0.1 | 17.87 | 0.1 | 19.14 | 0.1 | 16.11 | 4.49 | 11.23 | 4.88 | 4.88 | 6.05 | 4.88 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
85.152 | 0.29 | 0.1 | 1.46 | 0.1 | 7.81 | 0.1 | 18.07 | 4.88 | 17.68 | 4.79 | 11.43 | 4.59 | 12.01 | 9.57 | 9.86 | 11.52 | 9.86 | 21.88 | 0 | 18.46 | 0 | 19.73 | 0 | 16.8 | 4.49 | 12.4 | 4.79 | 4.88 | 5.47 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
85.184 | 0.2 | 0.1 | 1.27 | 0.1 | 2.25 | 0.1 | 18.07 | 4.88 | 18.16 | 4.79 | 12.01 | 4.59 | 13.18 | 9.57 | 9.86 | 10.84 | 9.86 | 22.07 | 0 | 19.24 | 0 | 20.31 | 0 | 17.29 | 4.49 | 13.18 | 4.69 | 4.88 | 5.27 | 4.88 | 24.22 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
85.216 | 0.2 | 0.1 | 1.17 | 0.1 | 2.05 | 0.1 | 9.47 | 4.88 | 18.56 | 4.88 | 12.6 | 4.69 | 14.06 | 9.67 | 9.86 | 10.25 | 9.86 | 21.88 | 0.1 | 19.82 | 0.1 | 20.9 | 0.1 | 18.07 | 4.59 | 14.06 | 4.79 | 4.88 | 5.08 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
85.248 | 0.2 | 0.1 | 0.98 | 0.1 | 1.86 | 0.1 | 3.03 | 4.88 | 19.04 | 4.88 | 13.09 | 4.69 | 14.65 | 9.67 | 9.86 | 9.86 | 9.86 | 21.68 | 0.1 | 20.31 | 0.1 | 21.49 | 0.1 | 18.75 | 4.59 | 14.75 | 4.59 | 4.88 | 5.08 | 4.88 | 12.6 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86
85.28 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.13 | 4.88 | 19.53 | 4.88 | 13.57 | 4.69 | 15.23 | 9.67 | 9.86 | 9.96 | 9.86 | 21.68 | 0.1 | 21 | 0.1 | 22.07 | 0.1 | 19.34 | 4.59 | 15.53 | 4.59 | 4.88 | 5.08 | 4.88 | 5.57 | 9.86 | 24.22 | 9.86 | 17.68 | 9.86
85.312 | 0.1 | 0.1 | 0.78 | 0.1 | 1.37 | 0 | 3.91 | 4.88 | 20.02 | 4.88 | 14.16 | 4.69 | 15.82 | 9.67 | 9.77 | 10.16 | 9.86 | 21.49 | 0 | 21.68 | 0 | 22.66 | 0 | 20.12 | 4.59 | 16.21 | 4.49 | 4.88 | 5.08 | 4.79 | 8.4 | 9.86 | 15.33 | 9.86 | 6.64 | 9.86
85.344 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 4.88 | 4.88 | 20.22 | 4.88 | 14.65 | 4.69 | 16.31 | 9.67 | 10.06 | 10.16 | 9.86 | 21.49 | 0 | 22.07 | 0 | 18.75 | 0 | 20.61 | 4.69 | 16.8 | 4.49 | 4.88 | 5.08 | 4.88 | 11.91 | 9.86 | 6.93 | 9.86 | 7.03 | 9.86
85.376 | 0.1 | 0.78 | 0.68 | 0.1 | 1.17 | 0.1 | 5.86 | 4.88 | 20.8 | 4.88 | 15.14 | 4.79 | 16.8 | 9.67 | 10.94 | 9.96 | 9.86 | 21.49 | 0 | 22.66 | 0.1 | 18.26 | 0 | 21.29 | 4.69 | 17.58 | 4.59 | 4.88 | 4.98 | 4.88 | 11.52 | 9.86 | 7.03 | 9.86 | 10.74 | 9.86
85.408 | 0 | 1.56 | 0.59 | 0.1 | 1.07 | 0.1 | 5.76 | 4.88 | 20.41 | 4.88 | 15.63 | 4.79 | 17.29 | 9.67 | 11.91 | 9.86 | 9.86 | 21.49 | 0 | 19.24 | 0.1 | 18.07 | 0.1 | 21.78 | 4.69 | 18.26 | 4.59 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 11.04 | 9.86 | 12.01 | 9.86
85.44 | 0 | 2.73 | 0.49 | 0.1 | 0.88 | 0.1 | 5.27 | 4.88 | 18.26 | 4.88 | 16.11 | 4.79 | 17.77 | 9.67 | 13.28 | 9.96 | 9.86 | 21.49 | 0.1 | 18.56 | 0.1 | 17.87 | 0.1 | 22.36 | 4.69 | 18.85 | 4.59 | 4.88 | 4.88 | 4.88 | 10.16 | 9.86 | 12.01 | 9.86 | 10.74 | 9.86
85.472 | 0 | 3.91 | 0.49 | 0 | 0.78 | 0 | 4.98 | 4.88 | 18.07 | 4.88 | 16.7 | 4.79 | 18.16 | 9.67 | 14.16 | 9.77 | 9.86 | 21.39 | 0 | 17.97 | 0 | 17.97 | 0 | 20.61 | 4.69 | 19.53 | 4.49 | 5.66 | 4.69 | 4.79 | 10.25 | 9.86 | 10.55 | 9.86 | 9.96 | 9.86
85.504 | 0 | 4.59 | 0.39 | 0.1 | 0.68 | 0.1 | 4.79 | 4.88 | 18.36 | 4.88 | 17.09 | 4.79 | 18.65 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 0 | 17.97 | 0 | 17.87 | 0 | 18.07 | 4.69 | 20.12 | 4.49 | 6.64 | 4.79 | 4.79 | 10.35 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
85.536 | 0.1 | 5.18 | 0.29 | 0.1 | 0.68 | 0.1 | 4.98 | 4.88 | 18.26 | 4.88 | 17.58 | 4.79 | 19.14 | 9.67 | 15.33 | 9.77 | 9.86 | 21.39 | 0.1 | 17.97 | 0.1 | 17.77 | 0.1 | 18.07 | 4.79 | 20.7 | 4.49 | 8.3 | 4.69 | 4.88 | 10.35 | 9.86 | 10.25 | 9.86 | 10.35 | 9.86
85.568 | 0.1 | 5.66 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 17.97 | 4.88 | 19.53 | 9.67 | 15.82 | 9.77 | 9.86 | 13.28 | 0.1 | 17.77 | 0.1 | 17.77 | 0 | 17.97 | 4.79 | 21.29 | 4.59 | 10.35 | 4.59 | 4.88 | 10.25 | 9.86 | 10.35 | 9.86 | 10.45 | 9.86
85.6 | 0.1 | 6.25 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.46 | 4.79 | 20.02 | 9.77 | 16.21 | 9.77 | 9.86 | 7.03 | 0.1 | 17.77 | 0.1 | 17.77 | 0 | 17.87 | 4.79 | 21.88 | 4.59 | 11.62 | 4.59 | 4.88 | 10.06 | 9.86 | 10.45 | 9.86 | 10.25 | 9.86
85.632 | 0 | 6.93 | 0.2 | 0.1 | 0.49 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 18.95 | 4.79 | 20.31 | 9.67 | 16.8 | 9.67 | 9.86 | 8.2 | 0 | 17.68 | 0 | 17.77 | 0 | 17.77 | 4.79 | 22.46 | 4.59 | 12.6 | 4.59 | 4.79 | 10.25 | 9.86 | 10.25 | 9.86 | 10.16 | 9.86
85.664 | 0 | 7.52 | 0.1 | 0.1 | 0.49 | 0 | 4.69 | 4.88 | 18.16 | 4.88 | 19.34 | 4.79 | 20.8 | 9.67 | 17.19 | 9.67 | 9.86 | 11.04 | 0 | 17.68 | 0 | 17.68 | 0 | 17.77 | 4.79 | 22.95 | 4.59 | 13.48 | 4.59 | 4.88 | 10.06 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
85.696 | 0.1 | 8.11 | 0.1 | 0.1 | 0.39 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 21.19 | 9.77 | 17.77 | 9.67 | 9.86 | 11.23 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.54 | 4.59 | 14.26 | 4.59 | 4.88 | 9.96 | 9.86 | 10.06 | 9.86 | 10.16 | 9.86
85.728 | 0.1 | 8.79 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.22 | 4.88 | 21.58 | 9.77 | 18.16 | 9.67 | 9.86 | 10.64 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 24.02 | 4.59 | 15.04 | 4.59 | 4.79 | 9.96 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
85.76 | 0.1 | 9.38 | 0 | 0.2 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.61 | 4.88 | 21.97 | 9.77 | 18.65 | 9.67 | 9.86 | 9.86 | 0 | 17.58 | 0.1 | 13.87 | 0.1 | 17.77 | 4.88 | 24.51 | 4.69 | 15.72 | 4.49 | 4.88 | 9.96 | 9.86 | 10.06 | 9.86 | 9.86 | 9.86
85.792 | 0 | 9.96 | 0 | 1.27 | 0.2 | 0.1 | 4.49 | 4.88 | 18.07 | 4.88 | 21 | 4.79 | 22.46 | 9.77 | 19.14 | 9.57 | 9.86 | 9.86 | 0 | 17.58 | 0 | 3.03 | 0 | 17.68 | 4.88 | 25.1 | 4.69 | 16.41 | 4.49 | 4.79 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
85.824 | 0 | 10.45 | 0 | 2.25 | 0.2 | 1.17 | 4.49 | 4.88 | 18.07 | 4.88 | 18.95 | 4.79 | 22.75 | 9.77 | 19.43 | 9.57 | 9.86 | 10.06 | 0 | 9.08 | 0 | 2.15 | 0 | 17.68 | 4.88 | 25.59 | 4.69 | 17.09 | 4.49 | 4.79 | 9.77 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
85.856 | 0.1 | 11.13 | 0.1 | 3.52 | 0.1 | 1.95 | 4.59 | 4.88 | 8.4 | 4.88 | 18.85 | 4.88 | 23.24 | 9.77 | 20.02 | 9.67 | 9.86 | 10.16 | 0.1 | 2.34 | 0.1 | 2.15 | 0.1 | 17.68 | 4.88 | 26.08 | 4.69 | 17.77 | 4.59 | 4.79 | 9.77 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
85.888 | 0.1 | 11.72 | 0.1 | 4.39 | 0.1 | 3.32 | 4.49 | 4.88 | 2.73 | 4.88 | 18.56 | 4.88 | 23.63 | 9.77 | 20.41 | 9.67 | 9.86 | 9.96 | 0.1 | 2.25 | 0.1 | 1.76 | 0.1 | 17.68 | 4.88 | 21.49 | 4.69 | 18.36 | 4.59 | 4.79 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86
85.92 | 0.1 | 12.21 | 0.1 | 4.98 | 0 | 4.2 | 4.49 | 4.88 | 3.13 | 4.88 | 18.36 | 4.88 | 23.93 | 9.77 | 20.8 | 9.67 | 10.64 | 9.96 | 0.1 | 2.25 | 0.1 | 1.76 | 0 | 8.01 | 4.88 | 20.22 | 4.79 | 19.14 | 4.59 | 5.76 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77 | 9.86
85.952 | 0 | 12.79 | 0.1 | 5.47 | 0 | 4.88 | 4.49 | 4.88 | 4.3 | 4.88 | 18.26 | 4.88 | 24.32 | 9.77 | 21.19 | 9.57 | 11.43 | 9.77 | 0 | 1.76 | 0 | 1.46 | 0 | 1.66 | 4.88 | 21.19 | 4.69 | 19.73 | 4.59 | 6.54 | 9.67 | 10.84 | 9.77 | 9.77 | 9.67 | 9.86
85.984 | 0 | 13.28 | 0 | 5.96 | 0.1 | 5.37 | 4.49 | 4.88 | 5.27 | 4.88 | 18.16 | 4.88 | 20.9 | 9.77 | 21.58 | 9.67 | 12.6 | 9.86 | 0 | 1.76 | 0.1 | 1.56 | 0 | 2.25 | 4.88 | 21.19 | 4.79 | 20.31 | 4.59 | 8.01 | 9.67 | 11.23 | 9.77 | 10.74 | 9.67 | 9.86
86.016 | 0.1 | 13.87 | 0.1 | 6.64 | 0.1 | 5.96 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 21.09 | 9.77 | 21.97 | 9.67 | 13.67 | 9.77 | 0 | 1.56 | 0.1 | 1.37 | 0.1 | 2.25 | 4.88 | 21.19 | 4.79 | 20.9 | 4.69 | 10.16 | 9.67 | 12.79 | 9.77 | 11.52 | 9.67 | 10.74
86.048 | 0.1 | 14.36 | 0.1 | 7.23 | 0.1 | 6.54 | 4.59 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 21.29 | 9.77 | 22.36 | 9.67 | 14.36 | 9.67 | 0.1 | 1.46 | 0.1 | 1.27 | 0.1 | 1.66 | 4.88 | 21.19 | 4.79 | 21.49 | 4.69 | 11.62 | 9.67 | 14.84 | 9.67 | 13.09 | 9.67 | 11.82
86.08 | 0.1 | 14.84 | 0.1 | 7.81 | 0 | 7.13 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 22.75 | 9.67 | 14.94 | 9.77 | 0.1 | 1.37 | 0.1 | 1.27 | 0.1 | 1.76 | 4.88 | 21.29 | 4.79 | 22.07 | 4.69 | 12.5 | 9.67 | 16.11 | 9.67 | 15.04 | 9.67 | 13.48
86.112 | 0 | 15.43 | 0 | 8.4 | 0.1 | 7.72 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 23.15 | 9.67 | 15.53 | 9.67 | 0 | 1.27 | 0 | 1.17 | 0 | 1.46 | 4.88 | 21.29 | 4.79 | 22.66 | 4.69 | 13.38 | 9.67 | 17.09 | 9.67 | 16.41 | 9.57 | 15.33
86.144 | 0 | 15.92 | 0 | 8.98 | 0.1 | 8.4 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 23.54 | 9.67 | 15.92 | 9.67 | 0 | 1.07 | 0 | 0.88 | 0 | 1.46 | 4.88 | 21.19 | 4.79 | 23.24 | 4.69 | 14.16 | 9.57 | 17.87 | 9.67 | 17.19 | 9.67 | 16.41
86.176 | 0.1 | 16.41 | 0.1 | 9.67 | 0.1 | 8.98 | 4.69 | 4.98 | 4.98 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.93 | 9.67 | 16.6 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 23.73 | 4.79 | 14.94 | 9.67 | 18.56 | 9.67 | 18.07 | 9.67 | 17.38
86.208 | 0.1 | 16.8 | 0.1 | 10.25 | 0.1 | 9.57 | 4.69 | 5.76 | 4.79 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 24.22 | 9.67 | 16.99 | 9.67 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.19 | 4.79 | 24.22 | 4.79 | 15.63 | 9.57 | 19.24 | 9.67 | 18.75 | 9.67 | 18.07
86.24 | 0.1 | 17.38 | 0.1 | 10.74 | 0 | 10.16 | 4.69 | 6.74 | 4.79 | 4.88 | 9.86 | 4.88 | 21.39 | 9.86 | 20.9 | 9.77 | 17.58 | 9.67 | 0.1 | 0.88 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.81 | 4.79 | 16.41 | 9.57 | 19.92 | 9.67 | 19.34 | 9.67 | 18.75
86.272 | 0 | 15.23 | 0 | 11.43 | 0.1 | 10.64 | 4.69 | 8.2 | 4.69 | 4.88 | 3.03 | 4.88 | 21.49 | 9.77 | 21.39 | 9.77 | 17.97 | 9.57 | 1.17 | 0.78 | 0 | 0.68 | 0 | 0.98 | 4.88 | 21.19 | 4.79 | 25.29 | 4.79 | 16.99 | 9.57 | 20.7 | 9.57 | 20.12 | 9.57 | 19.43
86.304 | 0 | 15.14 | 0 | 12.01 | 0.1 | 11.33 | 4.69 | 9.08 | 4.79 | 4.88 | 3.13 | 4.88 | 21.39 | 9.77 | 21.29 | 9.77 | 18.46 | 9.57 | 2.44 | 0.68 | 0 | 0.59 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.68 | 4.79 | 17.68 | 9.57 | 21.19 | 9.67 | 20.61 | 9.57 | 20.12
86.336 | 0.1 | 14.84 | 0.1 | 12.5 | 0.1 | 11.91 | 4.79 | 9.86 | 4.69 | 4.88 | 3.91 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 18.75 | 9.57 | 4.2 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 6.54 | 4.79 | 26.27 | 4.79 | 18.36 | 9.67 | 21.68 | 9.67 | 21.29 | 9.67 | 20.61
86.368 | 0.1 | 14.75 | 0.1 | 13.09 | 0.1 | 12.4 | 4.79 | 10.45 | 4.59 | 4.88 | 5.08 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 19.24 | 9.57 | 5.66 | 0.59 | 0.1 | 0.49 | 0 | 0.68 | 4.88 | 2.64 | 4.79 | 22.66 | 4.79 | 18.95 | 9.67 | 22.27 | 9.67 | 21.78 | 9.67 | 21.19
86.4 | 0.1 | 14.75 | 0.1 | 13.57 | 0.1 | 12.99 | 4.79 | 11.13 | 4.59 | 4.88 | 5.96 | 4.88 | 21.19 | 9.86 | 21.39 | 9.77 | 19.73 | 9.57 | 6.84 | 0.49 | 0.1 | 0.39 | 0.1 | 0.68 | 4.88 | 3.32 | 4.79 | 22.27 | 4.79 | 19.53 | 9.67 | 22.85 | 9.67 | 22.36 | 9.67 | 21.78
86.432 | 0 | 14.55 | 0 | 14.16 | 0.1 | 13.57 | 4.79 | 11.82 | 4.59 | 4.88 | 5.66 | 4.88 | 8.69 | 9.77 | 21.49 | 9.77 | 20.12 | 9.57 | 7.52 | 0.39 | 0.88 | 0.29 | 0 | 0.59 | 4.88 | 4.1 | 4.79 | 21.68 | 4.79 | 20.22 | 9.67 | 23.44 | 9.67 | 22.95 | 9.67 | 22.36
86.464 | 0 | 14.45 | 0 | 14.55 | 0.1 | 14.06 | 4.79 | 12.31 | 4.49 | 4.88 | 5.27 | 4.88 | 6.93 | 9.77 | 21.39 | 9.77 | 20.51 | 9.57 | 8.3 | 0.39 | 1.86 | 0.29 | 0.49 | 0.49 | 4.88 | 5.37 | 4.79 | 21.58 | 4.79 | 20.8 | 9.67 | 23.83 | 9.67 | 23.44 | 9.67 | 22.95
86.496 | 0.1 | 14.55 | 0.1 | 15.14 | 0.1 | 14.65 | 4.79 | 12.89 | 4.59 | 4.88 | 4.98 | 4.88 | 9.96 | 9.86 | 21.39 | 9.86 | 21 | 9.67 | 9.08 | 0.39 | 3.52 | 0.2 | 1.66 | 0.59 | 4.88 | 6.15 | 4.88 | 21.58 | 4.79 | 21.39 | 9.67 | 24.41 | 9.67 | 24.02 | 9.67 | 23.44
86.528 | 0.1 | 14.45 | 0.1 | 15.63 | 0.1 | 15.14 | 4.88 | 13.38 | 4.59 | 4.88 | 4.88 | 4.88 | 11.62 | 9.86 | 21.49 | 9.86 | 21.39 | 9.67 | 9.86 | 0.29 | 5.27 | 0.2 | 3.13 | 0.49 | 4.88 | 5.86 | 4.79 | 21.49 | 4.79 | 21.97 | 9.67 | 25 | 9.67 | 24.41 | 9.67 | 24.02
86.56 | 0.1 | 14.45 | 0.1 | 16.11 | 0.1 | 15.63 | 4.88 | 13.97 | 4.49 | 4.88 | 4.88 | 4.88 | 10.64 | 9.86 | 21.39 | 9.86 | 21.78 | 9.67 | 10.74 | 0.29 | 6.54 | 0.2 | 5.18 | 0.39 | 4.88 | 5.57 | 4.79 | 21.39 | 4.79 | 22.46 | 9.67 | 25.49 | 9.67 | 25 | 9.67 | 24.51
86.592 | 0 | 14.45 | 0 | 16.6 | 0.1 | 16.02 | 4.79 | 14.45 | 4.49 | 5.47 | 4.88 | 4.88 | 9.86 | 9.77 | 21.39 | 9.77 | 22.17 | 9.67 | 11.43 | 0.2 | 7.32 | 0.1 | 6.45 | 0.29 | 4.88 | 5.18 | 4.79 | 21.39 | 4.79 | 23.05 | 9.67 | 25.98 | 9.67 | 25.49 | 9.67 | 25.1
86.624 | 0 | 14.36 | 0 | 17.09 | 0.1 | 16.5 | 4.88 | 14.94 | 4.49 | 6.45 | 4.79 | 4.88 | 9.86 | 9.77 | 21.39 | 9.86 | 22.56 | 9.67 | 12.11 | 0.2 | 8.11 | 0.1 | 7.13 | 0.29 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 23.63 | 9.67 | 26.37 | 9.67 | 25.98 | 9.67 | 25.49
86.656 | 0.1 | 14.36 | 0.1 | 13.87 | 0.1 | 17.09 | 4.88 | 15.43 | 4.59 | 7.72 | 4.79 | 4.88 | 10.25 | 9.86 | 20.8 | 9.86 | 22.95 | 9.67 | 12.99 | 0.2 | 8.79 | 0.1 | 8.01 | 0.29 | 4.88 | 5.08 | 4.88 | 21.39 | 4.79 | 24.22 | 9.77 | 26.95 | 9.77 | 26.47 | 9.67 | 26.08
86.688 | 0.1 | 8.4 | 0.1 | 13.97 | 0.1 | 14.16 | 4.88 | 16.02 | 4.59 | 8.79 | 4.69 | 4.88 | 10.35 | 9.86 | 9.28 | 9.86 | 23.34 | 9.67 | 13.77 | 0.2 | 9.67 | 0.1 | 8.89 | 0.29 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 24.71 | 9.77 | 27.44 | 9.77 | 26.95 | 9.67 | 26.47
86.72 | 0.1 | 2.44 | 0.1 | 14.16 | 0 | 14.06 | 4.88 | 16.5 | 4.59 | 9.67 | 4.59 | 4.88 | 9.96 | 9.86 | 6.93 | 9.86 | 23.73 | 9.67 | 14.36 | 0.1 | 10.55 | 0.1 | 9.57 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 25.2 | 9.77 | 28.03 | 9.77 | 27.44 | 9.77 | 27.05
86.752 | 0 | 1.95 | 0.1 | 14.45 | 0.1 | 14.16 | 4.88 | 16.99 | 4.59 | 10.35 | 4.59 | 4.88 | 9.86 | 9.77 | 10.25 | 9.86 | 24.12 | 9.67 | 15.23 | 0.1 | 11.33 | 0 | 10.35 | 0.2 | 4.88 | 4.98 | 4.79 | 21.29 | 4.79 | 25.68 | 9.77 | 28.42 | 9.77 | 28.03 | 9.67 | 27.54
86.784 | 0 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 17.29 | 4.59 | 10.94 | 4.59 | 4.88 | 9.86 | 9.77 | 11.62 | 9.86 | 24.41 | 9.67 | 15.82 | 0.1 | 12.01 | 0.1 | 11.13 | 0.1 | 4.88 | 4.79 | 4.79 | 14.84 | 4.79 | 26.17 | 9.67 | 28.91 | 9.77 | 28.42 | 9.67 | 28.03
86.816 | 0.1 | 1.37 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 17.77 | 4.69 | 11.52 | 4.59 | 4.88 | 9.96 | 9.86 | 10.74 | 9.86 | 22.27 | 9.67 | 16.5 | 0.1 | 12.79 | 0.1 | 11.91 | 0.2 | 4.88 | 4.88 | 4.79 | 3.52 | 4.88 | 22.07 | 9.77 | 29.4 | 9.77 | 28.91 | 9.77 | 28.42
86.848 | 0.1 | 1.46 | 0.1 | 14.55 | 0.1 | 14.45 | 4.88 | 18.26 | 4.69 | 12.11 | 4.49 | 4.88 | 9.86 | 9.86 | 9.96 | 9.86 | 22.07 | 9.67 | 17.29 | 0.1 | 13.57 | 0.1 | 12.6 | 0.1 | 4.88 | 4.79 | 4.79 | 3.22 | 4.88 | 21.29 | 9.77 | 24.71 | 9.77 | 29.4 | 9.77 | 28.91
86.88 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.75 | 4.69 | 12.7 | 4.49 | 4.88 | 9.86 | 9.77 | 9.86 | 9.96 | 21.78 | 9.77 | 17.77 | 0.1 | 14.45 | 0.1 | 13.38 | 0.1 | 4.88 | 4.69 | 4.79 | 3.61 | 4.79 | 21.39 | 9.77 | 23.54 | 9.77 | 25.1 | 9.77 | 29.4
86.912 | 0 | 1.17 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.69 | 13.28 | 4.49 | 4.88 | 9.67 | 9.77 | 10.16 | 9.86 | 21.58 | 9.67 | 18.46 | 0.1 | 15.14 | 0.1 | 14.06 | 0.1 | 5.18 | 4.69 | 4.79 | 4.49 | 4.79 | 21.39 | 9.77 | 24.41 | 9.77 | 24.81 | 9.77 | 26.08
86.944 | 0 | 0.98 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.69 | 13.77 | 4.49 | 4.88 | 9.77 | 9.77 | 10.25 | 9.86 | 21.68 | 9.77 | 19.14 | 0.1 | 15.72 | 0.1 | 14.84 | 0.1 | 6.05 | 4.59 | 4.79 | 5.96 | 4.79 | 21.39 | 9.77 | 24.22 | 9.77 | 24.41 | 9.77 | 25.1
86.976 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.02 | 4.79 | 14.26 | 4.49 | 5.57 | 9.77 | 9.77 | 9.96 | 9.96 | 21.58 | 9.77 | 19.73 | 0.1 | 16.41 | 0.1 | 15.63 | 0.1 | 7.72 | 4.59 | 4.79 | 6.15 | 4.79 | 21.39 | 9.77 | 24.32 | 9.86 | 24.32 | 9.77 | 24.71
87.008 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.51 | 4.79 | 14.84 | 4.49 | 6.45 | 9.67 | 9.77 | 9.86 | 9.96 | 21.49 | 9.77 | 20.31 | 0.1 | 17.09 | 0.1 | 16.21 | 0.1 | 9.77 | 4.59 | 4.79 | 5.76 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.41 | 9.77 | 24.51
87.04 | 0.1 | 0.88 | 0.1 | 14.45 | 0 | 14.45 | 4.88 | 20.9 | 4.79 | 15.33 | 4.49 | 7.72 | 9.67 | 9.86 | 9.96 | 9.96 | 21.49 | 9.77 | 20.9 | 0.1 | 17.77 | 0.1 | 16.9 | 0.1 | 11.33 | 4.59 | 4.79 | 5.27 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.32 | 9.77 | 24.51
87.072 | 0 | 0.59 | 0 | 14.36 | 0.1 | 14.36 | 4.88 | 17.48 | 4.79 | 15.82 | 4.59 | 8.79 | 9.67 | 10.64 | 9.96 | 9.86 | 21.39 | 9.77 | 21.58 | 0.1 | 18.46 | 0 | 17.68 | 0.1 | 12.31 | 4.59 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.41 | 9.77 | 24.32 | 9.77 | 24.41
87.104 | 0 | 0.59 | 0 | 5.47 | 0.1 | 14.36 | 4.88 | 17.58 | 4.79 | 16.31 | 4.49 | 9.57 | 9.67 | 11.62 | 9.77 | 9.86 | 21.39 | 9.77 | 22.07 | 0.1 | 18.95 | 0.1 | 18.16 | 0.1 | 13.18 | 4.59 | 4.79 | 5.18 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22
87.136 | 0.1 | 0.49 | 0.1 | 1.76 | 0.1 | 5.76 | 4.88 | 17.87 | 4.79 | 16.8 | 4.59 | 10.25 | 9.67 | 12.99 | 9.77 | 9.96 | 21.49 | 9.77 | 20.31 | 0.1 | 19.73 | 0.1 | 18.95 | 0.1 | 14.06 | 4.59 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.32
87.168 | 0.1 | 0.49 | 0 | 1.95 | 0.1 | 1.66 | 4.88 | 18.16 | 4.79 | 17.29 | 4.59 | 10.94 | 9.57 | 13.87 | 9.77 | 9.96 | 21.39 | 9.77 | 17.09 | 0.1 | 20.22 | 0.1 | 19.53 | 0.1 | 14.84 | 4.59 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.32 | 9.77 | 24.22
87.2 | 0 | 0.49 | 0 | 1.76 | 0.1 | 1.95 | 4.88 | 18.07 | 4.88 | 17.68 | 4.59 | 11.52 | 9.57 | 14.55 | 9.77 | 9.96 | 21.39 | 9.77 | 17.58 | 0.1 | 20.8 | 0.1 | 20.12 | 0.1 | 15.53 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22
87.232 | 0 | 0.29 | 0 | 1.46 | 0.1 | 1.66 | 4.88 | 18.07 | 4.79 | 18.16 | 4.69 | 12.11 | 9.57 | 15.14 | 9.67 | 9.86 | 12.79 | 9.77 | 17.48 | 0.1 | 21.49 | 0 | 20.8 | 0.1 | 16.31 | 4.59 | 4.79 | 4.98 | 4.79 | 16.02 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22
87.264 | 0.1 | 0.29 | 0 | 1.46 | 0.1 | 1.37 | 4.88 | 18.07 | 4.79 | 18.56 | 4.59 | 12.7 | 9.57 | 15.63 | 9.67 | 9.86 | 6.93 | 9.77 | 17.58 | 0.1 | 22.07 | 0.1 | 21.39 | 0.1 | 16.9 | 4.49 | 4.79 | 4.98 | 4.79 | 3.91 | 9.77 | 24.02 | 9.77 | 24.12 | 9.77 | 24.12
87.296 | 0.78 | 0.29 | 0.1 | 1.17 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 19.14 | 4.69 | 13.18 | 9.67 | 16.21 | 9.67 | 9.96 | 8.2 | 9.77 | 17.58 | 0.1 | 22.66 | 0.1 | 21.97 | 0.1 | 17.58 | 4.59 | 4.79 | 4.88 | 4.79 | 3.22 | 9.77 | 8.59 | 9.86 | 20.9 | 9.86 | 24.12
87.328 | 1.56 | 0.2 | 0.1 | 1.17 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.53 | 4.69 | 13.77 | 9.67 | 16.6 | 9.67 | 9.96 | 11.04 | 9.86 | 17.58 | 0.1 | 19.43 | 0.1 | 22.46 | 0.1 | 18.26 | 4.59 | 5.66 | 4.79 | 4.79 | 3.42 | 9.77 | 6.25 | 9.86 | 7.62 | 9.86 | 16.9
87.36 | 2.83 | 0.2 | 0.1 | 0.98 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 19.92 | 4.69 | 14.26 | 9.67 | 17.19 | 9.57 | 9.96 | 11.23 | 9.86 | 17.58 | 0.1 | 18.65 | 0.1 | 18.26 | 0.1 | 18.95 | 4.59 | 6.74 | 4.79 | 4.79 | 4.39 | 9.77 | 9.86 | 9.86 | 6.84 | 9.77 | 6.93
87.392 | 3.91 | 0.1 | 0 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.31 | 4.69 | 14.84 | 9.67 | 17.68 | 9.57 | 9.86 | 10.45 | 9.77 | 17.48 | 0.1 | 18.26 | 0.1 | 17.87 | 0.1 | 19.43 | 4.59 | 8.5 | 4.69 | 4.79 | 5.76 | 9.77 | 12.21 | 9.77 | 10.55 | 9.77 | 7.23
87.424 | 4.49 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.79 | 20.8 | 4.69 | 15.33 | 9.67 | 18.07 | 9.57 | 9.96 | 9.77 | 9.86 | 17.48 | 0.1 | 17.97 | 0 | 17.77 | 0.1 | 20.22 | 4.59 | 10.25 | 4.69 | 4.79 | 6.15 | 9.77 | 11.04 | 9.86 | 12.11 | 9.77 | 10.94
87.456 | 5.18 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.7 | 4.79 | 15.82 | 9.67 | 18.56 | 9.57 | 9.96 | 9.96 | 9.86 | 17.38 | 0.1 | 17.97 | 0.1 | 17.68 | 0.1 | 20.8 | 4.69 | 11.62 | 4.69 | 4.79 | 5.76 | 9.77 | 9.96 | 9.86 | 11.04 | 9.86 | 11.91
87.488 | 5.66 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 18.07 | 4.88 | 18.65 | 4.79 | 16.31 | 9.67 | 18.95 | 9.57 | 9.96 | 10.16 | 9.86 | 17.48 | 0.1 | 17.87 | 0.1 | 17.77 | 0.1 | 21.39 | 4.69 | 12.6 | 4.69 | 4.79 | 5.47 | 9.77 | 10.06 | 9.86 | 10.25 | 9.86 | 10.64
87.52 | 6.35 | 0.1 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 6.45 | 4.88 | 18.56 | 4.79 | 16.8 | 9.67 | 19.43 | 9.57 | 9.96 | 10.16 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 21.88 | 4.69 | 13.48 | 4.59 | 4.79 | 5.18 | 9.77 | 10.35 | 9.86 | 9.96 | 9.86 | 10.06
87.552 | 6.93 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 2.93 | 4.79 | 18.46 | 4.79 | 17.19 | 9.67 | 19.82 | 9.57 | 9.86 | 9.86 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.77 | 0.1 | 22.46 | 4.79 | 14.26 | 4.59 | 4.79 | 5.18 | 9.77 | 10.45 | 9.77 | 10.35 | 9.77 | 9.96
87.584 | 7.52 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.88 | 3.42 | 4.79 | 18.26 | 4.79 | 17.68 | 9.67 | 20.22 | 9.57 | 9.86 | 9.86 | 9.86 | 7.42 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.05 | 4.79 | 14.94 | 4.59 | 4.79 | 5.18 | 9.77 | 10.35 | 9.86 | 10.45 | 9.77 | 10.45
87.616 | 8.11 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 4.39 | 4.88 | 18.26 | 4.79 | 18.16 | 9.77 | 20.7 | 9.67 | 9.96 | 9.96 | 9.86 | 1.95 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.54 | 4.79 | 15.72 | 4.49 | 5.96 | 5.08 | 9.77 | 10.06 | 9.86 | 10.35 | 9.86 | 10.45
87.648 | 8.79 | 0.1 | 0 | 0.29 | 0.1 | 0.29 | 4.88 | 5.57 | 4.88 | 18.16 | 4.79 | 18.56 | 9.77 | 21.09 | 9.67 | 9.96 | 9.86 | 9.86 | 2.25 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 24.12 | 4.79 | 16.41 | 4.59 | 6.84 | 4.98 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 10.16
87.68 | 9.38 | 0.1 | 0.49 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.16 | 4.79 | 19.14 | 9.77 | 21.49 | 9.67 | 9.96 | 9.77 | 9.86 | 1.95 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 24.71 | 4.79 | 17.09 | 4.59 | 8.69 | 4.88 | 9.77 | 10.25 | 9.86 | 10.06 | 9.86 | 10.06
87.712 | 9.96 | 0.1 | 1.27 | 0.2 | 0.1 | 0.29 | 4.88 | 5.47 | 4.79 | 18.16 | 4.79 | 19.43 | 9.77 | 21.97 | 9.67 | 10.16 | 9.67 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.2 | 4.79 | 17.77 | 4.59 | 10.45 | 4.98 | 9.77 | 10.16 | 9.86 | 10.16 | 9.77 | 10.16
87.744 | 10.45 | 0.1 | 2.25 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.16 | 4.79 | 19.82 | 9.77 | 22.36 | 9.67 | 10.74 | 9.77 | 9.86 | 1.56 | 0.1 | 11.72 | 0.1 | 17.68 | 0.1 | 25.68 | 4.79 | 18.46 | 4.59 | 11.72 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.77 | 10.16
87.776 | 11.13 | 0.1 | 3.71 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.31 | 9.77 | 22.66 | 9.67 | 11.82 | 9.67 | 9.96 | 1.46 | 0.1 | 2.83 | 0.1 | 16.99 | 0.1 | 26.08 | 4.79 | 19.04 | 4.59 | 12.79 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
87.808 | 11.72 | 0.1 | 4.49 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.8 | 9.77 | 23.05 | 9.67 | 13.28 | 9.67 | 9.86 | 1.37 | 0.1 | 2.25 | 0.1 | 4.2 | 0.1 | 21.49 | 4.79 | 19.73 | 4.69 | 13.57 | 4.79 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86 | 9.96
87.84 | 12.21 | 0.1 | 5.08 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.7 | 9.77 | 23.54 | 9.67 | 14.06 | 9.67 | 9.96 | 1.27 | 0.1 | 1.95 | 0.1 | 1.95 | 0.1 | 20.31 | 4.79 | 20.31 | 4.69 | 14.36 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
87.872 | 12.79 | 0.1 | 5.57 | 0.1 | 1.46 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 18.16 | 9.77 | 23.83 | 9.67 | 14.75 | 9.67 | 9.86 | 1.07 | 0.1 | 1.76 | 0 | 2.05 | 0.1 | 21.09 | 4.79 | 21 | 4.69 | 15.14 | 4.69 | 9.77 | 9.86 | 9.86 | 9.77 | 9.77 | 9.86
87.904 | 13.28 | 0.1 | 6.05 | 0.1 | 2.64 | 0 | 5.57 | 4.79 | 4.79 | 8.79 | 4.79 | 18.46 | 9.77 | 24.12 | 9.67 | 15.23 | 9.67 | 9.86 | 1.17 | 0.1 | 1.56 | 0 | 2.15 | 0.1 | 21.19 | 4.79 | 21.49 | 4.69 | 15.82 | 4.59 | 9.77 | 9.86 | 9.86 | 9.77 | 9.77 | 9.86
87.936 | 13.87 | 0.1 | 6.74 | 0.1 | 3.81 | 0.1 | 6.35 | 4.79 | 4.88 | 2.83 | 4.88 | 18.46 | 9.77 | 21.19 | 9.77 | 15.82 | 9.67 | 9.96 | 0.88 | 0.1 | 1.56 | 0.1 | 1.86 | 0.1 | 21.29 | 4.79 | 22.17 | 4.79 | 16.5 | 4.59 | 10.94 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
87.968 | 14.36 | 0.1 | 7.32 | 0.1 | 4.59 | 0.1 | 7.52 | 4.79 | 4.88 | 3.22 | 4.79 | 18.26 | 9.77 | 20.9 | 9.77 | 16.31 | 9.67 | 9.86 | 0.98 | 0.1 | 1.37 | 0.1 | 1.56 | 0.1 | 21.29 | 4.88 | 22.66 | 4.79 | 17.19 | 4.59 | 11.82 | 9.77 | 9.86 | 9.77 | 10.74 | 9.77
88 | 14.84 | 0.1 | 7.91 | 0.1 | 5.18 | 0.1 | 8.69 | 4.69 | 4.88 | 4.2 | 4.79 | 18.26 | 9.77 | 21.19 | 9.77 | 16.8 | 9.57 | 9.96 | 0.78 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 21.29 | 4.79 | 23.24 | 4.79 | 17.97 | 4.59 | 13.67 | 9.77 | 10.74 | 9.77 | 11.43 | 9.77
88.032 | 15.43 | 0.1 | 8.5 | 0.1 | 5.66 | 0 | 9.57 | 4.59 | 4.79 | 5.37 | 4.79 | 18.16 | 9.77 | 21.39 | 9.77 | 17.38 | 9.57 | 9.86 | 0.68 | 0.1 | 1.17 | 0.1 | 1.46 | 0.1 | 21.39 | 4.79 | 23.73 | 4.79 | 18.56 | 4.59 | 15.43 | 9.67 | 11.62 | 9.67 | 13.18 | 9.67
88.064 | 15.82 | 0 | 9.08 | 0.1 | 6.25 | 0.1 | 10.25 | 4.59 | 4.88 | 5.86 | 4.79 | 18.16 | 9.77 | 21.39 | 9.77 | 17.77 | 9.57 | 9.86 | 0.68 | 0.1 | 1.07 | 0.1 | 1.17 | 0.1 | 21.29 | 4.79 | 24.22 | 4.79 | 19.14 | 4.49 | 16.5 | 9.67 | 13.09 | 9.67 | 15.04 | 9.67
88.096 | 16.41 | 0.1 | 9.67 | 0.1 | 6.93 | 0.1 | 10.94 | 4.59 | 4.88 | 5.66 | 4.79 | 18.16 | 9.86 | 21.39 | 9.77 | 18.26 | 9.57 | 9.96 | 0.59 | 0.1 | 1.07 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.71 | 4.79 | 19.82 | 4.59 | 17.38 | 9.67 | 15.04 | 9.67 | 16.31 | 9.67
88.128 | 16.9 | 0.1 | 10.35 | 0.1 | 7.52 | 0.1 | 11.52 | 4.59 | 4.88 | 5.27 | 4.79 | 18.16 | 9.86 | 21.39 | 9.77 | 18.65 | 9.67 | 9.86 | 0.49 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 25.2 | 4.79 | 20.31 | 4.49 | 18.16 | 9.67 | 16.31 | 9.67 | 17.19 | 9.67
88.16 | 17.38 | 0.1 | 10.94 | 0.1 | 8.11 | 0.1 | 12.01 | 4.59 | 4.88 | 4.98 | 4.79 | 18.16 | 9.86 | 21.39 | 9.86 | 19.24 | 9.67 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.78 | 4.79 | 20.9 | 4.59 | 18.85 | 9.67 | 17.29 | 9.67 | 17.97 | 9.67
88.192 | 15.04 | 0.1 | 11.52 | 0.1 | 8.69 | 0 | 12.7 | 4.49 | 4.88 | 4.88 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 19.53 | 9.67 | 9.86 | 0.39 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 21.19 | 4.88 | 26.27 | 4.79 | 21.49 | 4.59 | 19.53 | 9.67 | 17.97 | 9.67 | 18.65 | 9.67
88.224 | 15.14 | 0.1 | 12.01 | 0.1 | 9.38 | 0.1 | 13.09 | 4.49 | 4.88 | 4.88 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 19.92 | 9.67 | 9.86 | 0.29 | 0.1 | 0.68 | 0 | 0.68 | 0.1 | 21.29 | 4.79 | 23.05 | 4.88 | 22.07 | 4.59 | 20.12 | 9.67 | 18.75 | 9.57 | 19.24 | 9.67
88.256 | 14.84 | 0.1 | 12.6 | 0.1 | 9.86 | 0.1 | 13.67 | 4.49 | 4.88 | 4.88 | 4.79 | 18.07 | 9.86 | 21.39 | 9.86 | 20.41 | 9.67 | 10.64 | 0.29 | 0.1 | 0.68 | 0.1 | 0.68 | 0.1 | 5.86 | 4.88 | 22.17 | 4.88 | 22.75 | 4.59 | 20.8 | 9.67 | 19.43 | 9.57 | 19.92 | 9.67
88.288 | 14.65 | 0.1 | 13.18 | 0.1 | 10.45 | 0.1 | 14.26 | 4.59 | 4.88 | 4.79 | 4.79 | 8.89 | 9.86 | 21.39 | 9.86 | 20.9 | 9.67 | 11.33 | 0.2 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 2.93 | 4.88 | 21.78 | 4.88 | 23.24 | 4.59 | 21.39 | 9.67 | 20.12 | 9.67 | 20.61 | 9.57
88.32 | 14.65 | 0.1 | 13.67 | 0.1 | 11.13 | 0.1 | 14.75 | 4.59 | 4.88 | 4.79 | 4.79 | 2.93 | 9.77 | 21.39 | 9.86 | 21.29 | 9.67 | 12.5 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.42 | 4.88 | 21.58 | 4.88 | 23.73 | 4.69 | 21.97 | 9.67 | 20.61 | 9.67 | 21.09 | 9.57
88.352 | 14.45 | 0.1 | 14.36 | 0.1 | 11.72 | 0 | 15.23 | 4.59 | 4.88 | 4.69 | 4.79 | 3.22 | 9.77 | 21.39 | 9.77 | 21.58 | 9.67 | 13.57 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4 | 4.88 | 21.49 | 4.88 | 24.32 | 4.69 | 22.66 | 9.67 | 21.19 | 9.57 | 21.78 | 9.67
88.384 | 14.45 | 0 | 14.75 | 0.1 | 12.31 | 0.1 | 15.72 | 4.59 | 4.88 | 4.69 | 4.79 | 4.1 | 9.77 | 11.04 | 9.86 | 21.97 | 9.67 | 14.36 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 5.37 | 4.79 | 21.39 | 4.88 | 24.81 | 4.69 | 23.15 | 9.67 | 21.78 | 9.67 | 22.36 | 9.67
88.416 | 14.55 | 0.1 | 15.23 | 0.1 | 12.7 | 0.1 | 16.21 | 4.69 | 4.88 | 4.69 | 4.79 | 5.37 | 9.86 | 6.54 | 9.86 | 22.46 | 9.67 | 14.94 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 6.35 | 4.88 | 21.29 | 4.88 | 25.39 | 4.69 | 23.73 | 9.67 | 22.36 | 9.67 | 22.85 | 9.67
88.448 | 14.45 | 0.1 | 15.72 | 0.1 | 13.28 | 0.1 | 16.7 | 4.69 | 4.88 | 4.59 | 4.79 | 5.96 | 9.86 | 9.28 | 9.86 | 22.85 | 9.67 | 15.53 | 0.1 | 0.1 | 0.2 | 0.68 | 0.29 | 0.1 | 5.96 | 4.88 | 21.29 | 4.88 | 25.88 | 4.79 | 24.22 | 9.67 | 22.85 | 9.67 | 23.44 | 9.67
88.48 | 14.45 | 0.1 | 16.21 | 0.1 | 13.87 | 0.1 | 17.19 | 4.69 | 4.88 | 4.59 | 4.79 | 5.66 | 9.86 | 11.52 | 9.86 | 23.15 | 9.67 | 16.02 | 0.1 | 0.1 | 0.2 | 1.66 | 0.29 | 0.1 | 5.47 | 4.88 | 21.19 | 4.88 | 26.37 | 4.79 | 24.71 | 9.67 | 23.54 | 9.67 | 23.93 | 9.67
88.512 | 14.36 | 0.1 | 16.7 | 0.1 | 14.36 | 0 | 17.68 | 4.69 | 4.88 | 4.49 | 4.79 | 5.18 | 9.77 | 10.74 | 9.86 | 23.54 | 9.67 | 16.5 | 0.1 | 1.07 | 0.2 | 3.22 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.88 | 22.75 | 4.79 | 25.2 | 9.77 | 23.93 | 9.67 | 24.41 | 9.67
88.544 | 14.36 | 0.1 | 17.19 | 0.1 | 14.94 | 0 | 18.07 | 4.69 | 5.57 | 4.59 | 4.79 | 4.88 | 9.77 | 10.06 | 9.86 | 23.93 | 9.67 | 16.9 | 0.1 | 2.05 | 0.1 | 5.18 | 0.2 | 0.1 | 5.08 | 4.79 | 21.19 | 4.88 | 22.17 | 4.79 | 25.68 | 9.67 | 24.51 | 9.67 | 25 | 9.67
88.576 | 14.36 | 0.1 | 14.16 | 0.1 | 15.43 | 0.1 | 18.56 | 4.79 | 6.35 | 4.59 | 4.79 | 4.98 | 9.86 | 9.96 | 9.86 | 24.32 | 9.77 | 17.38 | 0.1 | 3.71 | 0.1 | 6.45 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.88 | 21.78 | 4.79 | 26.17 | 9.77 | 25 | 9.67 | 25.39 | 9.67
88.608 | 8.11 | 0.1 | 14.75 | 0.1 | 15.92 | 0.1 | 18.95 | 4.79 | 7.72 | 4.59 | 4.79 | 4.98 | 9.86 | 10.06 | 9.86 | 20.9 | 9.77 | 17.97 | 0.1 | 5.47 | 0.1 | 7.23 | 0.2 | 1.07 | 5.08 | 4.88 | 21.19 | 4.88 | 21.58 | 4.79 | 26.76 | 9.77 | 25.49 | 9.67 | 25.88 | 9.67
88.64 | 2.25 | 0.1 | 14.55 | 0.1 | 16.41 | 0.1 | 19.43 | 4.79 | 8.79 | 4.59 | 4.79 | 4.88 | 9.86 | 10.16 | 9.86 | 20.9 | 9.77 | 18.46 | 0.1 | 6.54 | 0.1 | 8.01 | 0.1 | 2.15 | 4.98 | 4.88 | 12.99 | 4.88 | 21.58 | 4.79 | 27.05 | 9.77 | 25.98 | 9.67 | 26.47 | 9.67
88.672 | 2.05 | 0.1 | 14.55 | 0.1 | 16.8 | 0.1 | 19.92 | 4.79 | 9.67 | 4.59 | 4.79 | 4.79 | 9.77 | 10.06 | 9.86 | 21.29 | 9.77 | 18.85 | 0 | 7.42 | 0.1 | 8.79 | 0.1 | 4.1 | 4.98 | 4.88 | 3.03 | 4.88 | 21.39 | 4.79 | 27.64 | 9.77 | 26.47 | 9.67 | 26.95 | 9.67
88.704 | 1.76 | 0 | 14.55 | 0.1 | 17.29 | 0.1 | 20.22 | 4.79 | 10.25 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 21.39 | 9.77 | 19.24 | 0 | 8.3 | 0 | 9.67 | 0.1 | 5.57 | 4.88 | 4.79 | 3.22 | 4.88 | 21.29 | 4.79 | 28.13 | 9.77 | 27.05 | 9.67 | 27.44 | 9.67
88.736 | 1.46 | 0.1 | 14.45 | 0.1 | 15.43 | 0.1 | 20.7 | 4.79 | 10.94 | 4.59 | 4.79 | 4.79 | 9.86 | 9.86 | 9.86 | 21.39 | 9.77 | 19.73 | 0.1 | 8.98 | 0.1 | 10.45 | 0.1 | 6.74 | 4.79 | 4.88 | 3.52 | 4.88 | 21.29 | 4.88 | 28.61 | 9.77 | 27.54 | 9.77 | 28.03 | 9.77
88.768 | 1.37 | 0.1 | 14.55 | 0.1 | 15.23 | 0.1 | 21.09 | 4.88 | 11.52 | 4.59 | 4.79 | 4.69 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.12 | 0.1 | 9.77 | 0.1 | 11.23 | 0.1 | 7.52 | 4.69 | 4.88 | 4.59 | 4.88 | 21.19 | 4.88 | 29.1 | 9.77 | 28.03 | 9.77 | 28.32 | 9.77
88.8 | 1.27 | 0.1 | 14.45 | 0.1 | 14.84 | 0.1 | 19.24 | 4.88 | 12.11 | 4.69 | 4.79 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.61 | 0.1 | 10.55 | 0.1 | 12.01 | 0.1 | 8.4 | 4.69 | 4.88 | 5.96 | 4.88 | 21.29 | 4.79 | 29.49 | 9.77 | 28.42 | 9.77 | 28.91 | 9.77
88.832 | 1.07 | 0.1 | 14.36 | 0.1 | 14.75 | 0 | 18.75 | 4.79 | 12.6 | 4.69 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21 | 0 | 11.33 | 0 | 12.79 | 0 | 9.18 | 4.69 | 4.88 | 5.96 | 4.88 | 21.19 | 4.79 | 25.29 | 9.77 | 28.91 | 9.77 | 29.4 | 9.77
88.864 | 0.98 | 0.1 | 14.45 | 0.1 | 14.75 | 0.1 | 18.56 | 4.79 | 13.18 | 4.69 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.39 | 0 | 12.11 | 0 | 13.48 | 0 | 9.96 | 4.69 | 4.79 | 5.66 | 4.88 | 21.19 | 4.79 | 25.1 | 9.77 | 29.4 | 9.77 | 25.29 | 9.77
88.896 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.77 | 4.69 | 4.79 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 12.89 | 0.1 | 14.16 | 0.1 | 10.84 | 4.59 | 4.88 | 5.37 | 4.88 | 21.19 | 4.88 | 24.71 | 9.77 | 25.29 | 9.77 | 24.32 | 9.77
88.928 | 0.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 14.26 | 4.69 | 5.47 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.86 | 22.17 | 0.1 | 13.67 | 0.1 | 14.94 | 0.1 | 11.72 | 4.59 | 4.88 | 5.18 | 4.88 | 10.45 | 4.88 | 24.61 | 9.77 | 24.71 | 9.77 | 24.41 | 9.77
88.96 | 0.78 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.26 | 4.88 | 14.84 | 4.69 | 6.25 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 22.56 | 0 | 14.45 | 0.1 | 15.72 | 0.1 | 12.31 | 4.59 | 4.88 | 5.08 | 4.88 | 3.03 | 4.79 | 24.61 | 9.77 | 24.51 | 9.77 | 24.32 | 9.86
88.992 | 0.68 | 0.1 | 11.52 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.33 | 4.69 | 7.42 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 0 | 15.14 | 0 | 16.41 | 0 | 13.18 | 4.49 | 4.88 | 4.98 | 4.88 | 3.22 | 4.79 | 24.41 | 9.77 | 24.51 | 9.77 | 24.41 | 9.77
89.024 | 0.59 | 0.1 | 3.71 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.72 | 4.69 | 8.5 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 23.34 | 0.1 | 15.82 | 0 | 17.09 | 0 | 13.77 | 4.59 | 5.86 | 5.08 | 4.88 | 3.52 | 4.79 | 24.32 | 9.77 | 24.51 | 9.77 | 24.32 | 9.77
89.056 | 0.49 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.31 | 4.79 | 9.47 | 4.59 | 9.86 | 9.67 | 9.86 | 9.96 | 9.86 | 23.73 | 0.1 | 16.6 | 0.1 | 17.68 | 0.1 | 14.55 | 4.59 | 7.32 | 4.98 | 4.88 | 4.79 | 4.88 | 24.41 | 9.77 | 24.41 | 9.86 | 24.22 | 9.86
89.088 | 0.49 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.8 | 4.79 | 10.16 | 4.59 | 9.96 | 9.67 | 9.86 | 6.74 | 9.86 | 24.02 | 0.1 | 17.19 | 0.1 | 18.46 | 0.1 | 15.33 | 4.59 | 9.38 | 4.98 | 4.88 | 6.05 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.32 | 9.86
89.12 | 0.39 | 0.1 | 1.46 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 17.19 | 4.79 | 10.64 | 4.59 | 10.64 | 9.67 | 9.86 | 9.77 | 9.86 | 24.51 | 0.1 | 17.87 | 0.1 | 18.95 | 0.1 | 16.02 | 4.49 | 10.94 | 4.88 | 4.88 | 6.15 | 4.88 | 24.32 | 9.77 | 24.41 | 9.86 | 24.22 | 9.86
89.152 | 0.29 | 0.1 | 1.56 | 0.1 | 9.18 | 0 | 18.07 | 4.88 | 17.68 | 4.79 | 11.33 | 4.59 | 11.91 | 9.57 | 9.86 | 11.62 | 9.86 | 22.36 | 0 | 18.56 | 0 | 19.73 | 0 | 16.7 | 4.49 | 12.31 | 4.79 | 4.88 | 5.57 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
89.184 | 0.2 | 0 | 1.17 | 0.1 | 2.54 | 0.1 | 18.07 | 4.88 | 18.07 | 4.79 | 11.91 | 4.59 | 12.99 | 9.67 | 9.86 | 10.94 | 9.86 | 22.27 | 0 | 19.24 | 0 | 20.22 | 0 | 17.29 | 4.59 | 13.09 | 4.69 | 4.88 | 5.27 | 4.88 | 24.22 | 9.77 | 24.32 | 9.77 | 24.12 | 9.86
89.216 | 0.2 | 0.1 | 1.17 | 0.1 | 1.95 | 0.1 | 10.06 | 4.88 | 18.56 | 4.79 | 12.5 | 4.69 | 13.87 | 9.67 | 9.86 | 10.16 | 9.86 | 21.88 | 0.1 | 19.92 | 0.1 | 20.8 | 0.1 | 18.07 | 4.59 | 13.77 | 4.69 | 4.88 | 5.18 | 4.88 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
89.248 | 0.2 | 0.1 | 0.98 | 0.1 | 1.86 | 0.1 | 3.13 | 4.88 | 19.04 | 4.88 | 12.99 | 4.69 | 14.65 | 9.67 | 9.86 | 9.96 | 9.86 | 21.78 | 0 | 20.41 | 0 | 21.39 | 0.1 | 18.65 | 4.59 | 14.65 | 4.69 | 4.88 | 5.18 | 4.88 | 13.57 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
89.28 | 0.1 | 0.1 | 0.88 | 0.1 | 1.37 | 0.1 | 3.13 | 4.88 | 19.43 | 4.79 | 13.57 | 4.69 | 15.14 | 9.67 | 9.86 | 10.16 | 9.86 | 21.78 | 0.1 | 20.9 | 0.1 | 21.97 | 0 | 19.34 | 4.59 | 15.33 | 4.59 | 4.88 | 5.08 | 4.88 | 6.05 | 9.77 | 24.22 | 9.86 | 20.7 | 9.86
89.312 | 0.1 | 0.1 | 0.78 | 0.1 | 1.56 | 0 | 3.81 | 4.88 | 19.92 | 4.79 | 14.16 | 4.69 | 15.72 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 0 | 21.49 | 0 | 22.56 | 0 | 19.92 | 4.59 | 16.11 | 4.59 | 4.88 | 4.98 | 4.88 | 7.72 | 9.77 | 16.9 | 9.77 | 7.42 | 9.86
89.344 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0 | 4.98 | 4.88 | 20.31 | 4.88 | 14.55 | 4.69 | 16.21 | 9.67 | 9.96 | 10.06 | 9.86 | 21.49 | 0 | 22.17 | 0 | 18.65 | 0 | 20.51 | 4.59 | 16.6 | 4.49 | 4.88 | 4.98 | 4.88 | 11.62 | 9.77 | 7.32 | 9.77 | 6.54 | 9.86
89.376 | 0.1 | 0.68 | 0.68 | 0.1 | 1.07 | 0.1 | 5.86 | 4.88 | 20.7 | 4.88 | 15.14 | 4.79 | 16.7 | 9.67 | 10.74 | 10.06 | 9.86 | 21.58 | 0.1 | 22.66 | 0.1 | 17.58 | 0.1 | 21.19 | 4.69 | 17.38 | 4.59 | 4.88 | 4.98 | 4.88 | 11.72 | 9.86 | 6.74 | 9.86 | 10.45 | 9.86
89.408 | 0.1 | 1.37 | 0.59 | 0.1 | 1.17 | 0.1 | 5.76 | 4.88 | 20.8 | 4.88 | 15.63 | 4.79 | 17.19 | 9.67 | 11.72 | 9.96 | 9.86 | 21.49 | 0.1 | 19.34 | 0.1 | 17.87 | 0.1 | 21.68 | 4.69 | 18.07 | 4.49 | 4.88 | 4.88 | 4.88 | 10.35 | 9.77 | 10.55 | 9.86 | 12.21 | 9.86
89.44 | 0 | 2.64 | 0.49 | 0.1 | 0.88 | 0.1 | 5.18 | 4.88 | 18.65 | 4.88 | 16.02 | 4.79 | 17.68 | 9.67 | 13.09 | 9.96 | 9.86 | 21.49 | 0.1 | 18.65 | 0.1 | 17.77 | 0.1 | 22.27 | 4.69 | 18.65 | 4.49 | 4.88 | 4.79 | 4.88 | 10.06 | 9.77 | 12.11 | 9.86 | 10.94 | 9.86
89.472 | 0 | 3.81 | 0.39 | 0.1 | 0.78 | 0 | 4.98 | 4.88 | 18.26 | 4.88 | 16.6 | 4.79 | 18.07 | 9.67 | 14.06 | 9.86 | 9.86 | 21.39 | 0 | 18.07 | 0 | 17.87 | 0 | 22.17 | 4.69 | 19.34 | 4.49 | 5.86 | 4.69 | 4.79 | 10.25 | 9.86 | 10.74 | 9.77 | 9.96 | 9.86
89.504 | 0 | 4.49 | 0.39 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.46 | 4.79 | 16.99 | 4.79 | 18.65 | 9.67 | 14.65 | 9.77 | 9.86 | 21.39 | 0 | 17.97 | 0 | 17.77 | 0 | 19.24 | 4.79 | 20.02 | 4.49 | 6.64 | 4.69 | 4.88 | 10.45 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86
89.536 | 0.1 | 5.08 | 0.29 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.36 | 4.88 | 17.58 | 4.88 | 19.14 | 9.67 | 15.23 | 9.77 | 9.86 | 21.39 | 0.1 | 17.97 | 0.1 | 17.77 | 0 | 18.56 | 4.79 | 20.61 | 4.59 | 8.2 | 4.69 | 4.88 | 10.45 | 9.86 | 10.16 | 9.86 | 10.35 | 9.86
89.568 | 0.1 | 5.66 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.26 | 4.88 | 17.97 | 4.88 | 19.43 | 9.77 | 15.72 | 9.77 | 9.86 | 14.84 | 0.1 | 17.87 | 0.1 | 17.77 | 0.1 | 18.26 | 4.79 | 21 | 4.59 | 10.25 | 4.69 | 4.79 | 10.16 | 9.86 | 10.35 | 9.86 | 10.45 | 9.86
89.6 | 0.1 | 6.25 | 0.2 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.46 | 4.88 | 19.92 | 9.77 | 16.21 | 9.77 | 9.86 | 7.32 | 0.1 | 17.77 | 0.1 | 17.68 | 0 | 18.07 | 4.79 | 21.78 | 4.59 | 11.62 | 4.59 | 4.88 | 10.06 | 9.77 | 10.45 | 9.86 | 10.25 | 9.86
89.632 | 0 | 6.84 | 0.2 | 0.1 | 0.49 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 18.85 | 4.79 | 20.22 | 9.77 | 16.7 | 9.67 | 9.86 | 7.81 | 0 | 17.68 | 0 | 17.68 | 0 | 17.97 | 4.79 | 22.27 | 4.49 | 12.5 | 4.59 | 4.79 | 9.96 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86
89.664 | 0 | 7.42 | 0.1 | 0.1 | 0.49 | 0.1 | 4.69 | 4.88 | 18.16 | 4.79 | 19.24 | 4.79 | 20.7 | 9.77 | 17.19 | 9.67 | 9.86 | 10.74 | 0 | 17.68 | 0 | 17.68 | 0 | 17.87 | 4.79 | 22.85 | 4.59 | 13.48 | 4.59 | 4.88 | 10.06 | 9.77 | 10.06 | 9.77 | 10.06 | 9.86
89.696 | 0.1 | 8.11 | 0.1 | 0.1 | 0.39 | 0.1 | 4.59 | 4.88 | 18.16 | 4.88 | 19.73 | 4.88 | 21.19 | 9.77 | 17.58 | 9.67 | 9.86 | 11.43 | 0 | 17.68 | 0.1 | 17.68 | 0 | 17.87 | 4.88 | 23.44 | 4.59 | 14.16 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
89.728 | 0.1 | 8.69 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 21.58 | 9.77 | 18.16 | 9.67 | 9.86 | 10.64 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 17.77 | 4.88 | 24.02 | 4.59 | 14.94 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
89.76 | 0.1 | 9.28 | 0 | 0.2 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.61 | 4.88 | 21.88 | 9.77 | 18.56 | 9.67 | 9.86 | 9.96 | 0 | 17.58 | 0.1 | 15.43 | 0 | 17.77 | 4.88 | 24.51 | 4.69 | 15.72 | 4.49 | 4.79 | 9.86 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
89.792 | 0 | 9.86 | 0 | 1.17 | 0.2 | 0.1 | 4.49 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.36 | 9.77 | 18.95 | 9.67 | 9.86 | 9.86 | 0 | 17.58 | 0 | 3.52 | 0 | 17.77 | 4.79 | 25 | 4.69 | 16.41 | 4.49 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
89.824 | 0.1 | 10.45 | 0 | 1.95 | 0.2 | 1.07 | 4.49 | 4.79 | 18.07 | 4.79 | 18.95 | 4.88 | 22.66 | 9.77 | 19.43 | 9.67 | 9.86 | 10.16 | 0 | 9.67 | 0 | 2.15 | 0 | 17.68 | 4.88 | 25.49 | 4.69 | 17.09 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
89.856 | 0.1 | 11.04 | 0.1 | 3.42 | 0.1 | 1.95 | 4.49 | 4.88 | 8.89 | 4.88 | 18.95 | 4.88 | 23.15 | 9.86 | 19.92 | 9.67 | 9.86 | 10.16 | 0 | 2.44 | 0.1 | 2.25 | 0.1 | 17.77 | 4.88 | 26.08 | 4.69 | 17.77 | 4.49 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77 | 9.86
89.888 | 0.1 | 11.62 | 0.1 | 4.3 | 0.1 | 3.22 | 4.49 | 4.88 | 3.03 | 4.88 | 18.56 | 4.88 | 23.54 | 9.86 | 20.31 | 9.67 | 9.96 | 9.96 | 0 | 2.34 | 0.1 | 1.95 | 0.1 | 17.77 | 4.88 | 23.34 | 4.69 | 18.46 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
89.92 | 0.1 | 12.11 | 0 | 4.88 | 0 | 4.1 | 4.49 | 4.88 | 3.22 | 4.88 | 18.26 | 4.88 | 23.93 | 9.86 | 20.7 | 9.67 | 10.64 | 9.86 | 0.1 | 2.05 | 0.1 | 1.66 | 0.1 | 8.98 | 4.88 | 20.61 | 4.79 | 19.04 | 4.59 | 5.18 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
89.952 | 0 | 12.79 | 0.1 | 5.47 | 0 | 4.88 | 4.39 | 4.88 | 4.1 | 4.88 | 18.36 | 4.88 | 24.32 | 9.77 | 21.09 | 9.67 | 11.23 | 9.86 | 0 | 1.86 | 0 | 1.46 | 0 | 1.95 | 4.79 | 20.9 | 4.69 | 19.73 | 4.59 | 6.35 | 9.67 | 10.06 | 9.77 | 9.77 | 9.77 | 9.86
89.984 | 0 | 13.18 | 0 | 5.96 | 0.1 | 5.27 | 4.49 | 4.79 | 5.18 | 4.88 | 18.16 | 4.88 | 20.8 | 9.77 | 21.58 | 9.67 | 12.5 | 9.86 | 0 | 1.76 | 0.1 | 1.27 | 0 | 2.44 | 4.88 | 21.29 | 4.79 | 20.31 | 4.59 | 7.81 | 9.67 | 10.94 | 9.77 | 10.45 | 9.67 | 9.86
90.016 | 0.1 | 13.87 | 0.1 | 6.54 | 0.1 | 5.86 | 4.49 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 20.9 | 9.86 | 21.97 | 9.67 | 13.57 | 9.77 | 0.1 | 1.56 | 0.1 | 1.46 | 0.1 | 2.05 | 4.88 | 21.29 | 4.79 | 20.9 | 4.69 | 9.96 | 9.67 | 12.4 | 9.77 | 11.23 | 9.67 | 10.74
90.048 | 0.1 | 14.36 | 0.1 | 7.23 | 0.1 | 6.54 | 4.59 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 21.09 | 9.86 | 22.27 | 9.67 | 14.36 | 9.77 | 0 | 1.56 | 0.1 | 1.17 | 0.1 | 1.86 | 4.88 | 21.19 | 4.79 | 21.39 | 4.69 | 11.43 | 9.67 | 14.45 | 9.67 | 12.79 | 9.67 | 11.72
90.08 | 0.1 | 14.84 | 0 | 7.81 | 0.1 | 7.13 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 22.66 | 9.67 | 14.94 | 9.77 | 0 | 1.37 | 0.1 | 1.17 | 0.1 | 1.86 | 4.88 | 21.29 | 4.79 | 22.07 | 4.69 | 12.4 | 9.67 | 15.92 | 9.67 | 14.84 | 9.67 | 13.38
90.112 | 0 | 15.33 | 0 | 8.4 | 0.1 | 7.72 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 23.05 | 9.67 | 15.53 | 9.67 | 0 | 1.37 | 0 | 1.07 | 0 | 1.56 | 4.88 | 21.29 | 4.79 | 22.66 | 4.69 | 13.38 | 9.67 | 16.99 | 9.67 | 16.21 | 9.67 | 15.23
90.144 | 0 | 15.92 | 0 | 8.98 | 0.1 | 8.3 | 4.59 | 4.79 | 4.88 | 4.88 | 17.97 | 4.88 | 21.39 | 9.77 | 23.44 | 9.67 | 16.02 | 9.67 | 0 | 1.07 | 0 | 0.88 | 0 | 1.46 | 4.88 | 21.19 | 4.79 | 23.15 | 4.69 | 14.06 | 9.57 | 17.68 | 9.67 | 17.09 | 9.67 | 16.41
90.176 | 0.1 | 16.31 | 0.1 | 9.67 | 0.1 | 8.89 | 4.69 | 4.88 | 4.98 | 4.88 | 17.97 | 4.88 | 21.49 | 9.86 | 23.83 | 9.77 | 16.5 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 23.73 | 4.79 | 14.84 | 9.67 | 18.46 | 9.67 | 17.87 | 9.67 | 17.29
90.208 | 0.1 | 16.8 | 0.1 | 10.25 | 0.1 | 9.57 | 4.69 | 5.57 | 4.88 | 4.88 | 17.97 | 4.88 | 21.39 | 9.86 | 24.22 | 9.77 | 16.99 | 9.67 | 0 | 0.98 | 0.1 | 0.78 | 0.1 | 1.17 | 4.88 | 21.19 | 4.79 | 24.22 | 4.79 | 15.63 | 9.67 | 19.14 | 9.67 | 18.65 | 9.57 | 18.07
90.24 | 0.1 | 17.29 | 0 | 10.74 | 0.1 | 10.16 | 4.69 | 6.64 | 4.88 | 4.88 | 11.13 | 4.88 | 21.39 | 9.86 | 20.8 | 9.77 | 17.48 | 9.67 | 0.1 | 0.88 | 0 | 0.68 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.71 | 4.79 | 16.31 | 9.67 | 19.73 | 9.67 | 19.24 | 9.57 | 18.75
90.272 | 0 | 15.43 | 0 | 11.33 | 0.1 | 10.64 | 4.69 | 8.01 | 4.69 | 4.88 | 3.22 | 4.88 | 21.39 | 9.77 | 20.9 | 9.77 | 17.97 | 9.57 | 1.07 | 0.78 | 0 | 0.59 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.2 | 4.79 | 16.9 | 9.67 | 20.51 | 9.67 | 20.02 | 9.57 | 19.34
90.304 | 0 | 15.33 | 0 | 11.91 | 0.1 | 11.23 | 4.69 | 8.98 | 4.69 | 4.79 | 3.13 | 4.88 | 21.39 | 9.77 | 21.09 | 9.77 | 18.46 | 9.57 | 2.15 | 0.68 | 0 | 0.59 | 0 | 0.98 | 4.88 | 21.19 | 4.79 | 25.78 | 4.79 | 17.68 | 9.67 | 20.9 | 9.67 | 20.61 | 9.57 | 20.12
90.336 | 0.1 | 14.84 | 0.1 | 12.5 | 0.1 | 11.82 | 4.79 | 9.77 | 4.69 | 4.88 | 3.71 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 18.85 | 9.67 | 4.1 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 7.42 | 4.79 | 26.27 | 4.79 | 18.26 | 9.67 | 21.68 | 9.67 | 21.09 | 9.67 | 20.61
90.368 | 0.1 | 14.65 | 0 | 13.09 | 0.1 | 12.4 | 4.79 | 10.45 | 4.59 | 4.88 | 4.98 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 19.24 | 9.57 | 5.57 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 2.25 | 4.79 | 22.46 | 4.79 | 18.85 | 9.67 | 22.27 | 9.67 | 21.68 | 9.67 | 21.19
90.4 | 0.1 | 14.75 | 0.1 | 13.57 | 0.1 | 12.99 | 4.79 | 11.04 | 4.59 | 4.88 | 5.86 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 19.73 | 9.57 | 6.64 | 0.59 | 0.1 | 0.39 | 0 | 0.68 | 4.88 | 3.32 | 4.79 | 22.07 | 4.79 | 19.43 | 9.67 | 22.75 | 9.67 | 22.27 | 9.67 | 21.78
90.432 | 0.1 | 14.55 | 0 | 14.06 | 0.1 | 13.57 | 4.79 | 11.72 | 4.49 | 4.88 | 5.66 | 4.88 | 9.77 | 9.86 | 21.39 | 9.77 | 20.12 | 9.57 | 7.42 | 0.49 | 0.88 | 0.29 | 0 | 0.68 | 4.88 | 4.1 | 4.79 | 21.58 | 4.79 | 20.22 | 9.67 | 23.34 | 9.67 | 22.85 | 9.67 | 22.46
90.464 | 0 | 14.55 | 0 | 14.55 | 0.1 | 13.97 | 4.79 | 12.21 | 4.49 | 4.88 | 5.27 | 4.88 | 6.74 | 9.86 | 21.39 | 9.77 | 20.61 | 9.57 | 8.2 | 0.49 | 1.66 | 0.29 | 0.2 | 0.59 | 4.88 | 5.08 | 4.79 | 21.49 | 4.79 | 20.8 | 9.67 | 23.83 | 9.67 | 23.34 | 9.67 | 22.95
90.496 | 0.1 | 14.55 | 0 | 15.14 | 0.1 | 14.55 | 4.79 | 12.79 | 4.49 | 4.88 | 4.98 | 4.88 | 9.67 | 9.86 | 21.39 | 9.86 | 20.9 | 9.67 | 8.89 | 0.39 | 3.32 | 0.29 | 1.56 | 0.59 | 4.88 | 6.25 | 4.79 | 21.49 | 4.79 | 21.39 | 9.67 | 24.41 | 9.67 | 23.93 | 9.67 | 23.34
90.528 | 0.1 | 14.45 | 0 | 15.63 | 0.1 | 15.04 | 4.79 | 13.28 | 4.49 | 4.88 | 4.88 | 4.88 | 11.62 | 9.86 | 21.39 | 9.86 | 21.39 | 9.67 | 9.77 | 0.29 | 5.27 | 0.2 | 2.93 | 0.49 | 4.88 | 6.15 | 4.79 | 21.39 | 4.79 | 21.97 | 9.77 | 24.9 | 9.67 | 24.51 | 9.67 | 23.93
90.56 | 0.1 | 14.45 | 0 | 16.11 | 0 | 15.53 | 4.79 | 13.87 | 4.59 | 4.88 | 4.88 | 4.88 | 10.74 | 9.86 | 21.39 | 9.86 | 21.78 | 9.67 | 10.55 | 0.29 | 6.45 | 0.2 | 4.88 | 0.49 | 4.88 | 5.47 | 4.88 | 21.39 | 4.88 | 22.46 | 9.77 | 25.39 | 9.67 | 25 | 9.67 | 24.51
90.592 | 0 | 14.45 | 0 | 16.6 | 0.1 | 16.02 | 4.79 | 14.45 | 4.49 | 5.66 | 4.88 | 4.88 | 10.06 | 9.77 | 21.39 | 9.77 | 22.17 | 9.67 | 11.52 | 0.2 | 7.32 | 0.1 | 6.25 | 0.29 | 4.88 | 5.18 | 4.79 | 21.39 | 4.79 | 23.05 | 9.67 | 25.88 | 9.67 | 25.39 | 9.67 | 25.1
90.624 | 0 | 14.36 | 0 | 17.09 | 0.1 | 16.5 | 4.79 | 14.84 | 4.49 | 6.25 | 4.79 | 4.88 | 9.96 | 9.77 | 21.39 | 9.77 | 22.56 | 9.67 | 12.21 | 0.2 | 8.01 | 0.1 | 7.03 | 0.29 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 23.63 | 9.77 | 26.37 | 9.67 | 25.88 | 9.67 | 25.59
90.656 | 0.1 | 14.45 | 0 | 14.16 | 0.1 | 16.99 | 4.88 | 15.43 | 4.59 | 7.52 | 4.79 | 4.88 | 10.16 | 9.86 | 21.39 | 9.86 | 23.05 | 9.67 | 12.89 | 0.2 | 8.89 | 0.2 | 7.81 | 0.29 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 24.12 | 9.77 | 26.86 | 9.77 | 26.37 | 9.77 | 26.08
90.688 | 0.1 | 10.25 | 0 | 13.97 | 0.1 | 15.04 | 4.88 | 15.92 | 4.59 | 8.59 | 4.69 | 4.88 | 10.16 | 9.77 | 10.55 | 9.86 | 23.34 | 9.67 | 13.67 | 0.2 | 9.67 | 0.1 | 8.69 | 0.29 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 24.61 | 9.77 | 27.34 | 9.77 | 26.86 | 9.77 | 26.56
90.72 | 0.1 | 2.93 | 0 | 14.26 | 0.1 | 14.06 | 4.88 | 16.41 | 4.59 | 9.47 | 4.69 | 4.88 | 10.06 | 9.77 | 6.64 | 9.86 | 23.63 | 9.67 | 14.36 | 0.2 | 10.45 | 0.1 | 9.38 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.2 | 9.77 | 27.83 | 9.77 | 27.34 | 9.77 | 27.05
90.752 | 0 | 1.95 | 0 | 14.45 | 0.1 | 14.16 | 4.79 | 16.9 | 4.59 | 10.25 | 4.69 | 4.88 | 9.86 | 9.77 | 9.67 | 9.86 | 24.02 | 9.67 | 15.04 | 0.1 | 11.33 | 0.1 | 10.25 | 0.2 | 4.88 | 4.88 | 4.79 | 21.19 | 4.79 | 25.59 | 9.77 | 28.32 | 9.77 | 27.93 | 9.77 | 27.54
90.784 | 0 | 1.76 | 0 | 14.36 | 0.1 | 14.45 | 4.79 | 17.29 | 4.59 | 10.74 | 4.59 | 4.88 | 9.96 | 9.77 | 11.52 | 9.86 | 24.41 | 9.67 | 15.72 | 0.1 | 12.01 | 0.1 | 11.04 | 0.1 | 4.88 | 4.88 | 4.79 | 16.21 | 4.79 | 26.08 | 9.77 | 28.81 | 9.77 | 28.42 | 9.77 | 28.03
90.816 | 0.1 | 1.56 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 17.87 | 4.69 | 11.43 | 4.59 | 4.88 | 9.86 | 9.77 | 10.94 | 9.86 | 22.46 | 9.67 | 16.41 | 0.1 | 12.79 | 0.1 | 11.91 | 0.2 | 4.88 | 4.88 | 4.88 | 3.81 | 4.79 | 21.68 | 9.77 | 29.4 | 9.77 | 28.91 | 9.77 | 28.52
90.848 | 0.1 | 1.46 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 18.26 | 4.69 | 12.01 | 4.49 | 4.88 | 9.77 | 9.86 | 10.16 | 9.86 | 22.27 | 9.67 | 17.09 | 0.1 | 13.48 | 0.1 | 12.6 | 0.1 | 4.88 | 4.79 | 4.79 | 3.22 | 4.88 | 20.7 | 9.77 | 24.71 | 9.77 | 29.4 | 9.77 | 29.01
90.88 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.65 | 4.69 | 12.6 | 4.49 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 21.88 | 9.77 | 17.77 | 0.1 | 14.26 | 0.1 | 13.38 | 0.1 | 4.88 | 4.69 | 4.79 | 3.32 | 4.88 | 21.19 | 9.77 | 23.34 | 9.77 | 24.61 | 9.77 | 29.4
90.912 | 0 | 1.17 | 0 | 14.45 | 0.1 | 14.36 | 4.79 | 19.24 | 4.69 | 13.18 | 4.49 | 4.88 | 9.77 | 9.77 | 10.06 | 9.86 | 21.68 | 9.67 | 18.46 | 0.1 | 15.04 | 0.1 | 14.16 | 0.1 | 5.08 | 4.69 | 4.79 | 4.3 | 4.79 | 21.29 | 9.77 | 24.02 | 9.77 | 23.83 | 9.77 | 26.17
90.944 | 0 | 1.07 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.69 | 13.67 | 4.39 | 4.88 | 9.67 | 9.77 | 10.16 | 9.86 | 21.68 | 9.67 | 19.14 | 0.1 | 15.72 | 0.1 | 14.75 | 0.1 | 6.05 | 4.59 | 4.79 | 5.86 | 4.79 | 21.29 | 9.77 | 24.12 | 9.77 | 24.12 | 9.77 | 25
90.976 | 0.1 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 20.02 | 4.79 | 14.16 | 4.49 | 5.66 | 9.67 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 19.73 | 0.1 | 16.5 | 0.1 | 15.53 | 0.1 | 7.52 | 4.59 | 4.79 | 6.25 | 4.88 | 21.29 | 9.77 | 24.22 | 9.86 | 24.12 | 9.77 | 24.61
91.008 | 0.1 | 0.88 | 0 | 14.45 | 0.1 | 14.36 | 4.88 | 20.41 | 4.79 | 14.65 | 4.49 | 6.35 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.22 | 0.1 | 17.09 | 0.1 | 16.21 | 0.1 | 9.67 | 4.59 | 4.88 | 5.76 | 4.79 | 21.29 | 9.77 | 24.22 | 9.86 | 24.22 | 9.77 | 24.51
91.04 | 0.1 | 0.78 | 0 | 14.45 | 0.1 | 14.36 | 4.88 | 20.9 | 4.79 | 15.23 | 4.49 | 7.62 | 9.67 | 9.86 | 9.86 | 9.96 | 21.49 | 9.77 | 20.9 | 0.1 | 17.77 | 0.1 | 16.9 | 0.1 | 11.23 | 4.59 | 4.79 | 5.37 | 4.79 | 21.29 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.51
91.072 | 0 | 0.68 | 0 | 14.45 | 0.1 | 14.36 | 4.88 | 17.77 | 4.79 | 15.72 | 4.49 | 8.79 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 9.77 | 21.49 | 0.1 | 18.46 | 0 | 17.68 | 0.1 | 12.5 | 4.59 | 4.79 | 5.27 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77 | 24.41
91.104 | 0 | 0.59 | 0 | 5.66 | 0.1 | 14.36 | 4.88 | 17.77 | 4.79 | 16.21 | 4.49 | 9.57 | 9.67 | 11.52 | 9.77 | 9.86 | 21.39 | 9.77 | 21.97 | 0.1 | 19.04 | 0 | 18.16 | 0.1 | 13.18 | 4.49 | 4.79 | 5.27 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22
91.136 | 0.1 | 0.59 | 0 | 1.66 | 0.1 | 6.15 | 4.88 | 18.07 | 4.79 | 16.7 | 4.59 | 10.25 | 9.67 | 12.89 | 9.77 | 9.96 | 21.49 | 9.77 | 21.68 | 0.1 | 19.53 | 0.1 | 18.85 | 0.1 | 13.97 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77 | 24.32
91.168 | 0.1 | 0.49 | 0 | 2.05 | 0.1 | 1.76 | 4.88 | 18.26 | 4.79 | 17.09 | 4.59 | 10.84 | 9.57 | 13.77 | 9.77 | 9.96 | 21.39 | 9.77 | 18.07 | 0.1 | 20.22 | 0.1 | 19.43 | 0.1 | 14.75 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
91.2 | 0.1 | 0.39 | 0 | 1.76 | 0.1 | 1.95 | 4.88 | 18.16 | 4.79 | 17.58 | 4.59 | 11.43 | 9.57 | 14.55 | 9.77 | 9.96 | 21.39 | 9.77 | 17.87 | 0.1 | 20.9 | 0.1 | 20.12 | 0.1 | 15.53 | 4.59 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
91.232 | 0 | 0.29 | 0 | 1.46 | 0.1 | 1.66 | 4.88 | 18.07 | 4.79 | 18.07 | 4.69 | 12.11 | 9.57 | 15.04 | 9.67 | 9.86 | 13.38 | 9.77 | 17.58 | 0.1 | 21.39 | 0.1 | 20.9 | 0.1 | 16.21 | 4.59 | 4.79 | 4.98 | 4.79 | 17.97 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.12
91.264 | 0 | 0.29 | 0 | 1.37 | 0.1 | 1.37 | 4.88 | 18.16 | 4.79 | 18.46 | 4.69 | 12.6 | 9.57 | 15.53 | 9.67 | 9.86 | 7.32 | 9.77 | 17.68 | 0.1 | 21.97 | 0.1 | 21.39 | 0.1 | 16.9 | 4.49 | 4.88 | 4.88 | 4.79 | 4.39 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.12
91.296 | 0.78 | 0.29 | 0.1 | 1.27 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.95 | 4.69 | 13.18 | 9.67 | 16.11 | 9.67 | 9.96 | 8.01 | 9.77 | 17.58 | 0.1 | 22.56 | 0.1 | 21.88 | 0.1 | 17.58 | 4.59 | 4.88 | 4.88 | 4.88 | 3.03 | 9.77 | 10.45 | 9.86 | 23.93 | 9.86 | 24.12
91.328 | 1.56 | 0.2 | 0 | 1.17 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.34 | 4.69 | 13.67 | 9.67 | 16.6 | 9.67 | 9.96 | 10.84 | 9.77 | 17.58 | 0.1 | 19.14 | 0.1 | 22.46 | 0.1 | 18.26 | 4.59 | 5.66 | 4.79 | 4.79 | 3.32 | 9.77 | 5.86 | 9.86 | 8.79 | 9.86 | 17.97
91.36 | 2.73 | 0.2 | 0 | 1.07 | 0 | 1.17 | 4.88 | 18.16 | 4.88 | 19.82 | 4.69 | 14.16 | 9.67 | 17.09 | 9.67 | 9.96 | 11.43 | 9.86 | 17.58 | 0.1 | 18.46 | 0.1 | 17.87 | 0.1 | 18.85 | 4.59 | 6.54 | 4.79 | 4.79 | 4.2 | 9.77 | 9.28 | 9.86 | 6.25 | 9.86 | 7.13
91.392 | 3.91 | 0.1 | 0 | 0.88 | 0.1 | 0.98 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 14.75 | 9.67 | 17.58 | 9.57 | 9.86 | 10.35 | 9.77 | 17.48 | 0.1 | 18.16 | 0.1 | 16.99 | 0.1 | 19.53 | 4.59 | 8.4 | 4.69 | 4.79 | 5.76 | 9.77 | 12.21 | 9.86 | 10.16 | 9.77 | 7.23
91.424 | 4.59 | 0.1 | 0 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.79 | 20.61 | 4.69 | 15.23 | 9.67 | 17.97 | 9.57 | 9.86 | 9.86 | 9.86 | 17.48 | 0.1 | 17.87 | 0 | 17.48 | 0.1 | 20.12 | 4.59 | 10.25 | 4.69 | 4.79 | 6.45 | 9.77 | 11.33 | 9.86 | 12.4 | 9.77 | 10.84
91.456 | 5.18 | 0.1 | 0 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 21 | 4.79 | 15.63 | 9.67 | 18.46 | 9.57 | 9.96 | 9.86 | 9.86 | 17.38 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 20.8 | 4.69 | 11.62 | 4.59 | 4.79 | 5.86 | 9.77 | 10.25 | 9.86 | 11.13 | 9.86 | 11.82
91.488 | 5.76 | 0.1 | 0 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 16.21 | 9.67 | 18.95 | 9.57 | 9.96 | 10.25 | 9.86 | 17.48 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 21.29 | 4.69 | 12.5 | 4.59 | 4.79 | 5.37 | 9.77 | 9.96 | 9.86 | 10.25 | 9.86 | 10.64
91.52 | 6.35 | 0.1 | 0 | 0.68 | 0.1 | 0.68 | 4.88 | 7.32 | 4.88 | 18.85 | 4.79 | 16.7 | 9.67 | 19.34 | 9.57 | 9.96 | 10.16 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 21.88 | 4.69 | 13.38 | 4.59 | 4.79 | 5.18 | 9.77 | 10.25 | 9.86 | 9.86 | 9.86 | 10.16
91.552 | 6.93 | 0.1 | 0 | 0.49 | 0.1 | 0.59 | 4.79 | 2.73 | 4.79 | 18.46 | 4.79 | 17.19 | 9.67 | 19.82 | 9.57 | 9.86 | 9.86 | 9.86 | 17.38 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 22.46 | 4.79 | 14.26 | 4.59 | 4.79 | 5.18 | 9.77 | 10.45 | 9.86 | 10.35 | 9.77 | 10.06
91.584 | 7.42 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.88 | 3.42 | 4.79 | 18.36 | 4.79 | 17.58 | 9.67 | 20.22 | 9.57 | 9.86 | 9.86 | 9.86 | 8.01 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 23.05 | 4.69 | 14.94 | 4.59 | 4.79 | 5.37 | 9.77 | 10.35 | 9.86 | 10.45 | 9.77 | 10.25
91.616 | 8.11 | 0.1 | 0 | 0.39 | 0.1 | 0.49 | 4.88 | 4.3 | 4.88 | 18.46 | 4.79 | 18.07 | 9.77 | 20.7 | 9.67 | 9.96 | 9.96 | 9.86 | 1.95 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.63 | 4.79 | 15.72 | 4.59 | 5.76 | 5.08 | 9.77 | 10.16 | 9.86 | 10.35 | 9.86 | 10.45
91.648 | 8.79 | 0.1 | 0 | 0.39 | 0.1 | 0.39 | 4.88 | 5.47 | 4.88 | 18.16 | 4.79 | 18.46 | 9.77 | 21 | 9.67 | 9.96 | 9.96 | 9.86 | 2.05 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 24.02 | 4.79 | 16.41 | 4.59 | 6.84 | 5.08 | 9.77 | 10.16 | 9.86 | 10.16 | 9.77 | 10.25
91.68 | 9.38 | 0.1 | 0.29 | 0.29 | 0 | 0.29 | 4.88 | 5.96 | 4.88 | 18.16 | 4.79 | 18.85 | 9.77 | 21.49 | 9.67 | 9.96 | 9.77 | 9.86 | 1.95 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.71 | 4.79 | 17.09 | 4.59 | 8.5 | 4.98 | 9.77 | 10.25 | 9.86 | 10.06 | 9.86 | 9.96
91.712 | 9.86 | 0.1 | 1.27 | 0.2 | 0.1 | 0.2 | 4.79 | 5.66 | 4.79 | 18.16 | 4.79 | 19.43 | 9.77 | 21.97 | 9.67 | 9.96 | 9.67 | 9.86 | 1.66 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 25.2 | 4.79 | 17.87 | 4.59 | 10.45 | 4.88 | 9.77 | 10.06 | 9.86 | 10.06 | 9.77 | 10.16
91.744 | 10.45 | 0.1 | 2.25 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.79 | 18.07 | 4.79 | 19.82 | 9.77 | 22.27 | 9.67 | 10.64 | 9.67 | 9.86 | 1.66 | 0.1 | 12.99 | 0.1 | 17.68 | 0.1 | 25.68 | 4.79 | 18.46 | 4.59 | 11.72 | 4.88 | 9.77 | 9.96 | 9.86 | 10.06 | 9.77 | 10.06
91.776 | 11.13 | 0.1 | 3.61 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 20.22 | 9.77 | 22.66 | 9.67 | 11.62 | 9.77 | 9.86 | 1.46 | 0.1 | 3.22 | 0.1 | 17.58 | 0.1 | 26.27 | 4.79 | 19.04 | 4.59 | 12.6 | 4.79 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86 | 10.06
91.808 | 11.72 | 0.1 | 4.39 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.7 | 9.77 | 23.05 | 9.67 | 12.99 | 9.67 | 9.86 | 1.37 | 0.1 | 2.25 | 0.1 | 5.08 | 0.1 | 21.58 | 4.79 | 19.63 | 4.69 | 13.48 | 4.69 | 9.77 | 9.96 | 9.86 | 9.96 | 9.77 | 9.86
91.84 | 12.21 | 0.1 | 4.98 | 0.1 | 0.68 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21 | 9.77 | 23.44 | 9.67 | 13.87 | 9.67 | 9.86 | 1.46 | 0.1 | 2.05 | 0.1 | 1.86 | 0.1 | 20.31 | 4.79 | 20.31 | 4.59 | 14.36 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77 | 9.96
91.872 | 12.79 | 0.1 | 5.57 | 0.1 | 1.37 | 0.1 | 4.79 | 4.88 | 4.79 | 17.97 | 4.79 | 18.85 | 9.77 | 23.83 | 9.67 | 14.65 | 9.67 | 9.86 | 1.17 | 0.1 | 1.56 | 0 | 2.15 | 0.1 | 21 | 4.79 | 20.9 | 4.69 | 15.04 | 4.59 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77 | 9.96
91.904 | 13.28 | 0.1 | 6.05 | 0.1 | 2.64 | 0 | 5.57 | 4.79 | 4.79 | 10.35 | 4.79 | 18.85 | 9.77 | 24.22 | 9.67 | 15.14 | 9.57 | 9.86 | 0.98 | 0.1 | 1.86 | 0.1 | 2.05 | 0.1 | 21.19 | 4.79 | 21.49 | 4.69 | 15.72 | 4.59 | 9.77 | 9.86 | 9.86 | 9.77 | 9.77 | 9.77
91.936 | 13.87 | 0.1 | 6.64 | 0.1 | 3.81 | 0.1 | 6.25 | 4.79 | 4.88 | 3.32 | 4.88 | 18.56 | 9.86 | 22.17 | 9.67 | 15.72 | 9.67 | 9.96 | 0.98 | 0.1 | 1.56 | 0.1 | 1.66 | 0.1 | 21.29 | 4.88 | 21.97 | 4.69 | 16.41 | 4.59 | 10.74 | 9.86 | 9.86 | 9.77 | 9.86 | 9.77
91.968 | 14.36 | 0.1 | 7.23 | 0.1 | 4.59 | 0.1 | 7.32 | 4.79 | 4.88 | 3.22 | 4.88 | 18.46 | 9.77 | 21 | 9.67 | 16.21 | 9.67 | 9.86 | 0.88 | 0.1 | 1.37 | 0.1 | 1.56 | 0.1 | 21.19 | 4.88 | 22.56 | 4.69 | 17.19 | 4.59 | 11.82 | 9.77 | 9.86 | 9.77 | 10.25 | 9.86
92 | 14.84 | 0.1 | 7.91 | 0.1 | 5.18 | 0.1 | 8.59 | 4.69 | 4.88 | 3.91 | 4.79 | 18.46 | 9.77 | 21.29 | 9.67 | 16.7 | 9.57 | 9.96 | 0.78 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 21.29 | 4.88 | 23.15 | 4.69 | 17.87 | 4.59 | 13.48 | 9.67 | 10.84 | 9.77 | 11.43 | 9.77
92.032 | 15.43 | 0.1 | 8.5 | 0.1 | 5.66 | 0 | 9.47 | 4.59 | 4.79 | 5.27 | 4.79 | 18.16 | 9.77 | 21.39 | 9.77 | 17.19 | 9.57 | 9.86 | 0.68 | 0.1 | 1.27 | 0.1 | 1.37 | 0.1 | 21.29 | 4.88 | 23.73 | 4.69 | 18.46 | 4.49 | 15.23 | 9.67 | 11.52 | 9.67 | 12.89 | 9.67
92.064 | 15.82 | 0.1 | 8.98 | 0.1 | 6.25 | 0.1 | 10.16 | 4.59 | 4.79 | 5.96 | 4.79 | 18.16 | 9.77 | 21.39 | 9.77 | 17.68 | 9.57 | 9.86 | 0.59 | 0.1 | 1.07 | 0.1 | 1.37 | 0.1 | 21.29 | 4.79 | 24.12 | 4.79 | 19.14 | 4.49 | 16.41 | 9.67 | 13.09 | 9.67 | 14.84 | 9.67
92.096 | 16.41 | 0.1 | 9.67 | 0.1 | 6.84 | 0.1 | 10.74 | 4.59 | 4.88 | 5.66 | 4.88 | 18.16 | 9.86 | 21.39 | 9.77 | 18.07 | 9.57 | 9.96 | 0.49 | 0.1 | 1.07 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 24.71 | 4.79 | 19.73 | 4.59 | 17.29 | 9.67 | 15.14 | 9.67 | 16.11 | 9.67
92.128 | 16.8 | 0.1 | 10.25 | 0.1 | 7.42 | 0.1 | 11.33 | 4.59 | 4.88 | 5.27 | 4.88 | 18.16 | 9.77 | 21.39 | 9.77 | 18.56 | 9.57 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 25.2 | 4.79 | 20.31 | 4.59 | 18.07 | 9.67 | 16.31 | 9.67 | 17.09 | 9.67
92.16 | 17.29 | 0.1 | 10.84 | 0.1 | 8.01 | 0.1 | 11.91 | 4.59 | 4.79 | 5.08 | 4.88 | 18.16 | 9.77 | 21.39 | 9.77 | 18.95 | 9.67 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 25.78 | 4.79 | 20.9 | 4.59 | 18.85 | 9.67 | 17.19 | 9.67 | 17.87 | 9.67
92.192 | 15.23 | 0.1 | 11.43 | 0.1 | 8.69 | 0.1 | 12.6 | 4.49 | 4.79 | 4.88 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 19.43 | 9.57 | 9.86 | 0.29 | 0.1 | 0.78 | 0 | 0.78 | 0.1 | 21.19 | 4.88 | 26.37 | 4.79 | 21.39 | 4.59 | 19.53 | 9.57 | 17.97 | 9.67 | 18.65 | 9.67
92.224 | 15.14 | 0.1 | 12.01 | 0.1 | 9.28 | 0.1 | 13.09 | 4.49 | 4.79 | 4.98 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 19.82 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 21.19 | 4.79 | 22.66 | 4.79 | 22.07 | 4.59 | 20.12 | 9.67 | 18.65 | 9.67 | 19.24 | 9.57
92.256 | 14.84 | 0.1 | 12.5 | 0.1 | 9.86 | 0.1 | 13.67 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.77 | 20.22 | 9.67 | 10.55 | 0.29 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 6.64 | 4.88 | 21.97 | 4.79 | 22.66 | 4.59 | 20.7 | 9.67 | 19.34 | 9.67 | 20.02 | 9.67
92.288 | 14.65 | 0.1 | 13.09 | 0.1 | 10.45 | 0.1 | 14.16 | 4.59 | 4.79 | 4.88 | 4.88 | 10.16 | 9.86 | 21.39 | 9.77 | 20.7 | 9.67 | 11.23 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 2.54 | 4.79 | 21.68 | 4.79 | 23.24 | 4.59 | 21.29 | 9.67 | 20.02 | 9.67 | 20.61 | 9.67
92.32 | 14.75 | 0.1 | 13.67 | 0.1 | 11.04 | 0.1 | 14.65 | 4.59 | 4.88 | 4.79 | 4.88 | 3.32 | 9.86 | 21.39 | 9.77 | 21.19 | 9.67 | 12.31 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.52 | 4.88 | 21.49 | 4.79 | 23.73 | 4.59 | 21.78 | 9.67 | 20.61 | 9.67 | 21.09 | 9.67
92.352 | 14.45 | 0.1 | 14.16 | 0.1 | 11.62 | 0.1 | 15.23 | 4.59 | 4.79 | 4.69 | 4.79 | 3.22 | 9.77 | 21.29 | 9.77 | 21.49 | 9.67 | 13.57 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4 | 4.88 | 21.58 | 4.79 | 24.22 | 4.59 | 22.56 | 9.67 | 21.19 | 9.57 | 21.58 | 9.57
92.384 | 14.45 | 0.1 | 14.55 | 0.1 | 12.21 | 0.1 | 15.72 | 4.59 | 4.79 | 4.59 | 4.79 | 3.91 | 9.77 | 11.43 | 9.77 | 21.88 | 9.67 | 14.36 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 5.18 | 4.79 | 21.39 | 4.79 | 24.71 | 4.69 | 23.05 | 9.67 | 21.68 | 9.67 | 22.17 | 9.67
92.416 | 14.45 | 0.1 | 15.14 | 0.1 | 12.7 | 0.1 | 16.21 | 4.69 | 4.79 | 4.69 | 4.88 | 5.27 | 9.86 | 6.54 | 9.77 | 22.36 | 9.67 | 14.94 | 0.2 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 6.25 | 4.88 | 21.39 | 4.79 | 25.2 | 4.69 | 23.54 | 9.67 | 22.36 | 9.67 | 22.75 | 9.67
92.448 | 14.45 | 0.1 | 15.72 | 0.1 | 13.28 | 0.1 | 16.7 | 4.69 | 4.88 | 4.59 | 4.88 | 5.96 | 9.77 | 8.89 | 9.86 | 22.75 | 9.67 | 15.53 | 0.1 | 0.1 | 0.29 | 0.59 | 0.29 | 0.1 | 6.05 | 4.88 | 21.39 | 4.79 | 25.78 | 4.69 | 24.02 | 9.67 | 22.85 | 9.67 | 23.34 | 9.67
92.48 | 14.45 | 0.1 | 16.11 | 0.1 | 13.77 | 0.1 | 17.09 | 4.69 | 4.79 | 4.59 | 4.88 | 5.66 | 9.77 | 11.43 | 9.77 | 23.15 | 9.67 | 16.02 | 0.1 | 0.1 | 0.29 | 1.66 | 0.29 | 0.1 | 5.57 | 4.88 | 21.29 | 4.79 | 26.27 | 4.79 | 24.61 | 9.67 | 23.44 | 9.67 | 23.83 | 9.67
92.512 | 14.36 | 0.1 | 16.7 | 0.1 | 14.36 | 0 | 17.58 | 4.69 | 4.88 | 4.59 | 4.79 | 5.27 | 9.77 | 10.94 | 9.86 | 23.54 | 9.67 | 16.5 | 0 | 0.88 | 0.2 | 3.13 | 0.2 | 0.1 | 5.18 | 4.79 | 21.29 | 4.88 | 23.05 | 4.69 | 25.2 | 9.67 | 24.02 | 9.67 | 24.32 | 9.67
92.544 | 14.36 | 0.1 | 17.09 | 0.1 | 14.84 | 0.1 | 17.97 | 4.69 | 5.57 | 4.59 | 4.79 | 4.98 | 9.77 | 10.16 | 9.77 | 23.83 | 9.67 | 16.99 | 0.1 | 1.95 | 0.1 | 5.18 | 0.2 | 0.1 | 5.08 | 4.79 | 21.19 | 4.88 | 22.17 | 4.79 | 25.59 | 9.67 | 24.51 | 9.67 | 24.81 | 9.67
92.576 | 14.36 | 0.1 | 14.26 | 0.1 | 15.33 | 0.1 | 18.46 | 4.79 | 6.25 | 4.59 | 4.88 | 4.98 | 9.86 | 9.86 | 9.86 | 24.22 | 9.77 | 17.38 | 0.1 | 3.52 | 0.1 | 6.45 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 21.78 | 4.79 | 26.17 | 9.77 | 25 | 9.67 | 25.39 | 9.67
92.608 | 8.59 | 0.1 | 14.65 | 0.1 | 15.82 | 0.1 | 18.95 | 4.79 | 7.62 | 4.59 | 4.88 | 4.98 | 9.77 | 10.16 | 9.86 | 22.75 | 9.77 | 17.87 | 0.1 | 5.37 | 0.1 | 7.23 | 0.1 | 1.07 | 5.08 | 4.88 | 21.19 | 4.88 | 21.58 | 4.79 | 26.66 | 9.77 | 25.59 | 9.67 | 25.68 | 9.67
92.64 | 2.44 | 0.1 | 14.45 | 0.1 | 16.31 | 0.1 | 19.34 | 4.79 | 8.69 | 4.49 | 4.88 | 4.88 | 9.86 | 10.25 | 9.86 | 21.09 | 9.77 | 18.26 | 0.1 | 6.45 | 0.1 | 8.01 | 0.1 | 2.05 | 4.98 | 4.79 | 14.84 | 4.79 | 21.58 | 4.79 | 27.05 | 9.77 | 26.08 | 9.67 | 26.27 | 9.67
92.672 | 1.86 | 0.1 | 14.55 | 0.1 | 16.8 | 0 | 19.82 | 4.79 | 9.57 | 4.49 | 4.88 | 4.79 | 9.77 | 10.06 | 9.77 | 21.29 | 9.77 | 18.85 | 0 | 7.32 | 0.1 | 8.79 | 0.1 | 3.91 | 4.98 | 4.79 | 3.03 | 4.79 | 21.39 | 4.79 | 27.64 | 9.77 | 26.56 | 9.67 | 26.86 | 9.67
92.704 | 1.76 | 0 | 14.45 | 0.1 | 17.29 | 0.1 | 20.12 | 4.79 | 10.25 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 21.58 | 9.77 | 19.24 | 0 | 8.11 | 0 | 9.57 | 0 | 5.47 | 4.88 | 4.79 | 3.13 | 4.88 | 21.29 | 4.79 | 28.03 | 9.77 | 27.05 | 9.77 | 27.25 | 9.67
92.736 | 1.56 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 20.7 | 4.88 | 10.84 | 4.59 | 4.79 | 4.79 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 19.73 | 0.1 | 8.98 | 0.1 | 10.45 | 0.1 | 6.64 | 4.88 | 4.88 | 3.42 | 4.88 | 21.39 | 4.79 | 28.52 | 9.77 | 27.54 | 9.77 | 27.83 | 9.77
92.768 | 1.37 | 0.1 | 14.45 | 0.1 | 15.33 | 0.1 | 21.09 | 4.88 | 11.43 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.12 | 0.1 | 9.77 | 0.1 | 11.23 | 0.1 | 7.42 | 4.88 | 4.88 | 4.39 | 4.88 | 21.29 | 4.79 | 29.01 | 9.77 | 28.03 | 9.77 | 28.22 | 9.77
92.8 | 1.17 | 0.1 | 14.45 | 0.1 | 14.84 | 0.1 | 18.95 | 4.88 | 12.01 | 4.59 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.51 | 0.1 | 10.55 | 0.1 | 12.01 | 0.1 | 8.3 | 4.69 | 4.88 | 5.86 | 4.79 | 21.29 | 4.79 | 29.49 | 9.77 | 28.42 | 9.77 | 28.71 | 9.77
92.832 | 1.07 | 0.1 | 14.36 | 0.1 | 14.65 | 0 | 18.95 | 4.88 | 12.6 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 20.9 | 0 | 11.33 | 0 | 12.7 | 0 | 9.28 | 4.69 | 4.79 | 6.05 | 4.88 | 21.19 | 4.79 | 26.08 | 9.77 | 28.91 | 9.77 | 29.3 | 9.77
92.864 | 1.07 | 0.1 | 14.36 | 0.1 | 14.75 | 0.1 | 18.56 | 4.88 | 13.18 | 4.59 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.29 | 0 | 12.31 | 0 | 13.38 | 0 | 9.86 | 4.59 | 4.79 | 5.76 | 4.88 | 21.19 | 4.79 | 25.29 | 9.77 | 29.4 | 9.77 | 24.71 | 9.77
92.896 | 0.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 13.77 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.86 | 21.78 | 0.1 | 12.89 | 0.1 | 14.16 | 0.1 | 10.74 | 4.59 | 4.88 | 5.37 | 4.88 | 21.19 | 4.79 | 24.9 | 9.77 | 25.1 | 9.77 | 23.73 | 9.77
92.928 | 0.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.46 | 4.88 | 14.26 | 4.69 | 5.47 | 4.59 | 9.77 | 9.67 | 9.86 | 21.49 | 9.77 | 22.17 | 0.1 | 13.67 | 0.1 | 14.84 | 0.1 | 11.43 | 4.59 | 4.79 | 5.18 | 4.88 | 11.91 | 4.88 | 24.71 | 9.77 | 24.41 | 9.77 | 24.22 | 9.77
92.96 | 0.68 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.26 | 4.88 | 14.75 | 4.69 | 6.25 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 22.56 | 0 | 14.26 | 0.1 | 15.63 | 0.1 | 12.21 | 4.59 | 4.79 | 5.27 | 4.88 | 3.03 | 4.79 | 24.61 | 9.77 | 24.41 | 9.77 | 24.12 | 9.77
92.992 | 0.68 | 0.1 | 12.79 | 0.1 | 14.45 | 0 | 18.26 | 4.88 | 15.33 | 4.69 | 7.32 | 4.49 | 9.77 | 9.67 | 9.77 | 21.39 | 9.77 | 22.95 | 0 | 15.04 | 0 | 16.41 | 0 | 12.99 | 4.59 | 4.88 | 5.08 | 4.88 | 3.32 | 4.79 | 24.51 | 9.77 | 24.41 | 9.77 | 24.32 | 9.77
93.024 | 0.59 | 0 | 3.91 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.72 | 4.69 | 8.5 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 23.24 | 0.1 | 15.82 | 0 | 16.99 | 0 | 13.77 | 4.59 | 5.76 | 5.18 | 4.88 | 3.52 | 4.79 | 24.32 | 9.77 | 24.41 | 9.77 | 24.22 | 9.77
93.056 | 0.49 | 0.1 | 1.86 | 0.1 | 14.45 | 0.1 | 18.16 | 4.98 | 16.21 | 4.79 | 9.47 | 4.59 | 9.86 | 9.67 | 9.86 | 11.52 | 9.86 | 23.63 | 0.1 | 16.41 | 0.1 | 17.68 | 0.1 | 14.55 | 4.59 | 7.13 | 4.98 | 4.88 | 4.59 | 4.79 | 24.41 | 9.77 | 24.41 | 9.86 | 24.22 | 9.86
93.088 | 0.49 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.7 | 4.79 | 10.06 | 4.59 | 9.86 | 9.67 | 9.86 | 6.64 | 9.86 | 24.02 | 0.1 | 17.19 | 0.1 | 18.26 | 0.1 | 15.23 | 4.59 | 8.98 | 4.88 | 4.79 | 6.05 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
93.12 | 0.39 | 0.1 | 1.56 | 0.1 | 14.36 | 0.1 | 18.16 | 4.98 | 17.19 | 4.79 | 10.64 | 4.59 | 10.55 | 9.67 | 9.86 | 8.89 | 9.86 | 24.41 | 0.1 | 17.87 | 0.1 | 19.04 | 0.1 | 15.82 | 4.49 | 10.74 | 4.88 | 4.79 | 6.15 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
93.152 | 0.29 | 0.1 | 1.46 | 0.1 | 10.25 | 0 | 18.07 | 4.88 | 17.68 | 4.79 | 11.43 | 4.59 | 11.82 | 9.67 | 9.86 | 11.62 | 9.86 | 22.27 | 0 | 18.46 | 0 | 19.53 | 0 | 16.7 | 4.49 | 12.11 | 4.79 | 4.79 | 5.57 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
93.184 | 0.29 | 0 | 1.27 | 0.1 | 2.93 | 0 | 18.07 | 4.88 | 18.07 | 4.79 | 12.01 | 4.59 | 12.99 | 9.57 | 9.86 | 11.13 | 9.86 | 22.27 | 0 | 19.14 | 0 | 20.12 | 0 | 17.29 | 4.49 | 12.89 | 4.69 | 4.88 | 5.27 | 4.79 | 24.22 | 9.77 | 24.32 | 9.77 | 24.12 | 9.77
93.216 | 0.29 | 0.1 | 1.17 | 0.1 | 2.05 | 0.1 | 11.82 | 4.88 | 18.56 | 4.79 | 12.5 | 4.69 | 13.87 | 9.67 | 9.86 | 10.35 | 9.86 | 21.78 | 0.1 | 19.73 | 0.1 | 20.7 | 0.1 | 17.97 | 4.59 | 13.77 | 4.79 | 4.88 | 5.18 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
93.248 | 0.2 | 0.1 | 1.07 | 0.1 | 1.76 | 0.1 | 3.42 | 4.88 | 19.04 | 4.79 | 12.99 | 4.69 | 14.55 | 9.57 | 9.86 | 9.86 | 9.86 | 21.68 | 0.1 | 20.31 | 0.1 | 21.29 | 0.1 | 18.65 | 4.59 | 14.65 | 4.69 | 4.88 | 5.08 | 4.79 | 15.72 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
93.28 | 0.1 | 0.1 | 0.98 | 0.1 | 1.46 | 0.1 | 3.22 | 4.88 | 19.43 | 4.88 | 13.57 | 4.69 | 15.14 | 9.67 | 9.86 | 10.16 | 9.86 | 21.68 | 0 | 20.9 | 0.1 | 21.97 | 0.1 | 19.24 | 4.59 | 15.33 | 4.59 | 4.79 | 5.18 | 4.88 | 6.64 | 9.77 | 24.22 | 9.86 | 23.63 | 9.86
93.312 | 0.1 | 0.1 | 0.78 | 0.1 | 1.46 | 0.1 | 3.71 | 4.88 | 19.92 | 4.79 | 14.16 | 4.69 | 15.63 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 0 | 21.58 | 0 | 22.46 | 0 | 19.82 | 4.59 | 16.02 | 4.59 | 4.79 | 5.08 | 4.79 | 7.42 | 9.77 | 17.97 | 9.77 | 8.01 | 9.86
93.344 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0 | 4.79 | 4.88 | 20.22 | 4.88 | 14.65 | 4.69 | 16.11 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 0 | 22.07 | 0 | 18.16 | 0 | 20.41 | 4.69 | 16.7 | 4.49 | 4.88 | 4.98 | 4.79 | 11.13 | 9.77 | 7.81 | 9.77 | 6.35 | 9.86
93.376 | 0.1 | 0.68 | 0.59 | 0.1 | 1.27 | 0.1 | 5.86 | 4.88 | 20.7 | 4.88 | 15.14 | 4.79 | 16.6 | 9.67 | 10.55 | 9.96 | 9.86 | 21.58 | 0.1 | 22.66 | 0.1 | 17.38 | 0 | 21 | 4.69 | 17.38 | 4.59 | 4.88 | 4.98 | 4.79 | 11.82 | 9.77 | 6.64 | 9.86 | 10.16 | 9.86
93.408 | 0.1 | 1.37 | 0.59 | 0.1 | 1.07 | 0.1 | 5.66 | 4.88 | 21.09 | 4.88 | 15.63 | 4.79 | 17.19 | 9.67 | 11.52 | 9.96 | 9.86 | 21.49 | 0.1 | 19.43 | 0.1 | 17.77 | 0 | 21.58 | 4.69 | 17.97 | 4.59 | 4.88 | 4.88 | 4.79 | 10.55 | 9.77 | 10.35 | 9.86 | 12.31 | 9.86
93.44 | 0.1 | 2.44 | 0.49 | 0.1 | 0.88 | 0.1 | 5.37 | 4.88 | 18.95 | 4.88 | 16.11 | 4.79 | 17.58 | 9.67 | 12.89 | 9.96 | 9.86 | 21.49 | 0.1 | 18.65 | 0.1 | 17.68 | 0.1 | 22.17 | 4.79 | 18.65 | 4.59 | 4.79 | 4.79 | 4.79 | 9.96 | 9.77 | 12.11 | 9.77 | 11.04 | 9.86
93.472 | 0 | 3.81 | 0.39 | 0.1 | 0.88 | 0 | 4.98 | 4.88 | 18.46 | 4.88 | 16.6 | 4.79 | 18.07 | 9.67 | 13.77 | 9.77 | 9.86 | 21.39 | 0 | 18.16 | 0 | 17.87 | 0 | 22.66 | 4.69 | 19.34 | 4.49 | 5.76 | 4.69 | 4.79 | 10.25 | 9.77 | 10.94 | 9.77 | 9.96 | 9.86
93.504 | 0 | 4.49 | 0.39 | 0.1 | 0.88 | 0 | 4.88 | 4.88 | 18.46 | 4.79 | 16.99 | 4.79 | 18.46 | 9.67 | 14.55 | 9.77 | 9.86 | 21.39 | 0 | 17.97 | 0 | 17.68 | 0 | 19.53 | 4.79 | 19.92 | 4.49 | 6.45 | 4.69 | 4.79 | 10.45 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
93.536 | 0.1 | 5.18 | 0.39 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.36 | 4.88 | 17.58 | 4.88 | 18.95 | 9.67 | 15.14 | 9.86 | 9.86 | 21.39 | 0 | 17.97 | 0.1 | 17.77 | 0.1 | 18.75 | 4.79 | 20.61 | 4.59 | 7.91 | 4.69 | 4.79 | 10.45 | 9.77 | 10.25 | 9.86 | 10.45 | 9.86
93.568 | 0.1 | 5.57 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.36 | 4.88 | 17.97 | 4.88 | 19.43 | 9.67 | 15.63 | 9.77 | 9.86 | 17.38 | 0.1 | 17.87 | 0.1 | 17.77 | 0.1 | 18.26 | 4.79 | 21.19 | 4.59 | 10.06 | 4.69 | 4.79 | 10.16 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86
93.6 | 0.1 | 6.15 | 0.2 | 0.1 | 0.49 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 18.46 | 4.88 | 19.92 | 9.67 | 16.11 | 9.77 | 9.86 | 8.01 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 18.16 | 4.79 | 21.68 | 4.59 | 11.43 | 4.59 | 4.79 | 10.06 | 9.77 | 10.45 | 9.86 | 10.25 | 9.86
93.632 | 0 | 6.84 | 0.1 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.07 | 4.88 | 18.85 | 4.88 | 20.22 | 9.67 | 16.7 | 9.67 | 9.86 | 7.52 | 0 | 17.77 | 0 | 17.68 | 0 | 18.07 | 4.79 | 22.27 | 4.49 | 12.4 | 4.59 | 4.79 | 10.16 | 9.77 | 10.25 | 9.77 | 10.06 | 9.86
93.664 | 0 | 7.32 | 0.1 | 0.1 | 0.49 | 0 | 4.69 | 4.88 | 18.16 | 4.88 | 19.24 | 4.88 | 20.7 | 9.67 | 17.09 | 9.67 | 9.86 | 10.45 | 0 | 17.68 | 0 | 17.68 | 0 | 17.97 | 4.79 | 22.85 | 4.59 | 13.38 | 4.59 | 4.79 | 10.06 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86
93.696 | 0.1 | 8.01 | 0.1 | 0.1 | 0.29 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.73 | 4.88 | 21.09 | 9.77 | 17.58 | 9.67 | 9.86 | 11.62 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 23.44 | 4.59 | 14.06 | 4.59 | 4.88 | 10.06 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
93.728 | 0.1 | 8.5 | 0.1 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 21.49 | 9.77 | 18.07 | 9.67 | 9.86 | 10.64 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 23.83 | 4.69 | 14.84 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
93.76 | 0.1 | 9.28 | 0 | 0.1 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.51 | 4.88 | 21.88 | 9.77 | 18.46 | 9.67 | 9.86 | 9.86 | 0 | 17.58 | 0.1 | 16.99 | 0.1 | 17.77 | 4.88 | 24.32 | 4.69 | 15.63 | 4.59 | 4.79 | 9.96 | 9.77 | 10.06 | 9.86 | 9.86 | 9.86
93.792 | 0 | 9.77 | 0 | 1.07 | 0.2 | 0.1 | 4.49 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.27 | 9.77 | 18.95 | 9.67 | 9.86 | 9.86 | 0 | 17.58 | 0 | 3.71 | 0 | 17.77 | 4.88 | 24.9 | 4.69 | 16.21 | 4.49 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
93.824 | 0 | 10.35 | 0 | 1.86 | 0.2 | 0.88 | 4.49 | 4.88 | 18.07 | 4.88 | 18.85 | 4.88 | 22.56 | 9.77 | 19.34 | 9.67 | 9.86 | 10.16 | 0 | 10.64 | 0 | 2.05 | 0 | 17.77 | 4.88 | 25.39 | 4.69 | 16.99 | 4.49 | 4.79 | 9.86 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86
93.856 | 0.1 | 10.94 | 0 | 3.22 | 0.1 | 1.76 | 4.59 | 4.88 | 9.18 | 4.88 | 18.95 | 4.88 | 23.05 | 9.77 | 19.92 | 9.67 | 9.86 | 10.25 | 0 | 2.93 | 0.1 | 2.44 | 0 | 17.77 | 4.88 | 25.98 | 4.69 | 17.68 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
93.888 | 0.1 | 11.52 | 0 | 4.1 | 0.1 | 3.03 | 4.49 | 4.88 | 3.03 | 4.88 | 18.46 | 4.88 | 23.44 | 9.77 | 20.22 | 9.67 | 9.86 | 9.96 | 0.1 | 2.25 | 0.1 | 1.76 | 0.1 | 17.68 | 4.88 | 25.29 | 4.79 | 18.26 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
93.92 | 0.1 | 12.11 | 0 | 4.88 | 0 | 4 | 4.49 | 4.88 | 3.32 | 4.88 | 18.26 | 4.88 | 23.83 | 9.77 | 20.7 | 9.67 | 10.55 | 9.96 | 0 | 2.15 | 0.1 | 1.76 | 0 | 9.86 | 4.88 | 21.68 | 4.79 | 18.95 | 4.59 | 5.08 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86
93.952 | 0 | 12.7 | 0.1 | 5.37 | 0.1 | 4.79 | 4.39 | 4.88 | 4.1 | 4.88 | 18.36 | 4.88 | 24.32 | 9.77 | 21 | 9.67 | 11.33 | 9.86 | 0 | 1.76 | 0 | 1.56 | 0 | 2.05 | 4.88 | 21.19 | 4.69 | 19.43 | 4.59 | 6.35 | 9.67 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86
93.984 | 0 | 13.18 | 0 | 5.96 | 0.1 | 5.27 | 4.49 | 4.88 | 5.08 | 4.88 | 18.16 | 4.88 | 21.29 | 9.77 | 21.39 | 9.67 | 12.4 | 9.86 | 0 | 1.86 | 0 | 1.46 | 0 | 2.34 | 4.88 | 21.49 | 4.79 | 20.22 | 4.59 | 7.81 | 9.67 | 10.84 | 9.77 | 9.96 | 9.77 | 9.86
94.016 | 0.1 | 13.77 | 0 | 6.54 | 0.1 | 5.86 | 4.59 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 20.9 | 9.77 | 21.78 | 9.67 | 13.57 | 9.86 | 0.1 | 1.56 | 0.1 | 1.37 | 0.1 | 2.15 | 4.88 | 21.49 | 4.79 | 20.7 | 4.69 | 9.96 | 9.67 | 12.21 | 9.77 | 10.94 | 9.67 | 10.84
94.048 | 0.1 | 14.36 | 0 | 7.13 | 0.1 | 6.45 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 21.19 | 9.77 | 22.27 | 9.67 | 14.36 | 9.77 | 0 | 1.46 | 0.1 | 1.37 | 0.1 | 1.95 | 4.88 | 21.39 | 4.79 | 21.29 | 4.69 | 11.43 | 9.67 | 14.26 | 9.77 | 12.4 | 9.67 | 11.43
94.08 | 0.1 | 14.75 | 0 | 7.72 | 0.1 | 7.03 | 4.59 | 4.88 | 5.37 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 22.75 | 9.67 | 15.04 | 9.77 | 0.1 | 1.46 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.29 | 4.79 | 21.97 | 4.69 | 12.4 | 9.67 | 15.82 | 9.67 | 14.45 | 9.67 | 13.09
94.112 | 0 | 15.33 | 0 | 8.4 | 0.1 | 7.62 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.29 | 9.77 | 23.05 | 9.67 | 15.53 | 9.67 | 0 | 1.27 | 0 | 0.98 | 0 | 1.56 | 4.88 | 21.29 | 4.79 | 22.56 | 4.69 | 13.28 | 9.67 | 16.9 | 9.67 | 16.02 | 9.67 | 15.23
94.144 | 0 | 15.82 | 0 | 8.98 | 0.1 | 8.2 | 4.59 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 21.39 | 9.77 | 23.44 | 9.67 | 16.02 | 9.67 | 0 | 1.17 | 0 | 0.88 | 0 | 1.46 | 4.88 | 21.29 | 4.79 | 23.15 | 4.69 | 14.06 | 9.57 | 17.68 | 9.67 | 16.99 | 9.67 | 16.31
94.176 | 0.1 | 16.31 | 0 | 9.57 | 0.1 | 8.79 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.83 | 9.67 | 16.5 | 9.67 | 0.1 | 1.17 | 0.1 | 0.88 | 0 | 1.37 | 4.88 | 21.29 | 4.79 | 23.63 | 4.79 | 14.75 | 9.67 | 18.36 | 9.67 | 17.87 | 9.67 | 17.19
94.208 | 0.1 | 16.7 | 0 | 10.16 | 0.1 | 9.47 | 4.69 | 5.57 | 4.88 | 4.88 | 17.97 | 4.88 | 21.39 | 9.77 | 24.22 | 9.77 | 16.99 | 9.67 | 0 | 0.88 | 0.1 | 0.78 | 0.1 | 1.37 | 4.88 | 21.19 | 4.79 | 24.02 | 4.79 | 15.43 | 9.67 | 19.14 | 9.67 | 18.56 | 9.67 | 17.97
94.24 | 0.1 | 17.19 | 0 | 10.74 | 0 | 10.06 | 4.69 | 6.54 | 4.79 | 4.88 | 12.21 | 4.88 | 21.39 | 9.77 | 22.27 | 9.77 | 17.38 | 9.67 | 0 | 0.88 | 0.1 | 0.68 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.71 | 4.79 | 16.21 | 9.67 | 19.73 | 9.67 | 19.14 | 9.67 | 18.65
94.272 | 0 | 14.94 | 0 | 11.23 | 0.1 | 10.64 | 4.69 | 8.01 | 4.69 | 4.88 | 3.32 | 4.88 | 21.39 | 9.77 | 20.9 | 9.67 | 17.97 | 9.57 | 1.07 | 0.78 | 0 | 0.59 | 0 | 1.07 | 4.88 | 21.19 | 4.79 | 25.1 | 4.79 | 16.9 | 9.67 | 20.51 | 9.67 | 19.82 | 9.67 | 19.34
94.304 | 0 | 15.04 | 0 | 11.82 | 0.1 | 11.23 | 4.69 | 8.89 | 4.69 | 4.79 | 3.22 | 4.88 | 21.39 | 9.77 | 21.29 | 9.77 | 18.36 | 9.57 | 2.15 | 0.68 | 0 | 0.49 | 0 | 0.88 | 4.88 | 21.19 | 4.79 | 25.68 | 4.79 | 17.68 | 9.57 | 21 | 9.67 | 20.31 | 9.67 | 20.02
94.336 | 0.1 | 14.65 | 0 | 12.4 | 0.1 | 11.82 | 4.79 | 9.67 | 4.69 | 4.88 | 3.71 | 4.88 | 21.39 | 9.86 | 21.49 | 9.77 | 18.85 | 9.67 | 4 | 0.68 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 8.5 | 4.79 | 26.17 | 4.79 | 18.26 | 9.67 | 21.68 | 9.67 | 20.9 | 9.67 | 20.61
94.368 | 0.1 | 14.55 | 0 | 12.99 | 0.1 | 12.4 | 4.79 | 10.35 | 4.59 | 4.88 | 4.88 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 19.24 | 9.57 | 5.47 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 2.34 | 4.88 | 21.78 | 4.79 | 18.95 | 9.67 | 22.17 | 9.67 | 21.58 | 9.67 | 21.19
94.4 | 0.1 | 14.65 | 0 | 13.48 | 0.1 | 12.89 | 4.79 | 11.04 | 4.59 | 4.88 | 5.86 | 4.88 | 21.39 | 9.77 | 21.39 | 9.77 | 19.73 | 9.57 | 6.54 | 0.59 | 0.1 | 0.39 | 0.1 | 0.78 | 4.88 | 3.22 | 4.79 | 21.39 | 4.79 | 19.43 | 9.67 | 22.66 | 9.67 | 22.17 | 9.67 | 21.68
94.432 | 0 | 14.55 | 0 | 14.06 | 0.1 | 13.48 | 4.79 | 11.62 | 4.49 | 4.88 | 5.76 | 4.88 | 10.55 | 9.77 | 21.39 | 9.77 | 20.12 | 9.57 | 7.42 | 0.49 | 0.88 | 0.29 | 0 | 0.59 | 4.88 | 3.81 | 4.79 | 21.29 | 4.79 | 20.12 | 9.67 | 23.34 | 9.67 | 22.66 | 9.67 | 22.27
94.464 | 0 | 14.55 | 0 | 14.55 | 0.1 | 13.97 | 4.79 | 12.11 | 4.49 | 4.79 | 5.18 | 4.88 | 6.45 | 9.77 | 21.39 | 9.77 | 20.51 | 9.57 | 8.3 | 0.49 | 1.66 | 0.29 | 0.1 | 0.59 | 4.88 | 4.88 | 4.79 | 21.29 | 4.79 | 20.7 | 9.67 | 23.73 | 9.67 | 23.24 | 9.67 | 22.85
94.496 | 0.1 | 14.55 | 0.1 | 15.14 | 0.1 | 14.45 | 4.79 | 12.7 | 4.59 | 4.88 | 4.98 | 4.88 | 9.38 | 9.86 | 21.39 | 9.86 | 20.9 | 9.67 | 8.98 | 0.49 | 3.32 | 0.29 | 1.46 | 0.59 | 4.88 | 6.15 | 4.79 | 21.29 | 4.79 | 21.19 | 9.67 | 24.32 | 9.67 | 23.83 | 9.67 | 23.44
94.528 | 0.1 | 14.45 | 0 | 15.53 | 0.1 | 14.94 | 4.79 | 13.28 | 4.59 | 4.88 | 4.88 | 4.88 | 11.62 | 9.77 | 21.39 | 9.86 | 21.39 | 9.67 | 9.77 | 0.29 | 5.27 | 0.2 | 2.73 | 0.49 | 4.88 | 6.05 | 4.88 | 21.29 | 4.79 | 21.88 | 9.67 | 24.81 | 9.67 | 24.32 | 9.67 | 23.93
94.56 | 0.1 | 14.45 | 0 | 16.11 | 0 | 15.53 | 4.79 | 13.77 | 4.49 | 4.88 | 4.98 | 4.88 | 10.94 | 9.86 | 21.39 | 9.86 | 21.78 | 9.67 | 10.55 | 0.29 | 6.45 | 0.2 | 4.59 | 0.39 | 4.88 | 5.57 | 4.88 | 21.29 | 4.79 | 22.46 | 9.67 | 25.29 | 9.67 | 24.81 | 9.67 | 24.41
94.592 | 0 | 14.45 | 0 | 16.5 | 0.1 | 16.02 | 4.79 | 14.36 | 4.49 | 5.47 | 4.88 | 4.88 | 10.06 | 9.77 | 21.39 | 9.77 | 22.17 | 9.67 | 11.43 | 0.2 | 7.23 | 0.1 | 6.15 | 0.29 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 23.05 | 9.67 | 25.78 | 9.67 | 25.29 | 9.67 | 25
94.624 | 0 | 14.36 | 0 | 16.99 | 0.1 | 16.5 | 4.79 | 14.84 | 4.49 | 6.25 | 4.79 | 4.88 | 9.96 | 9.77 | 21.39 | 9.86 | 22.56 | 9.67 | 12.21 | 0.2 | 8.01 | 0.1 | 6.93 | 0.29 | 4.88 | 5.08 | 4.79 | 21.29 | 4.79 | 23.63 | 9.67 | 26.37 | 9.67 | 25.78 | 9.67 | 25.39
94.656 | 0.1 | 14.36 | 0 | 15.82 | 0.1 | 16.99 | 4.88 | 15.33 | 4.59 | 7.32 | 4.79 | 4.88 | 10.16 | 9.86 | 21.39 | 9.86 | 22.95 | 9.67 | 12.99 | 0.2 | 8.79 | 0.1 | 7.81 | 0.29 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 24.22 | 9.77 | 26.86 | 9.77 | 26.27 | 9.77 | 25.78
94.688 | 0.1 | 11.62 | 0 | 14.26 | 0.1 | 16.6 | 4.88 | 15.82 | 4.59 | 8.59 | 4.79 | 4.88 | 10.25 | 9.86 | 11.43 | 9.86 | 23.34 | 9.67 | 13.57 | 0.2 | 9.57 | 0.1 | 8.59 | 0.2 | 4.88 | 5.18 | 4.88 | 21.19 | 4.79 | 24.51 | 9.77 | 27.25 | 9.77 | 26.86 | 9.77 | 26.47
94.72 | 0.1 | 3.22 | 0 | 14.55 | 0 | 14.65 | 4.88 | 16.31 | 4.59 | 9.47 | 4.69 | 4.88 | 10.06 | 9.86 | 6.54 | 9.86 | 23.73 | 9.67 | 14.26 | 0.2 | 10.45 | 0.1 | 9.38 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 25 | 9.77 | 27.74 | 9.77 | 27.34 | 9.77 | 26.95
94.752 | 0 | 2.15 | 0 | 14.55 | 0.1 | 14.45 | 4.88 | 16.9 | 4.59 | 10.16 | 4.59 | 4.88 | 9.86 | 9.77 | 9.28 | 9.86 | 24.12 | 9.67 | 15.14 | 0.1 | 11.13 | 0.1 | 10.16 | 0.2 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 25.59 | 9.77 | 28.32 | 9.77 | 27.74 | 9.77 | 27.44
94.784 | 0 | 1.66 | 0 | 14.55 | 0.1 | 14.65 | 4.79 | 17.29 | 4.59 | 10.64 | 4.49 | 4.88 | 9.86 | 9.77 | 11.33 | 9.86 | 24.41 | 9.67 | 15.72 | 0.1 | 11.91 | 0.1 | 10.84 | 0.1 | 4.88 | 4.88 | 4.79 | 20.22 | 4.79 | 26.08 | 9.77 | 28.81 | 9.77 | 28.32 | 9.77 | 27.93
94.816 | 0.1 | 1.46 | 0 | 14.55 | 0.1 | 14.55 | 4.88 | 17.77 | 4.69 | 11.33 | 4.59 | 4.88 | 9.96 | 9.86 | 10.94 | 9.86 | 22.46 | 9.67 | 16.5 | 0.1 | 12.7 | 0.1 | 11.72 | 0.1 | 4.88 | 4.88 | 4.88 | 4.59 | 4.88 | 21.58 | 9.77 | 29.3 | 9.77 | 28.71 | 9.77 | 28.42
94.848 | 0.1 | 1.46 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 18.16 | 4.69 | 11.91 | 4.49 | 4.88 | 9.86 | 9.86 | 10.16 | 9.86 | 22.27 | 9.67 | 17.09 | 0.1 | 13.48 | 0.1 | 12.5 | 0.1 | 4.88 | 4.79 | 4.88 | 3.22 | 4.79 | 20.51 | 9.77 | 25 | 9.77 | 29.3 | 9.77 | 28.91
94.88 | 0.1 | 1.37 | 0 | 14.45 | 0 | 14.55 | 4.88 | 18.65 | 4.69 | 12.5 | 4.49 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 21.88 | 9.77 | 17.87 | 0.1 | 14.26 | 0.1 | 13.28 | 0.1 | 4.88 | 4.69 | 4.88 | 3.32 | 4.88 | 21.09 | 9.77 | 23.54 | 9.77 | 24.61 | 9.77 | 29.4
94.912 | 0 | 1.17 | 0 | 14.45 | 0.1 | 14.45 | 4.79 | 19.14 | 4.69 | 13.09 | 4.49 | 4.88 | 9.77 | 9.77 | 10.06 | 9.86 | 21.68 | 9.67 | 18.46 | 0.1 | 14.94 | 0.1 | 14.06 | 0.1 | 4.98 | 4.59 | 4.79 | 4.3 | 4.79 | 21.19 | 9.77 | 24.02 | 9.77 | 23.24 | 9.77 | 25.39
94.944 | 0 | 1.07 | 0 | 14.45 | 0.1 | 14.45 | 4.79 | 19.53 | 4.69 | 13.67 | 4.49 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.68 | 9.67 | 19.14 | 0.1 | 15.53 | 0.1 | 14.65 | 0.1 | 5.96 | 4.59 | 4.79 | 5.66 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 23.93 | 9.77 | 24.32
94.976 | 0.1 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.02 | 4.79 | 14.16 | 4.59 | 5.57 | 9.77 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 19.73 | 0.1 | 16.21 | 0.1 | 15.33 | 0.1 | 7.32 | 4.59 | 4.88 | 6.25 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.41
95.008 | 0.1 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 20.41 | 4.79 | 14.65 | 4.49 | 6.25 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.31 | 0.1 | 16.99 | 0.1 | 16.11 | 0.1 | 9.38 | 4.59 | 4.88 | 5.86 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
95.04 | 0.1 | 0.78 | 0 | 14.45 | 0.1 | 14.45 | 4.79 | 20.8 | 4.79 | 15.14 | 4.49 | 7.42 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.8 | 0.1 | 17.68 | 0.1 | 16.8 | 0.1 | 11.04 | 4.59 | 4.79 | 5.47 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.32
95.072 | 0 | 0.78 | 0 | 14.45 | 0.1 | 14.36 | 4.79 | 19.24 | 4.79 | 15.63 | 4.49 | 8.59 | 9.67 | 10.45 | 9.86 | 9.86 | 21.39 | 9.77 | 21.39 | 0.1 | 18.26 | 0.1 | 17.38 | 0.1 | 12.31 | 4.59 | 4.79 | 5.27 | 4.79 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
95.104 | 0 | 0.59 | 0 | 6.54 | 0.1 | 14.36 | 4.79 | 17.87 | 4.79 | 16.11 | 4.49 | 9.47 | 9.67 | 11.43 | 9.77 | 9.86 | 21.49 | 9.77 | 21.88 | 0.1 | 18.85 | 0.1 | 18.07 | 0.1 | 13.09 | 4.49 | 4.79 | 5.18 | 4.79 | 21.29 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
95.136 | 0.1 | 0.59 | 0 | 1.95 | 0.1 | 7.23 | 4.88 | 18.16 | 4.79 | 16.7 | 4.59 | 10.16 | 9.67 | 12.7 | 9.77 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 19.53 | 0.1 | 18.75 | 0.1 | 13.97 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
95.168 | 0.1 | 0.49 | 0 | 2.05 | 0.1 | 2.05 | 4.88 | 18.26 | 4.79 | 17.09 | 4.59 | 10.74 | 9.57 | 13.67 | 9.67 | 9.96 | 21.39 | 9.77 | 18.36 | 0.1 | 20.22 | 0.1 | 19.43 | 0.1 | 14.65 | 4.49 | 4.88 | 4.98 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
95.2 | 0.1 | 0.39 | 0 | 1.76 | 0.1 | 1.95 | 4.88 | 18.16 | 4.88 | 17.58 | 4.59 | 11.43 | 9.57 | 14.45 | 9.67 | 9.96 | 21.39 | 9.77 | 18.07 | 0.1 | 20.9 | 0.1 | 20.12 | 0.1 | 15.43 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
95.232 | 0 | 0.29 | 0 | 1.46 | 0.1 | 1.76 | 4.79 | 18.07 | 4.79 | 18.07 | 4.59 | 11.91 | 9.67 | 15.04 | 9.67 | 9.86 | 13.97 | 9.77 | 17.68 | 0.1 | 21.39 | 0.1 | 20.7 | 0.1 | 16.21 | 4.59 | 4.79 | 4.88 | 4.79 | 17.58 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.12
95.264 | 0 | 0.29 | 0 | 1.46 | 0.1 | 1.56 | 4.88 | 18.16 | 4.79 | 18.46 | 4.59 | 12.6 | 9.67 | 15.53 | 9.67 | 9.86 | 7.42 | 9.77 | 17.68 | 0.1 | 21.88 | 0 | 21.29 | 0.1 | 16.8 | 4.49 | 4.79 | 4.88 | 4.79 | 4.88 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.12
95.296 | 0.68 | 0.29 | 0 | 1.37 | 0.1 | 1.27 | 4.88 | 18.16 | 4.88 | 18.85 | 4.69 | 13.09 | 9.67 | 16.11 | 9.67 | 9.96 | 7.72 | 9.86 | 17.68 | 0.1 | 22.56 | 0.1 | 21.78 | 0.1 | 17.48 | 4.59 | 4.79 | 4.88 | 4.88 | 3.03 | 9.77 | 11.13 | 9.86 | 24.12 | 9.86 | 24.12
95.328 | 1.37 | 0.2 | 0 | 1.17 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.34 | 4.69 | 13.67 | 9.67 | 16.5 | 9.67 | 9.96 | 10.94 | 9.86 | 17.58 | 0.1 | 18.65 | 0.1 | 22.36 | 0.1 | 18.16 | 4.59 | 5.76 | 4.79 | 4.79 | 3.42 | 9.77 | 5.76 | 9.86 | 10.64 | 9.86 | 20.61
95.36 | 2.64 | 0.2 | 0 | 1.07 | 0 | 1.27 | 4.88 | 18.07 | 4.88 | 19.82 | 4.69 | 14.16 | 9.67 | 17.09 | 9.67 | 9.96 | 11.43 | 9.86 | 17.58 | 0.1 | 18.16 | 0.1 | 18.26 | 0.1 | 18.85 | 4.59 | 6.54 | 4.69 | 4.79 | 4.2 | 9.77 | 8.79 | 9.86 | 5.86 | 9.86 | 7.62
95.392 | 3.81 | 0.1 | 0 | 0.88 | 0.1 | 0.98 | 4.79 | 18.07 | 4.79 | 20.22 | 4.69 | 14.65 | 9.67 | 17.58 | 9.57 | 9.86 | 10.45 | 9.86 | 17.48 | 0.1 | 18.07 | 0.1 | 16.7 | 0.1 | 19.43 | 4.59 | 8.3 | 4.69 | 4.79 | 5.76 | 9.77 | 12.4 | 9.86 | 9.57 | 9.86 | 6.74
95.424 | 4.49 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 4.79 | 18.07 | 4.88 | 20.7 | 4.69 | 15.23 | 9.67 | 17.97 | 9.57 | 9.86 | 9.86 | 9.86 | 17.48 | 0.1 | 17.77 | 0 | 17.48 | 0.1 | 20.02 | 4.59 | 10.25 | 4.69 | 4.79 | 6.15 | 9.77 | 11.33 | 9.86 | 12.31 | 9.86 | 10.45
95.456 | 5.08 | 0.1 | 0 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 21 | 4.79 | 15.72 | 9.67 | 18.46 | 9.67 | 9.96 | 9.86 | 9.86 | 17.48 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 20.61 | 4.69 | 11.62 | 4.59 | 4.79 | 5.86 | 9.77 | 10.16 | 9.86 | 11.33 | 9.86 | 12.21
95.488 | 5.66 | 0.1 | 0 | 0.68 | 0.1 | 0.68 | 4.88 | 18.07 | 4.88 | 18.95 | 4.79 | 16.11 | 9.67 | 18.85 | 9.57 | 9.96 | 10.16 | 9.86 | 17.48 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 21.19 | 4.69 | 12.6 | 4.59 | 4.79 | 5.47 | 9.77 | 10.06 | 9.86 | 10.35 | 9.86 | 10.84
95.52 | 6.25 | 0.1 | 0 | 0.59 | 0.1 | 0.59 | 4.88 | 7.81 | 4.88 | 18.95 | 4.79 | 16.6 | 9.67 | 19.24 | 9.57 | 9.96 | 10.25 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 21.88 | 4.69 | 13.38 | 4.59 | 4.79 | 5.37 | 9.77 | 10.35 | 9.86 | 9.96 | 9.86 | 10.06
95.552 | 6.84 | 0.1 | 0 | 0.49 | 0 | 0.59 | 4.79 | 2.73 | 4.88 | 18.56 | 4.79 | 17.09 | 9.67 | 19.82 | 9.57 | 9.86 | 9.86 | 9.86 | 17.38 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 22.36 | 4.79 | 14.16 | 4.59 | 4.79 | 5.18 | 9.77 | 10.45 | 9.86 | 10.25 | 9.86 | 10.06
95.584 | 7.42 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.79 | 3.13 | 4.88 | 18.36 | 4.79 | 17.58 | 9.67 | 20.12 | 9.57 | 9.86 | 9.86 | 9.86 | 8.3 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 22.85 | 4.69 | 14.84 | 4.59 | 4.79 | 5.08 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86 | 10.35
95.616 | 8.01 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.88 | 4.2 | 4.88 | 18.46 | 4.79 | 17.97 | 9.67 | 20.7 | 9.57 | 9.96 | 9.96 | 9.86 | 2.05 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.44 | 4.79 | 15.72 | 4.59 | 5.66 | 5.08 | 9.77 | 10.16 | 9.86 | 10.25 | 9.86 | 10.45
95.648 | 8.59 | 0.1 | 0 | 0.39 | 0.1 | 0.39 | 4.88 | 5.57 | 4.88 | 18.26 | 4.79 | 18.46 | 9.77 | 21 | 9.67 | 9.96 | 9.96 | 9.86 | 2.15 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.02 | 4.79 | 16.41 | 4.59 | 6.74 | 5.08 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 10.35
95.68 | 9.28 | 0.1 | 0.29 | 0.29 | 0 | 0.29 | 4.88 | 5.96 | 4.88 | 18.26 | 4.79 | 18.95 | 9.77 | 21.49 | 9.67 | 9.86 | 9.77 | 9.96 | 1.95 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.51 | 4.79 | 16.99 | 4.59 | 8.5 | 4.98 | 9.77 | 10.16 | 9.86 | 10.25 | 9.86 | 9.96
95.712 | 9.86 | 0.1 | 1.27 | 0.29 | 0.1 | 0.2 | 4.79 | 5.57 | 4.88 | 18.16 | 4.79 | 19.34 | 9.77 | 21.88 | 9.67 | 9.96 | 9.77 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 25.1 | 4.79 | 17.68 | 4.59 | 10.35 | 4.88 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 9.96
95.744 | 10.45 | 0.1 | 2.15 | 0.2 | 0.1 | 0.2 | 4.79 | 5.27 | 4.79 | 18.07 | 4.79 | 19.82 | 9.77 | 22.17 | 9.67 | 10.55 | 9.67 | 9.86 | 1.76 | 0.1 | 13.48 | 0.1 | 17.58 | 0.1 | 25.59 | 4.79 | 18.36 | 4.59 | 11.62 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
95.776 | 11.04 | 0.1 | 3.61 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 20.12 | 9.77 | 22.56 | 9.67 | 11.62 | 9.77 | 9.96 | 1.66 | 0.1 | 3.71 | 0.1 | 17.68 | 0.1 | 25.98 | 4.79 | 19.04 | 4.59 | 12.6 | 4.79 | 9.77 | 9.96 | 9.86 | 10.06 | 9.86 | 10.06
95.808 | 11.62 | 0.1 | 4.39 | 0.1 | 0.1 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.61 | 9.77 | 22.95 | 9.67 | 12.89 | 9.67 | 9.96 | 1.46 | 0.1 | 2.25 | 0.1 | 6.15 | 0.1 | 23.34 | 4.79 | 19.63 | 4.59 | 13.48 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
95.84 | 12.21 | 0.1 | 4.98 | 0.1 | 0.68 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 21 | 9.77 | 23.34 | 9.67 | 13.87 | 9.67 | 9.96 | 1.27 | 0.1 | 2.05 | 0.1 | 1.66 | 0.1 | 20.7 | 4.79 | 20.12 | 4.59 | 14.26 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
95.872 | 12.7 | 0.1 | 5.57 | 0.1 | 1.27 | 0.1 | 4.79 | 4.88 | 4.79 | 18.07 | 4.79 | 19.14 | 9.77 | 23.73 | 9.67 | 14.55 | 9.67 | 9.86 | 1.17 | 0.1 | 1.86 | 0 | 2.15 | 0.1 | 21 | 4.79 | 20.8 | 4.59 | 15.04 | 4.69 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86 | 9.86
95.904 | 13.18 | 0 | 6.05 | 0.1 | 2.54 | 0 | 5.47 | 4.79 | 4.79 | 11.91 | 4.79 | 18.95 | 9.77 | 24.12 | 9.67 | 15.14 | 9.67 | 9.86 | 1.07 | 0.1 | 1.66 | 0.1 | 1.95 | 0.1 | 21.29 | 4.79 | 21.39 | 4.69 | 15.82 | 4.59 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77
95.936 | 13.77 | 0.1 | 6.64 | 0.1 | 3.81 | 0.1 | 6.15 | 4.79 | 4.88 | 3.52 | 4.79 | 18.56 | 9.77 | 23.63 | 9.67 | 15.72 | 9.67 | 9.96 | 1.07 | 0.1 | 1.46 | 0.1 | 1.76 | 0.1 | 21.29 | 4.88 | 21.97 | 4.69 | 16.41 | 4.59 | 10.64 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77
95.968 | 14.26 | 0.1 | 7.23 | 0.1 | 4.59 | 0.1 | 7.23 | 4.69 | 4.88 | 3.22 | 4.88 | 18.36 | 9.77 | 21.58 | 9.77 | 16.21 | 9.67 | 9.96 | 0.88 | 0.1 | 1.46 | 0.1 | 1.66 | 0.1 | 21.29 | 4.88 | 22.56 | 4.69 | 17.19 | 4.59 | 11.62 | 9.77 | 9.86 | 9.77 | 10.25 | 9.77
96 | 14.75 | 0.1 | 7.81 | 0.1 | 5.18 | 0.1 | 8.5 | 4.69 | 4.88 | 3.71 | 4.88 | 18.46 | 9.77 | 21.68 | 9.77 | 16.7 | 9.57 | 9.96 | 0.88 | 0.1 | 1.37 | 0.1 | 1.46 | 0.1 | 21.29 | 4.88 | 23.05 | 4.69 | 17.68 | 4.49 | 13.28 | 9.67 | 10.16 | 9.77 | 11.33 | 9.77
96.032 | 15.33 | 0.1 | 8.4 | 0.1 | 5.57 | 0 | 9.38 | 4.59 | 4.79 | 4.98 | 4.79 | 18.26 | 9.77 | 21.58 | 9.67 | 17.09 | 9.57 | 9.86 | 0.59 | 0.1 | 1.27 | 0.1 | 1.27 | 0.1 | 21.29 | 4.88 | 23.63 | 4.69 | 18.46 | 4.49 | 15.14 | 9.67 | 11.43 | 9.67 | 12.89 | 9.67
96.064 | 15.82 | 0.1 | 8.98 | 0.1 | 6.25 | 0.1 | 10.06 | 4.59 | 4.79 | 5.86 | 4.79 | 18.16 | 9.77 | 21.49 | 9.77 | 17.58 | 9.57 | 9.86 | 0.59 | 0.1 | 1.07 | 0 | 1.27 | 0.1 | 21.29 | 4.79 | 24.22 | 4.79 | 19.24 | 4.49 | 16.31 | 9.67 | 12.79 | 9.67 | 14.84 | 9.67
96.096 | 16.31 | 0.1 | 9.67 | 0.1 | 6.84 | 0.1 | 10.64 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 9.86 | 21.58 | 9.77 | 18.16 | 9.67 | 9.96 | 0.59 | 0.1 | 1.07 | 0.1 | 1.17 | 0.1 | 21.29 | 4.88 | 24.61 | 4.79 | 19.82 | 4.59 | 17.19 | 9.67 | 14.94 | 9.67 | 16.21 | 9.67
96.128 | 16.8 | 0.1 | 10.25 | 0.1 | 7.42 | 0.1 | 11.33 | 4.59 | 4.88 | 5.27 | 4.88 | 18.16 | 9.86 | 21.49 | 9.77 | 18.56 | 9.67 | 9.96 | 0.49 | 0.1 | 0.98 | 0.1 | 1.07 | 0.1 | 21.29 | 4.88 | 25.29 | 4.79 | 20.31 | 4.59 | 18.07 | 9.67 | 16.21 | 9.67 | 17.19 | 9.67
96.16 | 17.29 | 0.1 | 10.84 | 0.1 | 8.01 | 0.1 | 11.91 | 4.59 | 4.88 | 5.08 | 4.88 | 18.16 | 9.77 | 21.49 | 9.77 | 19.04 | 9.67 | 9.96 | 0.49 | 0.1 | 0.88 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.68 | 4.79 | 20.9 | 4.59 | 18.75 | 9.67 | 17.19 | 9.67 | 17.97 | 9.67
96.192 | 15.43 | 0.1 | 11.43 | 0.1 | 8.59 | 0.1 | 12.5 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 19.43 | 9.57 | 9.86 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.19 | 4.79 | 26.17 | 4.79 | 21.49 | 4.49 | 19.53 | 9.67 | 17.97 | 9.67 | 18.65 | 9.67
96.224 | 15.23 | 0.1 | 11.91 | 0.1 | 9.28 | 0.1 | 12.99 | 4.49 | 4.79 | 4.88 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 19.82 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0 | 0.68 | 0.1 | 21.19 | 4.79 | 22.36 | 4.79 | 22.07 | 4.59 | 20.02 | 9.57 | 18.65 | 9.57 | 19.24 | 9.67
96.256 | 14.84 | 0.1 | 12.5 | 0.1 | 9.86 | 0.1 | 13.48 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.77 | 20.31 | 9.67 | 10.74 | 0.29 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 7.42 | 4.88 | 21.78 | 4.79 | 22.56 | 4.59 | 20.61 | 9.67 | 19.34 | 9.57 | 20.02 | 9.67
96.288 | 14.65 | 0.1 | 13.09 | 0.1 | 10.45 | 0.1 | 14.06 | 4.59 | 4.88 | 4.88 | 4.88 | 11.91 | 9.86 | 21.39 | 9.77 | 20.7 | 9.67 | 11.13 | 0.29 | 0.1 | 0.59 | 0.1 | 0.59 | 0.1 | 2.25 | 4.88 | 21.58 | 4.79 | 23.15 | 4.59 | 21.19 | 9.67 | 20.02 | 9.57 | 20.51 | 9.67
96.32 | 14.75 | 0.1 | 13.57 | 0.1 | 11.04 | 0.1 | 14.65 | 4.59 | 4.88 | 4.79 | 4.88 | 3.61 | 9.86 | 21.39 | 9.77 | 21.09 | 9.67 | 12.11 | 0.2 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.32 | 4.88 | 21.49 | 4.79 | 23.73 | 4.59 | 21.88 | 9.67 | 20.41 | 9.57 | 21.19 | 9.67
96.352 | 14.55 | 0.1 | 14.26 | 0.1 | 11.62 | 0 | 15.04 | 4.59 | 4.79 | 4.69 | 4.79 | 3.22 | 9.86 | 21.29 | 9.77 | 21.58 | 9.67 | 13.57 | 0.2 | 0.1 | 0.39 | 0.1 | 0.39 | 0.1 | 4 | 4.79 | 21.49 | 4.79 | 24.22 | 4.59 | 22.46 | 9.67 | 21 | 9.57 | 21.78 | 9.67
96.384 | 14.45 | 0.1 | 14.65 | 0.1 | 12.11 | 0.1 | 15.53 | 4.59 | 4.79 | 4.69 | 4.88 | 3.81 | 9.77 | 12.5 | 9.77 | 21.97 | 9.67 | 14.26 | 0.1 | 0.1 | 0.29 | 0.1 | 0.49 | 0.1 | 4.98 | 4.79 | 21.39 | 4.79 | 24.71 | 4.69 | 22.85 | 9.67 | 21.68 | 9.67 | 22.27 | 9.57
96.416 | 14.55 | 0.1 | 15.14 | 0.1 | 12.7 | 0.1 | 16.11 | 4.69 | 4.88 | 4.69 | 4.88 | 4.98 | 9.86 | 6.74 | 9.86 | 22.36 | 9.67 | 14.84 | 0.2 | 0.1 | 0.29 | 0.1 | 0.29 | 0.1 | 6.25 | 4.88 | 21.29 | 4.79 | 25.2 | 4.69 | 23.63 | 9.67 | 22.27 | 9.67 | 22.85 | 9.67
96.448 | 14.45 | 0.1 | 15.63 | 0.1 | 13.28 | 0.1 | 16.5 | 4.69 | 4.88 | 4.59 | 4.88 | 6.05 | 9.86 | 8.69 | 9.86 | 22.66 | 9.67 | 15.43 | 0.1 | 0.1 | 0.29 | 0.29 | 0.29 | 0.1 | 6.05 | 4.79 | 21.39 | 4.79 | 25.78 | 4.69 | 24.12 | 9.67 | 22.85 | 9.67 | 23.44 | 9.67
96.48 | 14.45 | 0.1 | 16.11 | 0.1 | 13.87 | 0.1 | 16.99 | 4.69 | 4.88 | 4.59 | 4.88 | 5.76 | 9.86 | 11.33 | 9.86 | 23.05 | 9.77 | 15.92 | 0.1 | 0.1 | 0.29 | 1.56 | 0.29 | 0.1 | 5.57 | 4.79 | 21.29 | 4.79 | 26.27 | 4.69 | 24.51 | 9.67 | 23.34 | 9.67 | 23.93 | 9.67
96.512 | 14.36 | 0.1 | 16.6 | 0.1 | 14.36 | 0 | 17.58 | 4.69 | 4.79 | 4.59 | 4.88 | 5.37 | 9.77 | 11.04 | 9.86 | 23.54 | 9.67 | 16.41 | 0.1 | 0.88 | 0.2 | 3.03 | 0.2 | 0.1 | 5.18 | 4.79 | 21.29 | 4.79 | 23.15 | 4.69 | 25.1 | 9.67 | 23.93 | 9.67 | 24.51 | 9.67
96.544 | 14.36 | 0.1 | 17.09 | 0.1 | 14.84 | 0.1 | 17.97 | 4.69 | 5.47 | 4.59 | 4.79 | 4.98 | 9.77 | 10.25 | 9.86 | 23.83 | 9.67 | 16.9 | 0 | 1.86 | 0.2 | 4.88 | 0.1 | 0.1 | 5.08 | 4.79 | 21.19 | 4.88 | 22.17 | 4.79 | 25.49 | 9.67 | 24.41 | 9.67 | 25 | 9.67
96.576 | 14.36 | 0.1 | 13.77 | 0.1 | 15.33 | 0.1 | 18.46 | 4.79 | 6.25 | 4.59 | 4.88 | 4.98 | 9.86 | 10.16 | 9.86 | 24.22 | 9.77 | 17.38 | 0.1 | 3.32 | 0.2 | 6.25 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.79 | 21.78 | 4.79 | 25.98 | 9.77 | 24.9 | 9.67 | 25.59 | 9.67
96.608 | 9.96 | 0.1 | 14.16 | 0.1 | 15.82 | 0.1 | 18.85 | 4.79 | 7.42 | 4.49 | 4.88 | 4.98 | 9.86 | 10.06 | 9.86 | 22.95 | 9.77 | 17.87 | 0.1 | 5.27 | 0.1 | 7.13 | 0.1 | 0.98 | 5.08 | 4.79 | 21.19 | 4.79 | 21.58 | 4.79 | 26.56 | 9.77 | 25.49 | 9.67 | 25.98 | 9.67
96.64 | 2.83 | 0.1 | 14.26 | 0.1 | 16.31 | 0.1 | 19.24 | 4.79 | 8.59 | 4.49 | 4.88 | 4.98 | 9.86 | 10.16 | 9.86 | 21.19 | 9.77 | 18.26 | 0.1 | 6.35 | 0.1 | 8.01 | 0.1 | 1.95 | 5.08 | 4.88 | 16.21 | 4.79 | 21.58 | 4.79 | 27.05 | 9.77 | 26.08 | 9.67 | 26.47 | 9.67
96.672 | 1.86 | 0.1 | 14.36 | 0.1 | 16.8 | 0 | 19.73 | 4.79 | 9.47 | 4.49 | 4.88 | 4.79 | 9.77 | 10.16 | 9.86 | 21.39 | 9.77 | 18.75 | 0 | 7.23 | 0.1 | 8.79 | 0.1 | 3.81 | 4.98 | 4.79 | 3.32 | 4.79 | 21.39 | 4.79 | 27.64 | 9.77 | 26.56 | 9.67 | 26.95 | 9.67
96.704 | 1.76 | 0.1 | 14.36 | 0.1 | 17.29 | 0.1 | 20.12 | 4.79 | 10.16 | 4.49 | 4.79 | 4.79 | 9.77 | 9.96 | 9.86 | 21.58 | 9.77 | 19.24 | 0 | 8.01 | 0 | 9.47 | 0.1 | 5.47 | 4.98 | 4.79 | 3.03 | 4.88 | 21.29 | 4.79 | 28.03 | 9.77 | 27.05 | 9.67 | 27.44 | 9.67
96.736 | 1.46 | 0.1 | 14.45 | 0.1 | 15.04 | 0.1 | 20.61 | 4.88 | 10.74 | 4.59 | 4.88 | 4.79 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 19.73 | 0.1 | 8.89 | 0.1 | 10.25 | 0.1 | 6.64 | 4.79 | 4.88 | 3.42 | 4.88 | 21.39 | 4.79 | 28.52 | 9.77 | 27.44 | 9.77 | 27.93 | 9.77
96.768 | 1.37 | 0.1 | 14.45 | 0.1 | 15.23 | 0.1 | 21 | 4.88 | 11.43 | 4.59 | 4.88 | 4.79 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.12 | 0.1 | 9.67 | 0.1 | 11.04 | 0.1 | 7.42 | 4.79 | 4.79 | 4.3 | 4.79 | 21.29 | 4.79 | 29.01 | 9.77 | 28.03 | 9.77 | 28.42 | 9.77
96.8 | 1.27 | 0.1 | 14.45 | 0.1 | 14.84 | 0.1 | 18.16 | 4.88 | 11.91 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.41 | 0.1 | 10.45 | 0.1 | 11.82 | 0.1 | 8.3 | 4.79 | 4.88 | 5.86 | 4.79 | 21.29 | 4.79 | 29.49 | 9.77 | 28.42 | 9.77 | 28.91 | 9.77
96.832 | 1.07 | 0.1 | 14.36 | 0.1 | 14.65 | 0 | 18.85 | 4.88 | 12.5 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 20.9 | 0 | 11.23 | 0 | 12.6 | 0 | 9.28 | 4.69 | 4.79 | 6.15 | 4.79 | 21.19 | 4.79 | 26.17 | 9.77 | 28.91 | 9.77 | 29.4 | 9.77
96.864 | 1.07 | 0 | 14.36 | 0.1 | 14.65 | 0.1 | 18.26 | 4.88 | 13.09 | 4.59 | 4.79 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 0 | 12.21 | 0 | 13.28 | 0 | 9.77 | 4.59 | 4.79 | 5.76 | 4.79 | 21.19 | 4.79 | 25.1 | 9.77 | 29.3 | 9.77 | 24.71 | 9.77
96.896 | 0.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.26 | 4.88 | 13.57 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.68 | 0.1 | 12.89 | 0.1 | 14.06 | 0.1 | 10.64 | 4.59 | 4.79 | 5.37 | 4.88 | 21.19 | 4.79 | 24.81 | 9.77 | 24.71 | 9.77 | 23.54 | 9.77
96.928 | 0.88 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.26 | 4.88 | 14.16 | 4.69 | 5.66 | 4.59 | 9.86 | 9.77 | 9.86 | 21.39 | 9.86 | 22.07 | 0.1 | 13.48 | 0.1 | 14.75 | 0.1 | 11.43 | 4.59 | 4.79 | 5.27 | 4.79 | 13.09 | 4.79 | 24.61 | 9.77 | 23.54 | 9.77 | 24.12 | 9.77
96.96 | 0.78 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.26 | 4.88 | 14.65 | 4.69 | 6.05 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 22.46 | 0.1 | 14.26 | 0.1 | 15.63 | 0.1 | 12.31 | 4.59 | 4.79 | 5.08 | 4.79 | 3.13 | 4.79 | 24.61 | 9.77 | 24.12 | 9.77 | 24.22 | 9.77
96.992 | 0.59 | 0.1 | 13.97 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.23 | 4.69 | 7.23 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 0 | 15.14 | 0 | 16.21 | 0 | 13.09 | 4.59 | 4.88 | 5.08 | 4.79 | 3.13 | 4.79 | 24.51 | 9.77 | 24.22 | 9.77 | 24.32 | 9.77
97.024 | 0.59 | 0.1 | 4.3 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.63 | 4.69 | 8.4 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.24 | 0 | 15.72 | 0 | 16.9 | 0 | 13.77 | 4.59 | 5.86 | 4.98 | 4.79 | 3.52 | 4.79 | 24.41 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
97.056 | 0.49 | 0.1 | 1.76 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.21 | 4.79 | 9.38 | 4.59 | 9.86 | 9.67 | 9.86 | 11.91 | 9.86 | 23.54 | 0.1 | 16.41 | 0.1 | 17.58 | 0.1 | 14.45 | 4.59 | 7.03 | 4.98 | 4.79 | 4.39 | 4.79 | 24.41 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
97.088 | 0.49 | 0.1 | 1.86 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.6 | 4.79 | 10.06 | 4.59 | 9.86 | 9.67 | 9.86 | 6.54 | 9.86 | 23.93 | 0.1 | 16.99 | 0.1 | 18.36 | 0.1 | 15.23 | 4.49 | 8.89 | 4.98 | 4.79 | 5.96 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86
97.12 | 0.39 | 0.1 | 1.56 | 0.1 | 14.36 | 0.1 | 18.07 | 4.88 | 17.09 | 4.79 | 10.64 | 4.59 | 10.55 | 9.67 | 9.86 | 8.79 | 9.86 | 24.41 | 0.1 | 17.58 | 0.1 | 18.85 | 0 | 15.92 | 4.59 | 10.55 | 4.98 | 4.79 | 6.25 | 4.79 | 24.41 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
97.152 | 0.29 | 0.1 | 1.37 | 0.1 | 10.94 | 0 | 18.07 | 4.88 | 17.58 | 4.79 | 11.33 | 4.59 | 11.72 | 9.57 | 9.86 | 11.33 | 9.86 | 21.58 | 0 | 18.46 | 0 | 19.53 | 0 | 16.8 | 4.49 | 12.01 | 4.79 | 4.79 | 5.66 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
97.184 | 0.29 | 0.1 | 1.27 | 0.1 | 3.03 | 0.1 | 18.07 | 4.88 | 18.07 | 4.79 | 11.91 | 4.59 | 12.89 | 9.57 | 9.86 | 11.13 | 9.86 | 21.78 | 0 | 18.95 | 0 | 20.12 | 0 | 17.29 | 4.59 | 12.7 | 4.69 | 4.79 | 5.27 | 4.79 | 24.22 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
97.216 | 0.29 | 0.1 | 1.07 | 0.1 | 2.05 | 0.1 | 14.16 | 4.88 | 18.46 | 4.88 | 12.5 | 4.69 | 13.87 | 9.67 | 9.86 | 10.35 | 9.86 | 21.49 | 0.1 | 19.73 | 0.1 | 20.8 | 0.1 | 17.97 | 4.59 | 13.67 | 4.69 | 4.79 | 5.18 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
97.248 | 0.2 | 0.1 | 1.07 | 0.1 | 1.86 | 0.1 | 3.91 | 4.88 | 18.95 | 4.88 | 12.99 | 4.69 | 14.55 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 0.1 | 20.22 | 0.1 | 21.39 | 0.1 | 18.46 | 4.59 | 14.36 | 4.69 | 4.79 | 5.27 | 4.79 | 18.75 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
97.28 | 0.2 | 0.1 | 0.98 | 0.1 | 1.66 | 0.1 | 3.13 | 4.88 | 19.34 | 4.88 | 13.48 | 4.69 | 15.14 | 9.67 | 9.86 | 10.06 | 9.86 | 21.58 | 0.1 | 20.8 | 0.1 | 21.88 | 0 | 19.14 | 4.59 | 15.14 | 4.59 | 4.79 | 5.18 | 4.79 | 7.23 | 9.77 | 24.22 | 9.86 | 23.93 | 9.86
97.312 | 0.1 | 0.1 | 0.78 | 0.1 | 1.37 | 0 | 3.61 | 4.88 | 19.82 | 4.88 | 14.06 | 4.69 | 15.63 | 9.57 | 9.86 | 10.16 | 9.86 | 21.58 | 0 | 21.39 | 0 | 22.46 | 0 | 19.82 | 4.59 | 16.02 | 4.59 | 4.79 | 4.98 | 4.79 | 6.93 | 9.77 | 21 | 9.77 | 8.4 | 9.86
97.344 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0 | 4.59 | 4.88 | 20.22 | 4.88 | 14.55 | 4.69 | 16.11 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0 | 22.07 | 0 | 17.77 | 0 | 20.41 | 4.69 | 16.6 | 4.59 | 4.79 | 5.08 | 4.79 | 10.74 | 9.77 | 8.79 | 9.77 | 6.35 | 9.86
97.376 | 0.1 | 0.59 | 0.68 | 0.1 | 1.17 | 0.1 | 5.76 | 4.88 | 20.61 | 4.88 | 15.14 | 4.79 | 16.7 | 9.67 | 10.45 | 10.16 | 9.96 | 21.49 | 0.1 | 22.56 | 0.1 | 16.8 | 0 | 21 | 4.69 | 17.38 | 4.59 | 4.79 | 4.98 | 4.79 | 12.11 | 9.77 | 6.35 | 9.86 | 9.96 | 9.86
97.408 | 0 | 1.27 | 0.59 | 0.1 | 1.07 | 0.1 | 5.76 | 4.88 | 21 | 4.88 | 15.63 | 4.79 | 17.09 | 9.67 | 11.43 | 9.96 | 9.86 | 21.49 | 0.1 | 18.65 | 0.1 | 17.58 | 0.1 | 21.58 | 4.69 | 17.87 | 4.59 | 4.79 | 4.88 | 4.79 | 10.84 | 9.86 | 9.96 | 9.86 | 12.31 | 9.86
97.44 | 0.1 | 2.34 | 0.49 | 0.1 | 0.88 | 0.1 | 5.47 | 4.88 | 19.14 | 4.88 | 16.11 | 4.79 | 17.68 | 9.67 | 12.79 | 9.96 | 9.96 | 21.49 | 0 | 18.16 | 0.1 | 17.68 | 0 | 22.17 | 4.69 | 18.56 | 4.59 | 4.79 | 4.79 | 4.79 | 9.96 | 9.77 | 12.31 | 9.86 | 11.13 | 9.86
97.472 | 0 | 3.71 | 0.39 | 0.1 | 0.88 | 0 | 5.08 | 4.88 | 18.85 | 4.88 | 16.6 | 4.79 | 18.07 | 9.67 | 13.77 | 9.86 | 9.86 | 21.39 | 0 | 18.07 | 0 | 17.68 | 0 | 22.75 | 4.79 | 19.24 | 4.49 | 4.98 | 4.69 | 4.79 | 10.16 | 9.77 | 10.94 | 9.77 | 9.96 | 9.86
97.504 | 0 | 4.39 | 0.29 | 0.1 | 0.78 | 0 | 4.88 | 4.88 | 18.56 | 4.88 | 16.9 | 4.79 | 18.46 | 9.67 | 14.45 | 9.77 | 9.86 | 21.39 | 0 | 17.87 | 0 | 17.68 | 0 | 19.73 | 4.79 | 19.82 | 4.49 | 6.05 | 4.69 | 4.79 | 10.45 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86
97.536 | 0.1 | 5.08 | 0.29 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.46 | 4.88 | 17.48 | 4.88 | 18.95 | 9.67 | 14.94 | 9.77 | 9.96 | 21.39 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 18.75 | 4.79 | 20.41 | 4.49 | 7.62 | 4.69 | 4.79 | 10.55 | 9.77 | 9.96 | 9.86 | 10.35 | 9.86
97.568 | 0.1 | 5.57 | 0.29 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.46 | 4.88 | 17.97 | 4.88 | 19.34 | 9.67 | 15.53 | 9.77 | 9.96 | 18.85 | 0.1 | 17.77 | 0.1 | 17.77 | 0 | 18.36 | 4.79 | 21.09 | 4.59 | 9.77 | 4.69 | 4.79 | 10.25 | 9.77 | 10.35 | 9.77 | 10.45 | 9.86
97.6 | 0.1 | 6.15 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 18.46 | 4.79 | 19.82 | 9.67 | 16.02 | 9.77 | 9.96 | 8.4 | 0.1 | 17.68 | 0.1 | 17.68 | 0 | 18.07 | 4.79 | 21.68 | 4.59 | 11.33 | 4.59 | 4.79 | 10.06 | 9.77 | 10.45 | 9.77 | 10.25 | 9.86
97.632 | 0 | 6.74 | 0.1 | 0.1 | 0.49 | 0 | 4.69 | 4.88 | 18.16 | 4.88 | 18.85 | 4.79 | 20.31 | 9.67 | 16.6 | 9.67 | 9.86 | 7.52 | 0 | 17.68 | 0 | 17.68 | 0 | 18.07 | 4.79 | 22.17 | 4.49 | 12.31 | 4.59 | 4.79 | 10.06 | 9.77 | 10.35 | 9.77 | 9.96 | 9.86
97.664 | 0 | 7.32 | 0.1 | 0.1 | 0.49 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.24 | 4.88 | 20.7 | 9.67 | 17.09 | 9.67 | 9.86 | 10.25 | 0 | 17.68 | 0 | 17.68 | 0 | 17.97 | 4.79 | 22.75 | 4.59 | 13.18 | 4.59 | 4.79 | 10.06 | 9.77 | 10.16 | 9.77 | 10.06 | 9.86
97.696 | 0.1 | 8.01 | 0.1 | 0.1 | 0.39 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.73 | 4.88 | 21.19 | 9.77 | 17.58 | 9.67 | 9.96 | 11.62 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 23.34 | 4.59 | 13.97 | 4.59 | 4.79 | 10.16 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
97.728 | 0.1 | 8.59 | 0 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.88 | 21.58 | 9.77 | 17.97 | 9.67 | 9.96 | 10.74 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 17.87 | 4.88 | 23.83 | 4.59 | 14.65 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
97.76 | 0.1 | 9.08 | 0 | 0.1 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.51 | 4.88 | 21.88 | 9.77 | 18.46 | 9.67 | 9.96 | 10.06 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 24.32 | 4.69 | 15.43 | 4.49 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86
97.792 | 0 | 9.77 | 0 | 0.88 | 0.2 | 0.1 | 4.49 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.36 | 9.77 | 18.95 | 9.67 | 9.86 | 9.86 | 0 | 17.58 | 0 | 4.49 | 0 | 17.77 | 4.88 | 24.81 | 4.59 | 16.11 | 4.49 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
97.824 | 0 | 10.35 | 0 | 1.66 | 0.2 | 0.78 | 4.49 | 4.88 | 18.07 | 4.88 | 18.16 | 4.88 | 22.66 | 9.77 | 19.34 | 9.67 | 9.86 | 10.25 | 0 | 12.6 | 0 | 1.76 | 0.1 | 17.68 | 4.88 | 25.29 | 4.69 | 16.8 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
97.856 | 0.1 | 10.84 | 0 | 3.03 | 0.1 | 1.56 | 4.59 | 4.88 | 10.74 | 4.88 | 18.75 | 4.88 | 23.05 | 9.77 | 19.82 | 9.67 | 9.96 | 10.16 | 0.1 | 3.42 | 0.1 | 2.34 | 0 | 17.77 | 4.88 | 25.88 | 4.69 | 17.58 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
97.888 | 0.1 | 11.52 | 0 | 4 | 0.1 | 2.83 | 4.49 | 4.88 | 3.52 | 4.88 | 18.26 | 4.88 | 23.44 | 9.77 | 20.22 | 9.57 | 9.96 | 9.96 | 0 | 2.25 | 0.1 | 2.05 | 0 | 17.77 | 4.88 | 25.88 | 4.79 | 18.16 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
97.92 | 0.1 | 12.01 | 0 | 4.69 | 0 | 3.91 | 4.49 | 4.88 | 3.22 | 4.88 | 18.26 | 4.88 | 23.83 | 9.77 | 20.61 | 9.67 | 10.55 | 9.86 | 0 | 2.15 | 0.1 | 1.86 | 0 | 10.64 | 4.88 | 22.46 | 4.79 | 18.85 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
97.952 | 0 | 12.6 | 0.1 | 5.27 | 0.1 | 4.69 | 4.49 | 4.88 | 3.91 | 4.88 | 18.26 | 4.88 | 24.22 | 9.77 | 21 | 9.67 | 11.33 | 9.86 | 0 | 1.66 | 0 | 1.56 | 0 | 2.25 | 4.88 | 21.49 | 4.79 | 19.43 | 4.59 | 5.96 | 9.67 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86
97.984 | 0 | 13.18 | 0 | 5.86 | 0.1 | 5.18 | 4.49 | 4.88 | 4.88 | 4.88 | 18.16 | 4.88 | 22.07 | 9.77 | 21.39 | 9.67 | 12.31 | 9.86 | 0 | 1.86 | 0 | 1.46 | 0 | 2.34 | 4.88 | 21.58 | 4.79 | 20.12 | 4.59 | 7.42 | 9.67 | 10.74 | 9.77 | 9.96 | 9.77 | 9.86
98.016 | 0.1 | 13.67 | 0.1 | 6.45 | 0.1 | 5.76 | 4.49 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 21 | 9.77 | 21.78 | 9.67 | 13.57 | 9.77 | 0.1 | 1.66 | 0.1 | 1.46 | 0 | 2.15 | 4.88 | 21.49 | 4.79 | 20.61 | 4.69 | 9.57 | 9.67 | 12.21 | 9.77 | 10.94 | 9.77 | 10.25
98.048 | 0.1 | 14.26 | 0.1 | 7.03 | 0.1 | 6.35 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.29 | 9.86 | 22.27 | 9.67 | 14.36 | 9.77 | 0.1 | 1.37 | 0.1 | 1.27 | 0 | 1.66 | 4.88 | 21.49 | 4.79 | 21.29 | 4.69 | 11.13 | 9.67 | 14.26 | 9.77 | 12.21 | 9.67 | 11.23
98.08 | 0.1 | 14.75 | 0 | 7.62 | 0.1 | 6.93 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 22.66 | 9.67 | 14.94 | 9.77 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.76 | 4.88 | 21.39 | 4.79 | 21.78 | 4.69 | 12.21 | 9.67 | 15.82 | 9.67 | 14.26 | 9.67 | 12.79
98.112 | 0 | 15.23 | 0 | 8.2 | 0.1 | 7.52 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 23.05 | 9.67 | 15.43 | 9.77 | 0 | 1.27 | 0 | 0.88 | 0 | 1.46 | 4.88 | 21.29 | 4.79 | 22.36 | 4.69 | 13.09 | 9.67 | 16.9 | 9.67 | 15.82 | 9.67 | 15.04
98.144 | 0 | 15.82 | 0 | 8.89 | 0.1 | 8.2 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.77 | 23.34 | 9.67 | 15.92 | 9.67 | 0 | 1.27 | 0 | 0.88 | 0 | 1.46 | 4.88 | 21.29 | 4.79 | 22.95 | 4.69 | 13.87 | 9.57 | 17.58 | 9.67 | 16.9 | 9.67 | 16.21
98.176 | 0.1 | 16.21 | 0.1 | 9.47 | 0.1 | 8.79 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.73 | 9.67 | 16.5 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0 | 1.37 | 4.88 | 21.29 | 4.79 | 23.54 | 4.79 | 14.65 | 9.67 | 18.46 | 9.67 | 17.68 | 9.67 | 17.09
98.208 | 0.1 | 16.7 | 0.1 | 10.06 | 0.1 | 9.38 | 4.69 | 5.47 | 4.88 | 4.88 | 17.97 | 4.88 | 21.39 | 9.86 | 24.12 | 9.67 | 16.99 | 9.67 | 0.1 | 0.98 | 0.1 | 0.78 | 0.1 | 1.27 | 4.88 | 21.19 | 4.88 | 24.02 | 4.79 | 15.43 | 9.67 | 19.14 | 9.67 | 18.46 | 9.67 | 17.87
98.24 | 0.1 | 17.19 | 0.1 | 10.55 | 0.1 | 9.96 | 4.69 | 6.54 | 4.79 | 4.88 | 14.06 | 4.88 | 21.39 | 9.86 | 23.15 | 9.67 | 17.48 | 9.67 | 0.1 | 0.78 | 0.1 | 0.78 | 0 | 1.07 | 4.88 | 21.19 | 4.79 | 24.41 | 4.79 | 16.11 | 9.67 | 19.82 | 9.67 | 19.14 | 9.67 | 18.65
98.272 | 0 | 14.45 | 0 | 11.23 | 0.1 | 10.55 | 4.69 | 7.81 | 4.79 | 4.88 | 3.61 | 4.88 | 21.39 | 9.77 | 21 | 9.67 | 17.87 | 9.57 | 1.07 | 0.78 | 0 | 0.59 | 0 | 0.98 | 4.88 | 21.19 | 4.79 | 25 | 4.79 | 16.8 | 9.67 | 20.41 | 9.57 | 19.82 | 9.57 | 19.34
98.304 | 0.1 | 14.55 | 0 | 14.06 | 0.1 | 13.48 | 4.79 | 11.62 | 4.59 | 4.88 | 5.86 | 4.88 | 10.55 | 9.86 | 21.39 | 9.77 | 20.12 | 9.67 | 7.42 | 0.49 | 0.88 | 0.29 | 0.1 | 0.59 | 4.88 | 3.91 | 4.79 | 21 | 4.88 | 20.12 | 9.67 | 23.34 | 9.67 | 22.75 | 9.67 | 22.27
98.336 | 0 | 14.55 | 0.1 | 14.55 | 0.1 | 13.97 | 4.79 | 12.21 | 4.59 | 4.88 | 5.27 | 4.88 | 6.64 | 9.86 | 21.39 | 9.77 | 20.61 | 9.67 | 8.2 | 0.39 | 1.76 | 0.29 | 0.2 | 0.59 | 4.88 | 4.98 | 4.79 | 21.19 | 4.88 | 20.7 | 9.67 | 23.83 | 9.67 | 23.24 | 9.67 | 22.85
98.368 | 0 | 14.55 | 0 | 15.04 | 0.1 | 14.55 | 4.79 | 12.79 | 4.49 | 4.88 | 5.08 | 4.88 | 9.47 | 9.86 | 21.39 | 9.77 | 20.9 | 9.67 | 9.08 | 0.29 | 3.32 | 0.2 | 1.46 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 21.29 | 9.77 | 24.32 | 9.67 | 23.83 | 9.67 | 23.54
98.4 | 0.1 | 14.36 | 0 | 15.63 | 0 | 15.04 | 4.79 | 13.38 | 4.49 | 4.88 | 4.88 | 4.88 | 11.62 | 9.77 | 21.39 | 9.77 | 21.39 | 9.57 | 9.67 | 0.29 | 5.27 | 0.2 | 2.83 | 0.39 | 4.88 | 5.96 | 4.79 | 21.29 | 4.79 | 21.88 | 9.67 | 24.81 | 9.67 | 24.32 | 9.67 | 23.93
98.432 | 0.1 | 14.45 | 0 | 16.02 | 0 | 15.43 | 4.88 | 13.87 | 4.49 | 4.88 | 4.88 | 4.88 | 10.84 | 9.77 | 21.39 | 9.77 | 21.78 | 9.57 | 10.45 | 0.2 | 6.45 | 0.2 | 4.59 | 0.39 | 4.88 | 5.57 | 4.88 | 21.19 | 4.79 | 22.46 | 9.67 | 25.29 | 9.67 | 24.81 | 9.67 | 24.51
98.464 | 0.1 | 14.45 | 0 | 16.5 | 0.1 | 16.02 | 4.88 | 14.45 | 4.49 | 5.47 | 4.88 | 4.88 | 10.06 | 9.86 | 21.39 | 9.86 | 22.17 | 9.67 | 11.33 | 0.2 | 7.23 | 0.1 | 6.15 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.88 | 23.05 | 9.77 | 25.88 | 9.77 | 25.39 | 9.77 | 24.9
98.496 | 0 | 14.45 | 0 | 16.99 | 0.1 | 16.5 | 4.88 | 14.94 | 4.49 | 6.25 | 4.88 | 4.88 | 9.86 | 9.86 | 21.39 | 9.86 | 22.56 | 9.67 | 12.11 | 0.2 | 8.11 | 0.1 | 7.03 | 0.29 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 23.54 | 9.77 | 26.47 | 9.77 | 25.88 | 9.77 | 25.49
98.528 | 0 | 14.36 | 0 | 15.92 | 0.1 | 16.99 | 4.88 | 15.43 | 4.49 | 7.42 | 4.79 | 4.88 | 10.16 | 9.86 | 21.39 | 9.86 | 22.95 | 9.67 | 12.89 | 0.2 | 8.79 | 0.1 | 7.81 | 0.2 | 4.88 | 5.18 | 4.88 | 21.29 | 4.88 | 24.02 | 9.77 | 26.95 | 9.77 | 26.37 | 9.77 | 25.98
98.56 | 0.1 | 10.35 | 0 | 14.06 | 0 | 15.82 | 4.88 | 15.92 | 4.49 | 8.59 | 4.69 | 4.88 | 10.16 | 9.77 | 10.64 | 9.86 | 23.24 | 9.67 | 13.57 | 0.2 | 9.67 | 0.1 | 8.69 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.61 | 9.77 | 27.44 | 9.77 | 26.86 | 9.77 | 26.56
98.592 | 0.1 | 2.93 | 0 | 14.45 | 0.1 | 14.55 | 4.79 | 16.41 | 4.59 | 9.47 | 4.59 | 4.88 | 10.06 | 9.77 | 6.54 | 9.86 | 23.63 | 9.67 | 14.36 | 0.1 | 10.45 | 0.1 | 9.47 | 0.2 | 4.79 | 4.98 | 4.88 | 21.19 | 4.88 | 25 | 9.77 | 27.83 | 9.77 | 27.34 | 9.77 | 26.95
98.624 | 0.1 | 1.95 | 0 | 14.65 | 0.1 | 14.65 | 4.88 | 16.9 | 4.59 | 10.16 | 4.59 | 4.88 | 9.86 | 9.86 | 9.08 | 9.86 | 24.12 | 9.67 | 14.94 | 0.1 | 11.23 | 0.1 | 10.25 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.68 | 9.77 | 28.22 | 9.77 | 27.74 | 9.77 | 27.44
98.656 | 0.1 | 1.76 | 0 | 14.55 | 0.1 | 14.75 | 4.88 | 17.38 | 4.69 | 10.74 | 4.59 | 4.88 | 9.96 | 9.86 | 11.52 | 9.86 | 24.41 | 9.67 | 15.72 | 0.1 | 11.91 | 0.1 | 11.04 | 0.1 | 4.88 | 4.88 | 4.88 | 20.9 | 4.88 | 26.17 | 9.77 | 28.81 | 9.77 | 28.22 | 9.77 | 28.03
98.688 | 0 | 1.56 | 0 | 14.55 | 0.1 | 14.55 | 4.88 | 17.77 | 4.69 | 11.33 | 4.49 | 4.88 | 9.86 | 9.86 | 11.04 | 9.86 | 22.56 | 9.67 | 16.41 | 0.1 | 12.7 | 0 | 11.72 | 0.1 | 4.88 | 4.79 | 4.88 | 4.79 | 4.88 | 21.49 | 9.77 | 29.2 | 9.77 | 28.71 | 9.86 | 28.42
98.72 | 0.1 | 1.37 | 0 | 14.45 | 0 | 14.45 | 4.88 | 18.26 | 4.69 | 12.01 | 4.49 | 4.88 | 9.86 | 9.77 | 10.16 | 9.86 | 22.27 | 9.67 | 17.09 | 0.1 | 13.67 | 0.1 | 12.6 | 0.1 | 4.88 | 4.79 | 4.79 | 3.03 | 4.88 | 21.29 | 9.77 | 24.61 | 9.77 | 29.3 | 9.77 | 28.91
98.752 | 0.1 | 1.46 | 0 | 14.45 | 0 | 14.45 | 4.88 | 18.75 | 4.69 | 12.5 | 4.49 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 21.88 | 9.67 | 17.68 | 0.1 | 14.45 | 0.1 | 13.28 | 0.1 | 4.79 | 4.69 | 4.88 | 3.42 | 4.88 | 21 | 9.77 | 23.34 | 9.77 | 24.41 | 9.77 | 29.4
98.784 | 0.1 | 1.07 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.69 | 13.09 | 4.49 | 4.88 | 9.77 | 9.86 | 10.06 | 9.86 | 21.78 | 9.77 | 18.36 | 0.1 | 15.04 | 0.1 | 13.97 | 0.1 | 4.98 | 4.69 | 4.88 | 4.2 | 4.88 | 21.19 | 9.77 | 24.02 | 9.86 | 23.24 | 9.86 | 25.39
98.816 | 0.1 | 1.07 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.79 | 13.57 | 4.49 | 4.88 | 9.77 | 9.86 | 10.16 | 9.86 | 21.78 | 9.67 | 19.14 | 0.1 | 15.72 | 0.1 | 14.75 | 0.1 | 5.86 | 4.69 | 4.88 | 5.66 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 23.93 | 9.86 | 24.32
98.848 | 0 | 0.98 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 20.02 | 4.79 | 14.06 | 4.49 | 5.47 | 9.67 | 9.86 | 10.16 | 9.86 | 21.58 | 9.77 | 19.73 | 0.1 | 16.5 | 0.1 | 15.53 | 0.1 | 7.32 | 4.59 | 4.88 | 6.35 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86 | 24.41
98.88 | 0.1 | 0.78 | 0 | 14.45 | 0 | 14.45 | 4.88 | 20.41 | 4.79 | 14.65 | 4.49 | 6.25 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.77 | 20.31 | 0.1 | 17.09 | 0.1 | 16.31 | 0.1 | 9.57 | 4.59 | 4.88 | 5.66 | 4.79 | 21.29 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.41
98.912 | 0.1 | 0.78 | 0 | 14.45 | 0.1 | 14.36 | 4.88 | 20.8 | 4.79 | 15.14 | 4.49 | 7.32 | 9.67 | 9.86 | 9.77 | 9.86 | 21.49 | 9.67 | 20.8 | 0.1 | 17.68 | 0.1 | 16.8 | 0.1 | 10.94 | 4.49 | 4.88 | 5.27 | 4.88 | 21.19 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.32
98.944 | 0.1 | 0.68 | 0 | 14.45 | 0.1 | 14.36 | 4.88 | 18.85 | 4.79 | 15.63 | 4.49 | 8.59 | 9.67 | 10.45 | 9.86 | 9.86 | 21.49 | 9.77 | 21.49 | 0.1 | 18.36 | 0.1 | 17.58 | 0.1 | 12.11 | 4.59 | 4.88 | 5.27 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
98.976 | 0.1 | 0.59 | 0 | 6.64 | 0.1 | 14.36 | 4.88 | 17.77 | 4.79 | 16.11 | 4.59 | 9.47 | 9.67 | 11.43 | 9.86 | 9.86 | 21.49 | 9.77 | 21.97 | 0.1 | 19.04 | 0.1 | 18.16 | 0.1 | 13.09 | 4.49 | 4.88 | 5.18 | 4.88 | 21.29 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
99.008 | 0 | 0.59 | 0 | 1.95 | 0.1 | 7.23 | 4.88 | 18.07 | 4.88 | 16.7 | 4.59 | 10.16 | 9.67 | 12.7 | 9.77 | 9.86 | 21.49 | 9.77 | 21.97 | 0.1 | 19.63 | 0.1 | 18.75 | 0.1 | 13.97 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86 | 24.32
99.04 | 0.1 | 0.49 | 0 | 1.95 | 0 | 1.86 | 4.88 | 18.16 | 4.79 | 17.09 | 4.59 | 10.74 | 9.57 | 13.77 | 9.67 | 9.86 | 21.39 | 9.77 | 17.77 | 0.1 | 20.31 | 0 | 19.43 | 0.1 | 14.75 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
99.072 | 0.1 | 0.39 | 0 | 1.76 | 0.1 | 1.86 | 4.88 | 18.16 | 4.88 | 17.58 | 4.59 | 11.33 | 9.57 | 14.45 | 9.67 | 9.86 | 21.39 | 9.77 | 17.97 | 0.1 | 20.8 | 0.1 | 20.12 | 0.1 | 15.43 | 4.49 | 4.79 | 4.98 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
99.104 | 0.29 | 0.29 | 0 | 1.46 | 0.1 | 1.95 | 4.88 | 18.16 | 4.88 | 18.07 | 4.69 | 12.01 | 9.67 | 15.04 | 9.67 | 9.86 | 15.14 | 9.77 | 17.87 | 0.1 | 21.49 | 0.1 | 20.7 | 0.1 | 16.11 | 4.59 | 4.88 | 4.98 | 4.88 | 20.9 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
99.136 | 1.27 | 0.29 | 0 | 1.46 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.46 | 4.69 | 12.5 | 9.67 | 15.43 | 9.67 | 9.86 | 7.42 | 9.77 | 17.77 | 0.1 | 22.07 | 0.1 | 21.39 | 0.1 | 16.7 | 4.59 | 4.79 | 4.98 | 4.79 | 4.88 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
99.168 | 2.25 | 0.29 | 0 | 1.17 | 0.1 | 1.56 | 4.88 | 18.16 | 4.88 | 18.95 | 4.69 | 13.09 | 9.67 | 16.11 | 9.67 | 9.86 | 7.62 | 9.86 | 17.68 | 0.1 | 22.56 | 0.1 | 21.88 | 0.1 | 17.58 | 4.59 | 4.88 | 4.79 | 4.79 | 2.93 | 9.86 | 10.45 | 9.86 | 24.12 | 9.86 | 24.12
99.2 | 3.61 | 0.2 | 0 | 1.17 | 0 | 1.17 | 4.88 | 18.16 | 4.88 | 19.34 | 4.69 | 13.67 | 9.57 | 16.6 | 9.67 | 9.86 | 10.94 | 9.77 | 17.58 | 0.1 | 19.43 | 0.1 | 22.36 | 0.1 | 18.26 | 4.59 | 5.76 | 4.79 | 4.79 | 3.42 | 9.77 | 6.05 | 9.86 | 8.79 | 9.86 | 17.77
99.232 | 4.49 | 0.1 | 0 | 0.98 | 0 | 1.07 | 4.88 | 18.16 | 4.88 | 19.73 | 4.69 | 14.16 | 9.67 | 16.99 | 9.57 | 9.86 | 11.43 | 9.77 | 17.58 | 0.1 | 18.46 | 0.1 | 17.68 | 0.1 | 18.75 | 4.59 | 6.64 | 4.69 | 4.79 | 4.2 | 9.77 | 9.28 | 9.77 | 6.05 | 9.86 | 7.62
99.264 | 4.98 | 0.1 | 0 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 14.65 | 9.67 | 17.58 | 9.67 | 9.96 | 10.45 | 9.86 | 17.58 | 0.1 | 18.07 | 0.1 | 16.7 | 0.1 | 19.43 | 4.59 | 8.2 | 4.79 | 4.79 | 5.66 | 9.77 | 12.4 | 9.86 | 9.57 | 9.86 | 6.54
99.296 | 5.47 | 0.1 | 0 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 15.23 | 9.67 | 17.97 | 9.57 | 9.86 | 9.96 | 9.86 | 17.58 | 0.1 | 17.97 | 0.1 | 17.38 | 0.1 | 20.12 | 4.59 | 10.16 | 4.69 | 4.79 | 6.45 | 9.86 | 11.33 | 9.86 | 12.21 | 9.86 | 10.45
99.328 | 6.05 | 0.1 | 0 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 21 | 4.79 | 15.63 | 9.67 | 18.46 | 9.57 | 9.86 | 9.96 | 9.86 | 17.48 | 0.1 | 17.97 | 0.1 | 17.58 | 0.1 | 20.7 | 4.69 | 11.52 | 4.59 | 4.88 | 5.76 | 9.77 | 10.25 | 9.86 | 11.23 | 9.86 | 12.21
99.36 | 6.64 | 0 | 0 | 0.68 | 0 | 0.68 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 16.21 | 9.67 | 18.95 | 9.57 | 9.86 | 10.25 | 9.86 | 17.38 | 0.1 | 17.87 | 0.1 | 17.58 | 0.1 | 21.29 | 4.69 | 12.6 | 4.59 | 4.79 | 5.37 | 9.77 | 10.06 | 9.86 | 10.06 | 9.86 | 10.84
99.392 | 7.23 | 0.1 | 0 | 0.59 | 0.1 | 0.59 | 4.79 | 7.32 | 4.88 | 18.95 | 4.79 | 16.7 | 9.67 | 19.34 | 9.57 | 9.86 | 10.16 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.58 | 0.1 | 21.78 | 4.69 | 13.38 | 4.59 | 4.79 | 5.18 | 9.77 | 10.35 | 9.77 | 9.86 | 9.86 | 10.16
99.424 | 7.91 | 0.1 | 0 | 0.59 | 0.1 | 0.59 | 4.88 | 2.64 | 4.88 | 18.56 | 4.79 | 17.09 | 9.67 | 19.82 | 9.67 | 9.86 | 9.96 | 9.86 | 17.38 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 22.46 | 4.69 | 14.26 | 4.59 | 4.88 | 5.18 | 9.77 | 10.55 | 9.86 | 10.35 | 9.86 | 9.96
99.456 | 8.5 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.88 | 3.32 | 4.88 | 18.36 | 4.79 | 17.58 | 9.67 | 20.22 | 9.67 | 9.86 | 9.86 | 9.86 | 8.4 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.05 | 4.69 | 14.94 | 4.59 | 4.88 | 5.18 | 9.77 | 10.45 | 9.86 | 10.45 | 9.96 | 10.45
99.488 | 8.98 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.88 | 4.3 | 4.88 | 18.46 | 4.79 | 17.97 | 9.77 | 20.7 | 9.67 | 9.86 | 9.86 | 9.86 | 2.05 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 23.54 | 4.79 | 15.72 | 4.49 | 5.76 | 5.18 | 9.86 | 10.25 | 9.86 | 10.25 | 9.86 | 10.45
99.52 | 9.67 | 0.1 | 0 | 0.29 | 0 | 0.29 | 4.88 | 5.57 | 4.88 | 18.16 | 4.79 | 18.56 | 9.77 | 21 | 9.67 | 9.86 | 9.86 | 9.86 | 2.15 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 24.02 | 4.69 | 16.41 | 4.49 | 6.84 | 4.98 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.25
99.552 | 10.35 | 0.1 | 0.39 | 0.29 | 0.1 | 0.29 | 4.88 | 5.96 | 4.88 | 18.16 | 4.79 | 18.85 | 9.77 | 21.39 | 9.67 | 9.86 | 9.77 | 9.86 | 1.95 | 0.1 | 17.68 | 0.1 | 17.58 | 0.1 | 24.51 | 4.79 | 17.09 | 4.49 | 8.5 | 4.88 | 9.77 | 10.16 | 9.86 | 9.96 | 9.86 | 10.16
99.584 | 10.74 | 0.1 | 1.17 | 0.29 | 0.1 | 0.29 | 4.88 | 5.57 | 4.88 | 18.16 | 4.88 | 19.34 | 9.77 | 21.88 | 9.67 | 9.96 | 9.77 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.1 | 4.79 | 17.77 | 4.59 | 10.35 | 4.88 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.16
99.616 | 11.43 | 0.1 | 2.15 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.82 | 9.77 | 22.27 | 9.67 | 10.55 | 9.77 | 9.86 | 1.66 | 0.1 | 13.18 | 0.1 | 17.68 | 0.1 | 25.59 | 4.79 | 18.46 | 4.59 | 11.62 | 4.79 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86 | 10.16
99.648 | 11.91 | 0.1 | 3.61 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.22 | 9.77 | 22.66 | 9.67 | 11.52 | 9.77 | 9.86 | 1.46 | 0.1 | 3.22 | 0.1 | 17.58 | 0.1 | 26.08 | 4.79 | 18.95 | 4.59 | 12.6 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96
99.68 | 12.5 | 0.1 | 4.39 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.7 | 9.77 | 23.15 | 9.67 | 12.99 | 9.67 | 9.86 | 1.27 | 0.1 | 2.25 | 0.1 | 5.27 | 0.1 | 21.88 | 4.79 | 19.53 | 4.59 | 13.57 | 4.69 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86 | 9.96
99.712 | 13.09 | 0.1 | 4.98 | 0.1 | 0.68 | 0.1 | 4.79 | 5.08 | 4.88 | 18.07 | 4.79 | 21 | 9.77 | 23.34 | 9.67 | 13.77 | 9.67 | 9.86 | 1.27 | 0.1 | 2.05 | 0.1 | 1.86 | 0.1 | 20.7 | 4.79 | 20.22 | 4.59 | 14.36 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
99.744 | 13.57 | 0.1 | 5.47 | 0.1 | 1.37 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.95 | 9.77 | 23.73 | 9.77 | 14.55 | 9.67 | 9.96 | 1.17 | 0.1 | 1.86 | 0.1 | 2.25 | 0.1 | 21.29 | 4.79 | 20.9 | 4.69 | 15.04 | 4.69 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
99.776 | 14.26 | 0.1 | 6.05 | 0.1 | 2.44 | 0.1 | 5.47 | 4.79 | 4.88 | 11.13 | 4.88 | 18.95 | 9.77 | 24.12 | 9.67 | 15.04 | 9.67 | 9.96 | 1.17 | 0.1 | 1.66 | 0.1 | 2.05 | 0.1 | 21.39 | 4.79 | 21.49 | 4.69 | 15.72 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
99.808 | 14.65 | 0.1 | 6.64 | 0.1 | 3.81 | 0 | 6.15 | 4.69 | 4.88 | 3.61 | 4.88 | 18.56 | 9.77 | 23.73 | 9.77 | 15.63 | 9.67 | 9.96 | 0.88 | 0.1 | 1.46 | 0 | 1.86 | 0.1 | 21.29 | 4.79 | 21.97 | 4.69 | 16.5 | 4.59 | 10.64 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
99.84 | 15.14 | 0.1 | 7.23 | 0.1 | 4.49 | 0.1 | 7.42 | 4.79 | 4.88 | 3.22 | 4.79 | 18.46 | 9.77 | 21.09 | 9.77 | 16.11 | 9.67 | 9.86 | 0.88 | 0.1 | 1.56 | 0.1 | 1.56 | 0.1 | 21.29 | 4.79 | 22.66 | 4.69 | 17.19 | 4.59 | 11.82 | 9.67 | 9.86 | 9.77 | 10.74 | 9.77
99.872 | 15.72 | 0.1 | 7.81 | 0.1 | 5.08 | 0 | 8.5 | 4.69 | 4.88 | 3.71 | 4.79 | 18.46 | 9.77 | 21.68 | 9.67 | 16.7 | 9.57 | 9.86 | 0.78 | 0.1 | 1.27 | 0.1 | 1.37 | 0.1 | 21.29 | 4.79 | 23.15 | 4.69 | 17.87 | 4.49 | 13.38 | 9.67 | 10.16 | 9.67 | 11.23 | 9.77
99.904 | 16.11 | 0.1 | 8.5 | 0.1 | 5.66 | 0.1 | 9.47 | 4.69 | 4.88 | 4.98 | 4.88 | 18.26 | 9.86 | 21.68 | 9.77 | 17.19 | 9.57 | 9.96 | 0.78 | 0.1 | 1.27 | 0.1 | 1.37 | 0.1 | 21.39 | 4.88 | 23.73 | 4.79 | 18.56 | 4.59 | 15.14 | 9.77 | 11.23 | 9.67 | 12.79 | 9.77
99.936 | 16.7 | 0.1 | 9.08 | 0.1 | 6.15 | 0.1 | 10.06 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 9.77 | 21.58 | 9.77 | 17.58 | 9.57 | 9.96 | 0.68 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 21.29 | 4.88 | 24.12 | 4.79 | 19.14 | 4.49 | 16.41 | 9.67 | 12.7 | 9.67 | 14.94 | 9.67
99.968 | 17.09 | 0.1 | 9.67 | 0.1 | 6.74 | 0.1 | 10.64 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 9.86 | 21.58 | 9.77 | 18.07 | 9.57 | 9.96 | 0.59 | 0.1 | 1.17 | 0 | 1.17 | 0.1 | 21.29 | 4.88 | 24.71 | 4.79 | 19.73 | 4.59 | 17.29 | 9.67 | 14.75 | 9.67 | 16.21 | 9.67
100 | 13.77 | 0 | 10.25 | 0.1 | 7.42 | 0.1 | 11.33 | 4.59 | 4.88 | 5.27 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.57 | 9.86 | 0.49 | 0.1 | 0.88 | 0.1 | 0.98 | 0.1 | 21.29 | 4.79 | 25.2 | 4.79 | 20.31 | 4.59 | 18.16 | 9.67 | 16.31 | 9.67 | 17.19 | 9.67
100.032 | 14.16 | 0.1 | 10.74 | 0.1 | 8.01 | 0 | 11.91 | 4.59 | 4.88 | 4.98 | 4.79 | 18.07 | 9.77 | 21.39 | 9.77 | 18.95 | 9.57 | 9.86 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.79 | 25.68 | 4.79 | 20.9 | 4.49 | 18.85 | 9.67 | 17.09 | 9.67 | 17.77 | 9.67
100.064 | 14.26 | 0.1 | 11.43 | 0.1 | 8.59 | 0.1 | 12.5 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 19.34 | 9.67 | 9.96 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 26.27 | 4.79 | 21.49 | 4.59 | 19.43 | 9.67 | 17.87 | 9.67 | 18.65 | 9.67
100.096 | 14.45 | 0.1 | 11.91 | 0.1 | 9.18 | 0.1 | 12.99 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 19.92 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 21.29 | 4.88 | 22.36 | 4.79 | 22.17 | 4.59 | 20.12 | 9.57 | 18.65 | 9.57 | 19.34 | 9.67
100.128 | 14.36 | 0.1 | 12.5 | 0.1 | 9.77 | 0.1 | 13.57 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 20.22 | 9.67 | 10.74 | 0.29 | 0.1 | 0.59 | 0.1 | 0.68 | 0.1 | 7.62 | 4.88 | 21.78 | 4.79 | 22.56 | 4.59 | 20.7 | 9.67 | 19.24 | 9.57 | 19.92 | 9.57
100.16 | 14.45 | 0.1 | 13.09 | 0.1 | 10.45 | 0.1 | 14.16 | 4.49 | 4.88 | 4.88 | 4.79 | 10.84 | 9.86 | 21.39 | 9.77 | 20.7 | 9.67 | 11.23 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 2.44 | 4.79 | 21.68 | 4.79 | 23.15 | 4.59 | 21.29 | 9.67 | 19.92 | 9.67 | 20.61 | 9.57
100.192 | 14.45 | 0 | 13.57 | 0.1 | 10.94 | 0 | 14.65 | 4.59 | 4.88 | 4.79 | 4.79 | 3.32 | 9.77 | 21.39 | 9.77 | 21 | 9.67 | 12.11 | 0.29 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.22 | 4.79 | 21.49 | 4.79 | 23.73 | 4.59 | 21.88 | 9.67 | 20.51 | 9.67 | 21.09 | 9.57
100.224 | 14.45 | 0.1 | 14.16 | 0.1 | 11.52 | 0.1 | 15.14 | 4.59 | 4.88 | 4.69 | 4.88 | 3.32 | 9.86 | 21.39 | 9.86 | 21.49 | 9.67 | 13.48 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.81 | 4.88 | 21.49 | 4.79 | 24.22 | 4.69 | 22.46 | 9.67 | 21.19 | 9.67 | 21.78 | 9.67
100.256 | 14.45 | 0.1 | 14.65 | 0.1 | 12.21 | 0.1 | 15.63 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 9.86 | 12.7 | 9.86 | 21.88 | 9.67 | 14.26 | 0.2 | 0.1 | 0.39 | 0.1 | 0.39 | 0.1 | 5.18 | 4.88 | 21.39 | 4.79 | 24.71 | 4.69 | 23.05 | 9.67 | 21.78 | 9.67 | 22.27 | 9.67
100.288 | 14.45 | 0.1 | 15.14 | 0.1 | 12.7 | 0.1 | 16.11 | 4.59 | 4.88 | 4.59 | 4.88 | 4.88 | 9.86 | 6.74 | 9.86 | 22.27 | 9.67 | 14.94 | 0.1 | 0.1 | 0.29 | 0.1 | 0.39 | 0.1 | 6.25 | 4.88 | 21.39 | 4.79 | 25.2 | 4.69 | 23.54 | 9.67 | 22.27 | 9.67 | 22.75 | 9.67
100.32 | 14.36 | 0.1 | 15.63 | 0.1 | 13.18 | 0 | 16.6 | 4.59 | 4.88 | 4.59 | 4.79 | 5.96 | 9.77 | 8.59 | 9.86 | 22.66 | 9.67 | 15.43 | 0.1 | 0.1 | 0.29 | 0.29 | 0.29 | 0.1 | 5.96 | 4.79 | 21.39 | 4.79 | 25.78 | 4.69 | 24.22 | 9.67 | 22.85 | 9.67 | 23.34 | 9.67
100.352 | 14.45 | 0.1 | 16.11 | 0.1 | 13.77 | 0 | 17.09 | 4.69 | 4.88 | 4.59 | 4.88 | 5.76 | 9.86 | 11.23 | 9.86 | 23.15 | 9.67 | 15.92 | 0.1 | 0.1 | 0.2 | 1.46 | 0.2 | 0.1 | 5.57 | 4.79 | 21.29 | 4.79 | 26.27 | 4.69 | 24.61 | 9.67 | 23.44 | 9.67 | 23.93 | 9.67
100.384 | 14.36 | 0.1 | 16.6 | 0.1 | 14.26 | 0.1 | 17.58 | 4.69 | 4.88 | 4.59 | 4.88 | 5.37 | 9.86 | 11.13 | 9.86 | 23.44 | 9.77 | 16.41 | 0.1 | 0.88 | 0.2 | 3.03 | 0.2 | 0.1 | 5.27 | 4.88 | 21.29 | 4.79 | 23.15 | 4.79 | 25.1 | 9.67 | 24.02 | 9.67 | 24.41 | 9.67
100.416 | 13.87 | 0.1 | 17.09 | 0.1 | 14.75 | 0.1 | 17.97 | 4.79 | 5.57 | 4.59 | 4.88 | 5.08 | 9.86 | 10.25 | 9.86 | 23.83 | 9.77 | 16.9 | 0.1 | 1.76 | 0.2 | 5.08 | 0.2 | 0.1 | 5.18 | 4.88 | 21.29 | 4.79 | 22.27 | 4.79 | 25.59 | 9.67 | 24.51 | 9.67 | 25 | 9.67
100.448 | 4.2 | 0.1 | 13.77 | 0.1 | 15.23 | 0.1 | 18.46 | 4.79 | 6.35 | 4.59 | 4.88 | 4.88 | 9.86 | 9.86 | 9.86 | 24.22 | 9.77 | 17.38 | 0.1 | 3.32 | 0.1 | 6.35 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 21.78 | 4.79 | 26.17 | 9.77 | 24.9 | 9.67 | 25.49 | 9.67
100.48 | 1.76 | 0 | 14.36 | 0.1 | 15.82 | 0.1 | 18.85 | 4.69 | 7.62 | 4.49 | 4.79 | 4.98 | 9.77 | 10.06 | 9.86 | 22.17 | 9.67 | 17.97 | 0 | 5.18 | 0.1 | 7.23 | 0.1 | 1.07 | 4.98 | 4.79 | 21.19 | 4.88 | 21.58 | 4.79 | 26.66 | 9.67 | 25.59 | 9.67 | 26.08 | 9.67
100.512 | 1.86 | 0 | 14.26 | 0.1 | 16.21 | 0.1 | 19.34 | 4.79 | 8.69 | 4.49 | 4.79 | 4.98 | 9.77 | 10.25 | 9.77 | 21.19 | 9.77 | 18.26 | 0.1 | 6.35 | 0.1 | 8.01 | 0.1 | 1.95 | 4.98 | 4.79 | 15.92 | 4.79 | 21.58 | 4.79 | 27.15 | 9.77 | 25.98 | 9.67 | 26.47 | 9.67
100.544 | 1.56 | 0.1 | 14.36 | 0.1 | 16.7 | 0.1 | 19.73 | 4.79 | 9.47 | 4.49 | 4.88 | 4.98 | 9.86 | 10.16 | 9.86 | 21.49 | 9.77 | 18.75 | 0.1 | 7.32 | 0.1 | 8.69 | 0.1 | 3.71 | 4.98 | 4.88 | 3.91 | 4.88 | 21.49 | 4.79 | 27.64 | 9.77 | 26.37 | 9.67 | 26.95 | 9.77
100.576 | 1.37 | 0.1 | 14.36 | 0.1 | 17.19 | 0.1 | 20.12 | 4.79 | 10.25 | 4.59 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 21.58 | 9.77 | 19.24 | 0.1 | 8.11 | 0.1 | 9.57 | 0.1 | 5.37 | 4.98 | 4.88 | 3.03 | 4.79 | 21.39 | 4.79 | 28.03 | 9.77 | 26.95 | 9.77 | 27.44 | 9.77
100.608 | 1.27 | 0.1 | 14.36 | 0.1 | 14.55 | 0 | 20.61 | 4.88 | 10.84 | 4.59 | 4.88 | 4.79 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 19.53 | 0.1 | 8.79 | 0.1 | 10.45 | 0.1 | 6.54 | 4.88 | 4.79 | 3.61 | 4.79 | 21.39 | 4.79 | 28.61 | 9.77 | 27.44 | 9.77 | 27.83 | 9.77
100.64 | 1.27 | 0 | 14.45 | 0.1 | 15.23 | 0.1 | 21 | 4.79 | 11.43 | 4.49 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.12 | 0 | 9.67 | 0.1 | 11.13 | 0 | 7.52 | 4.79 | 4.79 | 4.39 | 4.79 | 21.29 | 4.79 | 29.01 | 9.77 | 28.03 | 9.67 | 28.52 | 9.77
100.672 | 1.07 | 0.1 | 14.45 | 0.1 | 14.65 | 0 | 18.56 | 4.88 | 12.01 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.41 | 0 | 10.35 | 0 | 11.91 | 0 | 8.2 | 4.69 | 4.88 | 5.86 | 4.79 | 21.29 | 4.79 | 29.59 | 9.77 | 28.42 | 9.67 | 28.91 | 9.77
100.704 | 1.07 | 0.1 | 14.36 | 0.1 | 14.65 | 0.1 | 18.85 | 4.88 | 12.6 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 20.9 | 0.1 | 11.13 | 0.1 | 12.7 | 0.1 | 9.08 | 4.79 | 4.88 | 6.15 | 4.79 | 21.19 | 4.79 | 26.17 | 9.77 | 28.91 | 9.77 | 29.3 | 9.77
100.736 | 0.78 | 0.1 | 14.45 | 0.1 | 14.65 | 0.1 | 18.46 | 4.88 | 13.18 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.29 | 0.1 | 12.01 | 0.1 | 13.48 | 0.1 | 9.77 | 4.69 | 4.79 | 5.76 | 4.79 | 21.19 | 4.79 | 25.29 | 9.77 | 29.3 | 9.77 | 24.61 | 9.77
100.768 | 0.78 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 18.36 | 4.88 | 13.67 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.86 | 21.68 | 0.1 | 12.79 | 0.1 | 14.26 | 0.1 | 10.74 | 4.59 | 4.79 | 5.37 | 4.79 | 21.19 | 4.79 | 24.81 | 9.77 | 24.61 | 9.77 | 23.54 | 9.77
100.8 | 0.68 | 0 | 14.36 | 0.1 | 14.55 | 0 | 18.46 | 4.88 | 14.16 | 4.69 | 5.47 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.07 | 0 | 13.48 | 0 | 14.84 | 0 | 11.43 | 4.59 | 4.79 | 5.08 | 4.79 | 11.52 | 4.79 | 24.71 | 9.77 | 23.83 | 9.77 | 24.32 | 9.77
100.832 | 0.59 | 0.1 | 14.36 | 0.1 | 14.45 | 0 | 18.26 | 4.88 | 14.75 | 4.69 | 6.05 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.46 | 0.1 | 14.16 | 0 | 15.63 | 0 | 12.11 | 4.59 | 4.79 | 5.08 | 4.79 | 3.03 | 4.79 | 24.61 | 9.77 | 24.12 | 9.77 | 24.12 | 9.77
100.864 | 0.59 | 0.1 | 14.16 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.23 | 4.69 | 7.23 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 0.1 | 15.04 | 0.1 | 16.21 | 0.1 | 12.99 | 4.59 | 4.88 | 5.27 | 4.79 | 3.22 | 4.79 | 24.61 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
100.896 | 0.49 | 0.1 | 4.2 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.72 | 4.79 | 8.5 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 0.1 | 15.63 | 0.1 | 16.99 | 0.1 | 13.77 | 4.59 | 5.76 | 5.08 | 4.79 | 3.42 | 4.79 | 24.41 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86
100.928 | 0.29 | 0.1 | 1.76 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.21 | 4.79 | 9.47 | 4.49 | 9.77 | 9.67 | 9.86 | 11.91 | 9.86 | 23.63 | 0.1 | 16.41 | 0.1 | 17.58 | 0.1 | 14.45 | 4.59 | 7.13 | 5.08 | 4.79 | 4.49 | 4.79 | 24.41 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
100.96 | 0.39 | 0 | 2.05 | 0.1 | 14.36 | 0 | 18.16 | 4.88 | 16.7 | 4.79 | 10.06 | 4.59 | 9.86 | 9.57 | 9.86 | 6.54 | 9.86 | 23.93 | 0 | 17.09 | 0.1 | 18.26 | 0.1 | 15.14 | 4.49 | 9.18 | 5.08 | 4.79 | 6.05 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.32 | 9.86
100.992 | 0.29 | 0.1 | 1.56 | 0.1 | 14.36 | 0 | 18.07 | 4.88 | 17.09 | 4.79 | 10.55 | 4.59 | 10.55 | 9.57 | 9.77 | 8.69 | 9.86 | 24.32 | 0 | 17.68 | 0 | 18.95 | 0.1 | 15.92 | 4.49 | 10.74 | 4.79 | 4.79 | 6.15 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
101.024 | 0.29 | 0.1 | 1.46 | 0.1 | 11.72 | 0.1 | 18.07 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 11.52 | 9.67 | 9.86 | 11.43 | 9.86 | 21.58 | 0.1 | 18.26 | 0.1 | 19.53 | 0.1 | 16.6 | 4.59 | 12.01 | 4.88 | 4.79 | 5.76 | 4.79 | 24.32 | 9.86 | 24.22 | 9.77 | 24.32 | 9.86
101.056 | 0.2 | 0.1 | 1.17 | 0.1 | 3.32 | 0.1 | 18.07 | 4.88 | 18.07 | 4.79 | 11.82 | 4.59 | 12.89 | 9.67 | 9.86 | 11.13 | 9.86 | 22.07 | 0.1 | 18.95 | 0.1 | 20.12 | 0.1 | 17.19 | 4.59 | 12.99 | 4.79 | 4.79 | 5.37 | 4.79 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
101.088 | 0.1 | 0.1 | 1.17 | 0.1 | 2.05 | 0.1 | 13.57 | 4.88 | 18.56 | 4.79 | 12.5 | 4.69 | 13.87 | 9.57 | 9.86 | 10.25 | 9.86 | 21.58 | 0.1 | 19.53 | 0.1 | 20.7 | 0.1 | 17.97 | 4.59 | 13.87 | 4.79 | 4.79 | 5.18 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
101.12 | 0.1 | 0 | 0.88 | 0.1 | 1.76 | 0 | 3.52 | 4.88 | 18.95 | 4.79 | 12.99 | 4.69 | 14.55 | 9.57 | 9.86 | 9.86 | 9.86 | 21.68 | 0 | 20.12 | 0 | 21.39 | 0 | 18.65 | 4.59 | 14.55 | 4.69 | 4.79 | 5.18 | 4.79 | 15.53 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
101.152 | 0.1 | 0 | 0.88 | 0.1 | 1.56 | 0 | 3.22 | 4.88 | 19.34 | 4.79 | 13.48 | 4.69 | 15.04 | 9.57 | 9.77 | 9.96 | 9.86 | 21.58 | 0 | 20.8 | 0.1 | 21.97 | 0 | 19.34 | 4.59 | 15.23 | 4.59 | 4.79 | 5.18 | 4.79 | 6.84 | 9.77 | 24.22 | 9.77 | 23.93 | 9.86
101.184 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.71 | 4.88 | 19.82 | 4.88 | 14.06 | 4.69 | 15.63 | 9.67 | 9.86 | 10.25 | 9.96 | 21.58 | 0.1 | 21.39 | 0.1 | 22.56 | 0.1 | 19.92 | 4.69 | 16.02 | 4.59 | 4.79 | 5.08 | 4.79 | 7.03 | 9.86 | 24.02 | 9.77 | 8.59 | 9.86
101.216 | 0 | 0.1 | 0.78 | 0.1 | 1.37 | 0.1 | 4.69 | 4.88 | 20.12 | 4.88 | 14.65 | 4.79 | 16.11 | 9.67 | 9.86 | 10.16 | 9.96 | 21.58 | 0.1 | 21.97 | 0.1 | 17.97 | 0.1 | 20.51 | 4.69 | 16.7 | 4.59 | 4.79 | 5.08 | 4.79 | 11.13 | 9.86 | 8.98 | 9.77 | 6.35 | 9.86
101.248 | 0 | 0.68 | 0.68 | 0.1 | 1.17 | 0.1 | 5.76 | 4.88 | 20.7 | 4.88 | 15.04 | 4.79 | 16.6 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 0.1 | 22.56 | 0.1 | 17.29 | 0.1 | 21 | 4.69 | 17.38 | 4.59 | 4.79 | 4.98 | 4.79 | 12.01 | 9.86 | 6.15 | 9.77 | 10.16 | 9.86
101.28 | 0.1 | 1.37 | 0.59 | 0.1 | 1.07 | 0.1 | 5.76 | 4.88 | 21 | 4.79 | 15.53 | 4.79 | 17.09 | 9.67 | 11.72 | 9.86 | 9.86 | 21.49 | 0 | 19.34 | 0.1 | 17.58 | 0 | 21.58 | 4.69 | 17.97 | 4.59 | 4.79 | 4.88 | 4.79 | 10.55 | 9.77 | 10.25 | 9.77 | 12.21 | 9.86
101.312 | 0.1 | 2.44 | 0.49 | 0.1 | 0.88 | 0 | 5.37 | 4.88 | 19.14 | 4.88 | 16.02 | 4.79 | 17.58 | 9.67 | 12.89 | 9.86 | 9.86 | 21.49 | 0 | 18.36 | 0 | 17.68 | 0 | 22.27 | 4.69 | 18.56 | 4.59 | 4.79 | 4.88 | 4.79 | 9.96 | 9.77 | 12.31 | 9.77 | 11.04 | 9.86
101.344 | 0.1 | 3.81 | 0.49 | 0.1 | 0.78 | 0.1 | 5.08 | 4.88 | 18.95 | 4.88 | 16.5 | 4.79 | 18.07 | 9.67 | 13.87 | 9.86 | 9.96 | 21.49 | 0.1 | 17.97 | 0.1 | 17.68 | 0.1 | 22.75 | 4.79 | 19.34 | 4.49 | 4.98 | 4.79 | 4.79 | 10.35 | 9.86 | 11.13 | 9.86 | 10.16 | 9.86
101.376 | 0.1 | 4.49 | 0.39 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.56 | 4.88 | 16.99 | 4.79 | 18.46 | 9.67 | 14.65 | 9.86 | 9.96 | 21.39 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 19.73 | 4.79 | 19.82 | 4.59 | 6.05 | 4.69 | 4.79 | 10.45 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86
101.408 | 0 | 5.08 | 0.29 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.46 | 4.88 | 17.58 | 4.88 | 18.95 | 9.67 | 15.14 | 9.77 | 9.96 | 21.39 | 0.1 | 17.87 | 0.1 | 17.77 | 0.1 | 18.75 | 4.79 | 20.51 | 4.59 | 7.62 | 4.79 | 4.79 | 10.45 | 9.86 | 9.96 | 9.86 | 10.45 | 9.86
101.44 | 0.1 | 5.66 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.36 | 4.79 | 17.97 | 4.88 | 19.43 | 9.67 | 15.63 | 9.77 | 9.96 | 16.02 | 0 | 17.77 | 0.1 | 17.77 | 0 | 18.26 | 4.79 | 21.09 | 4.49 | 9.77 | 4.59 | 4.79 | 10.16 | 9.77 | 10.35 | 9.77 | 10.45 | 9.86
101.472 | 0.1 | 6.15 | 0.2 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.26 | 4.79 | 18.46 | 4.79 | 19.82 | 9.67 | 16.11 | 9.77 | 9.86 | 8.2 | 0 | 17.68 | 0 | 17.68 | 0 | 18.07 | 4.79 | 21.68 | 4.59 | 11.13 | 4.59 | 4.79 | 10.06 | 9.77 | 10.55 | 9.77 | 10.25 | 9.86
101.504 | 0.1 | 6.84 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.85 | 4.88 | 20.31 | 9.77 | 16.7 | 9.67 | 9.96 | 7.42 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 18.07 | 4.79 | 22.27 | 4.59 | 12.31 | 4.59 | 4.79 | 10.16 | 9.86 | 10.35 | 9.86 | 10.16 | 9.86
101.536 | 0 | 7.32 | 0.1 | 0.1 | 0.49 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.24 | 4.88 | 20.7 | 9.77 | 17.19 | 9.67 | 9.96 | 10.35 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.97 | 4.79 | 22.85 | 4.59 | 13.18 | 4.59 | 4.79 | 10.16 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86
101.568 | 0 | 8.01 | 0.1 | 0.1 | 0.39 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 19.82 | 4.88 | 21.09 | 9.77 | 17.58 | 9.67 | 9.96 | 11.72 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 23.34 | 4.69 | 13.97 | 4.59 | 4.79 | 10.25 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
101.6 | 0.1 | 8.59 | 0 | 0.1 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.79 | 20.22 | 4.88 | 21.58 | 9.77 | 18.07 | 9.67 | 9.86 | 10.55 | 0 | 17.58 | 0.1 | 17.68 | 0 | 17.77 | 4.79 | 23.93 | 4.59 | 14.75 | 4.49 | 4.79 | 9.86 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
101.632 | 0.1 | 9.18 | 0 | 0.1 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.61 | 4.88 | 21.97 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 0 | 17.58 | 0 | 17.48 | 0 | 17.77 | 4.79 | 24.41 | 4.69 | 15.43 | 4.49 | 4.79 | 9.86 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86
101.664 | 0.1 | 9.77 | 0 | 0.88 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 21 | 4.88 | 22.36 | 9.77 | 19.04 | 9.67 | 9.96 | 9.86 | 0.1 | 17.58 | 0.1 | 4.69 | 0.1 | 17.77 | 4.88 | 25 | 4.69 | 16.11 | 4.49 | 4.79 | 9.96 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
101.696 | 0.1 | 10.45 | 0.1 | 1.66 | 0.2 | 0.78 | 4.49 | 4.88 | 18.07 | 4.88 | 18.16 | 4.88 | 22.66 | 9.77 | 19.34 | 9.67 | 9.96 | 10.06 | 0.1 | 13.38 | 0.1 | 1.86 | 0.1 | 17.77 | 4.88 | 25.39 | 4.69 | 16.8 | 4.59 | 4.79 | 9.86 | 9.77 | 9.86 | 9.77 | 10.06 | 9.86
101.728 | 0 | 11.04 | 0 | 3.03 | 0.1 | 1.56 | 4.49 | 4.88 | 10.25 | 4.88 | 18.75 | 4.88 | 23.05 | 9.77 | 19.82 | 9.57 | 9.96 | 10.16 | 0.1 | 3.32 | 0.1 | 2.15 | 0.1 | 17.77 | 4.88 | 25.98 | 4.79 | 17.58 | 4.59 | 4.79 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
101.76 | 0.1 | 11.52 | 0 | 4 | 0 | 2.93 | 4.49 | 4.88 | 3.22 | 4.79 | 18.46 | 4.88 | 23.44 | 9.77 | 20.22 | 9.57 | 9.86 | 9.96 | 0 | 2.15 | 0.1 | 1.86 | 0 | 17.68 | 4.88 | 24.02 | 4.69 | 18.16 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
101.792 | 0.1 | 12.11 | 0 | 4.69 | 0 | 3.81 | 4.39 | 4.88 | 3.32 | 4.88 | 18.26 | 4.88 | 23.83 | 9.77 | 20.7 | 9.67 | 10.55 | 9.86 | 0.1 | 2.05 | 0 | 1.86 | 0.1 | 10.64 | 4.79 | 21.58 | 4.79 | 18.85 | 4.59 | 4.98 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
101.824 | 0.1 | 12.7 | 0.1 | 5.27 | 0.1 | 4.59 | 4.49 | 4.88 | 3.81 | 4.88 | 18.26 | 4.88 | 24.22 | 9.86 | 21.09 | 9.67 | 11.33 | 9.86 | 0.1 | 1.76 | 0.1 | 1.56 | 0.1 | 2.54 | 4.88 | 21.58 | 4.79 | 19.53 | 4.59 | 5.96 | 9.77 | 9.86 | 9.77 | 9.77 | 9.77 | 9.86
101.856 | 0 | 13.18 | 0.1 | 5.86 | 0.1 | 5.18 | 4.49 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 22.27 | 9.86 | 21.49 | 9.67 | 12.4 | 9.86 | 0.1 | 1.95 | 0.1 | 1.46 | 0.1 | 2.34 | 4.88 | 21.49 | 4.79 | 20.12 | 4.59 | 7.42 | 9.67 | 10.74 | 9.77 | 9.96 | 9.77 | 9.86
101.888 | 0 | 13.67 | 0 | 6.45 | 0.1 | 5.76 | 4.49 | 4.88 | 6.05 | 4.88 | 18.16 | 4.88 | 21 | 9.86 | 21.97 | 9.67 | 13.57 | 9.77 | 0.1 | 1.56 | 0.1 | 1.37 | 0.1 | 2.05 | 4.88 | 21.39 | 4.79 | 20.61 | 4.69 | 9.67 | 9.67 | 12.21 | 9.77 | 10.94 | 9.77 | 10.35
101.92 | 0.1 | 14.26 | 0 | 7.03 | 0.1 | 6.35 | 4.49 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 21.19 | 9.86 | 22.27 | 9.67 | 14.26 | 9.77 | 0 | 1.46 | 0.1 | 1.17 | 0 | 1.66 | 4.88 | 21.19 | 4.79 | 21.29 | 4.59 | 11.43 | 9.67 | 14.36 | 9.67 | 12.6 | 9.67 | 11.43
101.952 | 0.1 | 14.75 | 0 | 7.62 | 0.1 | 6.93 | 4.49 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 22.66 | 9.67 | 14.94 | 9.67 | 0 | 1.27 | 0 | 1.17 | 0 | 1.76 | 4.88 | 21.29 | 4.79 | 21.88 | 4.69 | 12.21 | 9.67 | 15.72 | 9.67 | 14.45 | 9.67 | 12.89
101.984 | 0.1 | 15.23 | 0.1 | 8.2 | 0.1 | 7.52 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 23.05 | 9.67 | 15.53 | 9.77 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.46 | 4.88 | 21.29 | 4.88 | 22.56 | 4.69 | 13.09 | 9.67 | 16.8 | 9.67 | 15.92 | 9.67 | 15.04
102.016 | 0 | 15.72 | 0 | 8.79 | 0.1 | 8.11 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.39 | 9.86 | 23.44 | 9.67 | 15.92 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.29 | 4.88 | 23.05 | 4.69 | 13.97 | 9.67 | 17.68 | 9.67 | 16.99 | 9.67 | 16.31
102.048 | 0 | 16.31 | 0.1 | 9.47 | 0.1 | 8.69 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.83 | 9.67 | 16.5 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.29 | 4.88 | 23.54 | 4.79 | 14.65 | 9.67 | 18.36 | 9.67 | 17.77 | 9.67 | 17.19
102.08 | 0.1 | 16.7 | 0 | 10.06 | 0 | 9.47 | 4.69 | 5.57 | 4.88 | 4.79 | 17.97 | 4.88 | 21.39 | 9.86 | 24.22 | 9.67 | 16.99 | 9.67 | 0 | 0.88 | 0.1 | 0.78 | 0 | 1.17 | 4.88 | 21.19 | 4.79 | 24.02 | 4.69 | 15.43 | 9.57 | 19.04 | 9.67 | 18.46 | 9.67 | 17.97
102.112 | 0.1 | 17.19 | 0 | 10.55 | 0.1 | 9.86 | 4.69 | 6.45 | 4.79 | 4.88 | 13.67 | 4.88 | 21.49 | 9.77 | 22.07 | 9.67 | 17.48 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0 | 1.17 | 4.88 | 21.19 | 4.88 | 24.51 | 4.79 | 16.11 | 9.67 | 19.73 | 9.67 | 19.04 | 9.57 | 18.65
102.144 | 0.1 | 14.55 | 0.1 | 11.23 | 0.1 | 10.55 | 4.69 | 7.81 | 4.79 | 4.88 | 4.2 | 4.88 | 21.39 | 9.86 | 21 | 9.67 | 17.87 | 9.67 | 1.07 | 0.78 | 0.1 | 0.68 | 0 | 1.07 | 4.88 | 21.19 | 4.88 | 25.1 | 4.79 | 16.8 | 9.67 | 20.22 | 9.57 | 19.73 | 9.67 | 19.34
102.176 | 0.1 | 15.14 | 0 | 11.72 | 0.1 | 11.13 | 4.69 | 8.89 | 4.69 | 4.88 | 3.13 | 4.88 | 21.39 | 9.86 | 21.29 | 9.77 | 18.46 | 9.57 | 1.95 | 0.78 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 21.19 | 4.88 | 25.59 | 4.79 | 17.48 | 9.67 | 20.9 | 9.67 | 20.31 | 9.67 | 20.02
102.208 | 0 | 14.65 | 0 | 12.31 | 0.1 | 11.72 | 4.79 | 9.67 | 4.69 | 4.88 | 3.71 | 4.88 | 21.39 | 9.86 | 21.49 | 9.77 | 18.75 | 9.57 | 3.71 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 8.59 | 4.88 | 25.98 | 4.79 | 18.07 | 9.67 | 21.58 | 9.67 | 21 | 9.67 | 20.61
102.24 | 0.1 | 14.65 | 0 | 12.89 | 0 | 12.4 | 4.79 | 10.35 | 4.59 | 4.79 | 4.88 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 19.24 | 9.57 | 5.37 | 0.49 | 0 | 0.49 | 0 | 0.78 | 4.88 | 2.34 | 4.79 | 21.49 | 4.79 | 18.85 | 9.67 | 22.17 | 9.67 | 21.68 | 9.67 | 21.19
102.272 | 0.1 | 14.65 | 0 | 13.38 | 0.1 | 12.89 | 4.79 | 10.94 | 4.59 | 4.79 | 5.76 | 4.88 | 21.39 | 9.77 | 21.39 | 9.77 | 19.63 | 9.57 | 6.45 | 0.49 | 0 | 0.39 | 0 | 0.68 | 4.88 | 3.13 | 4.88 | 20.22 | 4.79 | 19.34 | 9.67 | 22.75 | 9.67 | 22.17 | 9.67 | 21.78
102.304 | 0.1 | 14.55 | 0 | 13.97 | 0.1 | 13.38 | 4.79 | 11.62 | 4.59 | 4.88 | 5.76 | 4.88 | 11.52 | 9.86 | 21.39 | 9.77 | 20.12 | 9.57 | 7.32 | 0.49 | 0.78 | 0.29 | 0.1 | 0.59 | 4.88 | 3.71 | 4.88 | 20.9 | 4.79 | 20.02 | 9.67 | 23.24 | 9.67 | 22.66 | 9.67 | 22.27
102.336 | 0 | 14.55 | 0 | 14.45 | 0.1 | 13.87 | 4.79 | 12.21 | 4.59 | 4.88 | 5.37 | 4.88 | 6.54 | 9.86 | 21.39 | 9.77 | 20.51 | 9.57 | 8.11 | 0.39 | 1.76 | 0.29 | 0.2 | 0.59 | 4.88 | 4.79 | 4.88 | 21.19 | 4.79 | 20.7 | 9.77 | 23.73 | 9.67 | 23.24 | 9.67 | 22.85
102.368 | 0 | 14.45 | 0.1 | 14.94 | 0.1 | 14.45 | 4.79 | 12.7 | 4.49 | 4.88 | 5.08 | 4.88 | 9.08 | 9.86 | 21.39 | 9.77 | 20.9 | 9.57 | 8.79 | 0.29 | 3.32 | 0.2 | 1.46 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.79 | 21.19 | 9.77 | 24.32 | 9.67 | 23.83 | 9.67 | 23.34
102.4 | 0.1 | 14.45 | 0 | 15.53 | 0 | 14.94 | 4.79 | 13.28 | 4.49 | 4.79 | 4.88 | 4.88 | 11.43 | 9.77 | 21.39 | 9.77 | 21.39 | 9.57 | 9.67 | 0.29 | 5.18 | 0.2 | 2.83 | 0.49 | 4.88 | 5.96 | 4.79 | 21.19 | 4.79 | 21.78 | 9.67 | 24.81 | 9.67 | 24.41 | 9.67 | 23.93
102.432 | 0.1 | 14.45 | 0 | 16.02 | 0.1 | 15.43 | 4.88 | 13.77 | 4.49 | 4.79 | 4.88 | 4.88 | 10.94 | 9.77 | 21.39 | 9.77 | 21.68 | 9.57 | 10.45 | 0.2 | 6.45 | 0.2 | 4.59 | 0.39 | 4.88 | 5.57 | 4.88 | 21.19 | 4.79 | 22.36 | 9.67 | 25.29 | 9.67 | 24.81 | 9.67 | 24.41
102.464 | 0.1 | 14.45 | 0 | 16.41 | 0.1 | 15.92 | 4.88 | 14.36 | 4.49 | 5.47 | 4.88 | 4.88 | 10.16 | 9.86 | 21.39 | 9.86 | 22.17 | 9.67 | 11.33 | 0.2 | 7.32 | 0.2 | 6.05 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.79 | 22.85 | 9.77 | 25.78 | 9.77 | 25.29 | 9.77 | 25
102.496 | 0 | 14.45 | 0.1 | 16.99 | 0.1 | 16.41 | 4.88 | 14.84 | 4.49 | 6.25 | 4.88 | 4.88 | 9.86 | 9.86 | 21.39 | 9.77 | 22.56 | 9.67 | 12.21 | 0.2 | 8.11 | 0.1 | 7.03 | 0.29 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 23.44 | 9.77 | 26.37 | 9.77 | 25.88 | 9.77 | 25.49
102.528 | 0 | 14.36 | 0 | 16.9 | 0.1 | 16.9 | 4.88 | 15.33 | 4.49 | 7.32 | 4.79 | 4.88 | 10.16 | 9.86 | 21.39 | 9.86 | 22.95 | 9.67 | 12.79 | 0.2 | 8.89 | 0.1 | 7.81 | 0.29 | 4.88 | 5.18 | 4.88 | 21.29 | 4.79 | 24.02 | 9.77 | 26.76 | 9.77 | 26.37 | 9.77 | 25.98
102.56 | 0.1 | 11.04 | 0 | 14.55 | 0 | 16.99 | 4.88 | 15.82 | 4.49 | 8.59 | 4.69 | 4.88 | 10.25 | 9.86 | 11.04 | 9.77 | 23.24 | 9.67 | 13.67 | 0.1 | 9.67 | 0.1 | 8.59 | 0.2 | 4.88 | 5.08 | 4.79 | 21.19 | 4.79 | 24.51 | 9.77 | 27.25 | 9.77 | 26.95 | 9.77 | 26.47
102.592 | 0.1 | 3.22 | 0 | 14.84 | 0.1 | 15.33 | 4.79 | 16.31 | 4.59 | 9.47 | 4.59 | 4.88 | 10.06 | 9.86 | 6.54 | 9.86 | 23.63 | 9.67 | 14.26 | 0.1 | 10.45 | 0.1 | 9.38 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 24.9 | 9.77 | 27.74 | 9.77 | 27.25 | 9.77 | 26.95
102.624 | 0.1 | 2.05 | 0.1 | 14.75 | 0.1 | 15.04 | 4.88 | 16.8 | 4.59 | 10.16 | 4.59 | 4.88 | 9.96 | 9.86 | 8.98 | 9.86 | 24.02 | 9.67 | 15.04 | 0.1 | 11.33 | 0.1 | 10.16 | 0.2 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 25.59 | 9.77 | 28.22 | 9.77 | 27.74 | 9.77 | 27.34
102.656 | 0 | 1.95 | 0 | 14.65 | 0.1 | 14.84 | 4.88 | 17.29 | 4.59 | 10.74 | 4.59 | 4.88 | 9.86 | 9.86 | 11.43 | 9.86 | 24.41 | 9.67 | 15.72 | 0.1 | 12.01 | 0.1 | 10.94 | 0.1 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 26.08 | 9.77 | 28.71 | 9.77 | 28.13 | 9.86 | 27.83
102.688 | 0 | 1.46 | 0 | 14.65 | 0.1 | 14.65 | 4.88 | 17.77 | 4.69 | 11.43 | 4.59 | 4.88 | 9.96 | 9.86 | 11.04 | 9.86 | 22.27 | 9.67 | 16.5 | 0.1 | 12.79 | 0.1 | 11.72 | 0.1 | 4.88 | 4.88 | 4.88 | 5.96 | 4.79 | 21.49 | 9.77 | 29.3 | 9.77 | 28.71 | 9.86 | 28.42
102.72 | 0.1 | 1.46 | 0 | 14.45 | 0 | 14.65 | 4.88 | 18.16 | 4.69 | 12.01 | 4.49 | 4.88 | 9.86 | 9.77 | 10.16 | 9.86 | 22.27 | 9.67 | 17.09 | 0.1 | 13.57 | 0.1 | 12.5 | 0.1 | 4.88 | 4.79 | 4.79 | 2.93 | 4.79 | 20.51 | 9.77 | 24.61 | 9.77 | 29.3 | 9.86 | 28.91
102.752 | 0.1 | 1.27 | 0 | 14.45 | 0.1 | 14.55 | 4.88 | 18.56 | 4.69 | 12.5 | 4.49 | 4.88 | 9.77 | 9.86 | 9.86 | 9.86 | 21.78 | 9.67 | 17.58 | 0.1 | 14.26 | 0.1 | 13.18 | 0.1 | 4.88 | 4.69 | 4.88 | 3.32 | 4.79 | 20.9 | 9.77 | 23.63 | 9.77 | 24.61 | 9.86 | 29.3
102.784 | 0.1 | 1.27 | 0 | 14.55 | 0.1 | 14.45 | 4.88 | 19.14 | 4.79 | 13.09 | 4.49 | 4.88 | 9.77 | 9.86 | 10.06 | 9.86 | 21.78 | 9.77 | 18.46 | 0.1 | 14.94 | 0.1 | 13.97 | 0.1 | 4.88 | 4.69 | 4.88 | 4.1 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 23.24 | 9.86 | 24.81
102.816 | 0 | 1.17 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.79 | 13.57 | 4.49 | 4.88 | 9.77 | 9.86 | 10.25 | 9.86 | 21.68 | 9.77 | 19.04 | 0.1 | 15.63 | 0.1 | 14.75 | 0.1 | 5.86 | 4.59 | 4.88 | 5.47 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 23.93 | 9.86 | 23.63
102.848 | 0 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.92 | 4.79 | 14.16 | 4.49 | 5.57 | 9.77 | 9.86 | 10.16 | 9.86 | 21.58 | 9.77 | 19.53 | 0.1 | 16.31 | 0.1 | 15.43 | 0.1 | 7.23 | 4.59 | 4.88 | 6.25 | 4.79 | 21.09 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
102.88 | 0.1 | 0.78 | 0 | 14.45 | 0 | 14.45 | 4.88 | 20.31 | 4.79 | 14.65 | 4.49 | 6.25 | 9.67 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.12 | 0.1 | 17.19 | 0.1 | 16.31 | 0.1 | 9.38 | 4.59 | 4.79 | 5.86 | 4.79 | 21.19 | 9.77 | 24.12 | 9.86 | 24.12 | 9.86 | 24.22
102.912 | 0.1 | 0.78 | 0 | 14.36 | 0.1 | 14.36 | 4.88 | 20.8 | 4.79 | 15.14 | 4.49 | 7.23 | 9.67 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.9 | 0 | 17.77 | 0.1 | 16.8 | 0.1 | 10.74 | 4.49 | 4.79 | 5.37 | 4.88 | 21.19 | 9.77 | 24.32 | 9.86 | 24.12 | 9.86 | 24.22
102.944 | 0.1 | 0.68 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 20.31 | 4.79 | 15.63 | 4.49 | 8.59 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 9.77 | 21.49 | 0.1 | 18.46 | 0.1 | 17.48 | 0.1 | 12.01 | 4.59 | 4.88 | 5.27 | 4.79 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
102.976 | 0.1 | 0.59 | 0.1 | 7.52 | 0.1 | 14.36 | 4.88 | 17.87 | 4.79 | 16.11 | 4.59 | 9.47 | 9.67 | 11.33 | 9.86 | 9.86 | 21.49 | 9.77 | 21.97 | 0.1 | 19.04 | 0.1 | 18.16 | 0.1 | 12.99 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
103.008 | 0 | 0.49 | 0.1 | 2.05 | 0.1 | 8.2 | 4.88 | 18.36 | 4.88 | 16.7 | 4.59 | 10.06 | 9.67 | 12.6 | 9.77 | 9.86 | 21.49 | 9.77 | 22.46 | 0.1 | 19.63 | 0 | 18.75 | 0.1 | 13.87 | 4.49 | 4.88 | 5.08 | 4.79 | 21.19 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86 | 24.32
103.04 | 0.1 | 0.49 | 0 | 1.95 | 0 | 1.95 | 4.88 | 18.26 | 4.79 | 17.09 | 4.59 | 10.74 | 9.57 | 13.67 | 9.77 | 9.86 | 21.39 | 9.77 | 18.95 | 0.1 | 20.22 | 0.1 | 19.43 | 0.1 | 14.65 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
103.072 | 0.1 | 0.49 | 0 | 1.76 | 0.1 | 1.86 | 4.88 | 18.26 | 4.79 | 17.58 | 4.59 | 11.33 | 9.67 | 14.36 | 9.67 | 9.86 | 21.39 | 9.77 | 18.46 | 0.1 | 20.9 | 0.1 | 20.02 | 0.1 | 15.14 | 4.49 | 4.79 | 4.98 | 4.79 | 21.09 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
103.104 | 0.2 | 0.39 | 0 | 1.56 | 0.1 | 1.76 | 4.88 | 18.26 | 4.88 | 17.97 | 4.69 | 11.91 | 9.67 | 14.94 | 9.67 | 9.86 | 15.92 | 9.86 | 18.07 | 0.1 | 21.49 | 0.1 | 20.7 | 0.1 | 16.02 | 4.59 | 4.88 | 4.98 | 4.88 | 21.19 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86 | 24.12
103.136 | 1.07 | 0.29 | 0.1 | 1.37 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.46 | 4.69 | 12.4 | 9.57 | 15.53 | 9.67 | 9.86 | 7.62 | 9.86 | 17.87 | 0.1 | 22.07 | 0.1 | 21.19 | 0.1 | 16.7 | 4.49 | 4.88 | 4.98 | 4.79 | 6.35 | 9.86 | 24.22 | 9.86 | 24.12 | 9.96 | 24.12
103.168 | 2.05 | 0.29 | 0 | 1.27 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.85 | 4.69 | 12.99 | 9.67 | 16.02 | 9.67 | 9.86 | 7.62 | 9.86 | 17.87 | 0.1 | 22.56 | 0 | 21.68 | 0.1 | 17.38 | 4.59 | 4.88 | 4.79 | 4.79 | 2.64 | 9.77 | 11.43 | 9.86 | 24.12 | 9.96 | 24.12
103.2 | 3.42 | 0.2 | 0 | 1.17 | 0 | 1.27 | 4.88 | 18.16 | 4.88 | 19.34 | 4.69 | 13.57 | 9.57 | 16.5 | 9.67 | 9.86 | 10.74 | 9.86 | 17.68 | 0.1 | 19.24 | 0.1 | 22.46 | 0.1 | 18.07 | 4.59 | 5.76 | 4.79 | 4.79 | 3.32 | 9.77 | 5.66 | 9.86 | 9.38 | 9.86 | 21.19
103.232 | 4.3 | 0.2 | 0 | 0.98 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 19.82 | 4.69 | 13.97 | 9.67 | 17.09 | 9.67 | 9.86 | 11.43 | 9.86 | 17.58 | 0.1 | 18.26 | 0.1 | 18.26 | 0.1 | 18.65 | 4.59 | 6.45 | 4.69 | 4.79 | 4.1 | 9.77 | 8.59 | 9.77 | 5.76 | 9.86 | 8.4
103.264 | 4.88 | 0.1 | 0 | 0.88 | 0.1 | 0.98 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 14.65 | 9.67 | 17.58 | 9.57 | 9.86 | 10.55 | 9.86 | 17.58 | 0.1 | 17.97 | 0.1 | 16.8 | 0.1 | 19.34 | 4.69 | 8.01 | 4.69 | 4.79 | 5.47 | 9.77 | 12.21 | 9.86 | 9.38 | 9.96 | 6.45
103.296 | 5.47 | 0.1 | 0 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 15.14 | 9.67 | 17.97 | 9.57 | 9.86 | 9.96 | 9.86 | 17.58 | 0.1 | 17.87 | 0.1 | 17.58 | 0.1 | 20.02 | 4.69 | 9.96 | 4.69 | 4.79 | 6.35 | 9.77 | 11.52 | 9.86 | 12.31 | 9.96 | 10.06
103.328 | 5.96 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 15.53 | 9.67 | 18.46 | 9.57 | 9.86 | 9.96 | 9.86 | 17.58 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 20.7 | 4.69 | 11.43 | 4.69 | 4.79 | 5.86 | 9.86 | 10.25 | 9.86 | 11.23 | 9.96 | 12.21
103.36 | 6.64 | 0 | 0 | 0.59 | 0.1 | 0.68 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 16.11 | 9.67 | 18.85 | 9.57 | 9.86 | 10.25 | 9.86 | 17.48 | 0.1 | 17.87 | 0.1 | 17.58 | 0 | 21.19 | 4.69 | 12.5 | 4.59 | 4.79 | 5.37 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86 | 11.04
103.392 | 7.23 | 0 | 0 | 0.49 | 0 | 0.68 | 4.88 | 8.3 | 4.88 | 18.95 | 4.79 | 16.6 | 9.67 | 19.24 | 9.57 | 9.86 | 10.25 | 9.86 | 17.48 | 0.1 | 17.77 | 0.1 | 17.58 | 0.1 | 21.78 | 4.69 | 13.28 | 4.59 | 4.79 | 5.18 | 9.77 | 10.25 | 9.86 | 9.96 | 9.86 | 10.25
103.424 | 7.81 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 2.64 | 4.88 | 18.56 | 4.79 | 17.09 | 9.67 | 19.82 | 9.57 | 9.86 | 9.96 | 9.86 | 17.48 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 22.27 | 4.69 | 14.16 | 4.59 | 4.88 | 5.18 | 9.77 | 10.45 | 9.86 | 10.35 | 9.96 | 9.77
103.456 | 8.4 | 0.1 | 0 | 0.49 | 0.1 | 0.49 | 4.88 | 3.22 | 4.88 | 18.36 | 4.79 | 17.58 | 9.67 | 20.22 | 9.57 | 9.86 | 9.86 | 9.86 | 9.57 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.85 | 4.69 | 14.94 | 4.59 | 4.79 | 5.37 | 9.77 | 10.45 | 9.86 | 10.35 | 9.96 | 10.35
103.488 | 8.98 | 0.1 | 0 | 0.39 | 0.1 | 0.49 | 4.88 | 4.1 | 4.88 | 18.46 | 4.79 | 17.97 | 9.77 | 20.61 | 9.67 | 9.86 | 9.96 | 9.86 | 2.44 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 23.44 | 4.79 | 15.63 | 4.59 | 5.76 | 4.98 | 9.86 | 10.16 | 9.86 | 10.35 | 9.86 | 10.45
103.52 | 9.67 | 0 | 0 | 0.29 | 0 | 0.39 | 4.88 | 5.37 | 4.88 | 18.16 | 4.79 | 18.46 | 9.67 | 21 | 9.67 | 9.86 | 9.86 | 9.86 | 2.05 | 0.1 | 17.68 | 0.1 | 17.68 | 0 | 24.02 | 4.69 | 16.41 | 4.49 | 6.64 | 5.08 | 9.77 | 10.06 | 9.86 | 10.06 | 9.86 | 10.25
103.552 | 10.25 | 0.1 | 0.1 | 0.29 | 0.1 | 0.29 | 4.79 | 6.05 | 4.88 | 18.16 | 4.79 | 18.85 | 9.77 | 21.39 | 9.57 | 9.86 | 9.86 | 9.86 | 2.05 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 24.51 | 4.79 | 16.99 | 4.49 | 8.2 | 4.88 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 9.96
103.584 | 10.74 | 0.1 | 1.17 | 0.29 | 0.1 | 0.29 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 19.34 | 9.77 | 21.88 | 9.67 | 9.86 | 9.67 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25 | 4.79 | 17.77 | 4.59 | 10.16 | 4.88 | 9.77 | 10.16 | 9.86 | 10.16 | 9.96 | 10.16
103.616 | 11.43 | 0.1 | 2.05 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.07 | 4.79 | 19.73 | 9.77 | 22.27 | 9.67 | 10.55 | 9.77 | 9.86 | 1.76 | 0.1 | 13.97 | 0.1 | 17.68 | 0.1 | 25.49 | 4.79 | 18.26 | 4.59 | 11.62 | 4.88 | 9.77 | 10.16 | 9.86 | 9.96 | 9.96 | 10.16
103.648 | 11.91 | 0.1 | 3.42 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.12 | 9.77 | 22.66 | 9.67 | 11.43 | 9.77 | 9.86 | 1.46 | 0.1 | 3.52 | 0 | 17.68 | 0.1 | 25.88 | 4.79 | 18.95 | 4.59 | 12.6 | 4.69 | 9.77 | 9.96 | 9.86 | 9.96 | 9.96 | 9.96
103.68 | 12.5 | 0 | 4.3 | 0.1 | 0 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.79 | 20.61 | 9.77 | 23.05 | 9.67 | 12.7 | 9.67 | 9.86 | 1.37 | 0 | 2.44 | 0 | 5.76 | 0 | 23.93 | 4.79 | 19.63 | 4.59 | 13.38 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
103.712 | 13.09 | 0.1 | 4.88 | 0.1 | 0.59 | 0.1 | 4.79 | 4.98 | 4.88 | 18.07 | 4.79 | 20.9 | 9.77 | 23.34 | 9.67 | 13.67 | 9.67 | 9.86 | 1.56 | 0.1 | 2.05 | 0.1 | 1.56 | 0.1 | 21.09 | 4.79 | 20.22 | 4.59 | 14.26 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
103.744 | 13.57 | 0.1 | 5.47 | 0.1 | 1.37 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.46 | 9.77 | 23.73 | 9.67 | 14.45 | 9.67 | 9.96 | 1.07 | 0.1 | 1.76 | 0.1 | 2.25 | 0.1 | 21.49 | 4.88 | 20.9 | 4.69 | 14.94 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
103.776 | 14.16 | 0.1 | 5.96 | 0.1 | 2.34 | 0 | 5.66 | 4.88 | 4.88 | 12.11 | 4.88 | 18.85 | 9.77 | 24.12 | 9.67 | 15.04 | 9.67 | 9.96 | 1.17 | 0.1 | 1.86 | 0.1 | 2.15 | 0.1 | 21.49 | 4.79 | 21.49 | 4.69 | 15.63 | 4.59 | 9.77 | 9.77 | 9.86 | 9.77 | 9.96 | 9.86
103.808 | 14.65 | 0.1 | 6.54 | 0.1 | 3.71 | 0 | 6.15 | 4.69 | 4.88 | 3.71 | 4.88 | 18.46 | 9.77 | 24.12 | 9.77 | 15.53 | 9.67 | 9.96 | 0.88 | 0.1 | 1.56 | 0 | 1.66 | 0.1 | 21.39 | 4.79 | 21.97 | 4.69 | 16.31 | 4.59 | 10.64 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
103.84 | 15.14 | 0 | 7.23 | 0.1 | 4.49 | 0.1 | 7.32 | 4.69 | 4.88 | 3.13 | 4.79 | 18.36 | 9.77 | 21.68 | 9.67 | 16.11 | 9.67 | 9.86 | 0.88 | 0 | 1.37 | 0.1 | 1.66 | 0.1 | 21.29 | 4.79 | 22.56 | 4.69 | 17.09 | 4.59 | 11.72 | 9.67 | 9.86 | 9.77 | 10.25 | 9.77
103.872 | 15.63 | 0.1 | 7.72 | 0.1 | 5.08 | 0 | 8.5 | 4.69 | 4.88 | 3.71 | 4.79 | 18.26 | 9.77 | 21.88 | 9.67 | 16.5 | 9.57 | 9.86 | 0.88 | 0 | 1.27 | 0.1 | 1.46 | 0.1 | 21.29 | 4.79 | 23.05 | 4.69 | 17.68 | 4.49 | 13.18 | 9.67 | 10.06 | 9.67 | 11.04 | 9.77
103.904 | 16.11 | 0.1 | 8.4 | 0.1 | 5.57 | 0.1 | 9.38 | 4.69 | 4.88 | 4.88 | 4.88 | 18.26 | 9.77 | 21.78 | 9.77 | 17.09 | 9.57 | 9.86 | 0.78 | 0.1 | 1.37 | 0.1 | 1.37 | 0.1 | 21.29 | 4.88 | 23.63 | 4.79 | 18.36 | 4.59 | 15.14 | 9.67 | 11.04 | 9.67 | 12.5 | 9.77
103.936 | 16.7 | 0.1 | 8.98 | 0.1 | 6.05 | 0.1 | 10.06 | 4.59 | 4.88 | 5.96 | 4.88 | 18.16 | 9.77 | 21.68 | 9.77 | 17.58 | 9.57 | 9.86 | 0.68 | 0.1 | 1.27 | 0.1 | 1.37 | 0.1 | 21.29 | 4.88 | 24.12 | 4.79 | 18.95 | 4.49 | 16.31 | 9.67 | 12.6 | 9.67 | 14.55 | 9.67
103.968 | 17.09 | 0.1 | 9.57 | 0.1 | 6.74 | 0.1 | 10.64 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 9.86 | 21.58 | 9.77 | 17.97 | 9.57 | 9.96 | 0.59 | 0.1 | 1.07 | 0 | 1.17 | 0.1 | 21.29 | 4.79 | 24.71 | 4.79 | 19.73 | 4.59 | 17.29 | 9.67 | 14.65 | 9.67 | 16.02 | 9.67
104 | 13.67 | 0 | 10.16 | 0.1 | 7.32 | 0.1 | 11.33 | 4.59 | 4.88 | 5.27 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.57 | 9.86 | 0.49 | 0 | 0.98 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 25.1 | 4.79 | 20.22 | 4.59 | 17.97 | 9.67 | 16.21 | 9.57 | 17.09 | 9.67
104.032 | 13.97 | 0.1 | 10.74 | 0.1 | 8.01 | 0 | 11.91 | 4.59 | 4.88 | 5.08 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.85 | 9.57 | 9.86 | 0.49 | 0 | 0.78 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 25.78 | 4.79 | 20.8 | 4.49 | 18.75 | 9.67 | 17.09 | 9.67 | 17.87 | 9.67
104.064 | 14.16 | 0.1 | 11.43 | 0.1 | 8.5 | 0.1 | 12.5 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 19.34 | 9.67 | 9.86 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 26.17 | 4.79 | 21.49 | 4.59 | 19.34 | 9.67 | 17.77 | 9.67 | 18.56 | 9.67
104.096 | 14.45 | 0.1 | 11.91 | 0.1 | 9.08 | 0.1 | 13.09 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 19.82 | 9.67 | 9.86 | 0.39 | 0.1 | 0.68 | 0.1 | 0.88 | 0.1 | 21.19 | 4.88 | 21.78 | 4.79 | 21.97 | 4.59 | 20.12 | 9.67 | 18.56 | 9.57 | 19.24 | 9.67
104.128 | 14.36 | 0.1 | 12.5 | 0.1 | 9.67 | 0.1 | 13.57 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 20.22 | 9.67 | 10.74 | 0.29 | 0.1 | 0.59 | 0 | 0.78 | 0.1 | 8.01 | 4.88 | 20.9 | 4.79 | 22.66 | 4.59 | 20.7 | 9.67 | 19.24 | 9.67 | 19.82 | 9.67
104.16 | 14.45 | 0 | 12.99 | 0.1 | 10.35 | 0.1 | 14.16 | 4.49 | 4.88 | 4.88 | 4.79 | 12.31 | 9.77 | 21.39 | 9.77 | 20.61 | 9.67 | 11.13 | 0.2 | 0 | 0.59 | 0.1 | 0.59 | 0.1 | 2.34 | 4.79 | 21.39 | 4.79 | 23.15 | 4.59 | 21.19 | 9.57 | 19.92 | 9.67 | 20.41 | 9.67
104.192 | 14.45 | 0.1 | 13.57 | 0.1 | 10.94 | 0 | 14.55 | 4.59 | 4.88 | 4.79 | 4.88 | 3.61 | 9.86 | 21.39 | 9.77 | 21 | 9.67 | 12.11 | 0.2 | 0 | 0.49 | 0.1 | 0.59 | 0.1 | 3.22 | 4.79 | 21.29 | 4.79 | 23.54 | 4.59 | 21.88 | 9.67 | 20.51 | 9.57 | 21 | 9.67
104.224 | 14.45 | 0.1 | 14.06 | 0.1 | 11.52 | 0.1 | 15.14 | 4.59 | 4.88 | 4.69 | 4.88 | 3.22 | 9.86 | 21.39 | 9.86 | 21.39 | 9.67 | 13.48 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 4 | 4.88 | 21.29 | 4.79 | 24.22 | 4.69 | 22.36 | 9.67 | 21 | 9.67 | 21.68 | 9.67
104.256 | 14.45 | 0.1 | 14.55 | 0.1 | 12.11 | 0.1 | 15.63 | 4.59 | 4.88 | 4.69 | 4.88 | 3.81 | 9.86 | 12.89 | 9.86 | 21.88 | 9.67 | 14.16 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 5.08 | 4.88 | 21.29 | 4.88 | 24.71 | 4.69 | 23.05 | 9.67 | 21.68 | 9.57 | 22.27 | 9.67
104.288 | 14.36 | 0.1 | 15.04 | 0.1 | 12.6 | 0.1 | 16.11 | 4.59 | 4.88 | 4.59 | 4.88 | 4.79 | 9.86 | 6.84 | 9.86 | 22.27 | 9.67 | 14.84 | 0.2 | 0.1 | 0.29 | 0 | 0.39 | 0.1 | 6.05 | 4.88 | 21.29 | 4.79 | 25.29 | 4.69 | 23.54 | 9.67 | 22.17 | 9.67 | 22.75 | 9.67
104.32 | 14.36 | 0 | 15.63 | 0.1 | 13.18 | 0 | 16.7 | 4.59 | 4.88 | 4.59 | 4.79 | 5.96 | 9.77 | 8.59 | 9.86 | 22.66 | 9.67 | 15.33 | 0.1 | 0 | 0.29 | 0.29 | 0.29 | 0.1 | 6.05 | 4.79 | 21.29 | 4.79 | 25.68 | 4.69 | 24.12 | 9.67 | 22.75 | 9.67 | 23.34 | 9.67
104.352 | 14.45 | 0 | 16.11 | 0.1 | 13.67 | 0 | 17.09 | 4.69 | 4.88 | 4.59 | 4.79 | 5.76 | 9.86 | 11.23 | 9.77 | 23.05 | 9.67 | 15.82 | 0.1 | 0.1 | 0.2 | 1.56 | 0.2 | 0.1 | 5.57 | 4.79 | 21.29 | 4.79 | 26.17 | 4.69 | 24.61 | 9.67 | 23.34 | 9.67 | 23.93 | 9.67
104.384 | 14.36 | 0.1 | 16.6 | 0.1 | 14.26 | 0.1 | 17.58 | 4.69 | 4.88 | 4.49 | 4.88 | 5.47 | 9.86 | 11.23 | 9.86 | 23.44 | 9.77 | 16.41 | 0.1 | 0.78 | 0.2 | 2.83 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.88 | 22.36 | 4.79 | 25.2 | 9.77 | 23.93 | 9.67 | 24.51 | 9.67
104.416 | 14.16 | 0.1 | 16.99 | 0.1 | 14.75 | 0.1 | 17.97 | 4.69 | 5.57 | 4.59 | 4.88 | 4.98 | 9.86 | 10.25 | 9.86 | 23.73 | 9.77 | 16.8 | 0.1 | 1.76 | 0.2 | 4.88 | 0.2 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 21.78 | 4.69 | 25.59 | 9.67 | 24.32 | 9.67 | 24.9 | 9.67
104.448 | 4.98 | 0.1 | 13.77 | 0.1 | 15.23 | 0.1 | 18.46 | 4.69 | 6.25 | 4.49 | 4.88 | 4.88 | 9.77 | 9.86 | 9.86 | 24.12 | 9.77 | 17.29 | 0.1 | 3.32 | 0.2 | 6.35 | 0.2 | 0.1 | 5.08 | 4.88 | 21.29 | 4.88 | 21.58 | 4.79 | 26.08 | 9.77 | 24.9 | 9.67 | 25.49 | 9.67
104.48 | 1.76 | 0 | 13.97 | 0.1 | 15.72 | 0 | 18.95 | 4.69 | 7.52 | 4.49 | 4.79 | 4.88 | 9.77 | 10.16 | 9.86 | 23.73 | 9.77 | 17.87 | 0 | 5.27 | 0.1 | 7.13 | 0.1 | 0.88 | 5.08 | 4.79 | 21.19 | 4.79 | 21.49 | 4.69 | 26.66 | 9.67 | 25.49 | 9.67 | 25.98 | 9.67
104.512 | 1.95 | 0 | 14.06 | 0 | 16.21 | 0 | 19.34 | 4.79 | 8.59 | 4.49 | 4.79 | 4.88 | 9.77 | 10.25 | 9.86 | 21.97 | 9.77 | 18.26 | 0.1 | 6.45 | 0.1 | 7.81 | 0.1 | 1.76 | 5.18 | 4.79 | 19.73 | 4.79 | 21.49 | 4.69 | 27.25 | 9.77 | 25.88 | 9.67 | 26.47 | 9.67
104.544 | 1.66 | 0.1 | 14.45 | 0.1 | 16.7 | 0.1 | 19.82 | 4.79 | 9.57 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 21.88 | 9.77 | 18.65 | 0.1 | 7.23 | 0.1 | 8.69 | 0.1 | 3.42 | 4.98 | 4.88 | 4.3 | 4.79 | 21.39 | 4.79 | 27.54 | 9.77 | 26.47 | 9.67 | 26.86 | 9.67
104.576 | 1.46 | 0.1 | 14.36 | 0.1 | 17.19 | 0.1 | 20.12 | 4.79 | 10.25 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.78 | 9.77 | 19.24 | 0.1 | 8.11 | 0.1 | 9.57 | 0.1 | 5.37 | 4.88 | 4.88 | 3.13 | 4.88 | 21.29 | 4.79 | 28.13 | 9.77 | 26.95 | 9.67 | 27.44 | 9.67
104.608 | 1.27 | 0.1 | 14.36 | 0.1 | 14.06 | 0.1 | 20.7 | 4.79 | 10.74 | 4.59 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.68 | 9.77 | 19.53 | 0.1 | 8.79 | 0.1 | 10.25 | 0.1 | 6.54 | 4.88 | 4.88 | 3.42 | 4.79 | 21.39 | 4.79 | 28.52 | 9.77 | 27.44 | 9.77 | 27.83 | 9.77
104.64 | 1.17 | 0 | 14.45 | 0.1 | 14.84 | 0.1 | 21 | 4.79 | 11.43 | 4.49 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.12 | 0 | 9.67 | 0.1 | 11.13 | 0.1 | 7.42 | 4.79 | 4.79 | 4.2 | 4.79 | 21.29 | 4.79 | 29.01 | 9.77 | 27.93 | 9.67 | 28.42 | 9.77
104.672 | 1.17 | 0.1 | 14.36 | 0.1 | 14.45 | 0 | 18.26 | 4.88 | 12.01 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.77 | 21.49 | 9.77 | 20.41 | 0.1 | 10.35 | 0 | 11.91 | 0 | 8.11 | 4.69 | 4.79 | 5.57 | 4.79 | 21.29 | 4.79 | 29.49 | 9.77 | 28.32 | 9.77 | 28.81 | 9.77
104.704 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.85 | 4.88 | 12.5 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.86 | 20.9 | 0.1 | 11.23 | 0.1 | 12.6 | 0.1 | 8.98 | 4.69 | 4.88 | 6.35 | 4.88 | 21.29 | 4.79 | 25.98 | 9.77 | 28.91 | 9.77 | 29.3 | 9.77
104.736 | 0.88 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.46 | 4.79 | 13.09 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.29 | 0.1 | 11.91 | 0.1 | 13.38 | 0.1 | 9.77 | 4.69 | 4.79 | 5.86 | 4.79 | 21.19 | 4.79 | 25.1 | 9.77 | 29.4 | 9.77 | 24.71 | 9.77
104.768 | 0.68 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.36 | 4.88 | 13.57 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.68 | 0.1 | 12.7 | 0.1 | 14.26 | 0.1 | 10.55 | 4.59 | 4.88 | 5.47 | 4.79 | 21.19 | 4.79 | 24.81 | 9.77 | 24.51 | 9.77 | 23.34 | 9.77
104.8 | 0.68 | 0 | 14.36 | 0.1 | 14.55 | 0 | 18.46 | 4.88 | 14.16 | 4.69 | 5.57 | 4.49 | 9.77 | 9.77 | 9.86 | 21.39 | 9.86 | 22.07 | 0.1 | 13.48 | 0.1 | 14.94 | 0.1 | 11.43 | 4.59 | 4.79 | 5.18 | 4.79 | 14.55 | 4.79 | 24.71 | 9.77 | 23.44 | 9.77 | 24.02 | 9.77
104.832 | 0.59 | 0.1 | 14.36 | 0.1 | 14.45 | 0 | 18.26 | 4.88 | 14.65 | 4.69 | 5.96 | 4.49 | 9.77 | 9.67 | 9.86 | 21.49 | 9.86 | 22.46 | 0 | 14.16 | 0.1 | 15.53 | 0.1 | 12.21 | 4.59 | 4.79 | 5.08 | 4.79 | 3.91 | 4.79 | 24.61 | 9.77 | 23.93 | 9.77 | 24.12 | 9.77
104.864 | 0.59 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.14 | 4.69 | 7.03 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.85 | 0.1 | 14.94 | 0.1 | 16.31 | 0.1 | 12.89 | 4.59 | 4.79 | 5.18 | 4.79 | 3.13 | 4.79 | 24.51 | 9.86 | 24.12 | 9.77 | 24.22 | 9.77
104.896 | 0.49 | 0.1 | 5.57 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.72 | 4.79 | 8.5 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.24 | 0.1 | 15.53 | 0.1 | 16.99 | 0.1 | 13.67 | 4.59 | 5.96 | 5.18 | 4.79 | 3.52 | 4.79 | 24.41 | 9.86 | 24.22 | 9.77 | 24.22 | 9.77
104.928 | 0.39 | 0.1 | 1.66 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.21 | 4.79 | 9.38 | 4.59 | 9.77 | 9.67 | 9.86 | 12.99 | 9.86 | 23.54 | 0.1 | 16.31 | 0.1 | 17.58 | 0.1 | 14.36 | 4.59 | 6.93 | 5.08 | 4.79 | 4.3 | 4.79 | 24.41 | 9.86 | 24.22 | 9.77 | 24.22 | 9.77
104.96 | 0.39 | 0 | 1.86 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.6 | 4.79 | 9.96 | 4.49 | 9.86 | 9.67 | 9.86 | 6.74 | 9.86 | 24.02 | 0.1 | 17.09 | 0.1 | 18.16 | 0.1 | 15.14 | 4.59 | 8.98 | 4.88 | 4.79 | 5.96 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.32 | 9.77
104.992 | 0.29 | 0 | 1.76 | 0.1 | 14.36 | 0 | 18.07 | 4.88 | 17.09 | 4.79 | 10.55 | 4.59 | 10.45 | 9.57 | 9.86 | 8.2 | 9.86 | 24.32 | 0.1 | 17.68 | 0 | 18.85 | 0.1 | 15.72 | 4.49 | 10.55 | 4.88 | 4.79 | 6.15 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
105.024 | 0.29 | 0.1 | 1.46 | 0.1 | 13.87 | 0.1 | 18.16 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 11.43 | 9.67 | 9.86 | 11.13 | 9.86 | 21.39 | 0.1 | 18.26 | 0.1 | 19.53 | 0.1 | 16.6 | 4.59 | 11.91 | 4.88 | 4.79 | 5.76 | 4.79 | 24.32 | 9.86 | 24.22 | 9.77 | 24.32 | 9.86
105.056 | 0.2 | 0.1 | 1.27 | 0.1 | 3.81 | 0 | 18.07 | 4.88 | 18.07 | 4.79 | 11.82 | 4.59 | 12.7 | 9.57 | 9.86 | 11.13 | 9.86 | 21.78 | 0.1 | 18.95 | 0.1 | 20.12 | 0.1 | 17.19 | 4.59 | 12.79 | 4.79 | 4.79 | 5.37 | 4.79 | 24.32 | 9.86 | 24.32 | 9.77 | 24.22 | 9.86
105.088 | 0.1 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 14.36 | 4.88 | 18.46 | 4.79 | 12.4 | 4.69 | 13.67 | 9.57 | 9.86 | 10.35 | 9.96 | 21.49 | 0.1 | 19.63 | 0.1 | 20.7 | 0.1 | 17.87 | 4.59 | 13.67 | 4.69 | 4.79 | 5.18 | 4.79 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86
105.12 | 0.1 | 0 | 1.07 | 0 | 1.86 | 0 | 3.91 | 4.88 | 18.95 | 4.79 | 12.99 | 4.69 | 14.36 | 9.57 | 9.86 | 9.77 | 9.86 | 21.58 | 0 | 20.12 | 0.1 | 21.29 | 0 | 18.56 | 4.59 | 14.45 | 4.69 | 4.79 | 5.08 | 4.79 | 16.6 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
105.152 | 0.1 | 0 | 0.88 | 0.1 | 1.56 | 0 | 3.13 | 4.88 | 19.34 | 4.79 | 13.48 | 4.69 | 15.04 | 9.57 | 9.86 | 9.96 | 9.86 | 21.58 | 0.1 | 20.8 | 0.1 | 21.88 | 0 | 19.24 | 4.59 | 15.14 | 4.69 | 4.79 | 5.08 | 4.79 | 7.03 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
105.184 | 0.1 | 0.1 | 0.88 | 0.1 | 1.56 | 0.1 | 3.52 | 4.88 | 19.82 | 4.88 | 14.06 | 4.69 | 15.53 | 9.67 | 9.86 | 10.16 | 9.96 | 21.49 | 0.1 | 21.29 | 0.1 | 22.36 | 0.1 | 19.82 | 4.59 | 15.82 | 4.69 | 4.79 | 5.08 | 4.79 | 6.84 | 9.86 | 24.22 | 9.86 | 10.45 | 9.86
105.216 | 0.1 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 4.59 | 4.88 | 20.22 | 4.88 | 14.55 | 4.69 | 16.11 | 9.67 | 9.86 | 10.16 | 9.96 | 21.49 | 0.1 | 21.88 | 0.1 | 17.68 | 0.1 | 20.31 | 4.69 | 16.41 | 4.59 | 4.79 | 5.08 | 4.79 | 10.74 | 9.77 | 9.86 | 9.77 | 5.96 | 9.86
105.248 | 0 | 0.39 | 0.68 | 0.1 | 1.17 | 0.1 | 5.96 | 4.88 | 20.7 | 4.88 | 15.04 | 4.69 | 16.6 | 9.67 | 10.45 | 9.96 | 9.96 | 21.49 | 0.1 | 22.46 | 0.1 | 16.9 | 0.1 | 20.9 | 4.69 | 17.19 | 4.59 | 4.79 | 4.98 | 4.79 | 12.11 | 9.86 | 6.05 | 9.77 | 9.47 | 9.86
105.28 | 0.1 | 1.27 | 0.59 | 0.1 | 1.07 | 0 | 5.76 | 4.88 | 21.09 | 4.79 | 15.63 | 4.79 | 17.09 | 9.67 | 11.62 | 9.86 | 9.86 | 21.49 | 0 | 19.14 | 0.1 | 17.38 | 0.1 | 21.58 | 4.69 | 17.97 | 4.59 | 4.79 | 4.88 | 4.79 | 10.55 | 9.77 | 10.16 | 9.77 | 12.31 | 9.86
105.312 | 0.1 | 2.25 | 0.49 | 0 | 0.88 | 0 | 5.27 | 4.88 | 19.14 | 4.88 | 16.02 | 4.79 | 17.58 | 9.67 | 12.7 | 9.86 | 9.86 | 21.49 | 0 | 18.16 | 0 | 17.58 | 0.1 | 22.17 | 4.69 | 18.46 | 4.59 | 4.79 | 4.88 | 4.79 | 9.96 | 9.77 | 12.4 | 9.77 | 11.23 | 9.86
105.344 | 0.1 | 3.71 | 0.49 | 0.1 | 0.78 | 0.1 | 5.08 | 4.88 | 18.95 | 4.88 | 16.5 | 4.79 | 18.07 | 9.67 | 13.77 | 9.86 | 9.96 | 21.39 | 0.1 | 17.97 | 0.1 | 17.68 | 0.1 | 22.75 | 4.79 | 19.24 | 4.59 | 4.98 | 4.79 | 4.79 | 10.25 | 9.86 | 11.23 | 9.77 | 10.16 | 9.86
105.376 | 0.1 | 4.49 | 0.39 | 0.1 | 0.88 | 0.1 | 4.98 | 4.88 | 18.56 | 4.88 | 16.99 | 4.79 | 18.46 | 9.67 | 14.45 | 9.86 | 9.96 | 21.39 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 19.34 | 4.79 | 19.82 | 4.59 | 5.96 | 4.79 | 4.79 | 10.55 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86
105.408 | 0 | 5.08 | 0.39 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.46 | 4.88 | 17.48 | 4.79 | 18.95 | 9.67 | 15.14 | 9.77 | 9.96 | 21.39 | 0.1 | 17.87 | 0.1 | 17.68 | 0.1 | 18.56 | 4.79 | 20.41 | 4.59 | 7.52 | 4.69 | 4.79 | 10.45 | 9.77 | 9.96 | 9.86 | 10.45 | 9.86
105.44 | 0.1 | 5.57 | 0.29 | 0.1 | 0.68 | 0.1 | 4.79 | 4.88 | 18.36 | 4.88 | 17.87 | 4.79 | 19.34 | 9.67 | 15.53 | 9.77 | 9.96 | 17.29 | 0 | 17.77 | 0.1 | 17.68 | 0.1 | 18.26 | 4.79 | 20.9 | 4.59 | 9.67 | 4.59 | 4.79 | 10.16 | 9.77 | 10.35 | 9.77 | 10.55 | 9.86
105.472 | 0.1 | 6.05 | 0.2 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.26 | 4.88 | 18.36 | 4.79 | 19.82 | 9.67 | 16.02 | 9.67 | 9.86 | 8.59 | 0 | 17.68 | 0 | 17.68 | 0 | 18.07 | 4.79 | 21.58 | 4.59 | 11.13 | 4.59 | 4.79 | 10.06 | 9.77 | 10.45 | 9.77 | 10.35 | 9.86
105.504 | 0.1 | 6.64 | 0.2 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.85 | 4.88 | 20.22 | 9.77 | 16.6 | 9.67 | 9.96 | 7.13 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 18.07 | 4.88 | 22.17 | 4.59 | 12.21 | 4.59 | 4.79 | 10.16 | 9.86 | 10.35 | 9.77 | 10.16 | 9.86
105.536 | 0 | 7.32 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 19.24 | 4.88 | 20.7 | 9.77 | 17.09 | 9.67 | 9.96 | 10.16 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.97 | 4.88 | 22.75 | 4.69 | 13.09 | 4.59 | 4.79 | 10.25 | 9.77 | 9.96 | 9.77 | 10.16 | 9.86
105.568 | 0 | 7.91 | 0.1 | 0.1 | 0.39 | 0.1 | 4.69 | 4.88 | 18.07 | 4.88 | 19.73 | 4.88 | 21 | 9.77 | 17.58 | 9.67 | 9.96 | 11.62 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 23.24 | 4.69 | 13.97 | 4.59 | 4.79 | 10.06 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86
105.6 | 0.1 | 8.5 | 0 | 0.1 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.12 | 4.79 | 21.49 | 9.77 | 18.07 | 9.67 | 9.96 | 10.64 | 0 | 17.58 | 0.1 | 17.68 | 0 | 17.87 | 4.79 | 23.83 | 4.59 | 14.75 | 4.59 | 4.79 | 10.06 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
105.632 | 0.1 | 9.08 | 0 | 0 | 0.2 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 21.88 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0 | 17.68 | 0 | 17.77 | 4.79 | 24.22 | 4.69 | 15.43 | 4.49 | 4.79 | 9.96 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86
105.664 | 0.1 | 9.67 | 0.1 | 0.78 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 22.27 | 9.77 | 18.95 | 9.67 | 9.96 | 9.77 | 0.1 | 17.58 | 0.1 | 5.47 | 0.1 | 17.77 | 4.88 | 24.71 | 4.69 | 16.21 | 4.59 | 4.79 | 9.96 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
105.696 | 0.1 | 10.35 | 0 | 1.56 | 0.2 | 0.68 | 4.49 | 4.88 | 18.07 | 4.88 | 17.87 | 4.88 | 22.66 | 9.77 | 19.34 | 9.57 | 9.96 | 10.25 | 0.1 | 13.87 | 0.1 | 1.86 | 0.1 | 17.77 | 4.88 | 25.29 | 4.79 | 16.8 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.96 | 9.86
105.728 | 0 | 10.94 | 0.1 | 2.73 | 0.1 | 1.46 | 4.59 | 4.88 | 10.94 | 4.88 | 18.36 | 4.88 | 23.05 | 9.77 | 19.82 | 9.57 | 9.96 | 10.16 | 0.1 | 3.52 | 0.1 | 2.15 | 0.1 | 17.77 | 4.88 | 25.88 | 4.79 | 17.58 | 4.59 | 4.79 | 9.86 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86
105.76 | 0.1 | 11.43 | 0 | 3.91 | 0.1 | 2.73 | 4.49 | 4.88 | 3.22 | 4.88 | 18.16 | 4.88 | 23.44 | 9.77 | 20.22 | 9.57 | 9.86 | 9.96 | 0 | 2.44 | 0 | 2.05 | 0.1 | 17.68 | 4.88 | 25.68 | 4.79 | 18.16 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
105.792 | 0.1 | 12.01 | 0 | 4.59 | 0.1 | 3.81 | 4.49 | 4.88 | 3.32 | 4.88 | 18.16 | 4.88 | 23.83 | 9.77 | 20.61 | 9.57 | 10.55 | 9.77 | 0.1 | 2.05 | 0 | 1.66 | 0 | 12.7 | 4.79 | 22.75 | 4.79 | 18.75 | 4.59 | 4.88 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
105.824 | 0.1 | 12.6 | 0.1 | 5.18 | 0.1 | 4.59 | 4.49 | 4.88 | 3.81 | 4.88 | 18.16 | 4.88 | 24.12 | 9.77 | 21.09 | 9.57 | 11.23 | 9.86 | 0.1 | 1.76 | 0.1 | 1.66 | 0.1 | 3.03 | 4.88 | 22.07 | 4.79 | 19.53 | 4.59 | 5.86 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
105.856 | 0 | 13.09 | 0.1 | 5.76 | 0.1 | 5.18 | 4.49 | 4.88 | 4.88 | 4.88 | 18.16 | 4.88 | 23.73 | 9.77 | 21.49 | 9.67 | 12.31 | 9.86 | 0.1 | 1.76 | 0.1 | 1.46 | 0.1 | 2.34 | 4.88 | 21.68 | 4.79 | 20.12 | 4.59 | 7.23 | 9.77 | 10.64 | 9.77 | 9.86 | 9.77 | 9.86
105.888 | 0 | 13.67 | 0.1 | 6.35 | 0.1 | 5.76 | 4.59 | 4.88 | 5.96 | 4.88 | 18.16 | 4.88 | 21.58 | 9.86 | 21.78 | 9.67 | 13.57 | 9.77 | 0.1 | 1.46 | 0.1 | 1.37 | 0.1 | 2.05 | 4.88 | 21.58 | 4.79 | 20.61 | 4.69 | 9.38 | 9.67 | 12.01 | 9.77 | 10.74 | 9.77 | 10.16
105.92 | 0.1 | 14.16 | 0 | 6.93 | 0 | 6.35 | 4.49 | 4.88 | 5.66 | 4.79 | 18.16 | 4.88 | 21.49 | 9.77 | 22.27 | 9.67 | 14.36 | 9.67 | 0 | 1.46 | 0 | 1.17 | 0 | 1.66 | 4.88 | 21.39 | 4.79 | 21.29 | 4.69 | 11.33 | 9.67 | 14.06 | 9.77 | 12.31 | 9.67 | 11.43
105.952 | 0.1 | 14.65 | 0 | 7.52 | 0.1 | 6.93 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.68 | 9.77 | 22.56 | 9.67 | 14.94 | 9.77 | 0 | 1.37 | 0 | 1.17 | 0 | 1.76 | 4.88 | 21.39 | 4.79 | 21.78 | 4.69 | 12.11 | 9.67 | 15.43 | 9.67 | 14.16 | 9.67 | 12.79
105.984 | 0.1 | 15.23 | 0.1 | 8.11 | 0.1 | 7.52 | 4.59 | 4.88 | 4.98 | 4.88 | 18.16 | 4.88 | 21.58 | 9.86 | 23.05 | 9.67 | 15.43 | 9.77 | 0.1 | 1.27 | 0.1 | 1.07 | 0.1 | 1.66 | 4.88 | 21.29 | 4.88 | 22.36 | 4.69 | 12.99 | 9.67 | 16.7 | 9.67 | 15.82 | 9.67 | 14.84
106.016 | 0 | 15.63 | 0.1 | 8.69 | 0.1 | 8.11 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 23.34 | 9.67 | 16.02 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.66 | 4.88 | 21.29 | 4.88 | 22.95 | 4.69 | 13.77 | 9.67 | 17.58 | 9.67 | 16.8 | 9.67 | 16.21
106.048 | 0 | 16.21 | 0.1 | 9.38 | 0.1 | 8.69 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.73 | 9.67 | 16.5 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.29 | 4.88 | 23.44 | 4.79 | 14.75 | 9.67 | 18.26 | 9.67 | 17.58 | 9.67 | 17.19
106.08 | 0.1 | 16.7 | 0 | 9.96 | 0.1 | 9.38 | 4.69 | 5.47 | 4.88 | 4.79 | 17.97 | 4.88 | 21.39 | 9.86 | 24.22 | 9.67 | 16.99 | 9.67 | 0 | 0.98 | 0 | 0.78 | 0 | 1.27 | 4.88 | 21.19 | 4.79 | 24.02 | 4.79 | 15.43 | 9.57 | 18.95 | 9.67 | 18.46 | 9.67 | 17.97
106.112 | 0.1 | 17.19 | 0.1 | 10.45 | 0.1 | 9.86 | 4.69 | 6.35 | 4.88 | 4.88 | 16.31 | 4.88 | 21.49 | 9.77 | 23.54 | 9.67 | 17.38 | 9.67 | 0.1 | 0.78 | 0 | 0.78 | 0 | 1.17 | 4.88 | 21.19 | 4.88 | 24.41 | 4.79 | 16.02 | 9.67 | 19.63 | 9.67 | 19.14 | 9.67 | 18.65
106.144 | 0.1 | 14.06 | 0.1 | 11.13 | 0.1 | 10.55 | 4.69 | 7.72 | 4.79 | 4.88 | 4.39 | 4.88 | 21.49 | 9.86 | 21.58 | 9.67 | 17.87 | 9.67 | 0.88 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.19 | 4.88 | 25 | 4.79 | 16.7 | 9.67 | 20.22 | 9.67 | 19.73 | 9.67 | 19.24
106.176 | 0.1 | 14.36 | 0.1 | 11.72 | 0.1 | 11.13 | 4.69 | 8.79 | 4.79 | 4.88 | 2.93 | 4.88 | 21.39 | 9.86 | 21.68 | 9.67 | 18.46 | 9.67 | 1.86 | 0.68 | 0.1 | 0.59 | 0.1 | 0.98 | 4.88 | 21.19 | 4.88 | 25.59 | 4.79 | 17.38 | 9.67 | 20.9 | 9.67 | 20.22 | 9.67 | 19.92
106.208 | 0 | 14.45 | 0.1 | 12.21 | 0.1 | 11.72 | 4.79 | 9.57 | 4.69 | 4.88 | 3.61 | 4.88 | 21.39 | 9.86 | 21.68 | 9.77 | 18.85 | 9.57 | 3.61 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 9.86 | 4.88 | 25.98 | 4.79 | 17.97 | 9.67 | 21.49 | 9.67 | 20.9 | 9.67 | 20.51
106.24 | 0.1 | 14.45 | 0 | 12.79 | 0.1 | 12.31 | 4.79 | 10.35 | 4.59 | 4.88 | 4.79 | 4.88 | 21.39 | 9.77 | 21.49 | 9.67 | 19.34 | 9.57 | 5.27 | 0.49 | 0 | 0.49 | 0 | 0.68 | 4.88 | 2.44 | 4.79 | 21.58 | 4.79 | 18.65 | 9.67 | 22.17 | 9.67 | 21.49 | 9.67 | 21.19
106.272 | 0.1 | 14.45 | 0 | 13.28 | 0.1 | 12.79 | 4.79 | 10.94 | 4.59 | 4.88 | 5.86 | 4.88 | 21.39 | 9.77 | 21.58 | 9.77 | 19.73 | 9.57 | 6.35 | 0.49 | 0 | 0.39 | 0 | 0.68 | 4.88 | 3.03 | 4.88 | 20.22 | 4.79 | 19.24 | 9.67 | 22.66 | 9.67 | 21.97 | 9.67 | 21.78
106.304 | 0.1 | 14.45 | 0.1 | 13.87 | 0.1 | 13.38 | 4.79 | 11.52 | 4.59 | 4.88 | 5.76 | 4.88 | 12.4 | 9.86 | 21.49 | 9.77 | 20.12 | 9.57 | 7.23 | 0.49 | 0.49 | 0.29 | 0.1 | 0.59 | 4.88 | 3.71 | 4.88 | 20.9 | 4.79 | 20.02 | 9.67 | 23.15 | 9.67 | 22.66 | 9.67 | 22.27
106.336 | 0 | 14.55 | 0.1 | 14.45 | 0.1 | 13.87 | 4.79 | 12.11 | 4.59 | 4.88 | 5.37 | 4.88 | 6.74 | 9.86 | 21.49 | 9.77 | 20.51 | 9.57 | 8.01 | 0.39 | 1.66 | 0.29 | 0.1 | 0.59 | 4.88 | 4.79 | 4.88 | 21.19 | 4.79 | 20.61 | 9.77 | 23.73 | 9.67 | 23.24 | 9.67 | 22.85
106.368 | 0 | 14.45 | 0 | 14.94 | 0.1 | 14.36 | 4.79 | 12.6 | 4.59 | 4.88 | 5.08 | 4.88 | 8.5 | 9.86 | 21.49 | 9.77 | 20.9 | 9.67 | 8.79 | 0.29 | 3.13 | 0.2 | 1.17 | 0.49 | 4.88 | 6.15 | 4.88 | 21.19 | 4.79 | 21.19 | 9.77 | 24.22 | 9.67 | 23.83 | 9.67 | 23.44
106.4 | 0.1 | 14.45 | 0 | 15.43 | 0 | 14.84 | 4.79 | 13.18 | 4.49 | 4.79 | 4.98 | 4.88 | 11.43 | 9.77 | 21.49 | 9.77 | 21.39 | 9.57 | 9.57 | 0.29 | 5.08 | 0.2 | 2.64 | 0.49 | 4.88 | 5.96 | 4.79 | 21.19 | 4.79 | 21.78 | 9.77 | 24.71 | 9.67 | 24.22 | 9.77 | 23.93
106.432 | 0.1 | 14.45 | 0 | 15.92 | 0.1 | 15.33 | 4.79 | 13.77 | 4.49 | 4.79 | 4.88 | 4.88 | 11.23 | 9.77 | 21.39 | 9.77 | 21.68 | 9.57 | 10.45 | 0.2 | 6.45 | 0.2 | 4.39 | 0.39 | 4.88 | 5.66 | 4.88 | 21.19 | 4.79 | 22.27 | 9.77 | 25.2 | 9.67 | 24.71 | 9.67 | 24.41
106.464 | 0.1 | 14.45 | 0.1 | 16.41 | 0.1 | 15.92 | 4.88 | 14.36 | 4.59 | 5.66 | 4.88 | 4.88 | 10.25 | 9.86 | 21.39 | 9.77 | 22.17 | 9.67 | 11.23 | 0.29 | 7.23 | 0.2 | 5.86 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.88 | 22.85 | 9.77 | 25.68 | 9.77 | 25.29 | 9.77 | 25
106.496 | 0.1 | 14.45 | 0.1 | 16.9 | 0.1 | 16.41 | 4.88 | 14.75 | 4.59 | 6.15 | 4.88 | 4.88 | 9.86 | 9.86 | 21.39 | 9.77 | 22.46 | 9.67 | 12.01 | 0.2 | 8.01 | 0.2 | 6.93 | 0.29 | 4.88 | 5.27 | 4.88 | 21.19 | 4.79 | 23.34 | 9.77 | 26.27 | 9.77 | 25.68 | 9.77 | 25.49
106.528 | 0 | 14.36 | 0.1 | 17.29 | 0.1 | 16.9 | 4.88 | 15.23 | 4.59 | 7.23 | 4.79 | 4.88 | 10.16 | 9.86 | 21.39 | 9.77 | 22.95 | 9.67 | 12.79 | 0.2 | 8.69 | 0.1 | 7.62 | 0.29 | 4.88 | 5.08 | 4.88 | 21.29 | 4.79 | 23.83 | 9.77 | 26.66 | 9.77 | 26.27 | 9.77 | 25.88
106.56 | 0.1 | 12.4 | 0 | 15.23 | 0.1 | 17.29 | 4.79 | 15.72 | 4.49 | 8.59 | 4.69 | 4.88 | 10.25 | 9.77 | 11.82 | 9.77 | 23.24 | 9.67 | 13.48 | 0.1 | 9.57 | 0.1 | 8.59 | 0.2 | 4.88 | 4.88 | 4.79 | 21.19 | 4.79 | 24.51 | 9.77 | 27.25 | 9.77 | 26.76 | 9.77 | 26.47
106.592 | 0.1 | 3.81 | 0 | 15.14 | 0.1 | 15.33 | 4.79 | 16.31 | 4.59 | 9.38 | 4.69 | 4.88 | 10.06 | 9.77 | 6.74 | 9.77 | 23.63 | 9.67 | 14.36 | 0.1 | 10.25 | 0.1 | 9.38 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 24.9 | 9.77 | 27.74 | 9.77 | 27.25 | 9.77 | 26.86
106.624 | 0.1 | 2.05 | 0.1 | 14.84 | 0.1 | 15.23 | 4.88 | 16.7 | 4.59 | 10.16 | 4.69 | 4.88 | 9.96 | 9.86 | 8.5 | 9.86 | 24.02 | 9.67 | 15.04 | 0.1 | 11.23 | 0.1 | 10.06 | 0.2 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 25.49 | 9.86 | 28.13 | 9.77 | 27.64 | 9.77 | 27.44
106.656 | 0.1 | 1.86 | 0.1 | 14.75 | 0.1 | 14.84 | 4.88 | 17.29 | 4.59 | 10.64 | 4.59 | 4.88 | 9.96 | 9.86 | 11.33 | 9.86 | 24.41 | 9.67 | 15.72 | 0.1 | 12.01 | 0.1 | 10.94 | 0.1 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 25.98 | 9.77 | 28.61 | 9.77 | 28.22 | 9.77 | 27.83
106.688 | 0 | 1.66 | 0.1 | 14.75 | 0.1 | 14.75 | 4.88 | 17.77 | 4.69 | 11.33 | 4.59 | 4.88 | 9.96 | 9.86 | 11.33 | 9.86 | 22.07 | 9.67 | 16.31 | 0.1 | 12.7 | 0.1 | 11.82 | 0.1 | 4.88 | 4.79 | 4.88 | 6.35 | 4.88 | 21.88 | 9.86 | 29.1 | 9.77 | 28.61 | 9.77 | 28.32
106.72 | 0.1 | 1.56 | 0 | 14.55 | 0 | 14.65 | 4.88 | 18.16 | 4.69 | 11.91 | 4.49 | 4.88 | 9.86 | 9.77 | 10.25 | 9.86 | 22.27 | 9.67 | 16.99 | 0.1 | 13.48 | 0.1 | 12.5 | 0.1 | 4.88 | 4.69 | 4.79 | 2.83 | 4.79 | 20.22 | 9.77 | 27.34 | 9.77 | 29.1 | 9.86 | 28.71
106.752 | 0.1 | 1.27 | 0 | 14.55 | 0.1 | 14.55 | 4.79 | 18.56 | 4.69 | 12.4 | 4.49 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 21.78 | 9.67 | 17.58 | 0.1 | 14.16 | 0.1 | 13.18 | 0.1 | 4.88 | 4.69 | 4.88 | 3.32 | 4.79 | 21 | 9.77 | 24.81 | 9.77 | 26.27 | 9.77 | 29.4
106.784 | 0.1 | 1.17 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 19.04 | 4.69 | 12.99 | 4.59 | 4.88 | 9.77 | 9.86 | 10.06 | 9.86 | 21.78 | 9.67 | 18.16 | 0.1 | 14.84 | 0.1 | 13.97 | 0.1 | 4.88 | 4.69 | 4.88 | 4.1 | 4.88 | 21.19 | 9.86 | 24.61 | 9.86 | 23.54 | 9.86 | 24.61
106.816 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.43 | 4.69 | 13.48 | 4.49 | 4.88 | 9.77 | 9.86 | 10.25 | 9.86 | 21.68 | 9.67 | 18.85 | 0.1 | 15.63 | 0.1 | 14.75 | 0.1 | 5.76 | 4.59 | 4.88 | 5.37 | 4.79 | 21.19 | 9.86 | 24.51 | 9.86 | 24.12 | 9.86 | 23.54
106.848 | 0 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.92 | 4.79 | 14.06 | 4.49 | 5.66 | 9.67 | 9.86 | 10.06 | 9.86 | 21.58 | 9.67 | 19.53 | 0.1 | 16.31 | 0 | 15.43 | 0.1 | 7.23 | 4.59 | 4.88 | 6.35 | 4.79 | 21.19 | 9.86 | 24.41 | 9.86 | 24.22 | 9.86 | 24.12
106.88 | 0.1 | 0.78 | 0 | 14.36 | 0 | 14.45 | 4.88 | 20.31 | 4.69 | 14.65 | 4.49 | 6.15 | 9.67 | 9.86 | 9.96 | 9.86 | 21.58 | 9.77 | 20.12 | 0 | 16.9 | 0.1 | 16.02 | 0.1 | 9.28 | 4.59 | 4.79 | 5.96 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.32
106.912 | 0.1 | 0.68 | 0 | 14.36 | 0 | 14.36 | 4.88 | 20.7 | 4.79 | 15.04 | 4.49 | 7.23 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.67 | 20.8 | 0 | 17.58 | 0.1 | 16.7 | 0.1 | 10.74 | 4.49 | 4.88 | 5.47 | 4.88 | 21.19 | 9.77 | 24.32 | 9.77 | 24.12 | 9.86 | 24.22
106.944 | 0.1 | 0.68 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 20.8 | 4.79 | 15.53 | 4.59 | 8.59 | 9.67 | 10.64 | 9.86 | 9.86 | 21.49 | 9.77 | 21.39 | 0.1 | 18.16 | 0.1 | 17.38 | 0.1 | 12.01 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
106.976 | 0.1 | 0.59 | 0.1 | 8.59 | 0.1 | 14.36 | 4.88 | 18.75 | 4.79 | 16.11 | 4.59 | 9.47 | 9.67 | 11.23 | 9.86 | 9.86 | 21.49 | 9.77 | 21.88 | 0.1 | 18.85 | 0.1 | 18.16 | 0.1 | 12.99 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
107.008 | 0 | 0.49 | 0.1 | 2.44 | 0.1 | 8.79 | 4.88 | 18.65 | 4.79 | 16.5 | 4.59 | 10.16 | 9.67 | 12.6 | 9.86 | 9.86 | 21.49 | 9.77 | 22.56 | 0.1 | 19.53 | 0.1 | 18.75 | 0.1 | 13.87 | 4.59 | 4.88 | 5.18 | 4.79 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
107.04 | 0.1 | 0.49 | 0 | 2.05 | 0 | 2.25 | 4.88 | 18.46 | 4.79 | 16.99 | 4.59 | 10.64 | 9.57 | 13.57 | 9.67 | 9.86 | 21.49 | 9.77 | 19.34 | 0 | 20.12 | 0.1 | 19.53 | 0 | 14.65 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
107.072 | 0.1 | 0.39 | 0 | 1.66 | 0.1 | 1.95 | 4.79 | 18.36 | 4.79 | 17.48 | 4.59 | 11.23 | 9.67 | 14.36 | 9.77 | 9.86 | 21.39 | 9.77 | 18.46 | 0 | 20.7 | 0.1 | 20.02 | 0.1 | 15.33 | 4.49 | 4.79 | 4.98 | 4.79 | 21.09 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
107.104 | 0.2 | 0.39 | 0.1 | 1.56 | 0.1 | 1.86 | 4.88 | 18.36 | 4.88 | 17.97 | 4.69 | 11.91 | 9.67 | 14.94 | 9.67 | 9.86 | 16.9 | 9.77 | 18.07 | 0.1 | 21.39 | 0.1 | 20.7 | 0.1 | 16.11 | 4.49 | 4.88 | 4.98 | 4.88 | 21.09 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
107.136 | 1.07 | 0.29 | 0.1 | 1.46 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.46 | 4.69 | 12.4 | 9.67 | 15.53 | 9.67 | 9.86 | 7.81 | 9.77 | 17.87 | 0.1 | 21.97 | 0.1 | 21.29 | 0.1 | 16.8 | 4.59 | 4.88 | 4.88 | 4.88 | 7.03 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86 | 24.22
107.168 | 2.05 | 0.29 | 0.1 | 1.27 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.85 | 4.69 | 12.99 | 9.67 | 16.02 | 9.67 | 9.86 | 7.52 | 9.77 | 17.87 | 0.1 | 22.46 | 0 | 21.78 | 0.1 | 17.48 | 4.59 | 4.88 | 4.88 | 4.88 | 2.34 | 9.77 | 12.89 | 9.86 | 24.12 | 9.86 | 24.12
107.2 | 3.42 | 0.2 | 0 | 1.17 | 0 | 1.17 | 4.88 | 18.16 | 4.79 | 19.24 | 4.69 | 13.57 | 9.67 | 16.6 | 9.67 | 9.86 | 10.74 | 9.77 | 17.68 | 0 | 18.46 | 0 | 22.36 | 0.1 | 18.07 | 4.59 | 5.18 | 4.79 | 4.79 | 3.32 | 9.77 | 5.66 | 9.86 | 11.13 | 9.86 | 20.9
107.232 | 4.3 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.79 | 19.73 | 4.69 | 13.97 | 9.67 | 16.99 | 9.67 | 9.86 | 11.52 | 9.77 | 17.58 | 0.1 | 17.29 | 0.1 | 20.61 | 0.1 | 18.75 | 4.59 | 6.35 | 4.69 | 4.79 | 4.1 | 9.77 | 8.01 | 9.86 | 5.57 | 9.86 | 8.69
107.264 | 4.98 | 0.1 | 0.1 | 0.98 | 0.1 | 0.98 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 14.55 | 9.67 | 17.48 | 9.67 | 9.86 | 10.55 | 9.77 | 17.58 | 0.1 | 17.77 | 0.1 | 17.38 | 0.1 | 19.34 | 4.59 | 7.72 | 4.69 | 4.88 | 5.27 | 9.86 | 11.91 | 9.86 | 8.59 | 9.86 | 6.25
107.296 | 5.47 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.61 | 4.79 | 15.04 | 9.67 | 17.97 | 9.67 | 9.86 | 9.86 | 9.77 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 20.02 | 4.59 | 9.86 | 4.69 | 4.88 | 6.15 | 9.86 | 11.62 | 9.86 | 12.01 | 9.86 | 9.77
107.328 | 5.96 | 0.1 | 0.1 | 0.88 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 21 | 4.79 | 15.53 | 9.67 | 18.46 | 9.57 | 9.86 | 9.86 | 9.86 | 17.58 | 0.1 | 17.77 | 0 | 17.87 | 0.1 | 20.61 | 4.69 | 11.33 | 4.69 | 4.88 | 5.86 | 9.77 | 10.45 | 9.86 | 11.43 | 9.86 | 12.5
107.36 | 6.64 | 0 | 0 | 0.68 | 0 | 0.68 | 4.79 | 18.07 | 4.88 | 18.56 | 4.79 | 16.11 | 9.67 | 18.85 | 9.57 | 9.86 | 10.16 | 9.77 | 17.48 | 0.1 | 17.77 | 0.1 | 17.77 | 0.1 | 21.19 | 4.69 | 12.4 | 4.59 | 4.79 | 5.47 | 9.77 | 10.25 | 9.86 | 10.25 | 9.86 | 10.94
107.392 | 7.13 | 0 | 0 | 0.49 | 0.1 | 0.68 | 4.79 | 8.79 | 4.88 | 18.75 | 4.79 | 16.6 | 9.67 | 19.24 | 9.57 | 9.86 | 10.16 | 9.77 | 17.48 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 21.58 | 4.69 | 13.18 | 4.59 | 4.79 | 5.18 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.25
107.424 | 7.72 | 0.1 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 3.13 | 4.88 | 18.26 | 4.79 | 16.99 | 9.77 | 19.82 | 9.67 | 9.86 | 10.16 | 9.86 | 17.48 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.27 | 4.69 | 14.06 | 4.59 | 4.79 | 5.18 | 9.86 | 10.35 | 9.86 | 10.25 | 9.96 | 9.96
107.456 | 8.4 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 18.26 | 4.88 | 17.48 | 9.77 | 20.12 | 9.67 | 9.86 | 9.96 | 9.86 | 10.64 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.95 | 4.69 | 14.84 | 4.59 | 4.88 | 5.18 | 9.86 | 10.45 | 9.86 | 10.35 | 9.86 | 10.45
107.488 | 8.98 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 4.88 | 4 | 4.88 | 18.26 | 4.88 | 17.87 | 9.77 | 20.61 | 9.67 | 9.86 | 9.86 | 9.86 | 2.73 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.54 | 4.69 | 15.43 | 4.59 | 5.37 | 4.98 | 9.86 | 10.25 | 9.86 | 10.35 | 9.86 | 10.45
107.52 | 9.67 | 0 | 0 | 0.29 | 0.1 | 0.29 | 4.88 | 5.18 | 4.88 | 18.16 | 4.79 | 18.46 | 9.77 | 21 | 9.67 | 9.86 | 9.96 | 9.86 | 2.25 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 24.02 | 4.69 | 16.11 | 4.59 | 6.45 | 4.98 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.25
107.552 | 10.25 | 0 | 0.1 | 0.29 | 0.1 | 0.29 | 4.79 | 5.86 | 4.88 | 18.16 | 4.88 | 18.75 | 9.77 | 21.39 | 9.67 | 9.86 | 9.77 | 9.77 | 2.05 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 24.41 | 4.69 | 16.8 | 4.49 | 7.91 | 4.88 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.06
107.584 | 10.74 | 0.1 | 1.17 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.16 | 4.88 | 19.24 | 9.77 | 21.78 | 9.67 | 9.86 | 9.77 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25 | 4.79 | 17.58 | 4.59 | 10.16 | 4.88 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.06
107.616 | 11.33 | 0.1 | 1.95 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.73 | 9.77 | 22.17 | 9.67 | 10.64 | 9.77 | 9.86 | 1.76 | 0.1 | 17.29 | 0.1 | 17.68 | 0.1 | 25.49 | 4.79 | 18.26 | 4.59 | 11.52 | 4.79 | 9.86 | 9.96 | 9.86 | 10.06 | 9.86 | 10.06
107.648 | 11.91 | 0.1 | 3.32 | 0.1 | 0.1 | 0.2 | 4.88 | 5.08 | 4.88 | 18.16 | 4.88 | 20.12 | 9.77 | 22.56 | 9.67 | 11.33 | 9.77 | 9.86 | 1.56 | 0.1 | 4.3 | 0 | 17.68 | 0.1 | 25.98 | 4.79 | 18.85 | 4.59 | 12.4 | 4.79 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
107.68 | 12.5 | 0 | 4.2 | 0.1 | 0.1 | 0.1 | 4.79 | 4.88 | 4.88 | 18.07 | 4.79 | 20.61 | 9.77 | 22.95 | 9.67 | 12.7 | 9.67 | 9.86 | 1.37 | 0 | 2.05 | 0 | 6.84 | 0.1 | 24.22 | 4.79 | 19.43 | 4.59 | 13.38 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
107.712 | 12.99 | 0.1 | 4.88 | 0.1 | 0.2 | 0.1 | 4.79 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 9.77 | 23.24 | 9.67 | 13.57 | 9.67 | 9.86 | 1.27 | 0 | 2.25 | 0.1 | 1.66 | 0.1 | 21.68 | 4.79 | 20.12 | 4.59 | 14.16 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86
107.744 | 13.57 | 0.1 | 5.37 | 0.1 | 1.17 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 18.26 | 9.77 | 23.83 | 9.67 | 14.36 | 9.67 | 9.86 | 1.27 | 0.1 | 1.95 | 0.1 | 2.44 | 0.1 | 21.58 | 4.79 | 20.8 | 4.69 | 14.84 | 4.69 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
107.776 | 14.16 | 0.1 | 5.96 | 0.1 | 2.15 | 0 | 5.27 | 4.79 | 4.88 | 14.36 | 4.88 | 18.75 | 9.86 | 24.02 | 9.77 | 14.94 | 9.67 | 9.86 | 1.07 | 0.1 | 1.76 | 0.1 | 2.15 | 0.1 | 21.58 | 4.79 | 21.39 | 4.69 | 15.53 | 4.69 | 9.77 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
107.808 | 14.55 | 0.1 | 6.54 | 0.1 | 3.61 | 0 | 6.05 | 4.79 | 4.88 | 4.1 | 4.88 | 18.26 | 9.86 | 24.41 | 9.77 | 15.53 | 9.67 | 9.86 | 0.98 | 0.1 | 1.56 | 0 | 1.76 | 0.1 | 21.49 | 4.79 | 21.97 | 4.69 | 16.21 | 4.59 | 10.74 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
107.84 | 15.14 | 0 | 7.13 | 0.1 | 4.39 | 0.1 | 7.03 | 4.69 | 4.88 | 3.22 | 4.88 | 18.26 | 9.77 | 21.88 | 9.67 | 16.11 | 9.67 | 9.86 | 0.88 | 0 | 1.56 | 0.1 | 1.66 | 0.1 | 21.29 | 4.79 | 22.46 | 4.69 | 16.99 | 4.49 | 11.52 | 9.77 | 9.77 | 9.77 | 10.16 | 9.77
107.872 | 15.63 | 0.1 | 7.72 | 0.1 | 5.08 | 0 | 8.4 | 4.69 | 4.88 | 3.52 | 4.88 | 18.26 | 9.86 | 21.97 | 9.77 | 16.5 | 9.57 | 9.86 | 0.88 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 21.39 | 4.79 | 23.05 | 4.69 | 17.58 | 4.49 | 12.89 | 9.67 | 9.96 | 9.77 | 10.94 | 9.77
107.904 | 16.11 | 0.1 | 8.3 | 0.1 | 5.47 | 0.1 | 9.38 | 4.59 | 4.88 | 4.69 | 4.88 | 18.16 | 9.86 | 21.78 | 9.77 | 16.99 | 9.67 | 9.86 | 0.78 | 0.1 | 1.46 | 0.1 | 1.46 | 0.1 | 21.39 | 4.79 | 23.54 | 4.69 | 18.36 | 4.59 | 14.94 | 9.67 | 10.94 | 9.77 | 12.31 | 9.67
107.936 | 16.6 | 0.1 | 8.89 | 0.1 | 6.05 | 0.1 | 9.96 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.58 | 9.67 | 9.86 | 0.68 | 0.1 | 1.17 | 0.1 | 1.27 | 0.1 | 21.39 | 4.79 | 24.12 | 4.79 | 18.95 | 4.59 | 16.31 | 9.67 | 12.4 | 9.67 | 14.36 | 9.77
107.968 | 16.99 | 0.1 | 9.57 | 0.1 | 6.64 | 0.1 | 10.55 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.97 | 9.67 | 9.86 | 0.59 | 0.1 | 1.07 | 0 | 1.17 | 0.1 | 21.39 | 4.79 | 24.51 | 4.79 | 19.53 | 4.49 | 17.19 | 9.67 | 14.45 | 9.67 | 15.92 | 9.67
108 | 13.77 | 0 | 10.16 | 0 | 7.23 | 0 | 11.23 | 4.49 | 4.88 | 5.37 | 4.79 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.67 | 9.86 | 0.49 | 0 | 0.88 | 0.1 | 1.07 | 0 | 21.29 | 4.79 | 25.1 | 4.79 | 20.12 | 4.49 | 17.97 | 9.67 | 16.11 | 9.67 | 16.99 | 9.67
108.032 | 13.97 | 0.1 | 10.64 | 0 | 7.81 | 0 | 11.82 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.85 | 9.57 | 9.86 | 0.49 | 0 | 0.88 | 0.1 | 0.98 | 0.1 | 21.29 | 4.79 | 25.59 | 4.79 | 20.7 | 4.49 | 18.65 | 9.67 | 17.09 | 9.57 | 17.77 | 9.67
108.064 | 14.16 | 0.1 | 11.33 | 0.1 | 8.4 | 0.1 | 12.4 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.34 | 9.67 | 9.86 | 0.49 | 0.1 | 0.78 | 0.1 | 0.98 | 0.1 | 21.29 | 4.88 | 26.17 | 4.79 | 21.39 | 4.59 | 19.34 | 9.67 | 17.77 | 9.67 | 18.56 | 9.67
108.096 | 14.36 | 0.1 | 11.91 | 0.1 | 9.08 | 0.1 | 12.89 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.73 | 9.67 | 9.86 | 0.39 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 21.29 | 4.88 | 21.49 | 4.79 | 21.78 | 4.59 | 20.02 | 9.67 | 18.46 | 9.67 | 19.14 | 9.67
108.128 | 14.36 | 0.1 | 12.5 | 0.1 | 9.67 | 0.1 | 13.48 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 20.22 | 9.67 | 10.25 | 0.29 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 8.59 | 4.88 | 20.41 | 4.79 | 22.46 | 4.59 | 20.61 | 9.67 | 19.14 | 9.67 | 19.82 | 9.67
108.16 | 14.36 | 0 | 12.99 | 0.1 | 10.25 | 0 | 14.06 | 4.49 | 4.88 | 4.79 | 4.88 | 12.89 | 9.77 | 21.39 | 9.77 | 20.7 | 9.67 | 11.04 | 0.29 | 0 | 0.49 | 0.1 | 0.59 | 0.1 | 2.34 | 4.79 | 21.19 | 4.79 | 23.05 | 4.59 | 21.19 | 9.67 | 19.92 | 9.67 | 20.41 | 9.57
108.192 | 14.45 | 0 | 13.48 | 0.1 | 10.84 | 0 | 14.45 | 4.49 | 4.88 | 4.79 | 4.88 | 3.91 | 9.86 | 21.39 | 9.77 | 21 | 9.67 | 12.01 | 0.29 | 0 | 0.49 | 0.1 | 0.59 | 0.1 | 3.13 | 4.79 | 21.19 | 4.79 | 23.54 | 4.59 | 21.78 | 9.57 | 20.41 | 9.57 | 20.9 | 9.67
108.224 | 14.45 | 0.1 | 14.06 | 0.1 | 11.43 | 0.1 | 15.04 | 4.59 | 4.88 | 4.79 | 4.88 | 3.13 | 9.86 | 21.39 | 9.86 | 21.49 | 9.67 | 13.38 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.71 | 4.88 | 21.19 | 4.79 | 24.12 | 4.69 | 22.36 | 9.67 | 21 | 9.67 | 21.58 | 9.67
108.256 | 14.45 | 0.1 | 14.55 | 0.1 | 12.01 | 0.1 | 15.63 | 4.59 | 4.88 | 4.69 | 4.88 | 3.61 | 9.86 | 13.18 | 9.86 | 21.78 | 9.67 | 14.16 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 21.29 | 4.79 | 24.61 | 4.69 | 22.95 | 9.67 | 21.58 | 9.67 | 22.07 | 9.67
108.288 | 14.45 | 0.1 | 15.04 | 0.1 | 12.5 | 0.1 | 16.11 | 4.59 | 4.88 | 4.59 | 4.88 | 4.88 | 9.86 | 6.93 | 9.86 | 22.27 | 9.67 | 14.75 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 6.15 | 4.88 | 21.29 | 4.88 | 25.1 | 4.69 | 23.34 | 9.67 | 22.17 | 9.67 | 22.66 | 9.67
108.32 | 14.36 | 0 | 15.53 | 0.1 | 13.09 | 0.1 | 16.5 | 4.59 | 4.88 | 4.59 | 4.79 | 5.86 | 9.86 | 8.3 | 9.86 | 22.56 | 9.67 | 15.33 | 0.1 | 0.1 | 0.29 | 0.2 | 0.39 | 0.1 | 5.96 | 4.79 | 21.29 | 4.79 | 25.59 | 4.69 | 24.12 | 9.67 | 22.75 | 9.67 | 23.34 | 9.67
108.352 | 14.36 | 0 | 16.02 | 0.1 | 13.67 | 0 | 16.9 | 4.69 | 4.88 | 4.59 | 4.88 | 5.76 | 9.86 | 11.04 | 9.86 | 22.95 | 9.67 | 15.82 | 0.1 | 0.1 | 0.2 | 1.56 | 0.29 | 0.1 | 5.57 | 4.79 | 21.29 | 4.79 | 26.08 | 4.69 | 24.61 | 9.67 | 23.24 | 9.67 | 23.83 | 9.67
108.384 | 14.36 | 0.1 | 16.6 | 0.1 | 14.16 | 0.1 | 17.58 | 4.69 | 4.88 | 4.59 | 4.88 | 5.37 | 9.86 | 11.23 | 9.86 | 23.34 | 9.77 | 16.31 | 0.1 | 0.59 | 0.2 | 2.83 | 0.29 | 0.1 | 5.27 | 4.88 | 21.29 | 4.88 | 21.78 | 4.69 | 25 | 9.67 | 23.83 | 9.67 | 24.32 | 9.67
108.416 | 14.36 | 0.1 | 16.99 | 0.1 | 14.65 | 0 | 17.97 | 4.69 | 5.57 | 4.59 | 4.88 | 5.08 | 9.86 | 10.35 | 9.86 | 23.73 | 9.77 | 16.9 | 0.1 | 1.66 | 0.2 | 4.79 | 0.2 | 0.1 | 5.27 | 4.88 | 21.29 | 4.88 | 20.8 | 4.69 | 25.39 | 9.77 | 24.32 | 9.67 | 24.81 | 9.67
108.448 | 5.18 | 0.1 | 14.16 | 0.1 | 15.23 | 0 | 18.46 | 4.79 | 6.15 | 4.49 | 4.88 | 4.98 | 9.86 | 9.96 | 9.86 | 24.12 | 9.77 | 17.29 | 0.1 | 3.13 | 0.2 | 6.25 | 0.2 | 0.1 | 5.08 | 4.79 | 21.29 | 4.88 | 21.29 | 4.79 | 25.98 | 9.77 | 24.9 | 9.67 | 25.29 | 9.67
108.48 | 1.66 | 0 | 13.87 | 0 | 15.63 | 0 | 18.85 | 4.69 | 7.32 | 4.39 | 4.88 | 4.88 | 9.86 | 10.06 | 9.86 | 24.12 | 9.77 | 17.77 | 0.1 | 5.08 | 0.1 | 7.03 | 0.1 | 0.88 | 4.98 | 4.79 | 21.19 | 4.88 | 21.29 | 4.79 | 26.56 | 9.67 | 25.39 | 9.67 | 25.88 | 9.67
108.512 | 1.95 | 0 | 14.26 | 0 | 16.11 | 0 | 19.24 | 4.79 | 8.59 | 4.49 | 4.88 | 4.79 | 9.86 | 10.16 | 9.86 | 22.36 | 9.77 | 18.16 | 0.1 | 6.25 | 0.1 | 7.81 | 0.1 | 1.76 | 5.18 | 4.79 | 20.9 | 4.79 | 21.39 | 4.69 | 26.95 | 9.77 | 25.78 | 9.67 | 26.37 | 9.67
108.544 | 1.76 | 0.1 | 14.45 | 0.1 | 16.7 | 0.1 | 19.73 | 4.79 | 9.47 | 4.49 | 4.88 | 4.79 | 9.86 | 10.16 | 9.86 | 22.17 | 9.77 | 18.65 | 0.1 | 7.13 | 0.1 | 8.69 | 0.1 | 3.32 | 4.88 | 4.79 | 5.08 | 4.88 | 21.29 | 4.79 | 27.54 | 9.77 | 26.27 | 9.77 | 26.76 | 9.77
108.576 | 1.46 | 0.1 | 14.45 | 0.1 | 17.19 | 0.1 | 20.12 | 4.79 | 10.16 | 4.49 | 4.88 | 4.79 | 9.86 | 10.06 | 9.86 | 21.88 | 9.77 | 19.24 | 0.1 | 8.01 | 0.1 | 9.47 | 0.1 | 5.27 | 4.98 | 4.88 | 2.83 | 4.79 | 21.29 | 4.79 | 28.13 | 9.77 | 26.76 | 9.77 | 27.34 | 9.77
108.608 | 1.27 | 0.1 | 14.45 | 0.1 | 13.77 | 0 | 20.61 | 4.79 | 10.64 | 4.59 | 4.88 | 4.79 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 19.53 | 0.1 | 8.69 | 0.1 | 10.25 | 0.1 | 6.45 | 4.88 | 4.88 | 3.42 | 4.88 | 21.29 | 4.79 | 28.52 | 9.77 | 27.34 | 9.77 | 27.74 | 9.77
108.64 | 1.27 | 0 | 14.45 | 0 | 14.06 | 0.1 | 20.9 | 4.79 | 11.33 | 4.49 | 4.79 | 4.69 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 20.02 | 0.1 | 9.57 | 0.1 | 10.94 | 0.1 | 7.23 | 4.79 | 4.79 | 4.3 | 4.79 | 21.19 | 4.79 | 29.01 | 9.77 | 27.83 | 9.77 | 28.32 | 9.77
108.672 | 1.07 | 0 | 14.45 | 0 | 14.16 | 0 | 17.68 | 4.79 | 11.91 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.41 | 0.1 | 10.25 | 0.1 | 11.82 | 0.1 | 8.01 | 4.79 | 4.79 | 5.57 | 4.79 | 21.29 | 4.79 | 29.4 | 9.77 | 28.22 | 9.77 | 28.71 | 9.77
108.704 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.07 | 4.79 | 12.5 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.9 | 0.1 | 11.04 | 0.1 | 12.6 | 0.1 | 8.89 | 4.79 | 4.79 | 6.35 | 4.88 | 21.19 | 4.79 | 25.2 | 9.77 | 28.71 | 9.77 | 29.2 | 9.77
108.736 | 0.88 | 0 | 14.45 | 0.1 | 14.36 | 0.1 | 18.07 | 4.88 | 12.99 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.19 | 0.1 | 11.82 | 0.1 | 13.38 | 0.1 | 9.67 | 4.69 | 4.79 | 5.96 | 4.88 | 21.19 | 4.79 | 24.41 | 9.77 | 29.3 | 9.77 | 25 | 9.77
108.768 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 13.57 | 4.69 | 4.88 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 21.58 | 0.1 | 12.5 | 0.1 | 14.16 | 0.1 | 10.55 | 4.69 | 4.79 | 5.47 | 4.88 | 21.19 | 4.79 | 24.51 | 9.86 | 24.51 | 9.77 | 23.54 | 9.77
108.8 | 0.68 | 0 | 14.36 | 0 | 14.45 | 0 | 18.26 | 4.79 | 14.16 | 4.69 | 5.66 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.07 | 0.1 | 13.38 | 0.1 | 14.75 | 0.1 | 11.23 | 4.59 | 4.79 | 5.18 | 4.79 | 18.46 | 4.79 | 24.51 | 9.77 | 23.34 | 9.77 | 24.02 | 9.77
108.832 | 0.59 | 0 | 14.36 | 0 | 14.45 | 0 | 18.16 | 4.88 | 14.65 | 4.69 | 6.05 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.77 | 22.36 | 0.1 | 14.16 | 0 | 15.43 | 0.1 | 12.01 | 4.49 | 4.79 | 5.18 | 4.79 | 4.69 | 4.79 | 24.51 | 9.77 | 23.93 | 9.77 | 24.22 | 9.77
108.864 | 0.59 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.23 | 4.69 | 7.03 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.75 | 0.1 | 14.75 | 0.1 | 16.21 | 0.1 | 12.7 | 4.59 | 4.79 | 5.18 | 4.88 | 2.93 | 4.79 | 24.41 | 9.86 | 24.12 | 9.77 | 24.22 | 9.77
108.896 | 0.49 | 0.1 | 5.86 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.63 | 4.79 | 8.4 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 0.1 | 15.53 | 0.1 | 16.8 | 0.1 | 13.38 | 4.59 | 5.76 | 5.08 | 4.79 | 3.42 | 4.79 | 24.41 | 9.86 | 24.12 | 9.77 | 24.22 | 9.86
108.928 | 0.39 | 0.1 | 1.76 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.11 | 4.79 | 9.28 | 4.49 | 9.86 | 9.67 | 9.86 | 13.57 | 9.86 | 23.54 | 0.1 | 16.21 | 0.1 | 17.48 | 0.1 | 14.26 | 4.59 | 6.84 | 5.08 | 4.79 | 4.39 | 4.79 | 24.41 | 9.86 | 24.12 | 9.77 | 24.22 | 9.86
108.96 | 0.39 | 0.1 | 1.76 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 16.7 | 4.79 | 9.96 | 4.49 | 9.86 | 9.67 | 9.86 | 6.93 | 9.86 | 23.93 | 0.1 | 16.99 | 0.1 | 18.16 | 0.1 | 15.14 | 4.49 | 8.59 | 4.88 | 4.79 | 5.57 | 4.79 | 24.32 | 9.77 | 24.22 | 9.86 | 24.32 | 9.77
108.992 | 0.29 | 0 | 1.76 | 0 | 14.36 | 0 | 18.07 | 4.88 | 17.09 | 4.79 | 10.55 | 4.49 | 10.45 | 9.57 | 9.86 | 8.01 | 9.77 | 24.32 | 0.1 | 17.58 | 0.1 | 18.85 | 0.1 | 15.82 | 4.49 | 10.35 | 4.98 | 4.79 | 6.35 | 4.79 | 24.41 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
109.024 | 0.29 | 0.1 | 1.46 | 0.1 | 14.36 | 0 | 18.16 | 4.88 | 17.58 | 4.79 | 11.23 | 4.59 | 11.33 | 9.67 | 9.86 | 11.04 | 9.86 | 21 | 0.1 | 18.26 | 0.1 | 19.34 | 0.1 | 16.6 | 4.59 | 11.72 | 4.88 | 4.79 | 5.86 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
109.056 | 0.2 | 0.1 | 1.37 | 0.1 | 4.79 | 0.1 | 18.07 | 4.88 | 18.07 | 4.79 | 11.82 | 4.59 | 12.6 | 9.57 | 9.86 | 11.33 | 9.86 | 21.29 | 0.1 | 18.85 | 0.1 | 20.02 | 0.1 | 17.19 | 4.59 | 12.6 | 4.79 | 4.79 | 5.37 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
109.088 | 0.1 | 0.1 | 1.17 | 0.1 | 1.86 | 0.1 | 17.58 | 4.88 | 18.46 | 4.79 | 12.4 | 4.59 | 13.67 | 9.57 | 9.86 | 10.45 | 9.86 | 21.39 | 0.1 | 19.34 | 0.1 | 20.7 | 0.1 | 17.77 | 4.59 | 13.57 | 4.79 | 4.79 | 5.18 | 4.79 | 24.32 | 9.86 | 24.12 | 9.86 | 24.22 | 9.86
109.12 | 0.1 | 0 | 1.07 | 0 | 1.86 | 0 | 4.59 | 4.88 | 18.85 | 4.79 | 12.89 | 4.59 | 14.36 | 9.57 | 9.86 | 9.77 | 9.86 | 21.49 | 0.1 | 20.12 | 0.1 | 21.29 | 0.1 | 18.46 | 4.59 | 14.36 | 4.69 | 4.79 | 5.18 | 4.79 | 18.16 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
109.152 | 0.1 | 0 | 0.98 | 0 | 1.86 | 0 | 3.03 | 4.88 | 19.34 | 4.79 | 13.48 | 4.69 | 14.94 | 9.57 | 9.86 | 10.06 | 9.86 | 21.49 | 0.1 | 20.7 | 0.1 | 21.88 | 0 | 19.14 | 4.59 | 15.04 | 4.69 | 4.79 | 5.08 | 4.79 | 7.72 | 9.86 | 24.12 | 9.77 | 24.12 | 9.86
109.184 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 3.52 | 4.88 | 19.82 | 4.88 | 14.06 | 4.69 | 15.53 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 21.19 | 0.1 | 22.36 | 0.1 | 19.82 | 4.69 | 15.72 | 4.69 | 4.79 | 5.18 | 4.79 | 6.74 | 9.86 | 24.12 | 9.86 | 11.23 | 9.86
109.216 | 0.1 | 0.1 | 0.88 | 0.1 | 1.27 | 0 | 4.59 | 4.88 | 20.22 | 4.79 | 14.55 | 4.69 | 16.02 | 9.67 | 9.86 | 10.25 | 9.86 | 21.49 | 0.1 | 21.88 | 0.1 | 17.77 | 0.1 | 20.31 | 4.69 | 16.5 | 4.69 | 4.79 | 5.18 | 4.79 | 10.55 | 9.86 | 10.45 | 9.86 | 5.66 | 9.86
109.248 | 0 | 0.2 | 0.68 | 0.1 | 1.07 | 0.1 | 5.66 | 4.88 | 20.61 | 4.88 | 15.04 | 4.69 | 16.5 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 0.1 | 22.46 | 0.1 | 16.7 | 0.1 | 20.9 | 4.69 | 17.19 | 4.69 | 4.79 | 4.98 | 4.79 | 12.31 | 9.86 | 6.05 | 9.86 | 8.98 | 9.86
109.28 | 0.1 | 1.17 | 0.59 | 0 | 1.07 | 0 | 5.86 | 4.88 | 21 | 4.79 | 15.53 | 4.69 | 16.99 | 9.67 | 11.52 | 9.86 | 9.86 | 21.49 | 0.1 | 18.65 | 0.1 | 17.38 | 0.1 | 21.58 | 4.69 | 17.87 | 4.59 | 4.79 | 4.88 | 4.79 | 10.64 | 9.77 | 9.86 | 9.77 | 12.31 | 9.86
109.312 | 0.1 | 2.15 | 0.49 | 0 | 0.88 | 0 | 5.47 | 4.88 | 18.65 | 4.79 | 16.02 | 4.79 | 17.48 | 9.67 | 12.6 | 9.86 | 9.86 | 21.39 | 0.1 | 17.19 | 0 | 17.58 | 0.1 | 22.07 | 4.69 | 18.46 | 4.59 | 4.79 | 4.88 | 4.79 | 9.96 | 9.77 | 12.31 | 9.77 | 11.52 | 9.86
109.344 | 0.1 | 3.61 | 0.49 | 0.1 | 0.88 | 0.1 | 5.08 | 4.88 | 18.95 | 4.88 | 16.5 | 4.79 | 17.97 | 9.67 | 13.67 | 9.86 | 9.96 | 21.39 | 0.1 | 17.68 | 0.1 | 17.58 | 0.1 | 22.56 | 4.79 | 19.14 | 4.59 | 4.88 | 4.79 | 4.79 | 10.25 | 9.86 | 11.43 | 9.86 | 10.45 | 9.86
109.376 | 0.1 | 4.39 | 0.49 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 18.46 | 4.88 | 16.99 | 4.79 | 18.46 | 9.67 | 14.36 | 9.77 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.58 | 0.1 | 18.56 | 4.79 | 19.73 | 4.59 | 5.86 | 4.79 | 4.79 | 10.45 | 9.86 | 10.16 | 9.77 | 9.96 | 9.86
109.408 | 0 | 4.98 | 0.39 | 0.1 | 0.68 | 0.1 | 4.98 | 4.88 | 18.36 | 4.88 | 17.48 | 4.79 | 18.85 | 9.67 | 15.04 | 9.77 | 9.96 | 21.39 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 18.16 | 4.79 | 20.31 | 4.59 | 7.23 | 4.69 | 4.79 | 10.45 | 9.86 | 9.96 | 9.77 | 10.25 | 9.86
109.44 | 0.1 | 5.47 | 0.29 | 0 | 0.59 | 0.1 | 4.88 | 4.88 | 18.46 | 4.79 | 17.87 | 4.79 | 19.34 | 9.67 | 15.53 | 9.77 | 9.96 | 19.14 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 18.16 | 4.79 | 21 | 4.49 | 9.57 | 4.69 | 4.79 | 10.25 | 9.77 | 10.45 | 9.77 | 10.64 | 9.86
109.472 | 0.1 | 5.96 | 0.29 | 0.1 | 0.59 | 0 | 4.79 | 4.88 | 18.26 | 4.79 | 18.26 | 4.79 | 19.82 | 9.67 | 16.02 | 9.67 | 9.86 | 8.69 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 17.87 | 4.79 | 21.58 | 4.59 | 10.94 | 4.59 | 4.79 | 10.25 | 9.77 | 10.35 | 9.77 | 10.35 | 9.86
109.504 | 0.1 | 6.54 | 0.2 | 0.1 | 0.59 | 0.1 | 4.79 | 4.88 | 18.26 | 4.88 | 18.75 | 4.79 | 20.12 | 9.77 | 16.5 | 9.67 | 9.96 | 7.13 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.97 | 4.88 | 22.07 | 4.59 | 12.11 | 4.59 | 4.79 | 10.35 | 9.86 | 10.25 | 9.86 | 10.25 | 9.86
109.536 | 0.1 | 7.23 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 20.61 | 9.77 | 17.09 | 9.67 | 9.96 | 9.96 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.79 | 22.66 | 4.69 | 12.99 | 4.59 | 4.79 | 10.16 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
109.568 | 0 | 7.81 | 0.1 | 0.1 | 0.39 | 0.1 | 4.69 | 4.88 | 18.16 | 4.88 | 19.73 | 4.88 | 21 | 9.77 | 17.48 | 9.67 | 9.96 | 11.52 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 23.15 | 4.69 | 13.97 | 4.59 | 4.79 | 10.06 | 9.86 | 9.96 | 9.86 | 10.16 | 9.86
109.6 | 0.1 | 8.4 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.79 | 20.12 | 4.88 | 21.39 | 9.67 | 17.97 | 9.67 | 9.86 | 10.74 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.88 | 23.73 | 4.59 | 14.65 | 4.59 | 4.79 | 10.06 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86
109.632 | 0.1 | 8.98 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.79 | 20.61 | 4.79 | 21.88 | 9.77 | 18.46 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.79 | 24.32 | 4.69 | 15.23 | 4.49 | 4.79 | 10.06 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86
109.664 | 0.1 | 9.67 | 0.1 | 0.68 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 22.27 | 9.77 | 18.85 | 9.67 | 9.96 | 9.86 | 0.1 | 17.58 | 0.1 | 6.05 | 0.1 | 17.77 | 4.88 | 24.81 | 4.69 | 16.02 | 4.59 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
109.696 | 0.1 | 10.25 | 0.1 | 1.37 | 0.2 | 0.68 | 4.59 | 4.88 | 18.07 | 4.88 | 17.77 | 4.88 | 22.66 | 9.77 | 19.34 | 9.57 | 9.96 | 10.06 | 0.1 | 16.5 | 0.1 | 1.66 | 0.1 | 17.77 | 4.88 | 25.29 | 4.79 | 16.7 | 4.59 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
109.728 | 0 | 10.74 | 0.1 | 2.64 | 0.1 | 1.46 | 4.59 | 4.88 | 13.28 | 4.88 | 18.16 | 4.88 | 22.95 | 9.77 | 19.82 | 9.57 | 9.86 | 10.16 | 0.1 | 4.1 | 0.1 | 2.44 | 0.1 | 17.77 | 4.88 | 25.88 | 4.79 | 17.38 | 4.59 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
109.76 | 0.1 | 11.43 | 0 | 3.81 | 0.1 | 2.73 | 4.49 | 4.88 | 3.61 | 4.79 | 18.16 | 4.88 | 23.44 | 9.77 | 20.22 | 9.57 | 9.86 | 9.96 | 0.1 | 2.15 | 0 | 2.05 | 0.1 | 17.68 | 4.88 | 26.37 | 4.79 | 18.16 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
109.792 | 0.1 | 11.91 | 0.1 | 4.49 | 0.1 | 3.81 | 4.49 | 4.88 | 3.13 | 4.79 | 18.07 | 4.88 | 23.73 | 9.77 | 20.61 | 9.57 | 10.64 | 9.86 | 0.1 | 2.15 | 0 | 1.66 | 0.1 | 14.06 | 4.88 | 23.24 | 4.79 | 18.65 | 4.59 | 4.88 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
109.824 | 0.1 | 12.5 | 0.1 | 5.18 | 0.1 | 4.49 | 4.49 | 4.88 | 3.71 | 4.88 | 18.16 | 4.88 | 24.12 | 9.77 | 21 | 9.67 | 11.13 | 9.86 | 0.1 | 1.86 | 0.1 | 1.66 | 0.1 | 3.32 | 4.88 | 22.27 | 4.79 | 19.43 | 4.69 | 5.86 | 9.77 | 9.86 | 9.86 | 9.77 | 9.77 | 9.86
109.856 | 0.1 | 13.09 | 0.1 | 5.76 | 0.1 | 5.18 | 4.49 | 4.88 | 4.79 | 4.88 | 18.16 | 4.88 | 24.12 | 9.77 | 21.49 | 9.67 | 12.11 | 9.96 | 0.1 | 1.86 | 0.1 | 1.46 | 0.1 | 2.25 | 4.88 | 21.78 | 4.79 | 20.02 | 4.69 | 7.23 | 9.77 | 10.94 | 9.77 | 9.86 | 9.77 | 9.86
109.888 | 0 | 13.57 | 0.1 | 6.25 | 0.1 | 5.66 | 4.59 | 4.88 | 5.86 | 4.88 | 18.16 | 4.88 | 22.07 | 9.77 | 21.88 | 9.67 | 13.48 | 9.86 | 0.1 | 1.56 | 0.1 | 1.37 | 0.1 | 2.25 | 4.88 | 21.58 | 4.79 | 20.61 | 4.69 | 9.38 | 9.67 | 12.01 | 9.77 | 10.74 | 9.77 | 9.96
109.92 | 0.1 | 14.16 | 0.1 | 6.84 | 0.1 | 6.25 | 4.59 | 4.88 | 5.76 | 4.79 | 18.16 | 4.79 | 21.68 | 9.77 | 22.27 | 9.67 | 14.26 | 9.67 | 0.1 | 1.46 | 0.1 | 1.27 | 0 | 1.66 | 4.88 | 21.49 | 4.79 | 21.19 | 4.69 | 11.23 | 9.67 | 13.87 | 9.77 | 12.11 | 9.67 | 11.04
109.952 | 0.1 | 14.65 | 0.1 | 7.42 | 0.1 | 6.84 | 4.49 | 4.88 | 5.37 | 4.79 | 18.07 | 4.79 | 21.78 | 9.77 | 22.66 | 9.67 | 14.84 | 9.67 | 0.1 | 1.37 | 0.1 | 1.37 | 0 | 1.76 | 4.88 | 21.49 | 4.79 | 21.68 | 4.69 | 12.01 | 9.67 | 15.43 | 9.67 | 13.87 | 9.67 | 12.5
109.984 | 0.1 | 15.14 | 0.1 | 8.11 | 0.1 | 7.52 | 4.59 | 4.88 | 5.08 | 4.88 | 18.16 | 4.88 | 21.68 | 9.86 | 22.95 | 9.67 | 15.33 | 9.77 | 0.1 | 1.27 | 0.1 | 0.98 | 0.1 | 1.66 | 4.88 | 21.29 | 4.88 | 22.36 | 4.79 | 12.99 | 9.67 | 16.6 | 9.67 | 15.63 | 9.67 | 14.55
110.016 | 0.1 | 15.63 | 0.1 | 8.69 | 0.1 | 8.11 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.68 | 9.86 | 23.34 | 9.67 | 15.82 | 9.67 | 0.1 | 1.07 | 0.1 | 0.98 | 0.1 | 1.46 | 4.88 | 21.39 | 4.88 | 22.95 | 4.79 | 13.77 | 9.67 | 17.48 | 9.67 | 16.7 | 9.67 | 16.02
110.048 | 0 | 16.11 | 0.1 | 9.28 | 0.1 | 8.69 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.86 | 23.83 | 9.67 | 16.41 | 9.67 | 0.1 | 0.98 | 0.1 | 0.78 | 0.1 | 1.56 | 4.88 | 21.29 | 4.88 | 23.44 | 4.79 | 14.55 | 9.67 | 18.26 | 9.67 | 17.58 | 9.67 | 16.99
110.08 | 0.1 | 16.6 | 0 | 9.86 | 0.1 | 9.38 | 4.69 | 5.57 | 4.88 | 4.79 | 18.07 | 4.79 | 21.49 | 9.86 | 24.22 | 9.67 | 16.99 | 9.67 | 0.1 | 0.88 | 0 | 0.88 | 0.1 | 1.27 | 4.88 | 21.19 | 4.79 | 24.02 | 4.79 | 15.43 | 9.67 | 19.04 | 9.67 | 18.36 | 9.67 | 17.77
110.112 | 0.1 | 17.09 | 0.1 | 10.45 | 0.1 | 9.86 | 4.69 | 6.25 | 4.88 | 4.88 | 16.41 | 4.79 | 21.49 | 9.77 | 24.41 | 9.67 | 17.38 | 9.67 | 0.1 | 0.78 | 0 | 0.68 | 0 | 1.07 | 4.88 | 21.19 | 4.88 | 24.41 | 4.79 | 16.02 | 9.67 | 19.63 | 9.57 | 18.95 | 9.67 | 18.56
110.144 | 0.1 | 13.77 | 0.1 | 11.04 | 0.1 | 10.45 | 4.69 | 7.62 | 4.79 | 4.88 | 4.98 | 4.88 | 21.49 | 9.86 | 22.27 | 9.67 | 17.87 | 9.57 | 0.88 | 0.88 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.19 | 4.88 | 25 | 4.79 | 16.8 | 9.67 | 20.31 | 9.67 | 19.63 | 9.67 | 19.24
110.176 | 0.1 | 13.97 | 0.1 | 11.62 | 0.1 | 11.04 | 4.79 | 8.69 | 4.79 | 4.88 | 2.93 | 4.79 | 21.39 | 9.86 | 22.07 | 9.77 | 18.26 | 9.57 | 1.86 | 0.68 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 21.19 | 4.88 | 25.49 | 4.79 | 17.38 | 9.67 | 20.9 | 9.67 | 20.22 | 9.67 | 19.82
110.208 | 0 | 14.16 | 0.1 | 12.21 | 0.1 | 11.62 | 4.79 | 9.57 | 4.69 | 4.88 | 3.52 | 4.79 | 21.39 | 9.86 | 21.88 | 9.77 | 18.65 | 9.57 | 3.52 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 11.62 | 4.88 | 25.98 | 4.79 | 18.07 | 9.67 | 21.49 | 9.67 | 20.7 | 9.67 | 20.31
110.24 | 0.1 | 14.45 | 0 | 12.7 | 0 | 12.31 | 4.79 | 10.25 | 4.59 | 4.79 | 4.79 | 4.79 | 21.39 | 9.77 | 21.68 | 9.67 | 19.24 | 9.57 | 5.27 | 0.49 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 2.73 | 4.88 | 21.78 | 4.79 | 18.75 | 9.67 | 22.17 | 9.67 | 21.49 | 9.67 | 21
110.272 | 0.1 | 14.36 | 0 | 13.28 | 0.1 | 12.7 | 4.79 | 10.94 | 4.59 | 4.79 | 5.76 | 4.79 | 21.39 | 9.77 | 21.68 | 9.77 | 19.53 | 9.57 | 6.35 | 0.49 | 0 | 0.29 | 0.1 | 0.78 | 4.88 | 3.13 | 4.88 | 20.51 | 4.79 | 19.24 | 9.67 | 22.56 | 9.67 | 22.07 | 9.67 | 21.58
110.304 | 0.1 | 14.36 | 0.1 | 13.77 | 0.1 | 13.28 | 4.79 | 11.43 | 4.59 | 4.88 | 6.05 | 4.88 | 12.89 | 9.86 | 21.58 | 9.77 | 20.02 | 9.57 | 7.32 | 0.49 | 0.29 | 0.29 | 0.1 | 0.59 | 4.88 | 3.61 | 4.88 | 21.19 | 4.79 | 19.92 | 9.77 | 23.05 | 9.67 | 22.66 | 9.67 | 22.17
110.336 | 0.1 | 14.45 | 0.1 | 14.36 | 0.1 | 13.77 | 4.79 | 12.01 | 4.59 | 4.79 | 5.37 | 4.79 | 6.93 | 9.86 | 21.49 | 9.77 | 20.51 | 9.57 | 8.11 | 0.39 | 1.46 | 0.29 | 0.1 | 0.59 | 4.88 | 4.59 | 4.88 | 21.29 | 4.88 | 20.51 | 9.67 | 23.63 | 9.67 | 23.15 | 9.67 | 22.66
110.368 | 0 | 14.45 | 0.1 | 14.94 | 0.1 | 14.36 | 4.79 | 12.6 | 4.59 | 4.88 | 5.18 | 4.79 | 8.2 | 9.86 | 21.49 | 9.77 | 20.9 | 9.57 | 8.79 | 0.29 | 3.03 | 0.2 | 1.17 | 0.59 | 4.88 | 5.96 | 4.88 | 21.29 | 4.88 | 21.09 | 9.77 | 24.22 | 9.67 | 23.73 | 9.67 | 23.24
110.4 | 0.1 | 14.36 | 0 | 15.43 | 0.1 | 14.94 | 4.79 | 13.18 | 4.49 | 4.79 | 4.98 | 4.79 | 11.33 | 9.77 | 21.49 | 9.77 | 21.29 | 9.57 | 9.67 | 0.2 | 5.08 | 0.2 | 2.34 | 0.49 | 4.88 | 5.96 | 4.79 | 21.19 | 4.79 | 21.78 | 9.77 | 24.71 | 9.67 | 24.22 | 9.67 | 23.73
110.432 | 0.1 | 14.36 | 0.1 | 15.82 | 0 | 15.33 | 4.88 | 13.67 | 4.49 | 4.79 | 4.98 | 4.79 | 11.13 | 9.77 | 21.39 | 9.77 | 21.68 | 9.57 | 10.45 | 0.2 | 6.25 | 0.2 | 4.1 | 0.39 | 4.88 | 5.66 | 4.88 | 21.19 | 4.79 | 22.36 | 9.77 | 25.2 | 9.67 | 24.61 | 9.67 | 24.22
110.464 | 0.1 | 14.36 | 0.1 | 16.31 | 0.1 | 15.92 | 4.88 | 14.26 | 4.49 | 5.18 | 4.88 | 4.88 | 10.25 | 9.86 | 21.49 | 9.86 | 22.07 | 9.67 | 11.23 | 0.2 | 7.13 | 0.1 | 5.66 | 0.29 | 4.88 | 5.47 | 4.88 | 21.29 | 4.88 | 22.85 | 9.77 | 25.68 | 9.67 | 25.29 | 9.77 | 24.9
110.496 | 0.1 | 14.36 | 0.1 | 16.8 | 0.1 | 16.41 | 4.88 | 14.75 | 4.59 | 6.05 | 4.79 | 4.79 | 9.86 | 9.86 | 21.49 | 9.86 | 22.46 | 9.67 | 11.91 | 0.2 | 8.01 | 0.1 | 6.64 | 0.29 | 4.88 | 5.08 | 4.88 | 21.29 | 4.88 | 23.34 | 9.77 | 26.17 | 9.77 | 25.59 | 9.77 | 25.39
110.528 | 0 | 14.36 | 0.1 | 17.29 | 0.1 | 16.8 | 4.88 | 15.23 | 4.49 | 7.03 | 4.79 | 4.79 | 10.06 | 9.77 | 21.39 | 9.86 | 22.85 | 9.67 | 12.7 | 0.2 | 8.69 | 0.1 | 7.62 | 0.29 | 4.88 | 4.98 | 4.88 | 21.29 | 4.88 | 23.93 | 9.77 | 26.76 | 9.77 | 26.17 | 9.77 | 25.88
110.56 | 0.1 | 14.16 | 0 | 15.43 | 0.1 | 17.38 | 4.88 | 15.72 | 4.59 | 8.4 | 4.79 | 4.79 | 10.25 | 9.77 | 12.89 | 9.77 | 23.24 | 9.67 | 13.57 | 0.1 | 9.47 | 0.1 | 8.3 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 24.41 | 9.77 | 27.25 | 9.77 | 26.76 | 9.77 | 26.27
110.592 | 0.1 | 4.69 | 0.1 | 15.33 | 0.1 | 15.43 | 4.88 | 16.21 | 4.59 | 9.38 | 4.69 | 4.79 | 10.16 | 9.77 | 7.13 | 9.86 | 23.54 | 9.67 | 14.26 | 0.1 | 10.25 | 0.1 | 9.18 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 25 | 9.86 | 27.64 | 9.77 | 27.25 | 9.77 | 26.86
110.624 | 0.1 | 1.76 | 0.1 | 14.84 | 0.1 | 15.23 | 4.88 | 16.7 | 4.59 | 9.96 | 4.59 | 4.79 | 9.96 | 9.77 | 8.2 | 9.86 | 23.93 | 9.67 | 15.04 | 0.1 | 11.04 | 0.1 | 9.96 | 0.2 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 25.39 | 9.86 | 28.13 | 9.77 | 27.64 | 9.77 | 27.34
110.656 | 0.1 | 1.95 | 0.1 | 14.75 | 0.1 | 14.84 | 4.88 | 17.19 | 4.69 | 10.55 | 4.59 | 4.79 | 9.86 | 9.86 | 11.13 | 9.86 | 24.32 | 9.67 | 15.72 | 0.1 | 11.82 | 0.1 | 10.74 | 0.1 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.88 | 9.86 | 28.52 | 9.77 | 28.13 | 9.77 | 27.83
110.688 | 0 | 1.66 | 0.1 | 14.75 | 0.1 | 14.65 | 4.88 | 17.68 | 4.69 | 11.13 | 4.59 | 4.79 | 9.96 | 9.86 | 11.33 | 9.86 | 21.39 | 9.67 | 16.41 | 0.1 | 12.6 | 0 | 11.62 | 0.1 | 4.88 | 4.88 | 4.88 | 7.32 | 4.88 | 23.15 | 9.86 | 29.1 | 9.77 | 28.61 | 9.77 | 28.22
110.72 | 0.1 | 1.37 | 0.1 | 14.55 | 0 | 14.65 | 4.88 | 18.07 | 4.69 | 11.82 | 4.49 | 4.79 | 9.86 | 9.77 | 10.35 | 9.86 | 22.17 | 9.67 | 16.99 | 0.1 | 13.38 | 0 | 12.4 | 0.1 | 4.88 | 4.69 | 4.79 | 2.34 | 4.88 | 20.22 | 9.77 | 28.61 | 9.77 | 29.1 | 9.77 | 28.71
110.752 | 0.1 | 1.37 | 0 | 14.55 | 0.1 | 14.55 | 4.88 | 18.56 | 4.69 | 12.31 | 4.49 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 21.49 | 9.67 | 17.58 | 0.1 | 14.06 | 0.1 | 13.09 | 0.1 | 4.88 | 4.69 | 4.88 | 3.42 | 4.88 | 21 | 9.86 | 25.49 | 9.77 | 28.32 | 9.77 | 29.1
110.784 | 0.1 | 1.17 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 19.14 | 4.79 | 12.89 | 4.59 | 4.79 | 9.77 | 9.86 | 9.96 | 9.86 | 21.58 | 9.67 | 18.26 | 0.1 | 14.75 | 0.1 | 13.87 | 0.1 | 4.88 | 4.69 | 4.88 | 3.91 | 4.88 | 21.19 | 9.86 | 24.9 | 9.86 | 24.71 | 9.86 | 24.61
110.816 | 0.1 | 0.98 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 19.43 | 4.79 | 13.38 | 4.49 | 4.79 | 9.77 | 9.86 | 10.25 | 9.86 | 21.58 | 9.67 | 18.95 | 0.1 | 15.63 | 0.1 | 14.65 | 0.1 | 6.05 | 4.69 | 4.88 | 5.27 | 4.88 | 21.19 | 9.86 | 24.61 | 9.86 | 24.51 | 9.86 | 23.24
110.848 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.92 | 4.79 | 13.97 | 4.49 | 5.18 | 9.77 | 9.86 | 10.16 | 9.86 | 21.49 | 9.77 | 19.53 | 0.1 | 16.21 | 0 | 15.33 | 0.1 | 6.93 | 4.59 | 4.88 | 6.15 | 4.88 | 21.19 | 9.86 | 24.61 | 9.77 | 24.41 | 9.77 | 23.93
110.88 | 0.1 | 0.88 | 0 | 14.36 | 0 | 14.45 | 4.88 | 20.22 | 4.79 | 14.55 | 4.49 | 5.96 | 9.67 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 20.22 | 0.1 | 16.99 | 0 | 16.02 | 0.1 | 8.98 | 4.49 | 4.79 | 5.96 | 4.88 | 21.19 | 9.86 | 24.41 | 9.86 | 24.22 | 9.86 | 24.12
110.912 | 0.1 | 0.78 | 0 | 14.45 | 0 | 14.36 | 4.88 | 20.7 | 4.79 | 15.04 | 4.49 | 6.93 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.67 | 20.8 | 0 | 17.58 | 0.1 | 16.6 | 0.1 | 10.55 | 4.49 | 4.88 | 5.47 | 4.88 | 21.19 | 9.86 | 24.41 | 9.77 | 24.32 | 9.86 | 24.12
110.944 | 0.1 | 0.68 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 21 | 4.79 | 15.43 | 4.59 | 8.3 | 9.67 | 10.55 | 9.96 | 9.86 | 21.49 | 9.77 | 21.39 | 0.1 | 18.46 | 0.1 | 17.48 | 0.1 | 11.91 | 4.59 | 4.88 | 5.27 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86 | 24.12
110.976 | 0.1 | 0.59 | 0.1 | 10.06 | 0.1 | 14.36 | 4.88 | 18.95 | 4.79 | 15.92 | 4.59 | 9.38 | 9.67 | 11.23 | 9.86 | 9.86 | 21.49 | 9.77 | 21.97 | 0.1 | 18.85 | 0.1 | 18.07 | 0.1 | 12.79 | 4.59 | 4.88 | 5.08 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.12
111.008 | 0 | 0.59 | 0.1 | 2.83 | 0.1 | 9.18 | 4.88 | 18.75 | 4.79 | 16.41 | 4.59 | 9.96 | 9.67 | 12.31 | 9.86 | 9.86 | 21.39 | 9.77 | 22.56 | 0.1 | 19.53 | 0 | 18.75 | 0.1 | 13.67 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22
111.04 | 0.1 | 0.49 | 0 | 1.86 | 0.1 | 2.25 | 4.88 | 18.46 | 4.79 | 16.99 | 4.59 | 10.55 | 9.57 | 13.57 | 9.77 | 9.86 | 21.39 | 9.77 | 19.43 | 0.1 | 20.12 | 0 | 19.24 | 0.1 | 14.55 | 4.49 | 4.79 | 5.08 | 4.88 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
111.072 | 0.1 | 0.39 | 0 | 1.76 | 0 | 1.95 | 4.88 | 18.36 | 4.79 | 17.38 | 4.59 | 11.13 | 9.57 | 14.26 | 9.77 | 9.86 | 21.39 | 9.77 | 18.46 | 0.1 | 20.7 | 0.1 | 19.92 | 0.1 | 15.23 | 4.49 | 4.88 | 4.98 | 4.88 | 21.09 | 9.86 | 24.22 | 9.77 | 24.12 | 9.86 | 24.22
111.104 | 0.1 | 0.39 | 0.1 | 1.46 | 0.1 | 1.76 | 4.88 | 18.36 | 4.88 | 17.77 | 4.69 | 11.82 | 9.67 | 14.94 | 9.77 | 9.86 | 19.14 | 9.77 | 18.07 | 0.1 | 21.39 | 0.1 | 20.61 | 0.1 | 16.02 | 4.49 | 4.88 | 4.88 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
111.136 | 1.07 | 0.29 | 0.1 | 1.46 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.26 | 4.69 | 12.31 | 9.67 | 15.43 | 9.67 | 9.96 | 8.69 | 9.77 | 17.87 | 0.1 | 21.97 | 0.1 | 21.19 | 0.1 | 16.7 | 4.49 | 4.88 | 4.98 | 4.88 | 7.32 | 9.86 | 24.22 | 9.86 | 24.12 | 9.86 | 24.12
111.168 | 1.86 | 0.29 | 0.1 | 1.27 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.75 | 4.79 | 12.89 | 9.67 | 15.92 | 9.67 | 9.86 | 7.23 | 9.77 | 17.87 | 0.1 | 22.46 | 0 | 21.68 | 0.1 | 17.29 | 4.59 | 4.88 | 4.88 | 4.88 | 2.34 | 9.86 | 13.18 | 9.86 | 24.12 | 9.86 | 24.12
111.2 | 3.22 | 0.2 | 0.1 | 1.17 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.24 | 4.69 | 13.48 | 9.67 | 16.41 | 9.67 | 9.86 | 10.45 | 9.77 | 17.77 | 0.1 | 18.07 | 0 | 22.36 | 0.1 | 17.97 | 4.59 | 5.08 | 4.79 | 4.79 | 3.32 | 9.86 | 5.76 | 9.77 | 11.82 | 9.86 | 24.02
111.232 | 4.1 | 0.1 | 0.1 | 1.07 | 0.1 | 1.07 | 4.88 | 18.07 | 4.88 | 19.53 | 4.79 | 13.97 | 9.67 | 16.9 | 9.67 | 9.86 | 11.52 | 9.77 | 17.58 | 0 | 16.8 | 0.1 | 22.17 | 0.1 | 18.56 | 4.59 | 6.25 | 4.69 | 4.88 | 4 | 9.86 | 7.81 | 9.77 | 5.86 | 9.86 | 10.25
111.264 | 4.88 | 0.1 | 0.1 | 0.98 | 0.1 | 0.98 | 4.88 | 18.07 | 4.88 | 20.02 | 4.79 | 14.45 | 9.67 | 17.38 | 9.67 | 9.96 | 10.64 | 9.77 | 17.68 | 0.1 | 17.58 | 0.1 | 18.75 | 0.1 | 19.24 | 4.59 | 7.62 | 4.79 | 4.88 | 5.18 | 9.86 | 11.62 | 9.86 | 8.01 | 9.86 | 6.05
111.296 | 5.37 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 15.04 | 9.67 | 17.87 | 9.67 | 9.86 | 9.96 | 9.77 | 17.58 | 0.1 | 17.58 | 0.1 | 18.36 | 0.1 | 19.92 | 4.59 | 9.77 | 4.69 | 4.88 | 6.35 | 9.86 | 11.72 | 9.77 | 11.82 | 9.86 | 9.28
111.328 | 5.86 | 0.1 | 0.1 | 0.88 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.8 | 4.79 | 15.53 | 9.67 | 18.36 | 9.67 | 9.86 | 9.86 | 9.77 | 17.58 | 0.1 | 17.68 | 0 | 18.07 | 0.1 | 20.51 | 4.69 | 11.33 | 4.59 | 4.88 | 5.86 | 9.86 | 10.45 | 9.86 | 11.72 | 9.86 | 12.31
111.36 | 6.54 | 0 | 0 | 0.59 | 0.1 | 0.68 | 4.88 | 18.07 | 4.88 | 17.87 | 4.79 | 16.02 | 9.67 | 18.75 | 9.57 | 9.86 | 10.16 | 9.77 | 17.58 | 0.1 | 17.68 | 0 | 17.97 | 0.1 | 21.19 | 4.69 | 12.4 | 4.59 | 4.79 | 5.57 | 9.86 | 10.06 | 9.77 | 10.16 | 9.86 | 11.23
111.392 | 7.13 | 0 | 0 | 0.59 | 0.1 | 0.68 | 4.88 | 9.38 | 4.88 | 17.97 | 4.79 | 16.5 | 9.67 | 19.24 | 9.57 | 9.86 | 10.16 | 9.77 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 21.68 | 4.69 | 13.09 | 4.59 | 4.88 | 5.18 | 9.86 | 10.35 | 9.77 | 9.86 | 9.86 | 10.25
111.424 | 7.72 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 3.03 | 4.88 | 18.07 | 4.88 | 16.9 | 9.67 | 19.73 | 9.67 | 9.96 | 9.96 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.27 | 4.69 | 14.06 | 4.59 | 4.88 | 5.18 | 9.86 | 10.55 | 9.86 | 10.25 | 9.86 | 9.96
111.456 | 8.3 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 18.07 | 4.88 | 17.38 | 9.67 | 20.12 | 9.67 | 9.96 | 9.96 | 9.86 | 11.62 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.85 | 4.69 | 14.84 | 4.59 | 4.88 | 5.18 | 9.86 | 10.25 | 9.86 | 10.45 | 9.86 | 10.16
111.488 | 8.89 | 0.1 | 0.1 | 0.39 | 0.1 | 0.49 | 4.88 | 3.91 | 4.88 | 18.07 | 4.88 | 17.87 | 9.77 | 20.51 | 9.67 | 9.86 | 9.96 | 9.86 | 2.93 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.34 | 4.69 | 15.63 | 4.59 | 5.08 | 5.18 | 9.86 | 10.16 | 9.86 | 10.45 | 9.86 | 10.45
111.52 | 9.57 | 0 | 0 | 0.29 | 0.1 | 0.39 | 4.88 | 5.27 | 4.88 | 18.07 | 4.79 | 18.26 | 9.67 | 20.9 | 9.67 | 9.86 | 9.86 | 9.77 | 2.25 | 0 | 17.68 | 0.1 | 17.77 | 0 | 23.93 | 4.69 | 16.21 | 4.59 | 6.35 | 4.98 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.35
111.552 | 10.16 | 0 | 0.2 | 0.29 | 0.1 | 0.29 | 4.88 | 5.96 | 4.88 | 18.07 | 4.79 | 18.75 | 9.77 | 21.29 | 9.67 | 9.96 | 9.86 | 9.77 | 2.34 | 0 | 17.58 | 0.1 | 17.77 | 0.1 | 24.32 | 4.69 | 16.9 | 4.49 | 7.72 | 4.98 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.16
111.584 | 10.64 | 0.1 | 1.07 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 19.24 | 9.77 | 21.78 | 9.67 | 9.96 | 9.77 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 24.9 | 4.79 | 17.58 | 4.59 | 9.86 | 4.98 | 9.86 | 10.16 | 9.86 | 10.06 | 9.86 | 10.06
111.616 | 11.33 | 0.1 | 1.95 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.63 | 9.77 | 22.07 | 9.67 | 10.64 | 9.77 | 9.86 | 1.86 | 0.1 | 17.38 | 0.1 | 17.68 | 0.1 | 25.49 | 4.79 | 18.26 | 4.59 | 11.33 | 4.79 | 9.86 | 9.96 | 9.86 | 10.16 | 9.86 | 10.06
111.648 | 11.82 | 0.1 | 3.22 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.12 | 9.77 | 22.46 | 9.67 | 11.43 | 9.77 | 9.86 | 1.46 | 0.1 | 4.59 | 0 | 17.68 | 0.1 | 25.98 | 4.79 | 18.85 | 4.59 | 12.4 | 4.79 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
111.68 | 12.4 | 0 | 4.3 | 0.1 | 0 | 0.1 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 20.51 | 9.77 | 22.95 | 9.67 | 12.7 | 9.67 | 9.86 | 1.37 | 0 | 2.05 | 0 | 7.81 | 0 | 25.68 | 4.79 | 19.53 | 4.59 | 13.38 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
111.712 | 12.99 | 0.1 | 4.88 | 0 | 0.2 | 0.1 | 4.88 | 4.79 | 4.88 | 17.97 | 4.88 | 20.8 | 9.77 | 23.24 | 9.67 | 13.67 | 9.67 | 9.86 | 1.37 | 0 | 2.15 | 0.1 | 1.95 | 0.1 | 22.85 | 4.79 | 20.12 | 4.59 | 14.16 | 4.59 | 9.86 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
111.744 | 13.48 | 0.1 | 5.37 | 0.1 | 1.17 | 0.1 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 17.68 | 9.77 | 23.63 | 9.77 | 14.36 | 9.67 | 9.86 | 1.27 | 0.1 | 1.95 | 0.1 | 2.34 | 0.1 | 22.07 | 4.79 | 20.7 | 4.69 | 14.84 | 4.69 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
111.776 | 14.16 | 0.1 | 5.96 | 0.1 | 2.05 | 0 | 5.08 | 4.88 | 4.88 | 17.77 | 4.88 | 17.97 | 9.77 | 24.02 | 9.77 | 15.04 | 9.67 | 9.86 | 1.17 | 0.1 | 1.76 | 0.1 | 2.34 | 0.1 | 21.78 | 4.79 | 21.29 | 4.69 | 15.53 | 4.59 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86
111.808 | 14.55 | 0.1 | 6.54 | 0.1 | 3.42 | 0 | 5.96 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 24.41 | 9.77 | 15.53 | 9.67 | 9.86 | 1.07 | 0.1 | 1.56 | 0 | 1.76 | 0.1 | 21.58 | 4.79 | 21.97 | 4.69 | 16.21 | 4.59 | 10.16 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
111.84 | 15.04 | 0 | 7.13 | 0 | 4.39 | 0.1 | 7.03 | 4.69 | 4.79 | 3.03 | 4.88 | 18.07 | 9.77 | 22.36 | 9.77 | 16.11 | 9.67 | 9.86 | 0.98 | 0 | 1.56 | 0 | 1.86 | 0.1 | 21.39 | 4.79 | 22.56 | 4.69 | 16.9 | 4.59 | 11.33 | 9.77 | 9.86 | 9.67 | 10.06 | 9.77
111.872 | 15.53 | 0 | 7.72 | 0.1 | 4.98 | 0 | 8.3 | 4.69 | 4.88 | 3.52 | 4.88 | 18.07 | 9.77 | 22.27 | 9.77 | 16.6 | 9.57 | 9.86 | 0.78 | 0 | 1.37 | 0.1 | 1.46 | 0.1 | 21.49 | 4.79 | 22.95 | 4.69 | 17.58 | 4.49 | 12.7 | 9.67 | 9.96 | 9.67 | 10.94 | 9.86
111.904 | 16.02 | 0.1 | 8.3 | 0.1 | 5.47 | 0.1 | 9.38 | 4.59 | 4.88 | 4.69 | 4.88 | 18.07 | 9.86 | 21.88 | 9.77 | 17.09 | 9.57 | 9.86 | 0.88 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 21.39 | 4.79 | 23.54 | 4.69 | 18.26 | 4.59 | 14.84 | 9.67 | 10.84 | 9.67 | 12.31 | 9.77
111.936 | 16.5 | 0.1 | 8.89 | 0.1 | 5.96 | 0.1 | 9.96 | 4.59 | 4.88 | 5.76 | 4.88 | 18.16 | 9.86 | 21.68 | 9.77 | 17.48 | 9.57 | 9.86 | 0.68 | 0.1 | 1.17 | 0.1 | 1.27 | 0.1 | 21.39 | 4.79 | 24.02 | 4.79 | 18.85 | 4.49 | 16.21 | 9.67 | 12.21 | 9.67 | 14.26 | 9.67
111.968 | 16.99 | 0.1 | 9.57 | 0.1 | 6.64 | 0.1 | 10.55 | 4.59 | 4.88 | 5.76 | 4.88 | 18.07 | 9.77 | 21.68 | 9.77 | 17.87 | 9.57 | 9.86 | 0.68 | 0.1 | 0.98 | 0 | 1.27 | 0.1 | 21.39 | 4.79 | 24.61 | 4.69 | 19.53 | 4.49 | 17.09 | 9.67 | 14.26 | 9.67 | 15.82 | 9.67
112 | 15.23 | 0 | 10.16 | 0 | 7.23 | 0 | 11.23 | 4.49 | 4.88 | 5.37 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.46 | 9.57 | 9.86 | 0.49 | 0 | 0.88 | 0 | 1.17 | 0 | 21.29 | 4.79 | 25 | 4.79 | 20.12 | 4.49 | 17.97 | 9.67 | 15.92 | 9.67 | 16.99 | 9.67
112.032 | 14.06 | 0.1 | 10.64 | 0.1 | 7.81 | 0 | 11.82 | 4.49 | 4.88 | 5.18 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.85 | 9.57 | 9.86 | 0.49 | 0 | 0.88 | 0 | 0.98 | 0.1 | 21.29 | 4.79 | 25.68 | 4.79 | 20.7 | 4.49 | 18.56 | 9.67 | 16.9 | 9.57 | 17.68 | 9.67
112.064 | 14.26 | 0.1 | 11.33 | 0.1 | 8.4 | 0.1 | 12.4 | 4.59 | 4.88 | 5.08 | 4.88 | 17.97 | 9.86 | 21.49 | 9.86 | 19.34 | 9.67 | 9.86 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.79 | 26.08 | 4.79 | 21.29 | 4.49 | 19.34 | 9.67 | 17.58 | 9.67 | 18.46 | 9.67
112.096 | 14.45 | 0.1 | 11.91 | 0.1 | 9.08 | 0.1 | 12.89 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.82 | 9.67 | 9.86 | 0.39 | 0.1 | 0.68 | 0.1 | 0.78 | 0.1 | 21.29 | 4.79 | 21.49 | 4.79 | 21.97 | 4.59 | 20.02 | 9.67 | 18.46 | 9.57 | 19.14 | 9.67
112.128 | 14.36 | 0.1 | 12.5 | 0.1 | 9.67 | 0.1 | 13.48 | 4.49 | 4.88 | 4.88 | 4.88 | 17.97 | 9.86 | 21.39 | 9.86 | 20.12 | 9.67 | 10.16 | 0.29 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 9.86 | 4.79 | 20.31 | 4.79 | 22.56 | 4.59 | 20.61 | 9.67 | 19.14 | 9.57 | 19.73 | 9.67
112.16 | 14.45 | 0 | 12.99 | 0 | 10.16 | 0 | 13.97 | 4.49 | 4.88 | 4.79 | 4.88 | 16.11 | 9.77 | 21.39 | 9.86 | 20.61 | 9.67 | 11.04 | 0.29 | 0 | 0.59 | 0 | 0.78 | 0.1 | 2.44 | 4.79 | 21.19 | 4.79 | 23.05 | 4.59 | 21.19 | 9.67 | 19.73 | 9.67 | 20.51 | 9.67
112.192 | 14.45 | 0 | 13.48 | 0.1 | 10.74 | 0 | 14.45 | 4.49 | 4.79 | 4.79 | 4.88 | 4.69 | 9.77 | 21.39 | 9.77 | 20.9 | 9.67 | 11.91 | 0.2 | 0 | 0.49 | 0.1 | 0.59 | 0 | 3.13 | 4.79 | 21.19 | 4.79 | 23.54 | 4.59 | 21.78 | 9.67 | 20.31 | 9.57 | 20.9 | 9.67
112.224 | 14.36 | 0.1 | 14.06 | 0.1 | 11.43 | 0.1 | 15.04 | 4.59 | 4.88 | 4.79 | 4.88 | 3.13 | 9.77 | 21.39 | 9.86 | 21.39 | 9.67 | 13.28 | 0.2 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 3.71 | 4.79 | 21.19 | 4.79 | 24.12 | 4.69 | 22.36 | 9.67 | 21 | 9.67 | 21.58 | 9.67
112.256 | 14.36 | 0.1 | 14.55 | 0.1 | 11.91 | 0 | 15.63 | 4.59 | 4.88 | 4.69 | 4.88 | 3.52 | 9.86 | 14.55 | 9.86 | 21.78 | 9.67 | 14.16 | 0.2 | 0.1 | 0.39 | 0.1 | 0.59 | 0.1 | 4.59 | 4.79 | 21.19 | 4.79 | 24.61 | 4.69 | 22.85 | 9.67 | 21.49 | 9.67 | 22.17 | 9.67
112.288 | 14.36 | 0 | 15.04 | 0.1 | 12.5 | 0 | 16.02 | 4.59 | 4.88 | 4.59 | 4.88 | 4.59 | 9.86 | 7.32 | 9.86 | 22.17 | 9.67 | 14.75 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 6.05 | 4.79 | 21.29 | 4.79 | 25.1 | 4.69 | 23.44 | 9.67 | 22.17 | 9.67 | 22.75 | 9.67
112.32 | 14.36 | 0 | 15.53 | 0 | 13.09 | 0 | 16.5 | 4.59 | 4.79 | 4.59 | 4.88 | 5.76 | 9.77 | 8.11 | 9.86 | 22.56 | 9.67 | 15.23 | 0.1 | 0 | 0.29 | 0.2 | 0.39 | 0 | 6.15 | 4.79 | 21.29 | 4.79 | 25.78 | 4.69 | 24.12 | 9.67 | 22.66 | 9.67 | 23.24 | 9.67
112.352 | 14.45 | 0 | 16.02 | 0.1 | 13.57 | 0 | 16.99 | 4.59 | 4.79 | 4.49 | 4.88 | 5.86 | 9.86 | 10.74 | 9.86 | 22.95 | 9.67 | 15.82 | 0.1 | 0 | 0.2 | 1.17 | 0.29 | 0.1 | 5.57 | 4.79 | 21.29 | 4.79 | 26.08 | 4.69 | 24.51 | 9.67 | 23.24 | 9.67 | 23.83 | 9.67
112.384 | 14.36 | 0.1 | 16.6 | 0.1 | 14.16 | 0.1 | 17.48 | 4.69 | 4.88 | 4.59 | 4.88 | 5.37 | 9.77 | 11.23 | 9.86 | 23.34 | 9.67 | 16.21 | 0.2 | 0.29 | 0.2 | 2.64 | 0.29 | 0.1 | 5.37 | 4.79 | 21.29 | 4.79 | 21.49 | 4.69 | 25 | 9.67 | 23.63 | 9.67 | 24.32 | 9.67
112.416 | 14.36 | 0.1 | 16.99 | 0.1 | 14.65 | 0 | 17.87 | 4.69 | 5.66 | 4.49 | 4.88 | 5.18 | 9.86 | 10.45 | 9.86 | 23.73 | 9.77 | 16.8 | 0.1 | 1.56 | 0.2 | 4.49 | 0.29 | 0.1 | 5.08 | 4.88 | 21.29 | 4.79 | 20.41 | 4.69 | 25.49 | 9.77 | 24.22 | 9.67 | 24.9 | 9.67
112.448 | 6.25 | 0.1 | 15.33 | 0.1 | 15.14 | 0 | 18.46 | 4.69 | 6.05 | 4.49 | 4.88 | 4.88 | 9.86 | 9.86 | 9.86 | 24.02 | 9.67 | 17.29 | 0.1 | 2.93 | 0.2 | 5.96 | 0.2 | 0.1 | 5.18 | 4.79 | 21.29 | 4.79 | 21 | 4.79 | 25.98 | 9.77 | 24.81 | 9.67 | 25.49 | 9.67
112.48 | 1.76 | 0 | 13.97 | 0 | 15.63 | 0 | 18.85 | 4.69 | 7.23 | 4.49 | 4.79 | 4.88 | 9.77 | 10.16 | 9.86 | 24.41 | 9.67 | 17.77 | 0.1 | 4.98 | 0.1 | 6.84 | 0.2 | 0.88 | 5.08 | 4.79 | 21.19 | 4.79 | 21.19 | 4.79 | 26.56 | 9.77 | 25.39 | 9.67 | 25.88 | 9.67
112.512 | 1.95 | 0 | 14.26 | 0.1 | 16.11 | 0 | 19.24 | 4.79 | 8.5 | 4.49 | 4.88 | 4.79 | 9.77 | 9.96 | 9.77 | 22.36 | 9.67 | 18.16 | 0.1 | 6.15 | 0.1 | 7.72 | 0.1 | 1.76 | 4.98 | 4.79 | 21.19 | 4.79 | 21.29 | 4.69 | 26.95 | 9.77 | 25.88 | 9.67 | 26.37 | 9.67
112.544 | 1.76 | 0.1 | 14.55 | 0.1 | 16.6 | 0.1 | 19.73 | 4.79 | 9.38 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 22.17 | 9.77 | 18.65 | 0.1 | 7.13 | 0.1 | 8.59 | 0.2 | 3.22 | 5.08 | 4.79 | 5.66 | 4.79 | 21.19 | 4.79 | 27.44 | 9.77 | 26.17 | 9.67 | 26.86 | 9.77
112.576 | 1.46 | 0.1 | 14.45 | 0.1 | 17.09 | 0.1 | 20.12 | 4.79 | 10.06 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.88 | 9.77 | 19.14 | 0.1 | 7.81 | 0.1 | 9.38 | 0.1 | 5.18 | 4.98 | 4.79 | 2.73 | 4.79 | 21.29 | 4.79 | 28.03 | 9.77 | 26.76 | 9.77 | 27.34 | 9.77
112.608 | 1.37 | 0.1 | 14.45 | 0.1 | 13.87 | 0 | 20.61 | 4.79 | 10.64 | 4.49 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.78 | 9.77 | 19.43 | 0.1 | 8.69 | 0.1 | 10.16 | 0.1 | 6.45 | 4.88 | 4.79 | 3.42 | 4.79 | 21.29 | 4.79 | 28.42 | 9.77 | 27.25 | 9.77 | 27.83 | 9.77
112.64 | 1.17 | 0 | 14.45 | 0 | 13.97 | 0.1 | 21 | 4.79 | 11.23 | 4.59 | 4.79 | 4.69 | 9.77 | 9.86 | 9.86 | 21.68 | 9.77 | 20.02 | 0.1 | 9.47 | 0.1 | 10.94 | 0.1 | 7.23 | 4.79 | 4.79 | 4.3 | 4.79 | 21.19 | 4.79 | 28.91 | 9.77 | 27.83 | 9.67 | 28.32 | 9.77
112.672 | 1.17 | 0 | 14.45 | 0 | 14.16 | 0 | 17.58 | 4.79 | 11.82 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.31 | 0.1 | 10.16 | 0 | 11.62 | 0.1 | 8.11 | 4.69 | 4.79 | 5.57 | 4.79 | 21.29 | 4.79 | 29.4 | 9.77 | 28.32 | 9.67 | 28.71 | 9.77
112.704 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 17.77 | 4.88 | 12.4 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.8 | 0.1 | 11.13 | 0.1 | 12.4 | 0.1 | 8.89 | 4.69 | 4.79 | 6.25 | 4.79 | 21.19 | 4.79 | 24.71 | 9.86 | 28.71 | 9.77 | 29.1 | 9.77
112.736 | 0.88 | 0.1 | 14.45 | 0.1 | 14.36 | 0 | 17.97 | 4.88 | 12.89 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 11.82 | 0.1 | 13.18 | 0.1 | 9.57 | 4.69 | 4.79 | 5.96 | 4.79 | 21.29 | 4.79 | 23.63 | 9.77 | 29.3 | 9.77 | 26.27 | 9.77
112.768 | 0.78 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 13.48 | 4.69 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.58 | 0.1 | 12.6 | 0.1 | 13.97 | 0.1 | 10.35 | 4.59 | 4.79 | 5.47 | 4.88 | 21.19 | 4.79 | 24.22 | 9.86 | 26.37 | 9.77 | 23.54 | 9.77
112.8 | 0.78 | 0 | 14.45 | 0 | 14.45 | 0 | 18.16 | 4.88 | 14.06 | 4.59 | 4.98 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 21.97 | 0.1 | 13.28 | 0.1 | 14.55 | 0.1 | 11.23 | 4.59 | 4.79 | 5.37 | 4.79 | 20.41 | 4.79 | 24.32 | 9.77 | 23.34 | 9.77 | 24.02 | 9.77
112.832 | 0.59 | 0 | 14.36 | 0 | 14.45 | 0 | 18.16 | 4.88 | 14.55 | 4.69 | 5.86 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.36 | 0.1 | 14.06 | 0 | 15.33 | 0.1 | 12.11 | 4.49 | 4.79 | 5.18 | 4.79 | 4.98 | 4.79 | 24.32 | 9.77 | 24.12 | 9.77 | 24.22 | 9.77
112.864 | 0.59 | 0.1 | 14.36 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.04 | 4.69 | 6.84 | 4.49 | 9.77 | 9.67 | 9.86 | 21.49 | 9.86 | 22.75 | 0.1 | 14.84 | 0.1 | 16.02 | 0.1 | 12.79 | 4.59 | 4.79 | 5.18 | 4.88 | 3.03 | 4.79 | 24.32 | 9.86 | 24.22 | 9.77 | 24.22 | 9.77
112.896 | 0.49 | 0.1 | 6.15 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.63 | 4.69 | 8.3 | 4.59 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 0.1 | 15.43 | 0.1 | 16.7 | 0.1 | 13.57 | 4.59 | 5.66 | 5.08 | 4.79 | 3.22 | 4.79 | 24.41 | 9.86 | 24.22 | 9.77 | 24.22 | 9.77
112.928 | 0.39 | 0.1 | 1.76 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 16.02 | 4.69 | 9.28 | 4.49 | 9.77 | 9.67 | 9.86 | 13.87 | 9.86 | 23.44 | 0.1 | 16.21 | 0.1 | 17.48 | 0.1 | 14.26 | 4.59 | 6.74 | 5.08 | 4.79 | 4.3 | 4.79 | 24.41 | 9.86 | 24.22 | 9.77 | 24.22 | 9.77
112.96 | 0.39 | 0 | 1.76 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.6 | 4.69 | 9.96 | 4.49 | 9.77 | 9.57 | 9.86 | 7.13 | 9.86 | 23.83 | 0.1 | 16.9 | 0.1 | 18.07 | 0.1 | 15.14 | 4.49 | 8.4 | 4.98 | 4.79 | 5.66 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
112.992 | 0.29 | 0 | 1.76 | 0 | 14.36 | 0 | 18.07 | 4.88 | 16.99 | 4.79 | 10.55 | 4.49 | 10.55 | 9.57 | 9.86 | 8.01 | 9.77 | 24.32 | 0.1 | 17.48 | 0 | 18.75 | 0.1 | 15.72 | 4.49 | 10.25 | 4.98 | 4.79 | 6.25 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
113.024 | 0.29 | 0 | 1.46 | 0.1 | 14.36 | 0 | 18.16 | 4.88 | 17.58 | 4.79 | 11.13 | 4.59 | 11.23 | 9.67 | 9.86 | 10.94 | 9.86 | 20.9 | 0.1 | 18.16 | 0.1 | 19.34 | 0.1 | 16.41 | 4.49 | 11.62 | 4.88 | 4.79 | 5.86 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
113.056 | 0.2 | 0.1 | 1.46 | 0.1 | 5.47 | 0 | 18.07 | 4.88 | 18.07 | 4.79 | 11.82 | 4.59 | 12.6 | 9.67 | 9.86 | 11.43 | 9.86 | 21 | 0.1 | 18.75 | 0.1 | 19.92 | 0.1 | 17.19 | 4.59 | 12.6 | 4.79 | 4.79 | 5.47 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
113.088 | 0.1 | 0.1 | 1.17 | 0.1 | 1.66 | 0 | 17.87 | 4.88 | 18.46 | 4.79 | 12.31 | 4.59 | 13.57 | 9.57 | 9.86 | 10.45 | 9.86 | 21.29 | 0.1 | 19.43 | 0.1 | 20.51 | 0.1 | 17.77 | 4.59 | 13.38 | 4.79 | 4.79 | 5.18 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
113.12 | 0.1 | 0 | 1.07 | 0 | 1.95 | 0 | 4.59 | 4.88 | 18.85 | 4.79 | 12.89 | 4.59 | 14.36 | 9.57 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 19.92 | 0.1 | 21.09 | 0.1 | 18.46 | 4.59 | 14.36 | 4.69 | 4.79 | 5.08 | 4.79 | 23.05 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
113.152 | 0.1 | 0 | 0.88 | 0.1 | 1.66 | 0 | 3.03 | 4.88 | 19.24 | 4.79 | 13.38 | 4.69 | 14.94 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 20.61 | 0.1 | 21.78 | 0.1 | 19.04 | 4.59 | 14.94 | 4.69 | 4.79 | 5.18 | 4.79 | 8.59 | 9.77 | 24.12 | 9.77 | 24.12 | 9.86
113.184 | 0.1 | 0.1 | 0.98 | 0.1 | 1.56 | 0.1 | 3.61 | 4.88 | 19.73 | 4.79 | 13.97 | 4.69 | 15.43 | 9.67 | 9.86 | 10.25 | 9.86 | 21.39 | 0.1 | 21.19 | 0.1 | 22.27 | 0.1 | 19.82 | 4.69 | 15.72 | 4.69 | 4.79 | 5.08 | 4.79 | 6.15 | 9.86 | 24.22 | 9.86 | 11.43 | 9.86
113.216 | 0 | 0.1 | 0.78 | 0.1 | 1.37 | 0 | 4.59 | 4.88 | 20.12 | 4.79 | 14.45 | 4.69 | 16.02 | 9.67 | 9.86 | 10.16 | 9.86 | 21.39 | 0.1 | 21.88 | 0.1 | 19.73 | 0.1 | 20.22 | 4.69 | 16.31 | 4.69 | 4.79 | 5.08 | 4.79 | 9.86 | 9.86 | 11.43 | 9.86 | 5.66 | 9.86
113.248 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 5.66 | 4.88 | 20.61 | 4.79 | 14.94 | 4.69 | 16.5 | 9.67 | 10.55 | 9.86 | 9.86 | 21.49 | 0.1 | 22.36 | 0.1 | 16.8 | 0.1 | 20.8 | 4.69 | 17.09 | 4.69 | 4.79 | 4.98 | 4.79 | 12.4 | 9.86 | 5.76 | 9.86 | 8.59 | 9.86
113.28 | 0.1 | 1.07 | 0.59 | 0 | 1.07 | 0 | 5.86 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 16.99 | 9.67 | 11.33 | 9.86 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.58 | 0.1 | 21.49 | 4.69 | 17.87 | 4.59 | 4.79 | 4.98 | 4.79 | 10.94 | 9.77 | 9.28 | 9.77 | 12.21 | 9.86
113.312 | 0.1 | 1.95 | 0.59 | 0.1 | 0.98 | 0 | 5.47 | 4.88 | 18.16 | 4.79 | 16.02 | 4.69 | 17.58 | 9.67 | 12.4 | 9.86 | 9.86 | 21.39 | 0.1 | 16.7 | 0 | 17.68 | 0.1 | 22.07 | 4.69 | 18.46 | 4.59 | 4.79 | 4.88 | 4.79 | 9.96 | 9.77 | 12.21 | 9.77 | 11.52 | 9.86
113.344 | 0.1 | 3.42 | 0.59 | 0.1 | 0.88 | 0.1 | 5.18 | 4.88 | 18.46 | 4.88 | 16.41 | 4.79 | 17.97 | 9.67 | 13.57 | 9.86 | 9.96 | 21.39 | 0.1 | 17.38 | 0.1 | 17.68 | 0.1 | 22.56 | 4.79 | 19.04 | 4.59 | 4.79 | 4.79 | 4.79 | 10.16 | 9.86 | 11.43 | 9.86 | 10.35 | 9.86
113.376 | 0.1 | 4.3 | 0.49 | 0.1 | 0.88 | 0 | 4.98 | 4.88 | 18.16 | 4.88 | 16.9 | 4.79 | 18.46 | 9.67 | 14.36 | 9.77 | 9.86 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 18.56 | 4.79 | 19.63 | 4.59 | 5.76 | 4.79 | 4.79 | 10.45 | 9.86 | 10.25 | 9.86 | 9.96 | 9.86
113.408 | 0 | 4.88 | 0.39 | 0.1 | 0.68 | 0 | 4.98 | 4.88 | 18.16 | 4.88 | 17.38 | 4.79 | 18.85 | 9.67 | 14.94 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.79 | 20.12 | 4.59 | 7.13 | 4.79 | 4.79 | 10.45 | 9.86 | 9.96 | 9.86 | 10.35 | 9.86
113.44 | 0.1 | 5.47 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.26 | 4.79 | 17.77 | 4.79 | 19.34 | 9.67 | 15.53 | 9.77 | 9.96 | 18.85 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 18.07 | 4.79 | 20.9 | 4.59 | 9.18 | 4.69 | 4.79 | 10.35 | 9.77 | 10.35 | 9.77 | 10.64 | 9.86
113.472 | 0.1 | 6.05 | 0.29 | 0.1 | 0.59 | 0 | 4.79 | 4.88 | 18.16 | 4.79 | 18.26 | 4.79 | 19.82 | 9.67 | 15.92 | 9.67 | 9.86 | 9.08 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.87 | 4.79 | 21.39 | 4.59 | 10.64 | 4.69 | 4.79 | 10.06 | 9.77 | 10.45 | 9.77 | 10.45 | 9.86
113.504 | 0.1 | 6.54 | 0.2 | 0.1 | 0.59 | 0 | 4.69 | 4.88 | 18.16 | 4.88 | 18.75 | 4.79 | 20.12 | 9.67 | 16.5 | 9.67 | 9.96 | 6.84 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.79 | 21.97 | 4.59 | 12.01 | 4.59 | 4.79 | 10.16 | 9.86 | 10.35 | 9.86 | 10.06 | 9.86
113.536 | 0.1 | 7.13 | 0.2 | 0.1 | 0.49 | 0 | 4.69 | 4.88 | 18.16 | 4.88 | 19.24 | 4.88 | 20.61 | 9.77 | 16.99 | 9.67 | 9.96 | 9.77 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 4.88 | 22.66 | 4.69 | 12.89 | 4.59 | 4.79 | 10.25 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
113.568 | 0 | 7.72 | 0.1 | 0.1 | 0.39 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 19.63 | 4.88 | 21 | 9.77 | 17.38 | 9.67 | 9.96 | 11.52 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.88 | 23.24 | 4.69 | 13.67 | 4.59 | 4.79 | 10.16 | 9.86 | 10.06 | 9.86 | 10.06 | 9.86
113.6 | 0.1 | 8.4 | 0.1 | 0 | 0.29 | 0 | 4.69 | 4.88 | 18.16 | 4.79 | 20.12 | 4.79 | 21.49 | 9.77 | 17.97 | 9.67 | 9.86 | 10.74 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 4.88 | 23.73 | 4.69 | 14.55 | 4.49 | 4.79 | 9.96 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86
113.632 | 0.1 | 8.98 | 0.1 | 0 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.79 | 20.51 | 4.79 | 21.78 | 9.77 | 18.36 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 4.79 | 24.32 | 4.69 | 15.23 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86
113.664 | 0.1 | 9.57 | 0.1 | 0.68 | 0.2 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 22.17 | 9.77 | 18.85 | 9.67 | 9.96 | 9.67 | 0.1 | 17.58 | 0.1 | 6.74 | 0.1 | 17.77 | 4.88 | 24.81 | 4.69 | 15.92 | 4.59 | 4.79 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
113.696 | 0.1 | 10.16 | 0.1 | 1.37 | 0.2 | 0.68 | 4.59 | 4.88 | 18.07 | 4.88 | 17.38 | 4.88 | 22.66 | 9.77 | 19.24 | 9.57 | 9.96 | 10.06 | 0.1 | 17.58 | 0.1 | 1.56 | 0.1 | 17.68 | 4.88 | 25.2 | 4.79 | 16.6 | 4.59 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86
113.728 | 0 | 10.74 | 0.1 | 2.44 | 0.1 | 1.27 | 4.59 | 4.88 | 16.21 | 4.88 | 17.68 | 4.88 | 23.05 | 9.77 | 19.73 | 9.67 | 9.96 | 10.16 | 0.1 | 5.47 | 0.1 | 2.15 | 0.1 | 17.77 | 4.88 | 25.68 | 4.79 | 17.38 | 4.59 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
113.76 | 0.1 | 11.33 | 0.1 | 3.81 | 0.1 | 2.54 | 4.49 | 4.88 | 4.2 | 4.79 | 17.87 | 4.88 | 23.34 | 9.77 | 20.12 | 9.57 | 9.86 | 9.96 | 0.1 | 1.86 | 0 | 2.15 | 0.1 | 17.68 | 4.88 | 26.17 | 4.79 | 18.07 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
113.792 | 0.1 | 11.91 | 0.1 | 4.49 | 0.1 | 3.71 | 4.49 | 4.88 | 3.03 | 4.88 | 18.07 | 4.88 | 23.73 | 9.77 | 20.51 | 9.57 | 10.74 | 9.77 | 0.1 | 2.25 | 0 | 1.86 | 0.1 | 15.14 | 4.88 | 22.95 | 4.79 | 18.56 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86
113.824 | 0.1 | 12.5 | 0.1 | 5.18 | 0.1 | 4.49 | 4.59 | 4.88 | 3.71 | 4.88 | 18.07 | 4.88 | 24.12 | 9.86 | 20.9 | 9.57 | 11.13 | 9.86 | 0.1 | 1.86 | 0.1 | 1.56 | 0.1 | 3.81 | 4.88 | 22.17 | 4.79 | 19.34 | 4.59 | 5.66 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
113.856 | 0 | 12.99 | 0.1 | 5.57 | 0.1 | 5.08 | 4.49 | 4.88 | 4.69 | 4.88 | 18.07 | 4.88 | 24.41 | 9.77 | 21.39 | 9.57 | 12.01 | 9.86 | 0.1 | 1.76 | 0.1 | 1.46 | 0.1 | 2.15 | 4.88 | 21.78 | 4.79 | 20.12 | 4.59 | 7.03 | 9.67 | 11.04 | 9.77 | 9.86 | 9.77 | 9.86
113.888 | 0 | 13.48 | 0.1 | 6.15 | 0.1 | 5.57 | 4.59 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 22.27 | 9.86 | 21.78 | 9.67 | 13.38 | 9.77 | 0.1 | 1.56 | 0.1 | 1.46 | 0.1 | 2.25 | 4.88 | 21.58 | 4.79 | 20.51 | 4.69 | 8.98 | 9.67 | 11.91 | 9.67 | 10.94 | 9.67 | 9.96
113.92 | 0.1 | 14.06 | 0.1 | 6.74 | 0.1 | 6.15 | 4.49 | 4.88 | 5.86 | 4.79 | 18.07 | 4.88 | 21.78 | 9.77 | 22.17 | 9.67 | 14.16 | 9.67 | 0.1 | 1.37 | 0.1 | 1.27 | 0.1 | 1.76 | 4.88 | 21.58 | 4.79 | 21.09 | 4.69 | 11.04 | 9.67 | 13.87 | 9.77 | 12.01 | 9.67 | 10.94
113.952 | 0.1 | 14.55 | 0.1 | 7.32 | 0.1 | 6.74 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.78 | 9.77 | 22.56 | 9.67 | 14.84 | 9.67 | 0.1 | 1.37 | 0 | 1.17 | 0.1 | 1.76 | 4.88 | 21.39 | 4.79 | 21.58 | 4.69 | 11.91 | 9.67 | 15.33 | 9.67 | 13.77 | 9.67 | 12.21
113.984 | 0.1 | 15.04 | 0.1 | 8.01 | 0.1 | 7.42 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.68 | 9.86 | 22.95 | 9.67 | 15.33 | 9.67 | 0.1 | 1.27 | 0.1 | 0.98 | 0.1 | 1.56 | 4.88 | 21.29 | 4.88 | 22.17 | 4.69 | 12.79 | 9.67 | 16.6 | 9.67 | 15.43 | 9.67 | 14.36
114.016 | 0.1 | 15.63 | 0.1 | 8.59 | 0.1 | 8.01 | 4.69 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.68 | 9.86 | 23.34 | 9.67 | 15.82 | 9.67 | 0.1 | 1.07 | 0.1 | 1.07 | 0.1 | 1.46 | 4.88 | 21.39 | 4.88 | 22.75 | 4.69 | 13.67 | 9.67 | 17.48 | 9.67 | 16.7 | 9.67 | 15.72
114.048 | 0 | 16.11 | 0.1 | 9.18 | 0.1 | 8.59 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.58 | 9.86 | 23.63 | 9.67 | 16.41 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.37 | 4.88 | 21.29 | 4.88 | 23.34 | 4.79 | 14.55 | 9.67 | 18.16 | 9.67 | 17.38 | 9.67 | 16.9
114.08 | 0.1 | 16.5 | 0.1 | 9.86 | 0.1 | 9.28 | 4.69 | 5.66 | 4.88 | 4.79 | 17.97 | 4.88 | 21.49 | 9.86 | 24.12 | 9.67 | 16.9 | 9.67 | 0.1 | 0.98 | 0 | 0.78 | 0.1 | 1.17 | 4.88 | 21.29 | 4.79 | 24.02 | 4.79 | 15.23 | 9.67 | 18.85 | 9.67 | 18.16 | 9.67 | 17.68
114.112 | 0.1 | 16.99 | 0.1 | 10.35 | 0.1 | 9.77 | 4.69 | 6.35 | 4.88 | 4.88 | 17.97 | 4.79 | 21.49 | 9.77 | 24.41 | 9.67 | 17.29 | 9.57 | 0.1 | 0.88 | 0 | 0.78 | 0.1 | 1.17 | 4.88 | 21.19 | 4.88 | 24.41 | 4.79 | 15.92 | 9.67 | 19.53 | 9.57 | 18.85 | 9.67 | 18.46
114.144 | 0.1 | 13.97 | 0.1 | 10.94 | 0.1 | 10.45 | 4.79 | 7.42 | 4.79 | 4.88 | 5.76 | 4.88 | 21.49 | 9.86 | 22.46 | 9.67 | 17.77 | 9.57 | 0.88 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.19 | 4.88 | 24.9 | 4.79 | 16.5 | 9.67 | 20.12 | 9.67 | 19.63 | 9.67 | 19.14
114.176 | 0.1 | 13.87 | 0.1 | 11.62 | 0.1 | 11.04 | 4.79 | 8.59 | 4.79 | 4.88 | 2.83 | 4.88 | 21.39 | 9.86 | 22.27 | 9.67 | 18.26 | 9.57 | 1.86 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 21.19 | 4.88 | 25.39 | 4.79 | 17.29 | 9.67 | 20.8 | 9.67 | 20.22 | 9.67 | 19.73
114.208 | 0 | 14.16 | 0.1 | 12.11 | 0.1 | 11.52 | 4.79 | 9.47 | 4.69 | 4.88 | 3.52 | 4.88 | 21.39 | 9.86 | 21.97 | 9.67 | 18.65 | 9.57 | 3.22 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 13.38 | 4.88 | 25.88 | 4.79 | 17.97 | 9.67 | 21.39 | 9.67 | 20.9 | 9.67 | 20.31
114.24 | 0.1 | 14.36 | 0.1 | 12.7 | 0.1 | 12.21 | 4.79 | 10.25 | 4.59 | 4.88 | 4.79 | 4.88 | 21.39 | 9.86 | 21.78 | 9.67 | 19.24 | 9.57 | 5.18 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 3.13 | 4.79 | 24.22 | 4.79 | 18.56 | 9.67 | 22.07 | 9.67 | 21.39 | 9.67 | 20.9
114.272 | 0.1 | 14.36 | 0.1 | 13.18 | 0.1 | 12.7 | 4.79 | 10.74 | 4.59 | 4.79 | 5.66 | 4.79 | 21.39 | 9.77 | 21.78 | 9.77 | 19.53 | 9.47 | 6.25 | 0.49 | 0 | 0.39 | 0.1 | 0.68 | 4.88 | 3.22 | 4.88 | 21.68 | 4.79 | 19.24 | 9.67 | 22.46 | 9.67 | 21.97 | 9.67 | 21.49
114.304 | 0.1 | 14.36 | 0.1 | 13.77 | 0.1 | 13.18 | 4.88 | 11.43 | 4.59 | 4.88 | 5.86 | 4.88 | 13.18 | 9.86 | 21.58 | 9.77 | 20.02 | 9.57 | 7.23 | 0.49 | 0.29 | 0.29 | 0.1 | 0.68 | 4.88 | 3.42 | 4.88 | 21.58 | 4.79 | 19.82 | 9.67 | 23.15 | 9.67 | 22.56 | 9.67 | 22.17
114.336 | 0.1 | 14.45 | 0.1 | 14.26 | 0.1 | 13.77 | 4.88 | 12.01 | 4.59 | 4.88 | 5.47 | 4.88 | 7.03 | 9.86 | 21.58 | 9.77 | 20.41 | 9.57 | 8.01 | 0.49 | 1.56 | 0.29 | 0.1 | 0.59 | 4.88 | 4.2 | 4.88 | 21.49 | 4.79 | 20.41 | 9.77 | 23.63 | 9.67 | 23.05 | 9.67 | 22.75
114.368 | 0 | 14.36 | 0.1 | 14.75 | 0.1 | 14.26 | 4.88 | 12.6 | 4.59 | 4.88 | 5.18 | 4.88 | 8.2 | 9.86 | 21.58 | 9.77 | 20.9 | 9.57 | 8.79 | 0.39 | 2.93 | 0.2 | 1.07 | 0.49 | 4.88 | 5.96 | 4.88 | 21.39 | 4.88 | 21 | 9.77 | 24.12 | 9.67 | 23.63 | 9.67 | 23.24
114.4 | 0.1 | 14.36 | 0.1 | 15.33 | 0.1 | 14.84 | 4.88 | 13.09 | 4.49 | 4.79 | 4.88 | 4.88 | 11.13 | 9.77 | 21.49 | 9.77 | 21.29 | 9.57 | 9.57 | 0.29 | 4.98 | 0.2 | 2.25 | 0.49 | 4.88 | 6.25 | 4.88 | 21.29 | 4.79 | 21.68 | 9.67 | 24.71 | 9.67 | 24.22 | 9.67 | 23.73
114.432 | 0.1 | 14.36 | 0.1 | 15.72 | 0.1 | 15.23 | 4.79 | 13.67 | 4.49 | 4.88 | 4.88 | 4.88 | 11.23 | 9.77 | 21.49 | 9.77 | 21.58 | 9.57 | 10.45 | 0.2 | 6.25 | 0.2 | 3.91 | 0.39 | 4.88 | 5.86 | 4.88 | 21.29 | 4.79 | 22.17 | 9.77 | 25.2 | 9.67 | 24.61 | 9.67 | 24.22
114.464 | 0.1 | 14.36 | 0.1 | 16.21 | 0.1 | 15.82 | 4.88 | 14.16 | 4.49 | 5.08 | 4.88 | 4.88 | 10.35 | 9.86 | 21.49 | 9.77 | 22.07 | 9.67 | 11.04 | 0.2 | 7.13 | 0.1 | 5.57 | 0.29 | 4.88 | 5.27 | 4.88 | 21.29 | 4.88 | 22.75 | 9.77 | 25.68 | 9.67 | 25.2 | 9.77 | 24.81
114.496 | 0.1 | 14.36 | 0.1 | 16.8 | 0.1 | 16.21 | 4.88 | 14.65 | 4.49 | 5.96 | 4.98 | 4.88 | 9.86 | 9.86 | 21.49 | 9.77 | 22.46 | 9.67 | 11.91 | 0.2 | 7.81 | 0.1 | 6.74 | 0.29 | 4.88 | 5.08 | 4.88 | 21.29 | 4.88 | 23.34 | 9.77 | 26.17 | 9.67 | 25.68 | 9.77 | 25.2
114.528 | 0 | 14.36 | 0.1 | 17.19 | 0.1 | 16.8 | 4.88 | 15.14 | 4.59 | 6.93 | 4.79 | 4.88 | 9.96 | 9.86 | 21.39 | 9.77 | 22.85 | 9.67 | 12.7 | 0.2 | 8.69 | 0.1 | 7.52 | 0.29 | 4.88 | 5.08 | 4.88 | 21.19 | 4.88 | 23.73 | 9.77 | 26.66 | 9.77 | 26.17 | 9.77 | 25.78
114.56 | 0.1 | 14.36 | 0.1 | 14.94 | 0.1 | 17.19 | 4.88 | 15.72 | 4.59 | 8.4 | 4.69 | 4.88 | 10.16 | 9.77 | 14.26 | 9.77 | 23.24 | 9.67 | 13.38 | 0.1 | 9.57 | 0.1 | 8.3 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 24.41 | 9.77 | 27.15 | 9.77 | 26.66 | 9.77 | 26.27
114.592 | 0.1 | 5.37 | 0.1 | 15.14 | 0.1 | 14.94 | 4.88 | 16.21 | 4.59 | 9.28 | 4.69 | 4.79 | 10.25 | 9.77 | 7.32 | 9.77 | 23.54 | 9.67 | 14.06 | 0.1 | 10.25 | 0.1 | 9.28 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 24.81 | 9.77 | 27.64 | 9.77 | 27.05 | 9.77 | 26.76
114.624 | 0.1 | 1.56 | 0.1 | 14.65 | 0.1 | 15.23 | 4.88 | 16.7 | 4.59 | 9.96 | 4.69 | 4.88 | 9.96 | 9.86 | 7.81 | 9.86 | 23.93 | 9.67 | 14.84 | 0.1 | 11.13 | 0.1 | 9.96 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.39 | 9.77 | 28.22 | 9.77 | 27.54 | 9.77 | 27.25
114.656 | 0.1 | 2.05 | 0.1 | 14.65 | 0.1 | 14.75 | 4.88 | 17.09 | 4.69 | 10.55 | 4.69 | 4.88 | 9.86 | 9.86 | 10.84 | 9.86 | 24.32 | 9.67 | 15.53 | 0.1 | 11.82 | 0.1 | 10.64 | 0.2 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 25.78 | 9.77 | 28.61 | 9.77 | 28.13 | 9.77 | 27.74
114.688 | 0 | 1.56 | 0.1 | 14.65 | 0.1 | 14.65 | 4.88 | 17.58 | 4.69 | 11.13 | 4.59 | 4.88 | 9.96 | 9.86 | 11.43 | 9.86 | 21.09 | 9.67 | 16.41 | 0.1 | 12.6 | 0 | 11.43 | 0.1 | 4.88 | 4.88 | 4.88 | 8.79 | 4.88 | 25.2 | 9.77 | 29.1 | 9.77 | 28.52 | 9.86 | 28.22
114.72 | 0.1 | 1.46 | 0.1 | 14.55 | 0.1 | 14.65 | 4.88 | 18.07 | 4.69 | 11.72 | 4.49 | 4.88 | 9.86 | 9.77 | 10.45 | 9.77 | 21.68 | 9.67 | 16.99 | 0 | 13.48 | 0 | 12.4 | 0.1 | 4.88 | 4.79 | 4.88 | 2.25 | 4.79 | 20.51 | 9.86 | 28.81 | 9.77 | 29.1 | 9.77 | 28.61
114.752 | 0.1 | 1.27 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.56 | 4.69 | 12.31 | 4.49 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 21.39 | 9.67 | 17.58 | 0.1 | 14.16 | 0.1 | 13.09 | 0.1 | 4.88 | 4.69 | 4.88 | 3.13 | 4.79 | 21.49 | 9.77 | 25.78 | 9.77 | 29.01 | 9.77 | 29.1
114.784 | 0.1 | 1.17 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 18.95 | 4.69 | 12.89 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 18.16 | 0.1 | 14.84 | 0.1 | 13.87 | 0.1 | 4.88 | 4.69 | 4.88 | 3.81 | 4.88 | 21.49 | 9.86 | 25 | 9.77 | 25.49 | 9.86 | 26.56
114.816 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.34 | 4.79 | 13.38 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 18.85 | 0.1 | 15.53 | 0.1 | 14.55 | 0.1 | 5.76 | 4.69 | 4.88 | 4.98 | 4.79 | 21.39 | 9.86 | 24.71 | 9.77 | 24.9 | 9.86 | 23.63
114.848 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.98 | 19.82 | 4.79 | 13.97 | 4.59 | 4.98 | 9.77 | 9.86 | 10.16 | 9.86 | 21.49 | 9.67 | 19.43 | 0.1 | 16.31 | 0 | 15.33 | 0.1 | 6.84 | 4.59 | 4.88 | 6.25 | 4.88 | 21.29 | 9.86 | 24.61 | 9.77 | 24.61 | 9.86 | 24.12
114.88 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 20.22 | 4.79 | 14.45 | 4.49 | 5.96 | 9.67 | 9.77 | 9.96 | 9.86 | 21.49 | 9.67 | 20.12 | 0.1 | 16.9 | 0.1 | 15.92 | 0 | 8.69 | 4.59 | 4.88 | 5.96 | 4.79 | 21.19 | 9.77 | 24.41 | 9.77 | 24.41 | 9.86 | 24.12
114.912 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 20.7 | 4.79 | 14.94 | 4.49 | 6.84 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.67 | 20.8 | 0.1 | 17.48 | 0.1 | 16.6 | 0.1 | 10.45 | 4.49 | 4.88 | 5.57 | 4.88 | 21.19 | 9.86 | 24.41 | 9.77 | 24.41 | 9.86 | 24.22
114.944 | 0.1 | 0.68 | 0.1 | 14.45 | 0.1 | 14.36 | 4.98 | 21.09 | 4.79 | 15.53 | 4.59 | 8.2 | 9.67 | 10.45 | 9.96 | 9.86 | 21.49 | 9.77 | 21.39 | 0.1 | 18.16 | 0 | 17.19 | 0.1 | 11.82 | 4.59 | 4.88 | 5.27 | 4.79 | 21.19 | 9.86 | 24.32 | 9.77 | 24.32 | 9.86 | 24.12
114.976 | 0.1 | 0.59 | 0.1 | 11.91 | 0.1 | 14.36 | 4.98 | 19.14 | 4.79 | 15.92 | 4.59 | 9.18 | 9.67 | 11.13 | 9.96 | 9.86 | 21.49 | 9.77 | 21.88 | 0.1 | 18.85 | 0.1 | 17.97 | 0.1 | 12.79 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.86 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
115.008 | 0 | 0.59 | 0.1 | 3.42 | 0.1 | 11.04 | 4.88 | 18.95 | 4.79 | 16.41 | 4.59 | 9.86 | 9.67 | 12.21 | 9.86 | 9.86 | 21.49 | 9.77 | 22.46 | 0.1 | 19.53 | 0 | 18.56 | 0.1 | 13.67 | 4.59 | 4.88 | 5.18 | 4.88 | 21.19 | 9.86 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
115.04 | 0.1 | 0.49 | 0.1 | 2.05 | 0.1 | 2.54 | 4.88 | 18.56 | 4.79 | 16.9 | 4.59 | 10.55 | 9.67 | 13.48 | 9.77 | 9.86 | 21.39 | 9.77 | 19.14 | 0.1 | 20.12 | 0 | 19.24 | 0.1 | 14.45 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
115.072 | 0.1 | 0.39 | 0.1 | 1.86 | 0.1 | 1.95 | 4.88 | 18.46 | 4.79 | 17.38 | 4.59 | 11.13 | 9.67 | 14.16 | 9.67 | 9.86 | 21.39 | 9.77 | 18.16 | 0.1 | 20.7 | 0.1 | 19.92 | 0.1 | 15.14 | 4.49 | 4.88 | 5.08 | 4.79 | 21.09 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
115.104 | 0.1 | 0.29 | 0.1 | 1.66 | 0.1 | 1.86 | 4.98 | 18.46 | 4.88 | 17.87 | 4.69 | 11.72 | 9.67 | 14.75 | 9.77 | 9.86 | 21 | 9.77 | 17.87 | 0.1 | 21.29 | 0.1 | 20.51 | 0.1 | 15.92 | 4.49 | 4.88 | 5.08 | 4.88 | 21.09 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
115.136 | 0.88 | 0.29 | 0.1 | 1.56 | 0.1 | 1.46 | 4.88 | 18.26 | 4.88 | 18.26 | 4.69 | 12.31 | 9.57 | 15.33 | 9.77 | 9.86 | 8.79 | 9.77 | 17.77 | 0.1 | 21.88 | 0.1 | 21.09 | 0.1 | 16.6 | 4.49 | 4.88 | 4.88 | 4.88 | 8.59 | 9.86 | 24.22 | 9.77 | 24.12 | 9.86 | 24.12
115.168 | 1.66 | 0.29 | 0.1 | 1.37 | 0.1 | 1.46 | 4.88 | 18.26 | 4.88 | 18.75 | 4.79 | 12.79 | 9.67 | 15.82 | 9.67 | 9.86 | 7.13 | 9.77 | 17.77 | 0.1 | 22.56 | 0 | 21.68 | 0.1 | 17.19 | 4.59 | 4.88 | 4.88 | 4.88 | 2.44 | 9.86 | 13.48 | 9.77 | 24.12 | 9.86 | 24.12
115.2 | 3.03 | 0.2 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 19.24 | 4.69 | 13.38 | 9.57 | 16.41 | 9.67 | 9.86 | 10.16 | 9.77 | 17.77 | 0.1 | 17.97 | 0 | 22.36 | 0.1 | 17.97 | 4.59 | 4.98 | 4.79 | 4.79 | 3.22 | 9.77 | 5.86 | 9.77 | 12.7 | 9.86 | 24.12
115.232 | 4 | 0.1 | 0.1 | 1.07 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 19.63 | 4.79 | 13.87 | 9.67 | 16.9 | 9.67 | 9.86 | 11.62 | 9.77 | 17.58 | 0.1 | 16.8 | 0.1 | 22.85 | 0.1 | 18.56 | 4.59 | 5.96 | 4.69 | 4.88 | 3.81 | 9.86 | 7.62 | 9.77 | 6.05 | 9.86 | 11.72
115.264 | 4.79 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 20.12 | 4.79 | 14.45 | 9.67 | 17.38 | 9.67 | 9.86 | 10.74 | 9.77 | 17.68 | 0.1 | 17.48 | 0.1 | 19.43 | 0.1 | 19.24 | 4.59 | 7.42 | 4.79 | 4.88 | 5.08 | 9.86 | 11.72 | 9.86 | 7.81 | 9.86 | 5.66
115.296 | 5.27 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.51 | 4.79 | 14.94 | 9.67 | 17.77 | 9.67 | 9.86 | 10.06 | 9.77 | 17.58 | 0.1 | 17.58 | 0.1 | 18.65 | 0.1 | 19.82 | 4.59 | 9.47 | 4.69 | 4.88 | 6.15 | 9.86 | 11.72 | 9.86 | 11.52 | 9.86 | 8.4
115.328 | 5.86 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 9.67 | 18.26 | 9.67 | 9.86 | 9.86 | 9.77 | 17.58 | 0.1 | 17.68 | 0 | 18.16 | 0.1 | 20.51 | 4.69 | 11.04 | 4.59 | 4.88 | 6.15 | 9.86 | 10.45 | 9.86 | 11.82 | 9.86 | 12.11
115.36 | 6.35 | 0 | 0.1 | 0.68 | 0 | 0.68 | 4.88 | 18.07 | 4.88 | 17.58 | 4.79 | 15.92 | 9.67 | 18.65 | 9.57 | 9.86 | 10.16 | 9.77 | 17.58 | 0 | 17.68 | 0.1 | 17.97 | 0.1 | 21.09 | 4.69 | 12.21 | 4.59 | 4.79 | 5.57 | 9.77 | 9.96 | 9.77 | 10.55 | 9.86 | 11.43
115.392 | 7.03 | 0 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 10.94 | 4.88 | 17.68 | 4.79 | 16.41 | 9.67 | 19.24 | 9.57 | 9.86 | 10.16 | 9.77 | 17.58 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 21.58 | 4.69 | 12.99 | 4.59 | 4.79 | 5.27 | 9.86 | 10.25 | 9.77 | 9.86 | 9.86 | 10.25
115.424 | 7.62 | 0 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 3.32 | 4.88 | 17.87 | 4.88 | 16.9 | 9.67 | 19.63 | 9.67 | 9.86 | 9.96 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.17 | 4.69 | 13.97 | 4.59 | 4.88 | 5.08 | 9.86 | 10.55 | 9.86 | 10.16 | 9.86 | 9.96
115.456 | 8.2 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.03 | 4.88 | 18.07 | 4.88 | 17.38 | 9.67 | 20.02 | 9.67 | 9.86 | 9.86 | 9.86 | 13.67 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.75 | 4.69 | 14.75 | 4.59 | 4.79 | 5.08 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86 | 10.25
115.488 | 8.79 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.81 | 4.88 | 18.07 | 4.88 | 17.77 | 9.67 | 20.51 | 9.67 | 9.86 | 9.96 | 9.86 | 3.32 | 0.1 | 17.68 | 0 | 17.87 | 0.1 | 23.34 | 4.69 | 15.33 | 4.59 | 5.08 | 5.08 | 9.77 | 10.16 | 9.86 | 10.45 | 9.86 | 10.35
115.52 | 9.47 | 0 | 0.1 | 0.29 | 0 | 0.39 | 4.88 | 4.98 | 4.88 | 18.07 | 4.79 | 18.26 | 9.67 | 20.9 | 9.67 | 9.86 | 9.86 | 9.77 | 2.15 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.83 | 4.69 | 16.21 | 4.59 | 6.25 | 5.08 | 9.86 | 10.06 | 9.86 | 10.06 | 9.86 | 10.35
115.552 | 9.96 | 0 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 18.65 | 9.77 | 21.19 | 9.67 | 9.86 | 9.86 | 9.86 | 2.25 | 0 | 17.68 | 0.1 | 17.68 | 0.1 | 24.32 | 4.79 | 16.8 | 4.49 | 7.72 | 5.08 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 9.96
115.584 | 10.55 | 0.1 | 0.88 | 0.29 | 0.1 | 0.29 | 4.88 | 5.66 | 4.88 | 18.07 | 4.88 | 19.14 | 9.77 | 21.68 | 9.67 | 9.86 | 9.77 | 9.86 | 1.86 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 24.9 | 4.79 | 17.48 | 4.59 | 9.77 | 4.98 | 9.86 | 10.16 | 9.86 | 10.16 | 9.86 | 10.06
115.616 | 11.23 | 0.1 | 1.76 | 0.2 | 0.1 | 0.2 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 21.97 | 9.67 | 10.55 | 9.77 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.49 | 4.79 | 18.16 | 4.59 | 11.33 | 4.79 | 9.86 | 10.06 | 9.86 | 9.96 | 9.86 | 10.16
115.648 | 11.72 | 0.1 | 3.03 | 0.2 | 0.1 | 0.2 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.02 | 9.77 | 22.36 | 9.67 | 11.33 | 9.77 | 9.86 | 1.56 | 0.1 | 5.18 | 0 | 17.68 | 0.1 | 25.88 | 4.79 | 18.75 | 4.59 | 12.31 | 4.79 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86 | 10.16
115.68 | 12.31 | 0 | 4.1 | 0.1 | 0.1 | 0.1 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 20.41 | 9.77 | 22.85 | 9.77 | 12.4 | 9.67 | 9.86 | 1.37 | 0.1 | 1.86 | 0 | 8.5 | 0.1 | 26.37 | 4.79 | 19.43 | 4.59 | 13.28 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
115.712 | 12.89 | 0 | 4.69 | 0 | 0.2 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.9 | 9.77 | 23.24 | 9.67 | 13.57 | 9.67 | 9.86 | 1.37 | 0 | 2.34 | 0 | 2.25 | 0.1 | 23.15 | 4.79 | 20.02 | 4.59 | 13.97 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86
115.744 | 13.38 | 0.1 | 5.27 | 0.1 | 1.17 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.38 | 9.77 | 23.63 | 9.77 | 14.36 | 9.67 | 9.86 | 1.27 | 0.1 | 1.95 | 0.1 | 2.25 | 0.1 | 22.36 | 4.79 | 20.7 | 4.69 | 14.75 | 4.69 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
115.776 | 14.06 | 0.1 | 5.86 | 0.1 | 2.05 | 0 | 5.08 | 4.79 | 4.88 | 17.77 | 4.88 | 17.58 | 9.77 | 23.93 | 9.77 | 14.84 | 9.67 | 9.86 | 1.27 | 0.1 | 1.76 | 0.1 | 1.95 | 0.1 | 21.88 | 4.79 | 21.19 | 4.69 | 15.53 | 4.59 | 9.86 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
115.808 | 14.45 | 0.1 | 6.45 | 0.1 | 3.42 | 0 | 5.86 | 4.79 | 4.88 | 5.27 | 4.88 | 17.77 | 9.77 | 24.32 | 9.77 | 15.43 | 9.67 | 9.86 | 1.07 | 0.1 | 1.56 | 0 | 1.66 | 0.1 | 21.68 | 4.79 | 21.78 | 4.69 | 16.11 | 4.59 | 9.96 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77
115.84 | 15.04 | 0 | 7.03 | 0 | 4.3 | 0 | 6.93 | 4.69 | 4.88 | 2.83 | 4.88 | 18.07 | 9.77 | 22.27 | 9.77 | 15.92 | 9.57 | 9.86 | 0.88 | 0 | 1.46 | 0 | 1.86 | 0 | 21.58 | 4.79 | 22.36 | 4.69 | 16.9 | 4.59 | 11.04 | 9.67 | 9.77 | 9.67 | 9.96 | 9.77
115.872 | 15.53 | 0 | 7.62 | 0.1 | 4.88 | 0 | 8.3 | 4.69 | 4.88 | 3.52 | 4.88 | 17.97 | 9.77 | 22.17 | 9.77 | 16.5 | 9.57 | 9.86 | 0.98 | 0 | 1.27 | 0.1 | 1.46 | 0.1 | 21.49 | 4.79 | 22.95 | 4.69 | 17.58 | 4.49 | 12.4 | 9.67 | 9.86 | 9.67 | 10.74 | 9.77
115.904 | 16.02 | 0.1 | 8.3 | 0.1 | 5.37 | 0.1 | 9.28 | 4.69 | 4.88 | 4.59 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 16.99 | 9.57 | 9.86 | 0.78 | 0.1 | 1.37 | 0 | 1.46 | 0.1 | 21.39 | 4.79 | 23.54 | 4.79 | 18.16 | 4.59 | 14.55 | 9.67 | 10.74 | 9.67 | 12.11 | 9.77
115.936 | 16.5 | 0.1 | 8.79 | 0.1 | 5.96 | 0.1 | 9.96 | 4.59 | 4.88 | 5.66 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 17.48 | 9.57 | 9.86 | 0.78 | 0.1 | 1.07 | 0.1 | 1.27 | 0.1 | 21.39 | 4.79 | 24.02 | 4.79 | 18.75 | 4.59 | 16.02 | 9.67 | 12.11 | 9.67 | 14.06 | 9.67
115.968 | 16.9 | 0.1 | 9.47 | 0.1 | 6.54 | 0.1 | 10.55 | 4.59 | 4.88 | 5.86 | 4.88 | 18.07 | 9.86 | 21.68 | 9.77 | 17.87 | 9.57 | 9.86 | 0.68 | 0.1 | 0.98 | 0 | 1.37 | 0.1 | 21.39 | 4.79 | 24.61 | 4.79 | 19.34 | 4.49 | 16.9 | 9.67 | 13.97 | 9.67 | 15.63 | 9.67
116 | 16.6 | 0 | 10.16 | 0.1 | 7.23 | 0 | 11.23 | 4.49 | 4.88 | 5.37 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.36 | 9.57 | 9.86 | 0.49 | 0 | 0.98 | 0 | 1.17 | 0 | 21.29 | 4.79 | 25.1 | 4.79 | 20.22 | 4.49 | 17.77 | 9.67 | 15.63 | 9.67 | 16.8 | 9.67
116.032 | 14.65 | 0 | 10.64 | 0 | 7.72 | 0 | 11.82 | 4.49 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 18.75 | 9.57 | 9.86 | 0.49 | 0 | 0.88 | 0.1 | 0.98 | 0.1 | 21.29 | 4.79 | 25.59 | 4.69 | 20.8 | 4.49 | 18.56 | 9.67 | 16.6 | 9.57 | 17.58 | 9.67
116.064 | 14.65 | 0.1 | 11.23 | 0.1 | 8.4 | 0 | 12.4 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.24 | 9.67 | 9.86 | 0.49 | 0.1 | 0.78 | 0.1 | 0.98 | 0.1 | 21.29 | 4.79 | 26.08 | 4.79 | 21.29 | 4.59 | 19.24 | 9.67 | 17.58 | 9.67 | 18.36 | 9.67
116.096 | 14.55 | 0.1 | 11.82 | 0.1 | 8.98 | 0 | 12.99 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 19.73 | 9.67 | 9.86 | 0.49 | 0.1 | 0.68 | 0.1 | 0.88 | 0.1 | 21.29 | 4.79 | 21.88 | 4.79 | 21.88 | 4.59 | 19.82 | 9.67 | 18.26 | 9.67 | 18.95 | 9.67
116.128 | 14.55 | 0.1 | 12.4 | 0.1 | 9.57 | 0 | 13.48 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 20.12 | 9.67 | 10.06 | 0.39 | 0.1 | 0.59 | 0 | 0.78 | 0.1 | 11.72 | 4.79 | 20.41 | 4.79 | 22.36 | 4.59 | 20.51 | 9.67 | 19.04 | 9.67 | 19.73 | 9.67
116.16 | 14.55 | 0 | 12.89 | 0 | 10.16 | 0 | 14.06 | 4.49 | 4.88 | 4.79 | 4.88 | 17.77 | 9.77 | 21.39 | 9.86 | 20.61 | 9.67 | 10.94 | 0.29 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 2.83 | 4.79 | 20.9 | 4.79 | 23.05 | 4.59 | 21 | 9.67 | 19.63 | 9.57 | 20.31 | 9.57
116.192 | 14.45 | 0 | 13.38 | 0 | 10.74 | 0 | 14.55 | 4.49 | 4.88 | 4.79 | 4.88 | 5.08 | 9.77 | 21.39 | 9.86 | 20.9 | 9.67 | 11.82 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.22 | 4.79 | 21.19 | 4.79 | 23.54 | 4.59 | 21.58 | 9.57 | 20.22 | 9.57 | 20.9 | 9.67
116.224 | 14.45 | 0.1 | 13.97 | 0.1 | 11.33 | 0.1 | 15.04 | 4.59 | 4.88 | 4.79 | 4.88 | 2.83 | 9.86 | 21.39 | 9.86 | 21.39 | 9.67 | 13.18 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.52 | 4.79 | 21.19 | 4.88 | 24.02 | 4.69 | 22.27 | 9.67 | 20.9 | 9.67 | 21.58 | 9.67
116.256 | 14.45 | 0.1 | 14.55 | 0.1 | 12.01 | 0 | 15.53 | 4.59 | 4.88 | 4.69 | 4.88 | 3.52 | 9.86 | 17.29 | 9.86 | 21.78 | 9.67 | 14.06 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.49 | 4.79 | 21.19 | 4.88 | 24.61 | 4.69 | 22.75 | 9.67 | 21.49 | 9.67 | 22.07 | 9.67
116.288 | 14.36 | 0.1 | 15.04 | 0.1 | 12.5 | 0 | 16.02 | 4.59 | 4.88 | 4.59 | 4.88 | 4.49 | 9.86 | 7.81 | 9.86 | 22.17 | 9.67 | 14.75 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 5.96 | 4.79 | 21.29 | 4.79 | 25.1 | 4.69 | 23.34 | 9.67 | 22.07 | 9.67 | 22.56 | 9.67
116.32 | 14.45 | 0 | 15.53 | 0 | 13.09 | 0 | 16.5 | 4.59 | 4.88 | 4.59 | 4.88 | 5.86 | 9.77 | 7.72 | 9.86 | 22.56 | 9.67 | 15.33 | 0.1 | 0 | 0.29 | 0.1 | 0.39 | 0 | 6.05 | 4.79 | 21.29 | 4.79 | 25.68 | 4.69 | 23.93 | 9.67 | 22.66 | 9.67 | 23.15 | 9.67
116.352 | 14.45 | 0 | 16.02 | 0.1 | 13.57 | 0 | 16.99 | 4.69 | 4.88 | 4.59 | 4.88 | 5.86 | 9.86 | 10.55 | 9.86 | 22.95 | 9.67 | 15.72 | 0.2 | 0.1 | 0.29 | 1.07 | 0.29 | 0.1 | 5.76 | 4.79 | 21.29 | 4.79 | 26.08 | 4.69 | 24.41 | 9.67 | 23.05 | 9.67 | 23.73 | 9.67
116.384 | 14.36 | 0.1 | 16.5 | 0.1 | 14.06 | 0 | 17.58 | 4.69 | 4.88 | 4.59 | 4.88 | 5.47 | 9.86 | 11.33 | 9.86 | 23.24 | 9.77 | 16.31 | 0.2 | 0.2 | 0.2 | 2.34 | 0.29 | 0.1 | 5.47 | 4.79 | 21.29 | 4.79 | 21.49 | 4.79 | 25 | 9.67 | 23.63 | 9.67 | 24.32 | 9.67
116.416 | 14.36 | 0.1 | 16.99 | 0.1 | 14.55 | 0.1 | 17.97 | 4.69 | 5.08 | 4.49 | 4.88 | 5.18 | 9.86 | 10.45 | 9.86 | 23.63 | 9.77 | 16.8 | 0.1 | 1.46 | 0.2 | 4.2 | 0.29 | 0.1 | 5.08 | 4.88 | 21.19 | 4.88 | 20.31 | 4.79 | 25.29 | 9.67 | 24.22 | 9.67 | 24.81 | 9.67
116.448 | 7.03 | 0 | 16.6 | 0.1 | 15.14 | 0 | 18.36 | 4.79 | 5.96 | 4.49 | 4.88 | 4.98 | 9.86 | 10.16 | 9.86 | 24.02 | 9.77 | 17.29 | 0.1 | 2.64 | 0.2 | 5.76 | 0.2 | 0.1 | 5.08 | 4.79 | 21.29 | 4.88 | 21 | 4.79 | 25.88 | 9.77 | 24.81 | 9.67 | 25.29 | 9.67
116.48 | 1.95 | 0 | 14.16 | 0.1 | 15.63 | 0 | 18.85 | 4.69 | 7.03 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 24.41 | 9.77 | 17.68 | 0.1 | 4.59 | 0.1 | 6.84 | 0.2 | 0.68 | 5.08 | 4.79 | 21.19 | 4.88 | 21.19 | 4.79 | 26.47 | 9.67 | 25.29 | 9.67 | 25.68 | 9.67
116.512 | 2.05 | 0 | 14.65 | 0 | 16.11 | 0 | 19.24 | 4.79 | 8.3 | 4.49 | 4.88 | 4.88 | 9.77 | 10.06 | 9.86 | 22.27 | 9.77 | 18.16 | 0.1 | 5.96 | 0.1 | 7.62 | 0.2 | 1.46 | 5.08 | 4.79 | 21.19 | 4.79 | 21.29 | 4.69 | 26.95 | 9.67 | 25.68 | 9.67 | 26.17 | 9.67
116.544 | 1.66 | 0.1 | 14.75 | 0.1 | 16.6 | 0 | 19.82 | 4.79 | 9.28 | 4.49 | 4.88 | 4.79 | 9.86 | 10.16 | 9.86 | 22.36 | 9.77 | 18.65 | 0.1 | 7.03 | 0.1 | 8.4 | 0.2 | 3.03 | 4.98 | 4.79 | 6.93 | 4.88 | 21.19 | 4.79 | 27.44 | 9.77 | 26.17 | 9.67 | 26.76 | 9.77
116.576 | 1.46 | 0 | 14.55 | 0.1 | 17.09 | 0 | 20.12 | 4.79 | 9.96 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.88 | 9.77 | 19.04 | 0.1 | 7.62 | 0.1 | 9.08 | 0.1 | 4.98 | 4.88 | 4.79 | 2.34 | 4.88 | 21.29 | 4.79 | 27.83 | 9.77 | 26.66 | 9.77 | 27.25 | 9.77
116.608 | 1.37 | 0.1 | 14.55 | 0.1 | 14.06 | 0 | 20.51 | 4.79 | 10.55 | 4.49 | 4.88 | 4.79 | 9.86 | 9.96 | 9.86 | 21.68 | 9.77 | 19.43 | 0.1 | 8.5 | 0.1 | 9.96 | 0.1 | 6.25 | 4.88 | 4.79 | 3.42 | 4.88 | 21.29 | 4.79 | 28.32 | 9.77 | 27.25 | 9.77 | 27.74 | 9.77
116.64 | 1.27 | 0 | 14.45 | 0.1 | 13.97 | 0 | 20.9 | 4.79 | 11.23 | 4.49 | 4.79 | 4.69 | 9.86 | 9.86 | 9.86 | 21.68 | 9.86 | 19.92 | 0.1 | 9.28 | 0.1 | 10.94 | 0.1 | 7.23 | 4.88 | 4.79 | 4 | 4.79 | 21.19 | 4.79 | 28.81 | 9.77 | 27.64 | 9.67 | 28.22 | 9.77
116.672 | 1.17 | 0 | 14.45 | 0 | 14.16 | 0 | 17.48 | 4.88 | 11.72 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.22 | 0.1 | 10.16 | 0.1 | 11.62 | 0.1 | 8.01 | 4.69 | 4.79 | 5.18 | 4.79 | 21.29 | 4.79 | 29.3 | 9.77 | 28.22 | 9.67 | 28.71 | 9.77
116.704 | 0.98 | 0.1 | 14.55 | 0.1 | 14.45 | 0.1 | 17.68 | 4.88 | 12.4 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.58 | 9.86 | 20.7 | 0.1 | 10.84 | 0.1 | 12.4 | 0.1 | 8.79 | 4.69 | 4.79 | 6.45 | 4.88 | 21.19 | 4.79 | 24.81 | 9.77 | 28.61 | 9.77 | 29.1 | 9.77
116.736 | 0.88 | 0 | 14.45 | 0.1 | 14.45 | 0 | 17.87 | 4.88 | 12.89 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.86 | 21.19 | 0.1 | 11.72 | 0.1 | 13.18 | 0.1 | 9.57 | 4.69 | 4.79 | 6.05 | 4.88 | 21.19 | 4.79 | 23.34 | 9.77 | 29.1 | 9.77 | 28.42 | 9.77
116.768 | 0.78 | 0.1 | 14.45 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 13.48 | 4.69 | 4.88 | 4.59 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 21.58 | 0.1 | 12.5 | 0.1 | 14.06 | 0.1 | 10.35 | 4.59 | 4.79 | 5.57 | 4.88 | 21.19 | 4.79 | 24.02 | 9.77 | 28.42 | 9.77 | 24.61 | 9.77
116.8 | 0.68 | 0 | 14.45 | 0 | 14.45 | 0 | 18.16 | 4.88 | 13.97 | 4.59 | 4.98 | 4.49 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 21.97 | 0.1 | 13.28 | 0.1 | 14.55 | 0.1 | 11.13 | 4.59 | 4.79 | 5.27 | 4.79 | 20.9 | 4.79 | 24.22 | 9.77 | 23.93 | 9.77 | 24.32 | 9.77
116.832 | 0.59 | 0 | 14.45 | 0 | 14.45 | 0 | 18.16 | 4.88 | 14.55 | 4.69 | 5.66 | 4.49 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 22.27 | 0.1 | 13.87 | 0 | 15.33 | 0.1 | 11.91 | 4.59 | 4.79 | 5.27 | 4.79 | 5.76 | 4.79 | 24.22 | 9.77 | 24.51 | 9.77 | 24.41 | 9.77
116.864 | 0.59 | 0.1 | 14.45 | 0.1 | 14.45 | 0.1 | 18.26 | 4.88 | 15.04 | 4.69 | 6.64 | 4.49 | 9.86 | 9.67 | 9.96 | 21.39 | 9.86 | 22.75 | 0.1 | 14.65 | 0.1 | 16.02 | 0.1 | 12.7 | 4.59 | 4.79 | 5.18 | 4.88 | 2.83 | 4.79 | 24.22 | 9.77 | 24.51 | 9.77 | 24.41 | 9.86
116.896 | 0.49 | 0.1 | 7.13 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.53 | 4.69 | 8.11 | 4.59 | 9.86 | 9.67 | 9.86 | 21.39 | 9.86 | 23.15 | 0.1 | 15.33 | 0.1 | 16.7 | 0.1 | 13.38 | 4.59 | 5.66 | 5.08 | 4.88 | 3.22 | 4.79 | 24.32 | 9.77 | 24.41 | 9.77 | 24.32 | 9.86
116.928 | 0.39 | 0 | 2.05 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 16.02 | 4.69 | 8.98 | 4.49 | 9.86 | 9.67 | 9.96 | 15.82 | 9.86 | 23.44 | 0.1 | 16.02 | 0.1 | 17.29 | 0.1 | 14.16 | 4.59 | 6.64 | 4.98 | 4.88 | 4.2 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.32 | 9.86
116.96 | 0.29 | 0 | 2.05 | 0 | 14.45 | 0 | 18.16 | 4.88 | 16.5 | 4.69 | 9.86 | 4.49 | 9.86 | 9.57 | 9.86 | 7.23 | 9.86 | 23.83 | 0.1 | 16.9 | 0.1 | 17.97 | 0.1 | 14.94 | 4.49 | 8.4 | 4.98 | 4.79 | 5.57 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.77
116.992 | 0.39 | 0 | 1.86 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.99 | 4.79 | 10.45 | 4.59 | 10.55 | 9.57 | 9.86 | 7.62 | 9.86 | 24.22 | 0.1 | 17.48 | 0 | 18.75 | 0.1 | 15.53 | 4.49 | 10.25 | 4.98 | 4.79 | 6.25 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
117.024 | 0.29 | 0.1 | 1.46 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 17.38 | 4.79 | 11.13 | 4.59 | 11.23 | 9.67 | 9.86 | 10.74 | 9.96 | 20.8 | 0.1 | 18.16 | 0.1 | 19.24 | 0.1 | 16.31 | 4.59 | 11.62 | 4.88 | 4.88 | 5.96 | 4.79 | 24.32 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86
117.056 | 0.2 | 0.1 | 1.37 | 0.1 | 5.76 | 0 | 18.16 | 4.88 | 17.87 | 4.79 | 11.72 | 4.59 | 12.4 | 9.57 | 9.86 | 11.43 | 9.86 | 20.9 | 0.1 | 18.85 | 0.1 | 20.02 | 0.1 | 17.09 | 4.59 | 12.6 | 4.79 | 4.88 | 5.47 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86
117.088 | 0.2 | 0.1 | 1.27 | 0.1 | 1.66 | 0 | 18.07 | 4.88 | 18.46 | 4.79 | 12.21 | 4.59 | 13.48 | 9.57 | 9.96 | 10.45 | 9.86 | 21.09 | 0.1 | 19.34 | 0.1 | 20.51 | 0.1 | 17.77 | 4.69 | 13.48 | 4.69 | 4.88 | 5.27 | 4.79 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
117.12 | 0.1 | 0 | 1.17 | 0 | 1.86 | 0 | 5.27 | 4.88 | 18.75 | 4.79 | 12.79 | 4.59 | 14.36 | 9.57 | 9.86 | 9.86 | 9.86 | 21.39 | 0.1 | 20.02 | 0.1 | 21.19 | 0.1 | 18.26 | 4.59 | 14.26 | 4.69 | 4.79 | 5.18 | 4.79 | 24.22 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
117.152 | 0.1 | 0 | 0.98 | 0 | 1.76 | 0 | 2.73 | 4.88 | 19.24 | 4.79 | 13.38 | 4.59 | 14.94 | 9.57 | 9.86 | 10.06 | 9.86 | 21.29 | 0.1 | 20.61 | 0.1 | 21.68 | 0.1 | 18.95 | 4.59 | 14.94 | 4.69 | 4.79 | 5.18 | 4.79 | 10.35 | 9.77 | 24.12 | 9.77 | 24.12 | 9.86
117.184 | 0.1 | 0.1 | 1.07 | 0.1 | 1.46 | 0 | 3.52 | 4.88 | 19.73 | 4.79 | 13.87 | 4.69 | 15.43 | 9.67 | 9.96 | 10.25 | 9.96 | 21.39 | 0.1 | 21 | 0.1 | 22.27 | 0.1 | 19.53 | 4.59 | 15.72 | 4.69 | 4.88 | 5.18 | 4.79 | 5.76 | 9.77 | 24.22 | 9.86 | 12.79 | 9.86
117.216 | 0.1 | 0.1 | 0.78 | 0.1 | 1.46 | 0 | 4.49 | 4.88 | 20.12 | 4.79 | 14.45 | 4.69 | 15.92 | 9.67 | 9.96 | 10.16 | 9.96 | 21.49 | 0.1 | 21.88 | 0.1 | 21.58 | 0.1 | 20.22 | 4.69 | 16.31 | 4.59 | 4.79 | 4.98 | 4.79 | 9.18 | 9.77 | 12.6 | 9.86 | 5.86 | 9.86
117.248 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0 | 5.66 | 4.88 | 20.51 | 4.79 | 14.94 | 4.69 | 16.41 | 9.67 | 10.55 | 9.96 | 9.96 | 21.39 | 0.1 | 22.36 | 0.1 | 17.29 | 0.1 | 20.9 | 4.69 | 16.99 | 4.59 | 4.88 | 4.98 | 4.79 | 12.5 | 9.77 | 5.76 | 9.86 | 8.01 | 9.86
117.28 | 0.1 | 1.07 | 0.59 | 0 | 1.07 | 0 | 5.86 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 16.99 | 9.67 | 11.33 | 9.96 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.97 | 0.1 | 21.39 | 4.69 | 17.87 | 4.59 | 4.88 | 4.88 | 4.79 | 11.13 | 9.77 | 8.59 | 9.77 | 11.91 | 9.86
117.312 | 0.1 | 1.95 | 0.59 | 0.1 | 0.88 | 0 | 5.47 | 4.88 | 17.58 | 4.79 | 15.92 | 4.79 | 17.48 | 9.67 | 12.31 | 9.86 | 9.86 | 21.39 | 0.1 | 16.8 | 0 | 17.87 | 0.1 | 21.97 | 4.69 | 18.46 | 4.59 | 4.79 | 4.79 | 4.79 | 10.06 | 9.77 | 11.72 | 9.77 | 11.62 | 9.86
117.344 | 0.1 | 3.42 | 0.59 | 0.1 | 0.78 | 0 | 5.18 | 4.88 | 17.77 | 4.88 | 16.41 | 4.79 | 17.87 | 9.67 | 13.48 | 9.86 | 9.96 | 21.39 | 0.1 | 17.38 | 0.1 | 17.87 | 0.1 | 22.46 | 4.79 | 19.14 | 4.59 | 4.88 | 4.79 | 4.79 | 10.25 | 9.77 | 11.62 | 9.86 | 10.45 | 9.86
117.376 | 0.1 | 4.2 | 0.49 | 0.1 | 0.88 | 0 | 4.98 | 4.88 | 17.97 | 4.88 | 16.9 | 4.79 | 18.46 | 9.67 | 14.36 | 9.86 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 17.97 | 4.79 | 19.63 | 4.59 | 5.86 | 4.79 | 4.79 | 10.25 | 9.77 | 10.35 | 9.86 | 9.96 | 9.86
117.408 | 0.1 | 4.88 | 0.39 | 0.1 | 0.68 | 0.1 | 4.88 | 4.88 | 18.16 | 4.88 | 17.38 | 4.79 | 18.85 | 9.67 | 14.94 | 9.77 | 9.96 | 21.39 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 16.99 | 4.79 | 20.22 | 4.59 | 7.03 | 4.79 | 4.79 | 10.45 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
117.44 | 0.1 | 5.37 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.16 | 4.79 | 17.87 | 4.79 | 19.24 | 9.67 | 15.53 | 9.77 | 9.96 | 21.09 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.79 | 20.9 | 4.59 | 8.98 | 4.59 | 4.79 | 10.25 | 9.77 | 10.25 | 9.77 | 10.45 | 9.86
117.472 | 0.1 | 5.96 | 0.29 | 0.1 | 0.49 | 0 | 4.88 | 4.88 | 18.07 | 4.79 | 18.16 | 4.79 | 19.73 | 9.67 | 16.02 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.68 | 4.79 | 21.49 | 4.59 | 10.55 | 4.59 | 4.79 | 10.25 | 9.77 | 10.25 | 9.77 | 10.35 | 9.86
117.504 | 0.1 | 6.54 | 0.29 | 0.1 | 0.59 | 0 | 4.79 | 4.88 | 18.16 | 4.88 | 18.75 | 4.88 | 20.12 | 9.67 | 16.5 | 9.67 | 9.96 | 6.84 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 22.17 | 4.59 | 11.91 | 4.59 | 4.79 | 10.06 | 9.77 | 10.35 | 9.86 | 10.25 | 9.86
117.536 | 0.1 | 7.13 | 0.2 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.16 | 4.88 | 19.24 | 4.79 | 20.61 | 9.67 | 16.99 | 9.67 | 9.96 | 9.47 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 22.66 | 4.59 | 12.99 | 4.59 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
117.568 | 0 | 7.72 | 0.1 | 0.1 | 0.39 | 0 | 4.79 | 4.88 | 18.07 | 4.88 | 19.63 | 4.88 | 21 | 9.67 | 17.48 | 9.67 | 9.86 | 11.52 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 23.15 | 4.69 | 13.77 | 4.59 | 4.79 | 10.16 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
117.6 | 0.1 | 8.3 | 0.1 | 0 | 0.29 | 0 | 4.69 | 4.88 | 18.16 | 4.79 | 20.02 | 4.79 | 21.49 | 9.67 | 17.97 | 9.67 | 9.86 | 10.74 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 23.63 | 4.69 | 14.55 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.77 | 10.06 | 9.86
117.632 | 0.1 | 8.89 | 0.1 | 0 | 0.2 | 0 | 4.59 | 4.88 | 18.07 | 4.79 | 20.31 | 4.79 | 21.78 | 9.77 | 18.36 | 9.67 | 9.86 | 10.16 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.79 | 24.12 | 4.69 | 15.33 | 4.49 | 4.79 | 9.86 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86
117.664 | 0.1 | 9.57 | 0.1 | 0.49 | 0.2 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.9 | 4.88 | 22.17 | 9.77 | 18.85 | 9.67 | 9.86 | 9.77 | 0.1 | 17.58 | 0.1 | 7.81 | 0.1 | 17.77 | 4.88 | 24.71 | 4.69 | 15.92 | 4.59 | 4.79 | 9.96 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
117.696 | 0.1 | 10.16 | 0.1 | 1.37 | 0.2 | 0.59 | 4.59 | 4.88 | 18.07 | 4.88 | 17.77 | 4.88 | 22.56 | 9.77 | 19.24 | 9.57 | 9.96 | 9.96 | 0.1 | 17.58 | 0.1 | 1.86 | 0.1 | 17.68 | 4.88 | 25.2 | 4.69 | 16.6 | 4.49 | 4.79 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
117.728 | 0 | 10.64 | 0.1 | 2.34 | 0.1 | 1.27 | 4.59 | 4.88 | 17.87 | 4.88 | 17.68 | 4.88 | 22.95 | 9.77 | 19.82 | 9.57 | 9.86 | 10.16 | 0.1 | 6.25 | 0.1 | 2.15 | 0.1 | 17.68 | 4.88 | 25.68 | 4.79 | 17.38 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86
117.76 | 0.1 | 11.33 | 0.1 | 3.71 | 0.1 | 2.44 | 4.59 | 4.88 | 4.88 | 4.79 | 17.77 | 4.88 | 23.34 | 9.77 | 20.12 | 9.57 | 9.86 | 9.96 | 0.1 | 1.66 | 0.1 | 2.05 | 0.1 | 17.68 | 4.79 | 26.17 | 4.79 | 17.97 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
117.792 | 0.1 | 11.91 | 0.1 | 4.49 | 0.1 | 3.71 | 4.49 | 4.88 | 2.93 | 4.79 | 18.07 | 4.88 | 23.63 | 9.77 | 20.51 | 9.67 | 10.16 | 9.86 | 0.1 | 2.34 | 0 | 1.76 | 0.1 | 17.48 | 4.79 | 22.85 | 4.79 | 18.65 | 4.59 | 4.79 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
117.824 | 0.1 | 12.5 | 0.1 | 5.08 | 0.1 | 4.49 | 4.59 | 4.88 | 3.52 | 4.88 | 18.07 | 4.88 | 24.02 | 9.77 | 20.9 | 9.67 | 11.04 | 9.86 | 0.1 | 1.95 | 0.1 | 1.86 | 0.1 | 4.49 | 4.88 | 22.07 | 4.79 | 19.24 | 4.59 | 5.86 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
117.856 | 0.1 | 12.99 | 0.1 | 5.57 | 0.1 | 5.08 | 4.59 | 4.88 | 4.49 | 4.88 | 18.07 | 4.88 | 24.41 | 9.77 | 21.39 | 9.67 | 11.91 | 9.86 | 0.1 | 1.66 | 0.1 | 1.46 | 0.1 | 2.15 | 4.88 | 21.68 | 4.79 | 19.92 | 4.59 | 6.84 | 9.67 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86
117.888 | 0 | 13.57 | 0.1 | 6.15 | 0.1 | 5.57 | 4.59 | 4.88 | 5.66 | 4.88 | 18.16 | 4.88 | 22.46 | 9.77 | 21.78 | 9.67 | 13.28 | 9.77 | 0.1 | 1.76 | 0.1 | 1.46 | 0.1 | 2.15 | 4.88 | 21.49 | 4.79 | 20.51 | 4.69 | 8.79 | 9.67 | 11.82 | 9.77 | 11.04 | 9.67 | 9.86
117.92 | 0.1 | 14.06 | 0.1 | 6.74 | 0.1 | 6.15 | 4.59 | 4.88 | 5.96 | 4.79 | 18.07 | 4.88 | 22.17 | 9.77 | 22.17 | 9.67 | 14.16 | 9.67 | 0.1 | 1.56 | 0 | 1.27 | 0.1 | 1.86 | 4.88 | 21.58 | 4.79 | 21.09 | 4.59 | 10.74 | 9.67 | 13.57 | 9.67 | 12.01 | 9.67 | 10.94
117.952 | 0.1 | 14.45 | 0.1 | 7.32 | 0.1 | 6.74 | 4.59 | 4.88 | 5.47 | 4.79 | 18.07 | 4.79 | 21.88 | 9.77 | 22.56 | 9.67 | 14.75 | 9.67 | 0.1 | 1.46 | 0.1 | 1.17 | 0.1 | 1.76 | 4.88 | 21.39 | 4.79 | 21.68 | 4.69 | 11.82 | 9.67 | 15.23 | 9.67 | 13.77 | 9.67 | 12.11
117.984 | 0.1 | 15.04 | 0.1 | 8.01 | 0.1 | 7.32 | 4.69 | 4.88 | 5.18 | 4.88 | 18.16 | 4.88 | 21.78 | 9.77 | 22.95 | 9.67 | 15.33 | 9.77 | 0.1 | 1.27 | 0.1 | 1.27 | 0.1 | 1.56 | 4.88 | 21.29 | 4.79 | 22.27 | 4.69 | 12.79 | 9.67 | 16.5 | 9.67 | 15.43 | 9.67 | 14.06
118.016 | 0.1 | 15.53 | 0.1 | 8.5 | 0.1 | 8.01 | 4.69 | 4.88 | 4.98 | 4.79 | 18.07 | 4.88 | 21.78 | 9.77 | 23.34 | 9.67 | 15.82 | 9.67 | 0.1 | 1.17 | 0.1 | 1.07 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 22.75 | 4.69 | 13.57 | 9.67 | 17.29 | 9.67 | 16.6 | 9.67 | 15.63
118.048 | 0 | 16.02 | 0.1 | 9.08 | 0.1 | 8.5 | 4.69 | 4.88 | 4.88 | 4.79 | 18.07 | 4.88 | 21.58 | 9.77 | 23.73 | 9.67 | 16.31 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.27 | 4.88 | 21.19 | 4.79 | 23.34 | 4.79 | 14.36 | 9.67 | 18.16 | 9.67 | 17.48 | 9.67 | 16.8
118.08 | 0.1 | 16.6 | 0.1 | 9.67 | 0 | 9.28 | 4.69 | 5.57 | 4.88 | 4.79 | 18.07 | 4.88 | 21.49 | 9.77 | 24.12 | 9.67 | 16.8 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 23.93 | 4.69 | 15.14 | 9.57 | 18.75 | 9.67 | 18.16 | 9.67 | 17.68
118.112 | 0.1 | 16.99 | 0.1 | 10.35 | 0.1 | 9.77 | 4.69 | 6.25 | 4.88 | 4.79 | 18.07 | 4.79 | 21.49 | 9.77 | 24.41 | 9.67 | 17.29 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0.1 | 1.07 | 4.88 | 21.19 | 4.79 | 24.41 | 4.69 | 15.72 | 9.57 | 19.43 | 9.57 | 18.95 | 9.67 | 18.46
118.144 | 0.1 | 13.97 | 0.1 | 10.94 | 0.1 | 10.35 | 4.79 | 7.32 | 4.79 | 4.88 | 6.93 | 4.88 | 21.49 | 9.77 | 22.27 | 9.67 | 17.77 | 9.57 | 0.59 | 0.78 | 0.1 | 0.59 | 0.1 | 1.07 | 4.88 | 21.19 | 4.88 | 24.9 | 4.79 | 16.5 | 9.67 | 20.12 | 9.57 | 19.53 | 9.67 | 19.14
118.176 | 0.1 | 13.97 | 0.1 | 11.52 | 0.1 | 10.94 | 4.79 | 8.59 | 4.79 | 4.88 | 2.64 | 4.88 | 21.49 | 9.86 | 22.36 | 9.77 | 18.26 | 9.57 | 1.66 | 0.78 | 0.1 | 0.59 | 0.1 | 0.88 | 4.88 | 21.19 | 4.88 | 25.39 | 4.79 | 17.19 | 9.67 | 20.61 | 9.67 | 20.22 | 9.67 | 19.53
118.208 | 0 | 14.16 | 0.1 | 12.11 | 0.1 | 11.52 | 4.79 | 9.47 | 4.69 | 4.79 | 3.42 | 4.88 | 21.49 | 9.86 | 21.88 | 9.77 | 18.65 | 9.57 | 3.13 | 0.59 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 14.45 | 4.88 | 25.78 | 4.79 | 17.87 | 9.67 | 21.39 | 9.67 | 20.8 | 9.67 | 20.31
118.24 | 0.1 | 14.36 | 0.1 | 12.7 | 0.1 | 12.11 | 4.79 | 10.06 | 4.59 | 4.79 | 4.49 | 4.88 | 21.39 | 9.77 | 21.78 | 9.77 | 19.24 | 9.57 | 5.18 | 0.49 | 0 | 0.49 | 0.1 | 0.78 | 4.88 | 3.03 | 4.79 | 25.49 | 4.79 | 18.56 | 9.67 | 21.97 | 9.67 | 21.49 | 9.67 | 20.8
118.272 | 0.1 | 14.36 | 0.1 | 13.18 | 0.1 | 12.6 | 4.79 | 10.64 | 4.59 | 4.79 | 5.57 | 4.79 | 21.39 | 9.77 | 21.68 | 9.77 | 19.53 | 9.57 | 6.35 | 0.49 | 0 | 0.39 | 0.1 | 0.68 | 4.88 | 3.22 | 4.88 | 22.46 | 4.79 | 19.24 | 9.67 | 22.46 | 9.67 | 22.07 | 9.67 | 21.49
118.304 | 0.1 | 14.36 | 0.1 | 13.67 | 0.1 | 13.18 | 4.88 | 11.33 | 4.59 | 4.88 | 5.96 | 4.88 | 14.75 | 9.86 | 21.68 | 9.77 | 20.02 | 9.57 | 7.23 | 0.49 | 0.2 | 0.39 | 0.1 | 0.59 | 4.88 | 3.42 | 4.88 | 21.97 | 4.79 | 19.82 | 9.67 | 23.05 | 9.67 | 22.66 | 9.67 | 22.07
118.336 | 0 | 14.45 | 0.1 | 14.16 | 0.1 | 13.77 | 4.88 | 11.91 | 4.59 | 4.79 | 5.66 | 4.88 | 7.42 | 9.86 | 21.58 | 9.77 | 20.41 | 9.57 | 7.91 | 0.49 | 1.37 | 0.29 | 0.1 | 0.49 | 4.88 | 4.39 | 4.88 | 21.68 | 4.79 | 20.41 | 9.67 | 23.54 | 9.67 | 23.15 | 9.67 | 22.56
118.368 | 0 | 14.36 | 0.1 | 14.75 | 0.1 | 14.26 | 4.88 | 12.5 | 4.59 | 4.79 | 5.18 | 4.88 | 7.91 | 9.86 | 21.58 | 9.77 | 20.8 | 9.57 | 8.69 | 0.29 | 2.64 | 0.2 | 1.07 | 0.59 | 4.88 | 5.86 | 4.88 | 21.58 | 4.79 | 21.09 | 9.77 | 24.12 | 9.67 | 23.63 | 9.67 | 23.05
118.4 | 0.1 | 14.36 | 0.1 | 15.23 | 0 | 14.75 | 4.88 | 13.09 | 4.59 | 4.79 | 5.08 | 4.79 | 10.94 | 9.77 | 21.49 | 9.77 | 21.29 | 9.57 | 9.47 | 0.29 | 4.69 | 0.2 | 2.15 | 0.49 | 4.88 | 6.15 | 4.88 | 21.39 | 4.79 | 21.58 | 9.77 | 24.51 | 9.67 | 24.12 | 9.67 | 23.73
118.432 | 0.1 | 14.36 | 0.1 | 15.72 | 0.1 | 15.23 | 4.88 | 13.57 | 4.49 | 4.79 | 4.98 | 4.79 | 11.23 | 9.77 | 21.49 | 9.77 | 21.58 | 9.57 | 10.35 | 0.2 | 5.96 | 0.2 | 3.81 | 0.39 | 4.88 | 5.66 | 4.88 | 21.39 | 4.79 | 22.17 | 9.77 | 25 | 9.67 | 24.61 | 9.67 | 24.12
118.464 | 0.1 | 14.36 | 0.1 | 16.21 | 0.1 | 15.72 | 4.88 | 14.16 | 4.59 | 4.98 | 4.98 | 4.88 | 10.45 | 9.86 | 21.49 | 9.77 | 22.07 | 9.67 | 11.23 | 0.2 | 7.03 | 0.1 | 5.47 | 0.39 | 4.88 | 5.37 | 4.88 | 21.29 | 4.79 | 22.66 | 9.77 | 25.59 | 9.67 | 25.2 | 9.77 | 24.81
118.496 | 0 | 14.36 | 0.1 | 16.7 | 0.1 | 16.21 | 4.88 | 14.65 | 4.49 | 5.76 | 4.88 | 4.88 | 9.77 | 9.86 | 21.49 | 9.77 | 22.36 | 9.67 | 11.91 | 0.2 | 7.81 | 0.1 | 6.64 | 0.29 | 4.88 | 5.18 | 4.88 | 21.29 | 4.79 | 23.24 | 9.77 | 26.08 | 9.77 | 25.68 | 9.77 | 25.29
118.528 | 0 | 14.36 | 0.1 | 17.19 | 0.1 | 16.7 | 4.88 | 15.14 | 4.59 | 6.84 | 4.88 | 4.88 | 10.06 | 9.86 | 21.39 | 9.77 | 22.75 | 9.67 | 12.7 | 0.2 | 8.59 | 0.1 | 7.52 | 0.29 | 4.88 | 5.08 | 4.88 | 21.29 | 4.79 | 23.73 | 9.77 | 26.66 | 9.77 | 26.17 | 9.77 | 25.68
118.56 | 0.1 | 14.36 | 0.1 | 14.36 | 0.1 | 17.19 | 4.88 | 15.63 | 4.49 | 8.3 | 4.69 | 4.88 | 10.16 | 9.77 | 14.94 | 9.77 | 23.15 | 9.67 | 13.38 | 0.1 | 9.57 | 0.1 | 8.4 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 24.32 | 9.77 | 27.05 | 9.77 | 26.66 | 9.77 | 26.17
118.592 | 0.1 | 5.66 | 0.1 | 14.75 | 0.1 | 14.75 | 4.88 | 16.11 | 4.59 | 9.08 | 4.79 | 4.88 | 10.16 | 9.77 | 7.52 | 9.77 | 23.44 | 9.67 | 14.26 | 0.1 | 10.16 | 0.1 | 9.28 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 24.81 | 9.77 | 27.54 | 9.67 | 27.25 | 9.77 | 26.76
118.624 | 0.1 | 1.66 | 0.1 | 14.55 | 0.1 | 15.04 | 4.88 | 16.6 | 4.59 | 9.86 | 4.69 | 4.88 | 9.96 | 9.86 | 7.72 | 9.86 | 23.93 | 9.67 | 14.84 | 0.1 | 10.84 | 0.1 | 9.86 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 25.29 | 9.77 | 28.13 | 9.77 | 27.54 | 9.77 | 27.25
118.656 | 0.1 | 1.86 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 17.09 | 4.69 | 10.45 | 4.69 | 4.88 | 9.86 | 9.86 | 10.74 | 9.86 | 24.32 | 9.67 | 15.63 | 0.1 | 11.72 | 0.1 | 10.64 | 0.1 | 4.88 | 4.88 | 4.88 | 21.19 | 4.88 | 25.78 | 9.77 | 28.42 | 9.77 | 28.22 | 9.77 | 27.74
118.688 | 0 | 1.76 | 0.1 | 14.55 | 0.1 | 14.55 | 4.88 | 17.58 | 4.69 | 11.13 | 4.59 | 4.88 | 9.96 | 9.86 | 11.43 | 9.86 | 20.9 | 9.67 | 16.21 | 0.1 | 12.5 | 0 | 11.43 | 0.1 | 4.88 | 4.79 | 4.88 | 9.38 | 4.79 | 26.27 | 9.77 | 29.01 | 9.77 | 28.61 | 9.77 | 28.13
118.72 | 0.1 | 1.37 | 0.1 | 14.45 | 0 | 14.65 | 4.88 | 18.07 | 4.69 | 11.72 | 4.59 | 4.88 | 9.86 | 9.77 | 10.45 | 9.86 | 21.39 | 9.67 | 17.09 | 0.1 | 13.28 | 0.1 | 12.21 | 0.1 | 4.88 | 4.69 | 4.79 | 2.54 | 4.79 | 21.58 | 9.77 | 29.59 | 9.77 | 29.01 | 9.77 | 28.61
118.752 | 0.1 | 1.37 | 0.1 | 14.55 | 0.1 | 14.45 | 4.88 | 18.46 | 4.69 | 12.31 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 21.29 | 9.67 | 17.58 | 0 | 13.97 | 0.1 | 12.89 | 0.1 | 4.88 | 4.69 | 4.88 | 3.32 | 4.79 | 21.97 | 9.86 | 26.27 | 9.77 | 29.4 | 9.77 | 29.1
118.784 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 18.95 | 4.79 | 12.79 | 4.59 | 4.88 | 9.77 | 9.86 | 9.96 | 9.86 | 21.39 | 9.67 | 18.26 | 0.1 | 14.75 | 0.1 | 13.77 | 0.1 | 4.88 | 4.69 | 4.88 | 3.71 | 4.79 | 21.68 | 9.86 | 25.29 | 9.77 | 25.78 | 9.86 | 28.52
118.816 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.34 | 4.79 | 13.38 | 4.59 | 4.88 | 9.77 | 9.86 | 10.25 | 9.86 | 21.49 | 9.67 | 18.85 | 0.1 | 15.53 | 0.1 | 14.45 | 0.1 | 5.76 | 4.69 | 4.88 | 4.88 | 4.79 | 21.58 | 9.86 | 24.81 | 9.77 | 25 | 9.86 | 24.71
118.848 | 0 | 0.98 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.82 | 4.79 | 13.87 | 4.59 | 4.98 | 9.77 | 9.86 | 10.25 | 9.86 | 21.49 | 9.67 | 19.43 | 0.1 | 16.21 | 0 | 15.23 | 0.1 | 6.84 | 4.69 | 4.88 | 6.15 | 4.88 | 21.49 | 9.86 | 24.61 | 9.77 | 24.61 | 9.86 | 24.51
118.88 | 0.1 | 0.88 | 0.1 | 14.36 | 0 | 14.45 | 4.88 | 20.22 | 4.79 | 14.45 | 4.49 | 5.86 | 9.67 | 9.86 | 9.96 | 9.86 | 21.49 | 9.67 | 20.12 | 0 | 16.8 | 0 | 15.92 | 0.1 | 8.5 | 4.59 | 4.79 | 6.05 | 4.79 | 21.29 | 9.77 | 24.51 | 9.86 | 24.41 | 9.86 | 24.32
118.912 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 20.7 | 4.79 | 14.94 | 4.59 | 6.74 | 9.67 | 9.77 | 9.86 | 9.86 | 21.49 | 9.67 | 20.7 | 0 | 17.48 | 0.1 | 16.6 | 0.1 | 10.35 | 4.49 | 4.88 | 5.57 | 4.79 | 21.19 | 9.86 | 24.51 | 9.77 | 24.41 | 9.86 | 24.41
118.944 | 0.1 | 0.68 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 21 | 4.79 | 15.43 | 4.59 | 8.2 | 9.67 | 10.64 | 9.96 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 18.16 | 0.1 | 17.19 | 0.1 | 11.62 | 4.59 | 4.88 | 5.27 | 4.79 | 21.29 | 9.86 | 24.32 | 9.86 | 24.41 | 9.86 | 24.32
118.976 | 0.1 | 0.59 | 0.1 | 12.99 | 0.1 | 14.36 | 4.88 | 18.65 | 4.79 | 15.92 | 4.59 | 9.08 | 9.67 | 11.04 | 9.96 | 9.86 | 21.49 | 9.67 | 21.78 | 0.1 | 18.85 | 0.1 | 17.87 | 0.1 | 12.7 | 4.59 | 4.88 | 5.37 | 4.88 | 21.19 | 9.86 | 24.32 | 9.77 | 24.32 | 9.86 | 24.32
119.008 | 0 | 0.59 | 0.1 | 3.91 | 0.1 | 11.72 | 4.88 | 18.95 | 4.79 | 16.41 | 4.59 | 9.86 | 9.67 | 12.01 | 9.86 | 9.86 | 21.49 | 9.77 | 22.36 | 0.1 | 19.34 | 0 | 18.56 | 0.1 | 13.57 | 4.59 | 4.88 | 5.08 | 4.79 | 21.19 | 9.86 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22
119.04 | 0.1 | 0.49 | 0.1 | 2.05 | 0.1 | 2.73 | 4.88 | 18.56 | 4.79 | 16.9 | 4.59 | 10.45 | 9.67 | 13.28 | 9.77 | 9.86 | 21.49 | 9.77 | 18.56 | 0.1 | 20.12 | 0 | 19.24 | 0.1 | 14.45 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
119.072 | 0.1 | 0.39 | 0.1 | 1.76 | 0.1 | 1.95 | 4.88 | 18.46 | 4.88 | 17.29 | 4.69 | 11.04 | 9.67 | 14.16 | 9.77 | 9.86 | 21.39 | 9.67 | 17.19 | 0 | 20.7 | 0.1 | 19.82 | 0.1 | 15.04 | 4.49 | 4.88 | 5.08 | 4.79 | 21.09 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
119.104 | 0.1 | 0.39 | 0.1 | 1.56 | 0.1 | 1.95 | 4.88 | 18.46 | 4.88 | 17.77 | 4.69 | 11.72 | 9.57 | 14.75 | 9.77 | 9.86 | 21.19 | 9.77 | 17.58 | 0.1 | 21.29 | 0.1 | 20.41 | 0.1 | 15.82 | 4.49 | 4.88 | 4.98 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.12
119.136 | 0.78 | 0.29 | 0.1 | 1.46 | 0.1 | 1.56 | 4.88 | 18.26 | 4.88 | 18.26 | 4.69 | 12.21 | 9.67 | 15.23 | 9.77 | 9.86 | 9.28 | 9.77 | 17.58 | 0.1 | 21.78 | 0.1 | 21 | 0.1 | 16.5 | 4.59 | 4.88 | 4.98 | 4.79 | 10.16 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
119.168 | 1.66 | 0.29 | 0.1 | 1.37 | 0.1 | 1.46 | 4.88 | 18.26 | 4.88 | 18.75 | 4.69 | 12.79 | 9.67 | 15.82 | 9.67 | 9.96 | 6.93 | 9.77 | 17.58 | 0.1 | 22.36 | 0 | 21.58 | 0.1 | 17.19 | 4.59 | 4.79 | 4.88 | 4.79 | 2.93 | 9.77 | 15.43 | 9.86 | 24.12 | 9.86 | 24.12
119.2 | 2.93 | 0.2 | 0.1 | 1.17 | 0.1 | 1.37 | 4.88 | 18.16 | 4.79 | 19.14 | 4.69 | 13.38 | 9.57 | 16.31 | 9.67 | 9.86 | 9.96 | 9.77 | 17.68 | 0 | 17.87 | 0 | 22.27 | 0.1 | 17.77 | 4.59 | 4.98 | 4.79 | 4.79 | 3.13 | 9.77 | 6.45 | 9.77 | 13.09 | 9.86 | 24.12
119.232 | 3.91 | 0.1 | 0.1 | 0.98 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 19.53 | 4.79 | 13.77 | 9.67 | 16.8 | 9.67 | 9.86 | 11.62 | 9.77 | 17.58 | 0 | 16.8 | 0.1 | 22.75 | 0.1 | 18.46 | 4.59 | 5.86 | 4.79 | 4.79 | 3.52 | 9.77 | 7.32 | 9.77 | 6.15 | 9.86 | 12.6
119.264 | 4.69 | 0.1 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.16 | 4.88 | 20.02 | 4.79 | 14.36 | 9.67 | 17.29 | 9.67 | 9.96 | 10.84 | 9.77 | 17.68 | 0.1 | 17.48 | 0.1 | 19.43 | 0.1 | 19.14 | 4.69 | 7.23 | 4.79 | 4.79 | 4.59 | 9.86 | 11.33 | 9.86 | 7.62 | 9.86 | 5.76
119.296 | 5.27 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.16 | 4.88 | 20.41 | 4.79 | 14.94 | 9.67 | 17.68 | 9.67 | 9.86 | 10.16 | 9.77 | 17.58 | 0.1 | 17.58 | 0.1 | 18.65 | 0.1 | 19.73 | 4.69 | 9.18 | 4.69 | 4.79 | 6.25 | 9.77 | 11.82 | 9.86 | 11.62 | 9.86 | 8.2
119.328 | 5.76 | 0.1 | 0.1 | 0.78 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.9 | 4.79 | 15.43 | 9.67 | 18.16 | 9.67 | 9.96 | 9.77 | 9.77 | 17.58 | 0.1 | 17.58 | 0 | 18.26 | 0.1 | 20.31 | 4.69 | 10.84 | 4.69 | 4.79 | 6.05 | 9.86 | 10.64 | 9.86 | 11.72 | 9.86 | 11.72
119.36 | 6.35 | 0 | 0.1 | 0.68 | 0 | 0.68 | 4.88 | 18.07 | 4.88 | 17.58 | 4.79 | 15.92 | 9.67 | 18.75 | 9.67 | 9.86 | 9.96 | 9.77 | 17.58 | 0 | 17.58 | 0 | 18.07 | 0.1 | 21 | 4.69 | 12.11 | 4.59 | 4.79 | 5.57 | 9.77 | 9.86 | 9.77 | 10.45 | 9.86 | 11.62
119.392 | 6.93 | 0 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 13.67 | 4.88 | 17.58 | 4.79 | 16.41 | 9.67 | 19.14 | 9.67 | 9.86 | 10.16 | 9.77 | 17.58 | 0 | 17.68 | 0.1 | 17.97 | 0.1 | 21.49 | 4.69 | 12.99 | 4.59 | 4.79 | 5.18 | 9.77 | 10.16 | 9.77 | 9.86 | 9.86 | 10.35
119.424 | 7.62 | 0.1 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 3.52 | 4.88 | 17.77 | 4.88 | 16.9 | 9.67 | 19.43 | 9.67 | 9.96 | 10.16 | 9.77 | 17.58 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 22.17 | 4.69 | 13.77 | 4.59 | 4.79 | 5.08 | 9.86 | 10.45 | 9.86 | 10.16 | 9.86 | 9.86
119.456 | 8.2 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.13 | 4.88 | 17.97 | 4.88 | 17.38 | 9.67 | 19.92 | 9.67 | 9.86 | 9.86 | 9.77 | 16.6 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.66 | 4.79 | 14.55 | 4.59 | 4.79 | 5.18 | 9.86 | 10.45 | 9.86 | 10.45 | 9.86 | 10.16
119.488 | 8.79 | 0.1 | 0.1 | 0.49 | 0.1 | 0.39 | 4.88 | 3.81 | 4.88 | 17.97 | 4.88 | 17.68 | 9.77 | 20.41 | 9.67 | 9.96 | 9.96 | 9.86 | 4 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.34 | 4.79 | 15.33 | 4.59 | 4.88 | 5.08 | 9.77 | 10.35 | 9.77 | 10.45 | 9.86 | 10.45
119.52 | 9.47 | 0 | 0.1 | 0.29 | 0 | 0.29 | 4.88 | 4.79 | 4.88 | 18.07 | 4.88 | 18.26 | 9.67 | 20.9 | 9.67 | 9.86 | 9.86 | 9.77 | 2.15 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.73 | 4.69 | 16.11 | 4.49 | 5.96 | 5.08 | 9.77 | 10.06 | 9.77 | 10.06 | 9.86 | 10.35
119.552 | 9.96 | 0 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 18.65 | 9.77 | 21.19 | 9.67 | 9.86 | 9.86 | 9.77 | 2.25 | 0 | 17.68 | 0 | 17.77 | 0.1 | 24.32 | 4.79 | 16.7 | 4.49 | 7.42 | 4.88 | 9.77 | 10.06 | 9.77 | 9.96 | 9.86 | 10.06
119.584 | 10.55 | 0.1 | 0.78 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 19.14 | 9.77 | 21.68 | 9.67 | 9.96 | 9.77 | 9.86 | 1.86 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 24.9 | 4.79 | 17.38 | 4.59 | 9.57 | 4.98 | 9.86 | 10.16 | 9.86 | 10.16 | 9.86 | 10.06
119.616 | 11.23 | 0.1 | 1.56 | 0.2 | 0.1 | 0.2 | 4.88 | 5.37 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 21.97 | 9.67 | 10.64 | 9.77 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.39 | 4.79 | 17.97 | 4.59 | 11.04 | 4.88 | 9.86 | 10.06 | 9.77 | 10.16 | 9.86 | 10.16
119.648 | 11.72 | 0.1 | 2.83 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 19.92 | 9.77 | 22.46 | 9.77 | 11.13 | 9.77 | 9.86 | 1.56 | 0.1 | 6.25 | 0 | 17.68 | 0.1 | 25.78 | 4.79 | 18.56 | 4.59 | 12.21 | 4.79 | 9.77 | 9.96 | 9.86 | 9.96 | 9.86 | 10.06
119.68 | 12.31 | 0 | 3.91 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 20.51 | 9.77 | 22.85 | 9.77 | 12.31 | 9.67 | 9.86 | 1.46 | 0 | 1.86 | 0 | 8.98 | 0 | 26.27 | 4.79 | 19.34 | 4.59 | 13.09 | 4.69 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86 | 9.96
119.712 | 12.79 | 0.1 | 4.69 | 0.1 | 0.1 | 0 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 20.9 | 9.77 | 23.15 | 9.67 | 13.48 | 9.67 | 9.77 | 1.37 | 0.1 | 2.25 | 0.1 | 2.34 | 0.1 | 23.05 | 4.79 | 19.92 | 4.59 | 13.87 | 4.69 | 9.77 | 9.86 | 9.77 | 9.77 | 9.86 | 9.96
119.744 | 13.38 | 0.1 | 5.27 | 0.1 | 0.98 | 0.1 | 4.88 | 4.98 | 4.88 | 17.97 | 4.88 | 17.68 | 9.77 | 23.63 | 9.77 | 14.36 | 9.67 | 9.86 | 1.56 | 0.1 | 2.05 | 0.1 | 2.25 | 0.1 | 22.27 | 4.79 | 20.61 | 4.69 | 14.65 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
119.776 | 13.97 | 0.1 | 5.76 | 0.1 | 1.86 | 0 | 4.98 | 4.88 | 4.88 | 17.97 | 4.88 | 17.58 | 9.77 | 23.93 | 9.77 | 14.84 | 9.67 | 9.86 | 1.17 | 0.1 | 1.76 | 0.1 | 2.05 | 0.1 | 21.78 | 4.79 | 21.09 | 4.69 | 15.33 | 4.69 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86 | 9.96
119.808 | 14.36 | 0.1 | 6.35 | 0.1 | 3.22 | 0 | 5.76 | 4.79 | 4.88 | 5.76 | 4.88 | 17.77 | 9.77 | 24.32 | 9.77 | 15.43 | 9.67 | 9.86 | 1.07 | 0.1 | 1.66 | 0 | 1.76 | 0.1 | 21.58 | 4.79 | 21.68 | 4.69 | 16.11 | 4.59 | 9.86 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
119.84 | 14.94 | 0 | 7.03 | 0 | 4.2 | 0 | 6.74 | 4.69 | 4.88 | 2.73 | 4.88 | 18.07 | 9.77 | 21.49 | 9.77 | 15.92 | 9.67 | 9.86 | 0.88 | 0.1 | 1.37 | 0 | 1.76 | 0 | 21.58 | 4.79 | 22.36 | 4.69 | 16.8 | 4.59 | 10.84 | 9.67 | 9.77 | 9.77 | 9.96 | 9.77
119.872 | 15.43 | 0 | 7.52 | 0.1 | 4.88 | 0 | 8.11 | 4.59 | 4.88 | 3.42 | 4.88 | 17.97 | 9.77 | 21.68 | 9.77 | 16.41 | 9.57 | 9.86 | 0.88 | 0 | 1.37 | 0 | 1.56 | 0.1 | 21.49 | 4.79 | 22.75 | 4.69 | 17.48 | 4.49 | 12.21 | 9.67 | 9.77 | 9.67 | 10.94 | 9.77
119.904 | 15.92 | 0.1 | 8.2 | 0.1 | 5.37 | 0.1 | 9.08 | 4.59 | 4.88 | 4.49 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 16.99 | 9.67 | 9.86 | 0.88 | 0.1 | 1.27 | 0.1 | 1.56 | 0.1 | 21.39 | 4.88 | 23.34 | 4.69 | 18.16 | 4.59 | 14.16 | 9.67 | 10.94 | 9.77 | 12.01 | 9.77
119.936 | 16.41 | 0.1 | 8.79 | 0.1 | 5.86 | 0 | 9.86 | 4.59 | 4.88 | 5.66 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 17.48 | 9.57 | 9.86 | 0.78 | 0.1 | 1.17 | 0.1 | 1.37 | 0.1 | 21.39 | 4.88 | 24.02 | 4.79 | 18.75 | 4.49 | 15.72 | 9.67 | 12.01 | 9.67 | 13.87 | 9.77
119.968 | 16.9 | 0.1 | 9.38 | 0.1 | 6.54 | 0.1 | 10.45 | 4.59 | 4.88 | 5.86 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 17.87 | 9.57 | 9.86 | 0.68 | 0.1 | 1.07 | 0 | 1.27 | 0.1 | 21.29 | 4.88 | 24.61 | 4.79 | 19.34 | 4.49 | 16.8 | 9.67 | 13.77 | 9.67 | 15.53 | 9.67
120 | 17.09 | 0 | 9.96 | 0 | 7.13 | 0 | 11.23 | 4.49 | 4.88 | 5.37 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 18.26 | 9.57 | 9.86 | 0.59 | 0 | 0.88 | 0 | 1.27 | 0.1 | 21.29 | 4.79 | 25 | 4.79 | 20.02 | 4.49 | 17.68 | 9.67 | 15.53 | 9.67 | 16.8 | 9.67
120.032 | 15.04 | 0 | 10.55 | 0.1 | 7.62 | 0 | 11.82 | 4.49 | 4.88 | 5.18 | 4.88 | 18.07 | 9.86 | 21.49 | 9.77 | 18.75 | 9.57 | 9.86 | 0.59 | 0 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 25.49 | 4.79 | 20.7 | 4.49 | 18.46 | 9.67 | 16.6 | 9.67 | 17.58 | 9.67
120.064 | 14.94 | 0.1 | 11.23 | 0.1 | 8.3 | 0.1 | 12.31 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 19.24 | 9.67 | 9.86 | 0.49 | 0.1 | 0.78 | 0 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 21.19 | 4.59 | 19.14 | 9.67 | 17.48 | 9.67 | 18.26 | 9.67
120.096 | 14.75 | 0.1 | 11.72 | 0.1 | 8.89 | 0 | 12.89 | 4.49 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.39 | 9.86 | 19.73 | 9.67 | 9.86 | 0.49 | 0.1 | 0.68 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 24.22 | 4.79 | 21.78 | 4.59 | 19.82 | 9.67 | 18.16 | 9.67 | 19.04 | 9.67
120.128 | 14.65 | 0 | 12.31 | 0.1 | 9.57 | 0 | 13.48 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.86 | 21.49 | 9.86 | 20.12 | 9.67 | 9.96 | 0.39 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 13.48 | 4.88 | 20.9 | 4.79 | 22.27 | 4.59 | 20.51 | 9.67 | 18.95 | 9.57 | 19.73 | 9.67
120.16 | 14.65 | 0 | 12.89 | 0.1 | 10.16 | 0 | 13.97 | 4.49 | 4.88 | 4.79 | 4.88 | 17.97 | 9.86 | 21.39 | 9.86 | 20.51 | 9.67 | 10.84 | 0.29 | 0.1 | 0.59 | 0 | 0.68 | 0.1 | 3.03 | 4.79 | 21.19 | 4.79 | 22.95 | 4.59 | 20.9 | 9.67 | 19.53 | 9.57 | 20.31 | 9.57
120.192 | 14.55 | 0 | 13.38 | 0 | 10.64 | 0 | 14.45 | 4.49 | 4.88 | 4.79 | 4.88 | 5.96 | 9.86 | 21.39 | 9.86 | 20.9 | 9.67 | 11.72 | 0.29 | 0.1 | 0.49 | 0.1 | 0.68 | 0.1 | 3.22 | 4.79 | 21.39 | 4.79 | 23.44 | 4.59 | 21.68 | 9.57 | 20.12 | 9.57 | 20.8 | 9.67
120.224 | 14.45 | 0.1 | 13.87 | 0.1 | 11.23 | 0.1 | 15.04 | 4.59 | 4.88 | 4.79 | 4.88 | 2.73 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 13.18 | 0.29 | 0.1 | 0.49 | 0 | 0.59 | 0.1 | 3.61 | 4.79 | 21.29 | 4.79 | 24.02 | 4.69 | 22.17 | 9.67 | 20.8 | 9.67 | 21.49 | 9.67
120.256 | 14.45 | 0.1 | 14.45 | 0.1 | 11.82 | 0 | 15.43 | 4.59 | 4.88 | 4.59 | 4.88 | 3.32 | 9.86 | 18.85 | 9.86 | 21.78 | 9.67 | 14.06 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.39 | 4.79 | 21.29 | 4.88 | 24.41 | 4.69 | 22.66 | 9.67 | 21.39 | 9.67 | 21.97 | 9.67
120.288 | 14.45 | 0.1 | 14.94 | 0.1 | 12.4 | 0 | 16.02 | 4.59 | 4.88 | 4.59 | 4.88 | 4.49 | 9.86 | 8.4 | 9.86 | 22.07 | 9.67 | 14.65 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 5.96 | 4.79 | 21.29 | 4.79 | 25 | 4.69 | 23.24 | 9.67 | 21.97 | 9.67 | 22.66 | 9.67
120.32 | 14.45 | 0 | 15.43 | 0 | 12.89 | 0 | 16.5 | 4.59 | 4.88 | 4.59 | 4.88 | 5.66 | 9.86 | 7.32 | 9.86 | 22.56 | 9.67 | 15.23 | 0.1 | 0 | 0.29 | 0.1 | 0.39 | 0 | 6.05 | 4.79 | 21.29 | 4.79 | 25.49 | 4.69 | 23.93 | 9.67 | 22.56 | 9.67 | 23.15 | 9.67
120.352 | 14.45 | 0 | 15.92 | 0 | 13.48 | 0 | 16.9 | 4.69 | 4.88 | 4.49 | 4.88 | 5.76 | 9.86 | 10.45 | 9.86 | 22.85 | 9.67 | 15.72 | 0.2 | 0 | 0.29 | 1.17 | 0.29 | 0.1 | 5.66 | 4.79 | 21.29 | 4.79 | 25.98 | 4.69 | 24.32 | 9.67 | 23.15 | 9.67 | 23.63 | 9.67
120.384 | 14.36 | 0.1 | 16.41 | 0.1 | 14.06 | 0 | 17.48 | 4.69 | 4.88 | 4.49 | 4.88 | 5.47 | 9.86 | 11.52 | 9.86 | 23.24 | 9.67 | 16.31 | 0.2 | 0.1 | 0.2 | 2.15 | 0.29 | 0.1 | 5.47 | 4.88 | 21.29 | 4.79 | 21.88 | 4.79 | 24.9 | 9.67 | 23.63 | 9.67 | 24.22 | 9.67
120.416 | 14.36 | 0.1 | 16.9 | 0.1 | 14.55 | 0 | 17.87 | 4.69 | 4.98 | 4.49 | 4.88 | 5.18 | 9.86 | 10.64 | 9.86 | 23.63 | 9.77 | 16.7 | 0.1 | 1.17 | 0.2 | 4 | 0.2 | 0.1 | 5.18 | 4.88 | 21.19 | 4.88 | 20.31 | 4.79 | 25.29 | 9.67 | 24.22 | 9.67 | 24.71 | 9.67
120.448 | 7.62 | 0.1 | 17.38 | 0.1 | 15.04 | 0.1 | 18.26 | 4.69 | 5.76 | 4.49 | 4.88 | 4.88 | 9.86 | 9.96 | 9.86 | 24.02 | 9.77 | 17.19 | 0.1 | 2.44 | 0.2 | 5.66 | 0.2 | 0.1 | 5.08 | 4.79 | 21.29 | 4.88 | 21 | 4.79 | 25.88 | 9.77 | 24.71 | 9.67 | 25.1 | 9.67
120.48 | 2.05 | 0 | 14.94 | 0.1 | 15.63 | 0 | 18.75 | 4.69 | 6.93 | 4.49 | 4.88 | 4.98 | 9.86 | 9.96 | 9.86 | 24.41 | 9.67 | 17.68 | 0.1 | 4.39 | 0.1 | 6.74 | 0.2 | 0.39 | 5.08 | 4.79 | 21.19 | 4.79 | 21.09 | 4.79 | 26.37 | 9.77 | 25.29 | 9.67 | 25.78 | 9.67
120.512 | 1.95 | 0 | 15.04 | 0 | 16.02 | 0 | 19.24 | 4.79 | 8.2 | 4.49 | 4.88 | 4.88 | 9.86 | 10.16 | 9.86 | 21.58 | 9.77 | 18.07 | 0.1 | 5.66 | 0.1 | 7.52 | 0.2 | 1.56 | 5.18 | 4.79 | 21.19 | 4.79 | 21.19 | 4.79 | 26.86 | 9.77 | 25.68 | 9.67 | 26.17 | 9.67
120.544 | 1.76 | 0.1 | 14.84 | 0.1 | 16.5 | 0 | 19.73 | 4.79 | 9.28 | 4.49 | 4.88 | 4.98 | 9.86 | 10.16 | 9.96 | 22.17 | 9.77 | 18.56 | 0.1 | 6.84 | 0.1 | 8.4 | 0.2 | 2.83 | 4.98 | 4.88 | 7.72 | 4.88 | 21.19 | 4.79 | 27.34 | 9.77 | 26.17 | 9.77 | 26.76 | 9.77
120.576 | 1.46 | 0.1 | 14.75 | 0.1 | 16.99 | 0.1 | 20.12 | 4.79 | 9.96 | 4.49 | 4.88 | 4.79 | 9.86 | 10.06 | 9.86 | 21.68 | 9.77 | 19.04 | 0.1 | 7.62 | 0.1 | 9.08 | 0.1 | 4.88 | 4.88 | 4.79 | 2.25 | 4.79 | 21.19 | 4.79 | 27.83 | 9.77 | 26.66 | 9.77 | 27.25 | 9.77
120.608 | 1.37 | 0.1 | 14.75 | 0.1 | 15.82 | 0 | 20.51 | 4.79 | 10.55 | 4.49 | 4.88 | 4.79 | 9.86 | 9.86 | 9.86 | 21.68 | 9.77 | 19.43 | 0.1 | 8.5 | 0.1 | 9.96 | 0.1 | 6.25 | 4.88 | 4.79 | 3.32 | 4.79 | 21.19 | 4.79 | 28.32 | 9.77 | 27.15 | 9.77 | 27.64 | 9.77
120.64 | 1.27 | 0 | 14.55 | 0 | 14.06 | 0.1 | 20.9 | 4.79 | 11.23 | 4.49 | 4.79 | 4.69 | 9.86 | 9.77 | 9.86 | 21.68 | 9.77 | 19.92 | 0.1 | 9.28 | 0.1 | 10.64 | 0.1 | 7.13 | 4.88 | 4.79 | 4 | 4.79 | 21.19 | 4.79 | 28.71 | 9.77 | 27.74 | 9.67 | 28.22 | 9.77
120.672 | 1.07 | 0 | 14.55 | 0 | 14.45 | 0 | 17.87 | 4.79 | 11.72 | 4.49 | 4.88 | 4.59 | 9.86 | 9.86 | 9.86 | 21.58 | 9.77 | 20.22 | 0.1 | 9.96 | 0 | 11.62 | 0.1 | 7.81 | 4.69 | 4.79 | 5.18 | 4.79 | 21.29 | 4.79 | 29.1 | 9.77 | 28.13 | 9.67 | 28.61 | 9.77
120.704 | 0.98 | 0.1 | 14.55 | 0.1 | 14.65 | 0 | 17.68 | 4.88 | 12.31 | 4.59 | 4.88 | 4.59 | 9.86 | 9.86 | 9.96 | 21.58 | 9.77 | 20.7 | 0.1 | 10.84 | 0.1 | 12.31 | 0.1 | 8.59 | 4.69 | 4.88 | 6.25 | 4.88 | 21.19 | 4.79 | 26.27 | 9.77 | 28.61 | 9.77 | 29.1 | 9.77
120.736 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 0 | 17.97 | 4.88 | 12.89 | 4.59 | 4.88 | 4.59 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.19 | 0.1 | 11.72 | 0.1 | 12.99 | 0.1 | 9.57 | 4.69 | 4.88 | 6.15 | 4.88 | 21.19 | 4.79 | 23.73 | 9.77 | 29.1 | 9.77 | 29.01 | 9.77
120.768 | 0.78 | 0 | 14.45 | 0.1 | 14.55 | 0 | 18.16 | 4.88 | 13.38 | 4.59 | 4.88 | 4.49 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 21.49 | 0.1 | 12.31 | 0.1 | 13.87 | 0.1 | 10.35 | 4.69 | 4.79 | 5.57 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 29.01 | 9.77 | 25.49 | 9.77
120.8 | 0.78 | 0 | 14.45 | 0.1 | 14.45 | 0 | 18.07 | 4.79 | 13.87 | 4.59 | 4.88 | 4.49 | 9.86 | 9.77 | 9.86 | 21.49 | 9.77 | 21.88 | 0.1 | 13.09 | 0.1 | 14.45 | 0.1 | 11.13 | 4.59 | 4.79 | 5.18 | 4.79 | 21.19 | 4.79 | 24.22 | 9.77 | 24.9 | 9.77 | 24.61 | 9.77
120.832 | 0.59 | 0 | 14.45 | 0 | 14.45 | 0 | 18.16 | 4.88 | 14.45 | 4.69 | 5.57 | 4.49 | 9.86 | 9.67 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 13.87 | 0 | 15.33 | 0.1 | 11.82 | 4.59 | 4.79 | 5.18 | 4.79 | 7.03 | 4.79 | 24.32 | 9.77 | 24.9 | 9.77 | 24.61 | 9.77
120.864 | 0.59 | 0.1 | 14.45 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.04 | 4.69 | 6.54 | 4.49 | 9.86 | 9.67 | 9.96 | 21.49 | 9.86 | 22.66 | 0.1 | 14.55 | 0.1 | 15.92 | 0.1 | 12.79 | 4.59 | 4.88 | 5.08 | 4.79 | 2.34 | 4.88 | 24.32 | 9.86 | 24.61 | 9.86 | 24.51 | 9.77
120.896 | 0.49 | 0.1 | 8.01 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.43 | 4.69 | 8.01 | 4.49 | 9.86 | 9.67 | 9.86 | 21.49 | 9.86 | 23.05 | 0.1 | 15.33 | 0.1 | 16.6 | 0.1 | 13.48 | 4.59 | 5.76 | 5.08 | 4.79 | 3.22 | 4.88 | 24.32 | 9.77 | 24.51 | 9.86 | 24.41 | 9.86
120.928 | 0.39 | 0.1 | 2.34 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 16.02 | 4.79 | 8.98 | 4.49 | 9.86 | 9.67 | 9.86 | 17.77 | 9.86 | 23.54 | 0.1 | 16.02 | 0.1 | 17.38 | 0.1 | 14.06 | 4.59 | 6.45 | 4.98 | 4.79 | 4.1 | 4.79 | 24.32 | 9.77 | 24.41 | 9.86 | 24.41 | 9.86
120.96 | 0.39 | 0 | 1.95 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.5 | 4.69 | 9.77 | 4.49 | 9.86 | 9.57 | 9.86 | 7.81 | 9.86 | 23.83 | 0.1 | 16.7 | 0.1 | 17.97 | 0.1 | 14.94 | 4.49 | 8.2 | 4.98 | 4.79 | 5.37 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
120.992 | 0.29 | 0 | 1.86 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.9 | 4.69 | 10.45 | 4.59 | 10.45 | 9.57 | 9.86 | 7.42 | 9.86 | 24.12 | 0.1 | 17.29 | 0 | 18.65 | 0.1 | 15.63 | 4.49 | 10.16 | 4.98 | 4.79 | 6.15 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
121.024 | 0.29 | 0.1 | 1.56 | 0.1 | 14.36 | 0 | 18.16 | 4.88 | 17.38 | 4.79 | 11.04 | 4.59 | 11.13 | 9.57 | 9.96 | 10.45 | 9.86 | 20.8 | 0.1 | 17.97 | 0.1 | 19.34 | 0.1 | 16.31 | 4.59 | 11.52 | 4.88 | 4.79 | 5.96 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
121.056 | 0.2 | 0 | 1.46 | 0.1 | 6.54 | 0 | 18.16 | 4.88 | 17.87 | 4.79 | 11.62 | 4.59 | 12.21 | 9.67 | 9.96 | 11.43 | 9.86 | 20.9 | 0.1 | 18.56 | 0.1 | 19.92 | 0.1 | 16.99 | 4.49 | 12.5 | 4.79 | 4.88 | 5.57 | 4.79 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
121.088 | 0.2 | 0.1 | 1.27 | 0.1 | 1.86 | 0 | 18.16 | 4.88 | 18.26 | 4.79 | 12.21 | 4.69 | 13.48 | 9.57 | 9.96 | 10.74 | 9.86 | 21.09 | 0.1 | 19.24 | 0.1 | 20.51 | 0.1 | 17.58 | 4.59 | 13.38 | 4.79 | 4.79 | 5.27 | 4.79 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
121.12 | 0.1 | 0 | 1.17 | 0 | 1.86 | 0 | 6.45 | 4.88 | 18.75 | 4.79 | 12.7 | 4.59 | 14.26 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 19.92 | 0.1 | 21 | 0.1 | 18.26 | 4.59 | 14.16 | 4.69 | 4.79 | 5.08 | 4.79 | 24.22 | 9.77 | 24.22 | 9.77 | 24.12 | 9.86
121.152 | 0.1 | 0 | 0.98 | 0 | 1.86 | 0 | 2.64 | 4.88 | 19.24 | 4.79 | 13.28 | 4.59 | 14.94 | 9.57 | 9.86 | 9.86 | 9.86 | 21.29 | 0.1 | 20.51 | 0.1 | 21.58 | 0.1 | 18.85 | 4.59 | 14.84 | 4.69 | 4.79 | 5.18 | 4.79 | 11.43 | 9.77 | 24.12 | 9.77 | 24.12 | 9.86
121.184 | 0.1 | 0.1 | 0.88 | 0.1 | 1.56 | 0 | 3.42 | 4.88 | 19.63 | 4.79 | 13.87 | 4.69 | 15.33 | 9.57 | 9.96 | 10.25 | 9.86 | 21.39 | 0.1 | 21.19 | 0.1 | 22.27 | 0.1 | 19.53 | 4.69 | 15.63 | 4.69 | 4.79 | 5.08 | 4.79 | 5.66 | 9.86 | 24.22 | 9.86 | 13.18 | 9.86
121.216 | 0 | 0.1 | 0.88 | 0.1 | 1.46 | 0 | 4.39 | 4.88 | 20.12 | 4.79 | 14.36 | 4.69 | 15.92 | 9.67 | 9.96 | 10.16 | 9.86 | 21.39 | 0.1 | 21.78 | 0.1 | 22.27 | 0.1 | 20.22 | 4.69 | 16.41 | 4.59 | 4.79 | 5.08 | 4.79 | 8.69 | 9.86 | 13.28 | 9.86 | 5.96 | 9.86
121.248 | 0 | 0.1 | 0.88 | 0.1 | 1.17 | 0 | 5.57 | 4.88 | 20.51 | 4.88 | 14.84 | 4.79 | 16.41 | 9.67 | 10.55 | 9.96 | 9.86 | 21.39 | 0.1 | 22.36 | 0.1 | 18.26 | 0.1 | 20.8 | 4.69 | 16.99 | 4.59 | 4.79 | 4.98 | 4.88 | 12.21 | 9.86 | 5.96 | 9.86 | 7.72 | 9.86
121.28 | 0.1 | 1.07 | 0.68 | 0 | 1.17 | 0 | 5.96 | 4.88 | 20.8 | 4.79 | 15.33 | 4.79 | 16.99 | 9.67 | 11.33 | 9.86 | 9.86 | 21.39 | 0.1 | 19.34 | 0.1 | 18.26 | 0.1 | 21.39 | 4.69 | 17.77 | 4.59 | 4.79 | 4.88 | 4.79 | 11.33 | 9.77 | 8.4 | 9.77 | 11.72 | 9.86
121.312 | 0.1 | 1.86 | 0.68 | 0 | 1.07 | 0 | 5.57 | 4.88 | 17.38 | 4.79 | 15.92 | 4.69 | 17.38 | 9.67 | 12.31 | 9.86 | 9.86 | 21.39 | 0.1 | 17.19 | 0 | 18.07 | 0.1 | 21.97 | 4.69 | 18.26 | 4.59 | 4.79 | 4.88 | 4.79 | 10.16 | 9.77 | 11.62 | 9.77 | 11.82 | 9.86
121.344 | 0.1 | 3.13 | 0.49 | 0.1 | 0.88 | 0 | 5.37 | 4.88 | 17.68 | 4.88 | 16.41 | 4.79 | 17.87 | 9.67 | 13.57 | 9.86 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.97 | 0.1 | 22.56 | 4.79 | 18.95 | 4.59 | 4.79 | 4.79 | 4.79 | 10.16 | 9.86 | 11.72 | 9.86 | 10.45 | 9.86
121.376 | 0.1 | 4 | 0.49 | 0.1 | 0.78 | 0 | 4.98 | 4.88 | 17.77 | 4.88 | 16.9 | 4.79 | 18.36 | 9.67 | 14.36 | 9.86 | 9.86 | 21.39 | 0.1 | 17.77 | 0.1 | 17.87 | 0.1 | 17.87 | 4.79 | 19.73 | 4.59 | 5.96 | 4.69 | 4.79 | 10.45 | 9.86 | 10.45 | 9.86 | 9.96 | 9.86
121.408 | 0 | 4.79 | 0.49 | 0.1 | 0.78 | 0 | 4.88 | 4.88 | 18.07 | 4.88 | 17.29 | 4.79 | 18.75 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 16.8 | 4.79 | 20.31 | 4.59 | 6.93 | 4.69 | 4.79 | 10.45 | 9.86 | 9.86 | 9.86 | 10.25 | 9.86
121.44 | 0.1 | 5.27 | 0.39 | 0 | 0.59 | 0 | 4.98 | 4.88 | 18.07 | 4.79 | 17.77 | 4.79 | 19.24 | 9.67 | 15.43 | 9.77 | 9.86 | 21.29 | 0.1 | 17.58 | 0 | 17.77 | 0.1 | 17.58 | 4.79 | 20.8 | 4.59 | 8.89 | 4.69 | 4.79 | 10.25 | 9.77 | 10.25 | 9.77 | 10.55 | 9.86
121.472 | 0.1 | 5.86 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.07 | 4.79 | 18.16 | 4.79 | 19.73 | 9.67 | 15.92 | 9.77 | 9.86 | 10.64 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.68 | 4.79 | 21.39 | 4.59 | 10.55 | 4.59 | 4.79 | 10.16 | 9.77 | 10.35 | 9.77 | 10.35 | 9.86
121.504 | 0.1 | 6.45 | 0.29 | 0.1 | 0.59 | 0 | 4.79 | 4.88 | 18.16 | 4.88 | 18.65 | 4.79 | 20.12 | 9.67 | 16.41 | 9.67 | 9.86 | 6.54 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 21.97 | 4.59 | 11.82 | 4.59 | 4.79 | 10.06 | 9.86 | 10.35 | 9.86 | 10.16 | 9.86
121.536 | 0 | 7.03 | 0.2 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.07 | 4.88 | 19.14 | 4.88 | 20.51 | 9.77 | 16.99 | 9.67 | 9.86 | 9.38 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.68 | 4.79 | 22.56 | 4.59 | 12.79 | 4.59 | 4.79 | 10.16 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
121.568 | 0 | 7.62 | 0.2 | 0.1 | 0.39 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 19.53 | 4.88 | 20.9 | 9.77 | 17.48 | 9.67 | 9.86 | 11.52 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 4.88 | 23.05 | 4.59 | 13.57 | 4.59 | 4.79 | 10.16 | 9.86 | 10.06 | 9.86 | 10.25 | 9.86
121.6 | 0.1 | 8.2 | 0.1 | 0.1 | 0.29 | 0 | 4.69 | 4.88 | 18.07 | 4.79 | 20.02 | 4.88 | 21.39 | 9.67 | 17.97 | 9.67 | 9.86 | 10.84 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 4.88 | 23.63 | 4.69 | 14.45 | 4.49 | 4.79 | 9.96 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86
121.632 | 0.1 | 8.89 | 0.1 | 0 | 0.2 | 0 | 4.69 | 4.88 | 18.07 | 4.79 | 20.31 | 4.79 | 21.68 | 9.77 | 18.36 | 9.67 | 9.86 | 10.06 | 0.1 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 4.79 | 24.12 | 4.69 | 15.23 | 4.59 | 4.79 | 9.86 | 9.77 | 10.06 | 9.77 | 9.86 | 9.86
121.664 | 0.1 | 9.47 | 0.1 | 0.49 | 0.29 | 0 | 4.59 | 4.88 | 18.07 | 4.88 | 20.8 | 4.88 | 22.17 | 9.77 | 18.85 | 9.67 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 8.5 | 0.1 | 17.77 | 4.88 | 24.61 | 4.69 | 15.92 | 4.59 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
121.696 | 0.1 | 10.06 | 0.1 | 1.37 | 0.2 | 0.49 | 4.59 | 4.88 | 18.07 | 4.88 | 19.34 | 4.88 | 22.56 | 9.77 | 19.24 | 9.67 | 9.86 | 9.96 | 0.1 | 17.58 | 0.1 | 1.86 | 0.1 | 17.77 | 4.88 | 25.1 | 4.69 | 16.6 | 4.59 | 4.79 | 9.96 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
121.728 | 0 | 10.64 | 0.1 | 2.34 | 0.2 | 1.27 | 4.59 | 4.88 | 18.07 | 4.79 | 17.77 | 4.88 | 22.95 | 9.77 | 19.73 | 9.67 | 9.86 | 10.16 | 0.1 | 7.23 | 0.1 | 2.25 | 0.1 | 17.77 | 4.88 | 25.68 | 4.69 | 17.19 | 4.49 | 4.79 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96 | 9.86
121.76 | 0.1 | 11.23 | 0.1 | 3.71 | 0.1 | 2.34 | 4.59 | 4.88 | 5.18 | 4.79 | 17.87 | 4.88 | 23.24 | 9.77 | 20.12 | 9.67 | 9.86 | 9.96 | 0.1 | 1.56 | 0.1 | 2.05 | 0.1 | 17.68 | 4.88 | 26.27 | 4.79 | 17.87 | 4.49 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
121.792 | 0.1 | 11.72 | 0.1 | 4.49 | 0 | 3.61 | 4.59 | 4.88 | 2.93 | 4.79 | 18.16 | 4.88 | 23.63 | 9.77 | 20.51 | 9.57 | 9.96 | 9.77 | 0.1 | 2.34 | 0 | 1.66 | 0.1 | 17.68 | 4.88 | 22.07 | 4.79 | 18.46 | 4.59 | 4.79 | 9.77 | 9.77 | 9.77 | 9.77 | 9.77 | 9.86
121.824 | 0.1 | 12.4 | 0.1 | 5.08 | 0.1 | 4.49 | 4.59 | 4.88 | 3.52 | 4.88 | 18.07 | 4.88 | 24.12 | 9.77 | 20.9 | 9.67 | 10.74 | 9.86 | 0.1 | 2.25 | 0.1 | 1.76 | 0.1 | 5.76 | 4.88 | 21.39 | 4.79 | 19.14 | 4.69 | 5.76 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
121.856 | 0.1 | 12.89 | 0.1 | 5.57 | 0.1 | 5.08 | 4.59 | 4.88 | 4.49 | 4.88 | 18.16 | 4.88 | 24.41 | 9.77 | 21.39 | 9.67 | 11.82 | 9.86 | 0.1 | 1.76 | 0.1 | 1.46 | 0.1 | 1.86 | 4.88 | 21.49 | 4.79 | 19.82 | 4.69 | 6.84 | 9.77 | 10.64 | 9.77 | 9.86 | 9.77 | 9.86
121.888 | 0 | 13.48 | 0.1 | 6.15 | 0.1 | 5.57 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 4.88 | 22.07 | 9.77 | 21.68 | 9.67 | 13.18 | 9.77 | 0.1 | 1.76 | 0.1 | 1.46 | 0.1 | 2.15 | 4.88 | 21.39 | 4.79 | 20.51 | 4.69 | 8.5 | 9.67 | 11.72 | 9.77 | 10.84 | 9.77 | 9.86
121.92 | 0.1 | 13.97 | 0.1 | 6.64 | 0.1 | 6.15 | 4.59 | 4.88 | 5.86 | 4.79 | 18.07 | 4.88 | 22.36 | 9.77 | 22.17 | 9.67 | 14.06 | 9.67 | 0.1 | 1.66 | 0.1 | 1.17 | 0.1 | 1.95 | 4.88 | 21.58 | 4.79 | 21.09 | 4.69 | 10.55 | 9.67 | 13.38 | 9.77 | 11.91 | 9.67 | 10.94
121.952 | 0.1 | 14.45 | 0.1 | 7.23 | 0.1 | 6.74 | 4.59 | 4.88 | 5.57 | 4.79 | 18.07 | 4.79 | 21.78 | 9.77 | 22.56 | 9.67 | 14.65 | 9.67 | 0.1 | 1.56 | 0.1 | 1.17 | 0 | 1.76 | 4.88 | 21.39 | 4.79 | 21.58 | 4.69 | 11.62 | 9.67 | 15.14 | 9.67 | 13.48 | 9.67 | 12.01
121.984 | 0.1 | 15.04 | 0.1 | 7.91 | 0.1 | 7.32 | 4.59 | 4.88 | 5.27 | 4.88 | 18.16 | 4.88 | 21.68 | 9.86 | 22.85 | 9.67 | 15.23 | 9.67 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.29 | 4.79 | 22.17 | 4.69 | 12.7 | 9.67 | 16.41 | 9.67 | 15.23 | 9.67 | 13.77
122.016 | 0.1 | 15.53 | 0.1 | 8.5 | 0.1 | 7.91 | 4.69 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.68 | 9.77 | 23.24 | 9.67 | 15.72 | 9.67 | 0.1 | 1.17 | 0.1 | 0.88 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 22.75 | 4.79 | 13.48 | 9.67 | 17.29 | 9.67 | 16.5 | 9.67 | 15.43
122.048 | 0 | 15.92 | 0.1 | 9.08 | 0.1 | 8.5 | 4.69 | 4.88 | 4.88 | 4.79 | 18.07 | 4.88 | 21.58 | 9.77 | 23.73 | 9.67 | 16.31 | 9.67 | 0.1 | 1.17 | 0.1 | 0.88 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 23.24 | 4.79 | 14.26 | 9.67 | 17.97 | 9.67 | 17.29 | 9.67 | 16.6
122.08 | 0.1 | 16.5 | 0.1 | 9.67 | 0.1 | 9.18 | 4.69 | 5.57 | 4.88 | 4.79 | 18.07 | 4.88 | 21.58 | 9.77 | 24.12 | 9.67 | 16.7 | 9.67 | 0.1 | 0.98 | 0 | 0.78 | 0.1 | 1.37 | 4.88 | 21.29 | 4.79 | 23.83 | 4.79 | 15.04 | 9.57 | 18.65 | 9.67 | 18.16 | 9.67 | 17.58
122.112 | 0.1 | 16.9 | 0.1 | 10.25 | 0.1 | 9.67 | 4.69 | 6.05 | 4.88 | 4.79 | 18.07 | 4.79 | 21.49 | 9.77 | 24.41 | 9.67 | 17.29 | 9.57 | 0.1 | 0.88 | 0 | 0.68 | 0.1 | 1.17 | 4.88 | 21.19 | 4.79 | 24.22 | 4.79 | 15.72 | 9.67 | 19.34 | 9.57 | 18.75 | 9.67 | 18.16
122.144 | 0.1 | 15.82 | 0.1 | 10.94 | 0.1 | 10.35 | 4.79 | 7.13 | 4.79 | 4.88 | 7.42 | 4.88 | 21.49 | 9.86 | 22.07 | 9.77 | 17.68 | 9.57 | 0.29 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.19 | 4.88 | 24.81 | 4.79 | 16.5 | 9.67 | 20.02 | 9.67 | 19.53 | 9.67 | 18.85
122.176 | 0.1 | 14.16 | 0.1 | 11.52 | 0.1 | 10.94 | 4.79 | 8.5 | 4.79 | 4.79 | 2.44 | 4.88 | 21.49 | 9.77 | 22.36 | 9.77 | 18.16 | 9.57 | 1.56 | 0.68 | 0.1 | 0.59 | 0.1 | 0.98 | 4.88 | 21.19 | 4.79 | 25.39 | 4.79 | 17.09 | 9.67 | 20.61 | 9.67 | 20.12 | 9.67 | 19.53
122.208 | 0 | 14.45 | 0.1 | 12.01 | 0.1 | 11.43 | 4.79 | 9.38 | 4.79 | 4.79 | 3.32 | 4.88 | 21.49 | 9.77 | 21.78 | 9.77 | 18.65 | 9.57 | 2.93 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 17.77 | 4.79 | 25.78 | 4.79 | 17.77 | 9.67 | 21.29 | 9.67 | 20.8 | 9.67 | 20.22
122.24 | 0.1 | 14.45 | 0.1 | 12.6 | 0 | 12.11 | 4.79 | 10.06 | 4.59 | 4.79 | 4.39 | 4.88 | 21.39 | 9.77 | 21.78 | 9.77 | 19.24 | 9.57 | 4.88 | 0.59 | 0.1 | 0.49 | 0.1 | 0.78 | 4.88 | 3.81 | 4.79 | 25.68 | 4.79 | 18.46 | 9.67 | 21.88 | 9.67 | 21.39 | 9.67 | 20.7
122.272 | 0.1 | 14.45 | 0.1 | 13.09 | 0.1 | 12.6 | 4.79 | 10.55 | 4.59 | 4.79 | 5.57 | 4.79 | 21.39 | 9.77 | 21.68 | 9.77 | 19.43 | 9.57 | 6.15 | 0.49 | 0 | 0.49 | 0.1 | 0.68 | 4.88 | 3.03 | 4.79 | 22.66 | 4.79 | 19.04 | 9.67 | 22.46 | 9.67 | 21.78 | 9.67 | 21.39
122.304 | 0.1 | 14.45 | 0.1 | 13.67 | 0.1 | 13.18 | 4.88 | 11.23 | 4.59 | 4.79 | 6.05 | 4.88 | 16.7 | 9.77 | 21.58 | 9.77 | 19.92 | 9.57 | 7.03 | 0.49 | 0.1 | 0.29 | 0.1 | 0.59 | 4.88 | 3.42 | 4.88 | 22.07 | 4.79 | 19.73 | 9.67 | 23.05 | 9.67 | 22.66 | 9.67 | 21.97
122.336 | 0.1 | 14.45 | 0.1 | 14.16 | 0.1 | 13.67 | 4.88 | 11.82 | 4.59 | 4.79 | 5.57 | 4.88 | 7.72 | 9.86 | 21.58 | 9.77 | 20.31 | 9.57 | 7.81 | 0.39 | 1.17 | 0.29 | 0.1 | 0.59 | 4.88 | 4.2 | 4.88 | 21.68 | 4.79 | 20.31 | 9.67 | 23.44 | 9.67 | 23.05 | 9.67 | 22.56
122.368 | 0 | 14.36 | 0.1 | 14.75 | 0.1 | 14.16 | 4.88 | 12.4 | 4.59 | 4.79 | 5.18 | 4.88 | 7.72 | 9.77 | 21.58 | 9.77 | 20.8 | 9.57 | 8.5 | 0.29 | 2.54 | 0.2 | 0.88 | 0.59 | 4.88 | 5.66 | 4.79 | 21.58 | 4.88 | 20.9 | 9.67 | 24.02 | 9.67 | 23.54 | 9.67 | 23.15
122.4 | 0.1 | 14.36 | 0.1 | 15.23 | 0 | 14.75 | 4.88 | 12.99 | 4.49 | 4.79 | 4.98 | 4.88 | 10.84 | 9.77 | 21.49 | 9.77 | 21.19 | 9.57 | 9.47 | 0.29 | 4.49 | 0.2 | 1.95 | 0.49 | 4.88 | 6.15 | 4.88 | 21.39 | 4.79 | 21.49 | 9.67 | 24.61 | 9.67 | 24.22 | 9.67 | 23.63
122.432 | 0.1 | 14.36 | 0.1 | 15.72 | 0.1 | 15.23 | 4.88 | 13.48 | 4.49 | 4.79 | 4.98 | 4.88 | 11.43 | 9.77 | 21.49 | 9.77 | 21.58 | 9.57 | 10.16 | 0.2 | 5.76 | 0.2 | 3.61 | 0.39 | 4.88 | 5.86 | 4.88 | 21.39 | 4.88 | 22.07 | 9.67 | 25 | 9.67 | 24.61 | 9.67 | 24.12
122.464 | 0.1 | 14.36 | 0.1 | 16.21 | 0.1 | 15.72 | 4.88 | 14.06 | 4.49 | 4.88 | 4.98 | 4.88 | 10.45 | 9.86 | 21.49 | 9.86 | 21.97 | 9.67 | 10.84 | 0.2 | 6.93 | 0.1 | 5.37 | 0.39 | 4.88 | 5.47 | 4.88 | 21.29 | 4.88 | 22.56 | 9.77 | 25.59 | 9.67 | 25.1 | 9.77 | 24.61
122.496 | 0 | 14.36 | 0.1 | 16.7 | 0.1 | 16.21 | 4.88 | 14.55 | 4.49 | 5.57 | 4.88 | 4.88 | 9.96 | 9.77 | 21.49 | 9.86 | 22.36 | 9.67 | 11.72 | 0.2 | 7.62 | 0.1 | 6.45 | 0.29 | 4.88 | 5.18 | 4.79 | 21.29 | 4.88 | 23.24 | 9.77 | 26.17 | 9.67 | 25.59 | 9.77 | 25.2
122.528 | 0 | 14.36 | 0.1 | 17.19 | 0.1 | 16.7 | 4.88 | 15.04 | 4.59 | 6.64 | 4.88 | 4.88 | 9.96 | 9.77 | 21.49 | 9.86 | 22.75 | 9.67 | 12.5 | 0.2 | 8.4 | 0.1 | 7.42 | 0.29 | 4.88 | 5.08 | 4.79 | 21.29 | 4.88 | 23.73 | 9.77 | 26.47 | 9.67 | 26.17 | 9.77 | 25.68
122.56 | 0.1 | 14.36 | 0.1 | 14.55 | 0.1 | 17.19 | 4.88 | 15.53 | 4.59 | 8.11 | 4.79 | 4.88 | 10.16 | 9.77 | 15.82 | 9.86 | 23.15 | 9.67 | 13.28 | 0.1 | 9.38 | 0.1 | 8.2 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 24.22 | 9.77 | 27.05 | 9.67 | 26.76 | 9.77 | 26.17
122.592 | 0.1 | 6.54 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 16.02 | 4.59 | 8.98 | 4.69 | 4.88 | 10.25 | 9.77 | 7.72 | 9.86 | 23.44 | 9.67 | 14.06 | 0.1 | 10.06 | 0.1 | 9.08 | 0.2 | 4.88 | 4.98 | 4.79 | 21.19 | 4.79 | 24.71 | 9.77 | 27.54 | 9.67 | 27.05 | 9.77 | 26.76
122.624 | 0.1 | 1.76 | 0.1 | 14.45 | 0.1 | 14.84 | 4.88 | 16.5 | 4.59 | 9.77 | 4.69 | 4.88 | 9.96 | 9.77 | 7.72 | 9.86 | 23.83 | 9.67 | 14.65 | 0.1 | 10.74 | 0.1 | 9.77 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.29 | 9.77 | 28.03 | 9.77 | 27.64 | 9.77 | 27.15
122.656 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 16.99 | 4.59 | 10.45 | 4.59 | 4.88 | 9.86 | 9.77 | 10.64 | 9.86 | 24.22 | 9.67 | 15.43 | 0.1 | 11.62 | 0.1 | 10.55 | 0.2 | 4.88 | 4.88 | 4.88 | 21.19 | 4.79 | 25.78 | 9.77 | 28.52 | 9.77 | 28.13 | 9.77 | 27.64
122.688 | 0 | 1.66 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 17.58 | 4.69 | 11.04 | 4.69 | 4.88 | 9.96 | 9.77 | 11.43 | 9.86 | 20.8 | 9.67 | 16.11 | 0.1 | 12.4 | 0 | 11.43 | 0.1 | 4.88 | 4.88 | 4.88 | 9.77 | 4.79 | 26.27 | 9.77 | 29.01 | 9.77 | 28.52 | 9.77 | 28.13
122.72 | 0.1 | 1.37 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 17.87 | 4.69 | 11.62 | 4.59 | 4.88 | 9.86 | 9.77 | 10.55 | 9.86 | 20.9 | 9.67 | 16.7 | 0.1 | 13.18 | 0 | 12.11 | 0.1 | 4.88 | 4.79 | 4.79 | 2.54 | 4.79 | 22.85 | 9.77 | 29.49 | 9.77 | 29.01 | 9.77 | 28.61
122.752 | 0.1 | 1.46 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.26 | 4.69 | 12.21 | 4.59 | 4.88 | 9.86 | 9.77 | 9.86 | 9.86 | 21.09 | 9.67 | 17.48 | 0.1 | 13.97 | 0.1 | 12.89 | 0.1 | 4.88 | 4.69 | 4.88 | 3.32 | 4.79 | 22.27 | 9.77 | 25.88 | 9.77 | 29.4 | 9.77 | 29.01
122.784 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.55 | 4.88 | 18.85 | 4.69 | 12.79 | 4.59 | 4.88 | 9.86 | 9.77 | 9.96 | 9.86 | 21.39 | 9.67 | 18.16 | 0.1 | 14.55 | 0.1 | 13.67 | 0.1 | 4.88 | 4.69 | 4.88 | 3.71 | 4.88 | 21.78 | 9.77 | 25.1 | 9.77 | 26.27 | 9.86 | 29.4
122.816 | 0.1 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.79 | 13.28 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.39 | 9.67 | 18.75 | 0.1 | 15.33 | 0.1 | 14.36 | 0.1 | 5.76 | 4.69 | 4.88 | 4.79 | 4.88 | 21.58 | 9.77 | 24.71 | 9.77 | 25.2 | 9.86 | 25.68
122.848 | 0 | 0.88 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.82 | 4.79 | 13.87 | 4.59 | 4.98 | 9.77 | 9.77 | 10.25 | 9.86 | 21.39 | 9.67 | 19.34 | 0.1 | 16.11 | 0 | 15.14 | 0.1 | 6.64 | 4.59 | 4.88 | 6.05 | 4.88 | 21.58 | 9.86 | 24.61 | 9.77 | 24.81 | 9.86 | 25
122.88 | 0.1 | 0.88 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.12 | 4.79 | 14.36 | 4.49 | 5.76 | 9.67 | 9.77 | 10.06 | 9.86 | 21.49 | 9.77 | 20.02 | 0 | 16.8 | 0.1 | 15.92 | 0.1 | 8.5 | 4.59 | 4.88 | 6.15 | 4.79 | 21.39 | 9.77 | 24.61 | 9.77 | 24.51 | 9.86 | 24.51
122.912 | 0.1 | 0.68 | 0.1 | 14.36 | 0 | 14.36 | 4.88 | 20.61 | 4.79 | 14.84 | 4.59 | 6.64 | 9.67 | 9.77 | 9.86 | 9.86 | 21.39 | 9.67 | 20.61 | 0.1 | 17.29 | 0.1 | 16.6 | 0.1 | 10.35 | 4.59 | 4.79 | 5.76 | 4.88 | 21.29 | 9.77 | 24.41 | 9.77 | 24.51 | 9.77 | 24.51
122.944 | 0.1 | 0.78 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 21 | 4.79 | 15.33 | 4.59 | 8.11 | 9.67 | 10.06 | 9.96 | 9.86 | 21.39 | 9.77 | 21.19 | 0.1 | 18.07 | 0 | 17.29 | 0.1 | 11.72 | 4.59 | 4.88 | 5.47 | 4.88 | 21.39 | 9.86 | 24.32 | 9.86 | 24.41 | 9.86 | 24.41
122.976 | 0.1 | 0.59 | 0.1 | 14.06 | 0.1 | 14.36 | 4.88 | 17.97 | 4.79 | 15.92 | 4.59 | 9.08 | 9.67 | 11.04 | 9.86 | 9.86 | 21.49 | 9.77 | 21.78 | 0.1 | 18.75 | 0.1 | 17.87 | 0.1 | 12.6 | 4.59 | 4.79 | 5.27 | 4.88 | 21.19 | 9.86 | 24.41 | 9.77 | 24.32 | 9.86 | 24.32
123.008 | 0 | 0.59 | 0.1 | 3.81 | 0.1 | 12.6 | 4.88 | 18.46 | 4.79 | 16.31 | 4.59 | 9.77 | 9.67 | 11.91 | 9.86 | 9.86 | 21.39 | 9.77 | 22.27 | 0.1 | 19.24 | 0 | 18.56 | 0.1 | 13.48 | 4.59 | 4.88 | 5.08 | 4.88 | 21.29 | 9.86 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22
123.04 | 0.1 | 0.49 | 0.1 | 1.95 | 0.1 | 3.03 | 4.88 | 18.26 | 4.79 | 16.9 | 4.59 | 10.45 | 9.67 | 13.28 | 9.77 | 9.86 | 21.39 | 9.77 | 17.77 | 0.1 | 20.02 | 0 | 19.24 | 0.1 | 14.26 | 4.49 | 4.79 | 5.18 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
123.072 | 0.1 | 0.39 | 0.1 | 1.86 | 0.1 | 1.86 | 4.88 | 18.16 | 4.79 | 17.29 | 4.69 | 11.04 | 9.57 | 14.06 | 9.77 | 9.86 | 21.39 | 9.77 | 16.6 | 0.1 | 20.51 | 0.1 | 19.73 | 0.1 | 15.04 | 4.49 | 4.88 | 4.98 | 4.88 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22
123.104 | 0.1 | 0.39 | 0.1 | 1.76 | 0.1 | 1.76 | 4.88 | 18.26 | 4.88 | 17.77 | 4.69 | 11.72 | 9.57 | 14.65 | 9.77 | 9.96 | 21.39 | 9.77 | 17.29 | 0.1 | 21.19 | 0.1 | 20.51 | 0.1 | 15.82 | 4.49 | 4.88 | 4.98 | 4.88 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
123.136 | 0.78 | 0.29 | 0.1 | 1.46 | 0.1 | 1.56 | 4.88 | 18.16 | 4.88 | 18.16 | 4.69 | 12.21 | 9.67 | 15.23 | 9.67 | 9.86 | 10.64 | 9.77 | 17.48 | 0.1 | 21.78 | 0 | 20.9 | 0.1 | 16.41 | 4.49 | 4.88 | 4.98 | 4.88 | 11.72 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
123.168 | 1.56 | 0.29 | 0.1 | 1.37 | 0.1 | 1.46 | 4.88 | 18.16 | 4.88 | 18.65 | 4.79 | 12.79 | 9.57 | 15.72 | 9.67 | 9.86 | 6.64 | 9.77 | 17.58 | 0.1 | 22.27 | 0 | 21.68 | 0.1 | 17.09 | 4.59 | 4.79 | 4.88 | 4.88 | 3.03 | 9.86 | 17.87 | 9.77 | 24.22 | 9.86 | 24.12
123.2 | 2.83 | 0.2 | 0.1 | 1.07 | 0.1 | 1.27 | 4.88 | 18.16 | 4.79 | 19.14 | 4.69 | 13.38 | 9.67 | 16.21 | 9.67 | 9.86 | 9.67 | 9.77 | 17.58 | 0.1 | 19.73 | 0 | 22.27 | 0 | 17.87 | 4.59 | 4.88 | 4.79 | 4.79 | 3.22 | 9.77 | 6.74 | 9.77 | 14.45 | 9.86 | 24.12
123.232 | 3.91 | 0.1 | 0.1 | 1.07 | 0.1 | 1.17 | 4.88 | 18.07 | 4.88 | 19.53 | 4.79 | 13.77 | 9.67 | 16.7 | 9.67 | 9.86 | 11.52 | 9.77 | 17.58 | 0 | 17.29 | 0.1 | 22.75 | 0.1 | 18.46 | 4.59 | 5.86 | 4.69 | 4.79 | 3.52 | 9.77 | 7.03 | 9.77 | 6.54 | 9.86 | 13.48
123.264 | 4.59 | 0.1 | 0.1 | 0.98 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 20.02 | 4.79 | 14.36 | 9.67 | 17.29 | 9.67 | 9.86 | 10.94 | 9.86 | 17.58 | 0.1 | 17.77 | 0.1 | 19.53 | 0.1 | 19.14 | 4.59 | 7.13 | 4.69 | 4.88 | 4.59 | 9.86 | 10.84 | 9.77 | 7.23 | 9.86 | 6.15
123.296 | 5.18 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.31 | 4.79 | 14.84 | 9.67 | 17.68 | 9.67 | 9.86 | 10.16 | 9.77 | 17.58 | 0.1 | 17.77 | 0 | 18.75 | 0.1 | 19.73 | 4.59 | 8.98 | 4.69 | 4.88 | 6.05 | 9.86 | 12.01 | 9.77 | 11.23 | 9.86 | 7.62
123.328 | 5.76 | 0.1 | 0.1 | 0.78 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 20.8 | 4.79 | 15.43 | 9.67 | 18.16 | 9.67 | 9.96 | 9.77 | 9.77 | 17.58 | 0.1 | 17.68 | 0 | 18.26 | 0.1 | 20.31 | 4.69 | 10.64 | 4.69 | 4.79 | 6.15 | 9.77 | 10.64 | 9.86 | 11.91 | 9.86 | 11.52
123.36 | 6.35 | 0 | 0.1 | 0.68 | 0.1 | 0.78 | 4.88 | 18.07 | 4.88 | 17.38 | 4.79 | 15.92 | 9.67 | 18.56 | 9.67 | 9.86 | 10.16 | 9.77 | 17.58 | 0 | 17.58 | 0 | 18.16 | 0.1 | 20.9 | 4.69 | 12.01 | 4.59 | 4.79 | 5.66 | 9.77 | 9.96 | 9.77 | 10.55 | 9.86 | 11.72
123.392 | 6.93 | 0 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 16.31 | 4.88 | 17.58 | 4.88 | 16.41 | 9.67 | 18.95 | 9.57 | 9.86 | 10.25 | 9.77 | 17.58 | 0.1 | 17.68 | 0.1 | 18.07 | 0.1 | 21.49 | 4.69 | 12.89 | 4.59 | 4.79 | 5.27 | 9.77 | 10.16 | 9.77 | 9.77 | 9.86 | 10.45
123.424 | 7.52 | 0 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 4.2 | 4.88 | 17.87 | 4.88 | 16.9 | 9.67 | 19.43 | 9.67 | 9.96 | 10.06 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.97 | 0.1 | 22.07 | 4.69 | 13.77 | 4.59 | 4.88 | 5.18 | 9.86 | 10.35 | 9.86 | 10.16 | 9.86 | 9.96
123.456 | 8.2 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.03 | 4.88 | 18.07 | 4.88 | 17.29 | 9.67 | 19.92 | 9.67 | 9.96 | 9.86 | 9.86 | 17.48 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.66 | 4.79 | 14.55 | 4.59 | 4.79 | 5.18 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86 | 10.25
123.488 | 8.79 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.52 | 4.88 | 17.97 | 4.88 | 17.68 | 9.67 | 20.31 | 9.67 | 9.96 | 9.86 | 9.86 | 5.08 | 0.1 | 17.68 | 0 | 17.87 | 0.1 | 23.15 | 4.69 | 15.23 | 4.59 | 4.88 | 5.18 | 9.77 | 10.35 | 9.77 | 10.45 | 9.86 | 10.45
123.52 | 9.47 | 0 | 0.1 | 0.29 | 0 | 0.29 | 4.88 | 4.59 | 4.88 | 17.97 | 4.79 | 18.26 | 9.67 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 1.86 | 0 | 17.68 | 0 | 17.77 | 0.1 | 23.73 | 4.69 | 16.02 | 4.49 | 5.86 | 4.98 | 9.77 | 10.16 | 9.77 | 10.16 | 9.86 | 10.35
123.552 | 9.96 | 0 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 18.65 | 9.67 | 21.09 | 9.67 | 9.96 | 9.77 | 9.77 | 2.25 | 0 | 17.68 | 0.1 | 17.77 | 0.1 | 24.32 | 4.69 | 16.6 | 4.49 | 7.23 | 4.88 | 9.77 | 10.06 | 9.77 | 10.16 | 9.86 | 10.25
123.584 | 10.55 | 0.1 | 0.78 | 0.29 | 0.1 | 0.29 | 4.88 | 5.76 | 4.88 | 18.07 | 4.88 | 19.24 | 9.77 | 21.58 | 9.67 | 9.86 | 9.77 | 9.86 | 2.15 | 0.1 | 17.58 | 0.1 | 17.77 | 0.1 | 24.71 | 4.79 | 17.38 | 4.49 | 9.28 | 4.98 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86 | 10.06
123.616 | 11.13 | 0.1 | 1.56 | 0.2 | 0.1 | 0.2 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 21.97 | 9.67 | 10.74 | 9.67 | 9.86 | 1.76 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.2 | 4.79 | 18.07 | 4.59 | 10.84 | 4.88 | 9.86 | 9.96 | 9.86 | 10.16 | 9.86 | 10.16
123.648 | 11.72 | 0.1 | 2.73 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 20.02 | 9.77 | 22.36 | 9.67 | 11.13 | 9.77 | 9.86 | 1.56 | 0.1 | 7.13 | 0 | 17.68 | 0.1 | 25.78 | 4.79 | 18.56 | 4.59 | 12.11 | 4.88 | 9.77 | 9.96 | 9.86 | 10.06 | 9.86 | 10.16
123.68 | 12.31 | 0 | 3.91 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.79 | 18.07 | 4.88 | 20.41 | 9.77 | 22.75 | 9.67 | 12.21 | 9.67 | 9.86 | 1.46 | 0 | 1.66 | 0 | 10.64 | 0.1 | 26.27 | 4.79 | 19.34 | 4.59 | 12.99 | 4.79 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.96
123.712 | 12.79 | 0 | 4.59 | 0.1 | 0.1 | 0 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.8 | 9.77 | 23.15 | 9.67 | 13.48 | 9.67 | 9.86 | 1.46 | 0 | 2.25 | 0.1 | 2.83 | 0.1 | 22.85 | 4.79 | 19.92 | 4.59 | 13.97 | 4.69 | 9.77 | 9.96 | 9.77 | 9.86 | 9.86 | 9.86
123.744 | 13.38 | 0.1 | 5.27 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 17.77 | 9.77 | 23.54 | 9.77 | 14.26 | 9.67 | 9.86 | 1.27 | 0.1 | 2.05 | 0.1 | 2.44 | 0.1 | 22.17 | 4.79 | 20.51 | 4.69 | 14.55 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.86
123.776 | 13.97 | 0.1 | 5.76 | 0.1 | 1.66 | 0 | 4.98 | 4.88 | 4.88 | 17.97 | 4.88 | 17.58 | 9.77 | 23.93 | 9.77 | 14.84 | 9.67 | 9.86 | 1.17 | 0.1 | 1.66 | 0.1 | 1.95 | 0.1 | 21.68 | 4.88 | 21 | 4.69 | 15.33 | 4.69 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
123.808 | 14.36 | 0.1 | 6.35 | 0.1 | 2.93 | 0 | 5.57 | 4.79 | 4.88 | 6.84 | 4.88 | 17.87 | 9.77 | 24.32 | 9.77 | 15.43 | 9.67 | 9.86 | 1.07 | 0.1 | 1.66 | 0 | 1.86 | 0.1 | 21.58 | 4.79 | 21.58 | 4.69 | 16.02 | 4.59 | 9.86 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86
123.84 | 14.94 | 0 | 6.93 | 0 | 4 | 0 | 6.64 | 4.79 | 4.88 | 2.64 | 4.88 | 18.07 | 9.77 | 21.19 | 9.77 | 15.92 | 9.67 | 9.86 | 0.88 | 0.1 | 1.37 | 0 | 1.86 | 0.1 | 21.68 | 4.79 | 22.27 | 4.69 | 16.7 | 4.59 | 10.74 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77
123.872 | 15.43 | 0 | 7.52 | 0.1 | 4.79 | 0 | 7.91 | 4.69 | 4.88 | 3.32 | 4.88 | 17.97 | 9.77 | 21.29 | 9.77 | 16.41 | 9.57 | 9.86 | 0.88 | 0 | 1.37 | 0.1 | 1.56 | 0.1 | 21.49 | 4.79 | 22.75 | 4.69 | 17.38 | 4.49 | 12.11 | 9.67 | 9.86 | 9.67 | 10.94 | 9.77
123.904 | 15.92 | 0.1 | 8.11 | 0.1 | 5.27 | 0 | 8.98 | 4.69 | 4.88 | 4.49 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 16.9 | 9.57 | 9.86 | 0.78 | 0.1 | 1.37 | 0.1 | 1.37 | 0.1 | 21.39 | 4.88 | 23.34 | 4.79 | 18.07 | 4.59 | 13.97 | 9.77 | 11.04 | 9.77 | 11.91 | 9.77
123.936 | 16.41 | 0 | 8.69 | 0.1 | 5.86 | 0.1 | 9.67 | 4.59 | 4.88 | 5.57 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 17.38 | 9.67 | 9.86 | 0.78 | 0.1 | 1.07 | 0.1 | 1.37 | 0.1 | 21.39 | 4.88 | 23.83 | 4.79 | 18.65 | 4.59 | 15.53 | 9.67 | 11.91 | 9.67 | 13.77 | 9.77
123.968 | 16.8 | 0.1 | 9.38 | 0.1 | 6.45 | 0 | 10.45 | 4.59 | 4.88 | 5.86 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 17.87 | 9.57 | 9.86 | 0.68 | 0.1 | 1.17 | 0 | 1.27 | 0.1 | 21.29 | 4.88 | 24.32 | 4.79 | 19.34 | 4.49 | 16.7 | 9.67 | 13.67 | 9.67 | 15.33 | 9.67
124 | 17.29 | 0 | 9.96 | 0 | 7.03 | 0 | 11.04 | 4.49 | 4.79 | 5.47 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 18.26 | 9.57 | 9.86 | 0.59 | 0 | 0.88 | 0 | 1.07 | 0.1 | 21.29 | 4.79 | 24.9 | 4.79 | 20.02 | 4.49 | 17.58 | 9.67 | 15.43 | 9.67 | 16.6 | 9.67
124.032 | 15.33 | 0 | 10.45 | 0.1 | 7.62 | 0 | 11.72 | 4.59 | 4.88 | 5.08 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 18.65 | 9.57 | 9.86 | 0.59 | 0 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 25.49 | 4.79 | 20.61 | 4.49 | 18.26 | 9.67 | 16.6 | 9.67 | 17.48 | 9.67
124.064 | 15.04 | 0.1 | 11.13 | 0.1 | 8.2 | 0 | 12.21 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 19.24 | 9.67 | 9.86 | 0.49 | 0.1 | 0.78 | 0.1 | 0.98 | 0.1 | 21.29 | 4.79 | 25.88 | 4.79 | 21.09 | 4.59 | 18.95 | 9.67 | 17.58 | 9.67 | 18.26 | 9.67
124.096 | 14.75 | 0.1 | 11.72 | 0.1 | 8.79 | 0 | 12.79 | 4.49 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 19.53 | 9.67 | 9.86 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.59 | 4.88 | 21.68 | 4.59 | 19.73 | 9.67 | 18.26 | 9.67 | 18.85 | 9.67
124.128 | 14.65 | 0.1 | 12.31 | 0.1 | 9.47 | 0.1 | 13.38 | 4.59 | 4.88 | 4.88 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 20.02 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 14.55 | 4.79 | 21.88 | 4.88 | 22.27 | 4.59 | 20.31 | 9.67 | 18.95 | 9.67 | 19.43 | 9.67
124.16 | 14.65 | 0 | 12.79 | 0.1 | 9.96 | 0 | 13.87 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 20.51 | 9.67 | 10.64 | 0.29 | 0 | 0.59 | 0 | 0.68 | 0.1 | 3.13 | 4.79 | 21.49 | 4.79 | 22.85 | 4.59 | 20.9 | 9.67 | 19.53 | 9.67 | 20.12 | 9.57
124.192 | 14.55 | 0 | 13.38 | 0.1 | 10.55 | 0 | 14.36 | 4.59 | 4.79 | 4.79 | 4.88 | 6.54 | 9.77 | 21.39 | 9.86 | 20.9 | 9.67 | 11.62 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.13 | 4.79 | 21.58 | 4.79 | 23.24 | 4.59 | 21.49 | 9.67 | 20.12 | 9.67 | 20.8 | 9.67
124.224 | 14.45 | 0.1 | 13.87 | 0.1 | 11.13 | 0 | 14.94 | 4.59 | 4.88 | 4.69 | 4.88 | 2.54 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 12.99 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.42 | 4.79 | 21.49 | 4.88 | 23.93 | 4.69 | 22.17 | 9.67 | 20.8 | 9.67 | 21.39 | 9.67
124.256 | 14.45 | 0.1 | 14.36 | 0.1 | 11.72 | 0 | 15.43 | 4.59 | 4.88 | 4.69 | 4.88 | 3.42 | 9.77 | 21 | 9.86 | 21.68 | 9.67 | 13.87 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.3 | 4.79 | 21.39 | 4.88 | 24.41 | 4.69 | 22.66 | 9.67 | 21.39 | 9.67 | 21.88 | 9.67
124.288 | 14.45 | 0 | 14.94 | 0.1 | 12.31 | 0 | 15.92 | 4.59 | 4.88 | 4.59 | 4.88 | 4.59 | 9.86 | 8.89 | 9.86 | 22.07 | 9.67 | 14.55 | 0.2 | 0.1 | 0.29 | 0 | 0.49 | 0.1 | 5.76 | 4.88 | 21.29 | 4.88 | 24.9 | 4.69 | 23.34 | 9.67 | 21.88 | 9.67 | 22.56 | 9.67
124.32 | 14.45 | 0 | 15.43 | 0 | 12.89 | 0 | 16.5 | 4.59 | 4.79 | 4.59 | 4.88 | 5.66 | 9.86 | 7.23 | 9.86 | 22.36 | 9.67 | 15.14 | 0.2 | 0 | 0.29 | 0.1 | 0.39 | 0 | 6.15 | 4.79 | 21.29 | 4.88 | 25.49 | 4.69 | 23.83 | 9.67 | 22.56 | 9.67 | 23.15 | 9.67
124.352 | 14.45 | 0 | 15.92 | 0.1 | 13.38 | 0 | 16.9 | 4.59 | 4.79 | 4.59 | 4.88 | 5.96 | 9.77 | 9.96 | 9.86 | 22.85 | 9.67 | 15.63 | 0.2 | 0 | 0.2 | 1.07 | 0.29 | 0.1 | 5.76 | 4.79 | 21.29 | 4.79 | 25.88 | 4.69 | 24.22 | 9.67 | 23.05 | 9.67 | 23.63 | 9.67
124.384 | 14.36 | 0.1 | 16.41 | 0.1 | 13.97 | 0 | 17.38 | 4.69 | 4.88 | 4.59 | 4.88 | 5.47 | 9.86 | 11.62 | 9.86 | 23.24 | 9.67 | 16.21 | 0.2 | 0.1 | 0.2 | 2.05 | 0.29 | 0.1 | 5.37 | 4.88 | 21.29 | 4.88 | 24.22 | 4.79 | 24.81 | 9.67 | 23.73 | 9.67 | 24.22 | 9.67
124.416 | 14.36 | 0.1 | 16.9 | 0.1 | 14.45 | 0 | 17.77 | 4.69 | 4.98 | 4.59 | 4.88 | 5.18 | 9.86 | 10.64 | 9.86 | 23.63 | 9.67 | 16.8 | 0.1 | 1.17 | 0.2 | 3.91 | 0.29 | 0.1 | 5.18 | 4.79 | 21.29 | 4.88 | 20.9 | 4.79 | 25.29 | 9.77 | 24.22 | 9.67 | 24.61 | 9.67
124.448 | 8.01 | 0 | 17.38 | 0.1 | 15.04 | 0 | 18.26 | 4.69 | 5.66 | 4.49 | 4.88 | 4.88 | 9.86 | 10.06 | 9.86 | 24.02 | 9.77 | 17.19 | 0.1 | 2.15 | 0.2 | 5.57 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.88 | 21.39 | 4.79 | 25.78 | 9.77 | 24.61 | 9.67 | 25.2 | 9.67
124.48 | 2.34 | 0 | 15.33 | 0 | 15.43 | 0 | 18.75 | 4.69 | 6.84 | 4.39 | 4.88 | 4.98 | 9.77 | 9.86 | 9.86 | 24.32 | 9.67 | 17.68 | 0.1 | 4.1 | 0.1 | 6.64 | 0.2 | 0.2 | 5.27 | 4.79 | 21.19 | 4.88 | 21.29 | 4.79 | 26.37 | 9.67 | 25.2 | 9.67 | 25.59 | 9.67
124.512 | 2.05 | 0 | 15.23 | 0.1 | 16.02 | 0 | 19.24 | 4.69 | 8.11 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 18.07 | 0.1 | 5.57 | 0.1 | 7.42 | 0.1 | 1.37 | 5.08 | 4.79 | 21.19 | 4.88 | 21.19 | 4.79 | 26.76 | 9.67 | 25.68 | 9.67 | 26.08 | 9.67
124.544 | 1.76 | 0.1 | 14.94 | 0.1 | 16.5 | 0 | 19.63 | 4.79 | 9.08 | 4.49 | 4.88 | 4.88 | 9.86 | 10.25 | 9.86 | 21.68 | 9.77 | 18.56 | 0.1 | 6.74 | 0.1 | 8.3 | 0.1 | 2.73 | 5.08 | 4.79 | 8.69 | 4.88 | 21.29 | 4.79 | 27.25 | 9.77 | 26.17 | 9.77 | 26.56 | 9.77
124.576 | 1.46 | 0 | 14.75 | 0.1 | 16.99 | 0 | 20.02 | 4.79 | 9.77 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 18.95 | 0.1 | 7.52 | 0.1 | 8.98 | 0.1 | 4.59 | 4.98 | 4.88 | 2.44 | 4.88 | 21.29 | 4.79 | 27.74 | 9.77 | 26.66 | 9.77 | 27.05 | 9.77
124.608 | 1.56 | 0.1 | 14.75 | 0.1 | 16.99 | 0 | 20.41 | 4.79 | 10.45 | 4.49 | 4.88 | 4.79 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 19.43 | 0.1 | 8.3 | 0.1 | 9.86 | 0.1 | 6.05 | 4.88 | 4.79 | 3.32 | 4.88 | 21.29 | 4.79 | 28.32 | 9.77 | 27.25 | 9.77 | 27.64 | 9.77
124.64 | 1.17 | 0 | 14.55 | 0.1 | 14.55 | 0 | 20.9 | 4.79 | 11.13 | 4.49 | 4.88 | 4.69 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 19.92 | 0.1 | 9.08 | 0.1 | 10.55 | 0.1 | 7.13 | 4.79 | 4.79 | 3.91 | 4.88 | 21.19 | 4.79 | 28.71 | 9.77 | 27.64 | 9.77 | 28.13 | 9.77
124.672 | 1.27 | 0 | 14.55 | 0 | 14.84 | 0 | 19.63 | 4.79 | 11.72 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.49 | 9.77 | 20.22 | 0.1 | 9.86 | 0 | 11.43 | 0 | 7.81 | 4.69 | 4.79 | 4.98 | 4.79 | 21.19 | 4.79 | 29.01 | 9.77 | 28.03 | 9.77 | 28.52 | 9.77
124.704 | 1.07 | 0.1 | 14.55 | 0.1 | 14.75 | 0 | 17.87 | 4.88 | 12.31 | 4.59 | 4.88 | 4.59 | 9.77 | 9.86 | 9.86 | 21.58 | 9.77 | 20.7 | 0.1 | 10.55 | 0.1 | 12.21 | 0.1 | 8.59 | 4.79 | 4.79 | 6.35 | 4.88 | 21.19 | 4.79 | 28.42 | 9.86 | 28.52 | 9.77 | 29.01 | 9.77
124.736 | 0.88 | 0 | 14.45 | 0.1 | 14.65 | 0 | 18.26 | 4.88 | 12.79 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21 | 0.1 | 11.43 | 0.1 | 12.89 | 0.1 | 9.38 | 4.69 | 4.79 | 6.05 | 4.88 | 21.19 | 4.79 | 24.71 | 9.77 | 29.01 | 9.77 | 29.49 | 9.77
124.768 | 0.78 | 0 | 14.45 | 0.1 | 14.65 | 0 | 18.36 | 4.88 | 13.38 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.49 | 0.1 | 12.31 | 0.1 | 13.87 | 0.1 | 10.06 | 4.59 | 4.79 | 5.66 | 4.88 | 21.19 | 4.79 | 24.61 | 9.77 | 29.49 | 9.77 | 25.78 | 9.77
124.8 | 0.68 | 0 | 14.45 | 0 | 14.45 | 0 | 18.16 | 4.79 | 13.87 | 4.69 | 4.98 | 4.59 | 9.77 | 9.67 | 9.86 | 21.49 | 9.77 | 21.97 | 0.1 | 13.09 | 0.1 | 14.45 | 0.1 | 11.04 | 4.59 | 4.79 | 5.27 | 4.88 | 21.19 | 4.79 | 24.41 | 9.77 | 25.2 | 9.77 | 24.81 | 9.77
124.832 | 0.68 | 0 | 14.45 | 0 | 14.45 | 0 | 18.16 | 4.88 | 14.45 | 4.69 | 5.57 | 4.49 | 9.77 | 9.67 | 9.86 | 21.49 | 9.77 | 22.27 | 0.1 | 13.87 | 0.1 | 15.14 | 0 | 11.91 | 4.59 | 4.79 | 5.18 | 4.79 | 8.01 | 4.79 | 24.41 | 9.77 | 25 | 9.77 | 24.61 | 9.77
124.864 | 0.59 | 0.1 | 14.45 | 0.1 | 14.55 | 0 | 18.26 | 4.88 | 14.94 | 4.69 | 6.64 | 4.59 | 9.77 | 9.67 | 9.86 | 21.49 | 9.86 | 22.66 | 0.1 | 14.55 | 0.1 | 15.82 | 0.1 | 12.6 | 4.59 | 4.79 | 5.08 | 4.88 | 2.25 | 4.88 | 24.41 | 9.77 | 24.71 | 9.86 | 24.61 | 9.77
124.896 | 0.59 | 0 | 8.89 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.43 | 4.69 | 7.91 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.86 | 23.05 | 0.1 | 15.14 | 0.1 | 16.6 | 0.1 | 13.28 | 4.59 | 5.08 | 5.08 | 4.88 | 3.32 | 4.79 | 24.41 | 9.77 | 24.51 | 9.86 | 24.51 | 9.77
124.928 | 0.49 | 0.1 | 2.54 | 0.1 | 14.45 | 0 | 18.16 | 4.88 | 15.92 | 4.79 | 8.89 | 4.49 | 9.77 | 9.67 | 9.86 | 19.34 | 9.86 | 23.44 | 0.1 | 15.92 | 0.1 | 17.29 | 0.1 | 14.06 | 4.49 | 6.25 | 4.98 | 4.88 | 4 | 4.88 | 24.32 | 9.86 | 24.51 | 9.86 | 24.41 | 9.86
124.96 | 0.39 | 0 | 1.95 | 0 | 14.45 | 0 | 18.16 | 4.88 | 16.41 | 4.69 | 9.77 | 4.49 | 9.77 | 9.57 | 9.86 | 8.01 | 9.77 | 23.83 | 0.1 | 16.6 | 0.1 | 17.87 | 0.1 | 14.75 | 4.49 | 7.81 | 4.98 | 4.88 | 5.27 | 4.79 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.77
124.992 | 0.29 | 0 | 1.76 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.9 | 4.79 | 10.45 | 4.59 | 10.45 | 9.57 | 9.86 | 7.23 | 9.86 | 24.22 | 0.1 | 17.19 | 0.1 | 18.56 | 0.1 | 15.53 | 4.49 | 9.86 | 4.98 | 4.79 | 6.25 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77
125.024 | 0.29 | 0.1 | 1.46 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 17.29 | 4.79 | 11.04 | 4.59 | 11.04 | 9.57 | 9.86 | 10.45 | 9.86 | 22.36 | 0.1 | 17.87 | 0.1 | 19.14 | 0.1 | 16.21 | 4.59 | 11.33 | 4.88 | 4.88 | 5.96 | 4.88 | 24.32 | 9.86 | 24.32 | 9.86 | 24.22 | 9.86
125.056 | 0.29 | 0 | 1.56 | 0.1 | 7.52 | 0 | 18.07 | 4.88 | 17.87 | 4.79 | 11.62 | 4.59 | 12.11 | 9.57 | 9.86 | 11.62 | 9.86 | 21 | 0.1 | 18.65 | 0.1 | 19.82 | 0.1 | 16.9 | 4.59 | 12.4 | 4.79 | 4.88 | 5.47 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
125.088 | 0.2 | 0.1 | 1.27 | 0.1 | 2.05 | 0 | 18.16 | 4.88 | 18.26 | 4.79 | 12.21 | 4.59 | 13.38 | 9.57 | 9.86 | 10.64 | 9.86 | 21.29 | 0.1 | 19.14 | 0.1 | 20.41 | 0.1 | 17.48 | 4.59 | 13.28 | 4.79 | 4.88 | 5.27 | 4.88 | 24.32 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86
125.12 | 0.2 | 0 | 1.17 | 0 | 1.95 | 0 | 7.32 | 4.79 | 18.75 | 4.79 | 12.79 | 4.59 | 14.16 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 19.82 | 0.1 | 21 | 0.1 | 18.16 | 4.59 | 14.06 | 4.69 | 4.79 | 5.08 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
125.152 | 0.1 | 0 | 1.07 | 0.1 | 1.76 | 0 | 2.73 | 4.88 | 19.14 | 4.79 | 13.28 | 4.59 | 14.75 | 9.57 | 9.86 | 9.96 | 9.86 | 21.39 | 0.1 | 20.51 | 0.1 | 21.68 | 0.1 | 18.85 | 4.59 | 14.75 | 4.69 | 4.79 | 5.18 | 4.79 | 12.5 | 9.77 | 24.12 | 9.77 | 24.12 | 9.86
125.184 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0 | 3.22 | 4.88 | 19.53 | 4.88 | 13.77 | 4.69 | 15.33 | 9.67 | 9.86 | 10.06 | 9.86 | 21.49 | 0.1 | 21 | 0.1 | 22.17 | 0.1 | 19.53 | 4.69 | 15.53 | 4.59 | 4.88 | 5.18 | 4.88 | 5.86 | 9.86 | 24.22 | 9.86 | 13.57 | 9.86
125.216 | 0 | 0.1 | 0.78 | 0.1 | 1.46 | 0 | 4.2 | 4.88 | 20.02 | 4.79 | 14.36 | 4.69 | 15.82 | 9.67 | 9.86 | 10.16 | 9.86 | 21.49 | 0.1 | 21.58 | 0.1 | 22.56 | 0.1 | 20.12 | 4.69 | 16.31 | 4.59 | 4.88 | 5.08 | 4.88 | 8.11 | 9.86 | 13.87 | 9.86 | 6.15 | 9.86
125.248 | 0 | 0.1 | 0.78 | 0.1 | 1.37 | 0 | 5.37 | 4.88 | 20.31 | 4.88 | 14.94 | 4.69 | 16.41 | 9.67 | 10.64 | 10.16 | 9.86 | 21.49 | 0.1 | 22.07 | 0.1 | 18.56 | 0.1 | 20.7 | 4.69 | 16.99 | 4.59 | 4.88 | 4.98 | 4.88 | 12.01 | 9.86 | 6.25 | 9.86 | 7.72 | 9.86
125.28 | 0.1 | 0.78 | 0.59 | 0 | 1.07 | 0 | 5.96 | 4.88 | 20.9 | 4.79 | 15.43 | 4.69 | 16.9 | 9.67 | 11.23 | 9.96 | 9.86 | 21.39 | 0.1 | 21.49 | 0 | 18.46 | 0.1 | 21.29 | 4.69 | 17.58 | 4.59 | 4.88 | 4.88 | 4.79 | 11.52 | 9.77 | 8.01 | 9.86 | 11.62 | 9.86
125.312 | 0.1 | 1.66 | 0.68 | 0.1 | 1.07 | 0 | 5.57 | 4.88 | 17.38 | 4.79 | 15.82 | 4.69 | 17.29 | 9.67 | 12.11 | 9.86 | 9.86 | 21.39 | 0.1 | 18.46 | 0 | 18.07 | 0.1 | 21.88 | 4.69 | 18.16 | 4.59 | 4.79 | 4.88 | 4.79 | 10.25 | 9.77 | 11.62 | 9.86 | 11.82 | 9.86
125.344 | 0.1 | 2.93 | 0.49 | 0.1 | 0.88 | 0 | 5.18 | 4.88 | 17.58 | 4.88 | 16.31 | 4.79 | 17.77 | 9.67 | 13.38 | 9.96 | 9.96 | 21.39 | 0.1 | 18.16 | 0.1 | 17.97 | 0.1 | 22.46 | 4.69 | 18.85 | 4.59 | 4.88 | 4.88 | 4.88 | 9.96 | 9.86 | 11.72 | 9.86 | 10.45 | 9.86
125.376 | 0.1 | 3.91 | 0.49 | 0.1 | 0.88 | 0 | 4.98 | 4.88 | 17.77 | 4.88 | 16.8 | 4.79 | 18.26 | 9.67 | 14.16 | 9.86 | 9.86 | 21.39 | 0.1 | 17.97 | 0.1 | 17.97 | 0.1 | 18.36 | 4.69 | 19.53 | 4.59 | 5.76 | 4.69 | 4.79 | 10.35 | 9.86 | 10.45 | 9.86 | 9.86 | 9.86
125.408 | 0.1 | 4.59 | 0.39 | 0.1 | 0.68 | 0 | 4.88 | 4.88 | 18.07 | 4.88 | 17.29 | 4.79 | 18.65 | 9.67 | 14.84 | 9.77 | 9.86 | 21.39 | 0.1 | 17.87 | 0.1 | 17.77 | 0.1 | 16.99 | 4.79 | 20.12 | 4.59 | 6.84 | 4.69 | 4.79 | 10.45 | 9.86 | 9.96 | 9.86 | 10.16 | 9.86
125.44 | 0.1 | 5.27 | 0.39 | 0 | 0.68 | 0 | 4.88 | 4.88 | 17.97 | 4.79 | 17.68 | 4.79 | 19.24 | 9.67 | 15.43 | 9.77 | 9.86 | 21.39 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 17.58 | 4.79 | 20.7 | 4.59 | 8.69 | 4.69 | 4.79 | 10.35 | 9.77 | 10.16 | 9.86 | 10.55 | 9.86
125.472 | 0.1 | 5.76 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.07 | 4.79 | 18.07 | 4.79 | 19.53 | 9.67 | 15.82 | 9.77 | 9.86 | 11.62 | 0.1 | 17.68 | 0 | 17.68 | 0.1 | 17.77 | 4.79 | 21.39 | 4.59 | 10.35 | 4.59 | 4.79 | 10.16 | 9.77 | 10.35 | 9.86 | 10.35 | 9.86
125.504 | 0.1 | 6.35 | 0.29 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.16 | 4.88 | 18.65 | 4.79 | 20.02 | 9.67 | 16.41 | 9.67 | 9.86 | 6.54 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 21.78 | 4.59 | 11.72 | 4.59 | 4.79 | 10.25 | 9.86 | 10.35 | 9.86 | 10.16 | 9.86
125.536 | 0.1 | 6.93 | 0.2 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.07 | 4.88 | 19.14 | 4.79 | 20.41 | 9.67 | 16.9 | 9.67 | 9.86 | 8.89 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.68 | 4.79 | 22.46 | 4.59 | 12.7 | 4.59 | 4.79 | 10.06 | 9.86 | 10.16 | 9.86 | 9.96 | 9.86
125.568 | 0 | 7.62 | 0.2 | 0.1 | 0.49 | 0 | 4.79 | 4.88 | 18.07 | 4.88 | 19.53 | 4.88 | 20.9 | 9.67 | 17.29 | 9.67 | 9.86 | 11.43 | 0.1 | 17.58 | 0 | 17.68 | 0.1 | 17.77 | 4.79 | 23.05 | 4.69 | 13.57 | 4.59 | 4.88 | 10.16 | 9.77 | 10.06 | 9.86 | 10.16 | 9.86
125.6 | 0.1 | 8.2 | 0.1 | 0 | 0.29 | 0 | 4.69 | 4.88 | 18.07 | 4.79 | 20.02 | 4.79 | 21.29 | 9.67 | 17.77 | 9.67 | 9.86 | 11.13 | 0.1 | 17.58 | 0.1 | 17.58 | 0.1 | 17.77 | 4.79 | 23.63 | 4.69 | 14.36 | 4.59 | 4.79 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16 | 9.77
125.632 | 0.1 | 8.79 | 0.1 | 0.1 | 0.29 | 0 | 4.69 | 4.88 | 18.07 | 4.88 | 20.41 | 4.79 | 21.68 | 9.67 | 18.26 | 9.67 | 9.86 | 10.25 | 0.1 | 17.58 | 0 | 17.58 | 0.1 | 17.77 | 4.79 | 24.12 | 4.69 | 15.04 | 4.59 | 4.79 | 9.86 | 9.77 | 10.16 | 9.86 | 9.86 | 9.86
125.664 | 0.1 | 9.47 | 0.1 | 0.2 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.8 | 4.88 | 22.17 | 9.77 | 18.65 | 9.67 | 9.96 | 9.77 | 0.1 | 17.58 | 0.1 | 8.89 | 0.1 | 17.77 | 4.88 | 24.71 | 4.69 | 15.82 | 4.59 | 4.79 | 9.86 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86
125.696 | 0.1 | 9.96 | 0.1 | 1.17 | 0.2 | 0.29 | 4.59 | 4.88 | 18.07 | 4.88 | 19.53 | 4.88 | 22.46 | 9.77 | 19.14 | 9.57 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 2.25 | 0.1 | 17.77 | 4.88 | 25.2 | 4.69 | 16.5 | 4.59 | 4.88 | 9.96 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
125.728 | 0 | 10.55 | 0.1 | 2.15 | 0.1 | 1.17 | 4.59 | 4.88 | 18.07 | 4.88 | 17.77 | 4.88 | 22.85 | 9.77 | 19.53 | 9.67 | 9.86 | 10.16 | 0.1 | 8.5 | 0.1 | 2.25 | 0.1 | 17.77 | 4.88 | 25.59 | 4.79 | 17.09 | 4.49 | 4.79 | 9.86 | 9.77 | 9.96 | 9.86 | 9.86 | 9.86
125.76 | 0.1 | 11.13 | 0.1 | 3.61 | 0.1 | 2.25 | 4.59 | 4.88 | 5.76 | 4.79 | 17.97 | 4.88 | 23.24 | 9.77 | 20.02 | 9.57 | 9.86 | 10.06 | 0.1 | 1.86 | 0 | 1.95 | 0.1 | 17.68 | 4.79 | 26.08 | 4.69 | 17.87 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.77
125.792 | 0.1 | 11.72 | 0.1 | 4.39 | 0.1 | 3.52 | 4.59 | 4.88 | 2.64 | 4.88 | 18.26 | 4.88 | 23.63 | 9.77 | 20.31 | 9.57 | 9.96 | 9.86 | 0.1 | 2.34 | 0 | 1.76 | 0 | 17.68 | 4.79 | 21.58 | 4.79 | 18.46 | 4.59 | 4.79 | 9.77 | 9.77 | 9.77 | 9.86 | 9.77 | 9.77
125.824 | 0.1 | 12.31 | 0.1 | 4.98 | 0.1 | 4.39 | 4.59 | 4.88 | 3.42 | 4.88 | 18.16 | 4.88 | 24.02 | 9.77 | 20.9 | 9.67 | 10.64 | 9.86 | 0.1 | 1.95 | 0.1 | 1.66 | 0.1 | 6.45 | 4.88 | 20.9 | 4.79 | 19.04 | 4.59 | 5.66 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
125.856 | 0.1 | 12.89 | 0.1 | 5.47 | 0.1 | 4.98 | 4.59 | 4.88 | 4.59 | 4.88 | 18.16 | 4.88 | 24.32 | 9.77 | 21.29 | 9.67 | 11.62 | 9.86 | 0.1 | 1.66 | 0.1 | 1.56 | 0.1 | 1.76 | 4.88 | 21.19 | 4.79 | 19.73 | 4.59 | 6.64 | 9.77 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86
125.888 | 0.1 | 13.38 | 0.1 | 6.05 | 0.1 | 5.47 | 4.59 | 4.88 | 5.47 | 4.88 | 18.16 | 4.88 | 21.49 | 9.77 | 21.68 | 9.67 | 12.99 | 9.77 | 0.1 | 1.76 | 0.1 | 1.37 | 0.1 | 2.25 | 4.88 | 21.19 | 4.79 | 20.31 | 4.69 | 8.3 | 9.67 | 11.52 | 9.77 | 10.74 | 9.67 | 9.86
125.92 | 0.1 | 13.87 | 0.1 | 6.64 | 0 | 6.05 | 4.59 | 4.88 | 5.76 | 4.79 | 18.07 | 4.88 | 22.17 | 9.77 | 21.97 | 9.67 | 13.97 | 9.67 | 0.1 | 1.37 | 0 | 1.37 | 0.1 | 2.05 | 4.88 | 21.29 | 4.79 | 21 | 4.69 | 10.45 | 9.67 | 13.18 | 9.67 | 11.91 | 9.67 | 10.94
125.952 | 0.1 | 14.45 | 0.1 | 7.23 | 0.1 | 6.54 | 4.59 | 4.88 | 5.47 | 4.79 | 18.16 | 4.79 | 21.58 | 9.77 | 22.46 | 9.67 | 14.55 | 9.67 | 0.1 | 1.46 | 0.1 | 1.17 | 0 | 1.66 | 4.88 | 21.19 | 4.79 | 21.49 | 4.69 | 11.62 | 9.67 | 15.04 | 9.67 | 13.48 | 9.67 | 11.91
125.984 | 0.1 | 14.94 | 0.1 | 7.81 | 0.1 | 7.32 | 4.59 | 4.88 | 5.18 | 4.88 | 18.16 | 4.88 | 21.58 | 9.77 | 22.75 | 9.67 | 15.23 | 9.67 | 0.1 | 1.27 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.29 | 4.79 | 22.17 | 4.69 | 12.6 | 9.67 | 16.31 | 9.67 | 15.23 | 9.67 | 13.48
126.016 | 0.1 | 15.43 | 0.1 | 8.5 | 0.1 | 7.91 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.58 | 9.77 | 23.24 | 9.67 | 15.72 | 9.67 | 0.1 | 1.17 | 0.1 | 1.07 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 22.75 | 4.69 | 13.57 | 9.67 | 17.19 | 9.67 | 16.5 | 9.67 | 15.23
126.048 | 0.1 | 15.92 | 0.1 | 8.98 | 0.1 | 8.5 | 4.69 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 23.54 | 9.67 | 16.21 | 9.67 | 0.1 | 1.07 | 0.1 | 0.88 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 23.24 | 4.69 | 14.26 | 9.57 | 17.97 | 9.67 | 17.38 | 9.67 | 16.5
126.08 | 0.1 | 16.41 | 0.1 | 9.67 | 0 | 9.08 | 4.69 | 5.18 | 4.98 | 4.79 | 18.07 | 4.79 | 21.49 | 9.77 | 24.02 | 9.67 | 16.7 | 9.67 | 0.1 | 0.98 | 0 | 0.78 | 0.1 | 1.27 | 4.88 | 21.29 | 4.79 | 24.02 | 4.69 | 15.04 | 9.57 | 18.65 | 9.67 | 18.07 | 9.67 | 17.38
126.112 | 0.1 | 16.9 | 0.1 | 10.25 | 0.1 | 9.67 | 4.69 | 5.96 | 4.79 | 4.88 | 18.07 | 4.79 | 21.49 | 9.77 | 24.32 | 9.67 | 17.29 | 9.57 | 0.1 | 0.88 | 0 | 0.78 | 0 | 1.17 | 4.88 | 21.19 | 4.79 | 24.41 | 4.79 | 15.72 | 9.67 | 19.34 | 9.57 | 18.75 | 9.67 | 18.16
126.144 | 0.1 | 16.9 | 0.1 | 10.84 | 0.1 | 10.25 | 4.69 | 6.84 | 4.88 | 4.88 | 7.72 | 4.88 | 21.39 | 9.77 | 21.29 | 9.67 | 17.68 | 9.57 | 0.2 | 0.78 | 0.1 | 0.68 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 24.9 | 4.79 | 16.41 | 9.67 | 19.92 | 9.67 | 19.43 | 9.67 | 18.75
126.176 | 0.1 | 14.94 | 0.1 | 11.43 | 0.1 | 10.84 | 4.79 | 8.3 | 4.79 | 4.88 | 2.64 | 4.88 | 21.49 | 9.77 | 21.78 | 9.67 | 18.07 | 9.57 | 1.46 | 0.78 | 0.1 | 0.59 | 0.1 | 0.98 | 4.88 | 21.19 | 4.79 | 25.29 | 4.79 | 17.09 | 9.67 | 20.61 | 9.67 | 20.12 | 9.67 | 19.43
126.208 | 0 | 14.84 | 0.1 | 12.01 | 0.1 | 11.43 | 4.79 | 9.28 | 4.69 | 4.88 | 3.32 | 4.88 | 21.39 | 9.77 | 21.49 | 9.67 | 18.56 | 9.57 | 2.64 | 0.68 | 0.1 | 0.49 | 0.1 | 0.88 | 4.88 | 21 | 4.88 | 25.68 | 4.79 | 17.68 | 9.67 | 21.19 | 9.67 | 20.7 | 9.67 | 20.12
126.24 | 0.1 | 14.65 | 0.1 | 12.5 | 0.1 | 12.01 | 4.79 | 9.96 | 4.59 | 4.79 | 4.39 | 4.88 | 21.39 | 9.77 | 21.58 | 9.67 | 18.95 | 9.57 | 4.59 | 0.59 | 0 | 0.49 | 0.1 | 0.78 | 4.88 | 4.69 | 4.79 | 26.37 | 4.79 | 18.46 | 9.67 | 21.97 | 9.67 | 21.39 | 9.67 | 20.8
126.272 | 0.1 | 14.55 | 0.1 | 13.09 | 0.1 | 12.5 | 4.79 | 10.55 | 4.59 | 4.79 | 5.47 | 4.79 | 21.39 | 9.77 | 21.58 | 9.77 | 19.34 | 9.57 | 5.96 | 0.49 | 0 | 0.49 | 0 | 0.68 | 4.88 | 2.73 | 4.88 | 23.15 | 4.79 | 19.04 | 9.67 | 22.27 | 9.57 | 21.88 | 9.67 | 21.29
126.304 | 0.1 | 14.55 | 0.1 | 13.57 | 0.1 | 13.09 | 4.79 | 11.23 | 4.59 | 4.88 | 5.96 | 4.88 | 17.87 | 9.77 | 21.49 | 9.77 | 19.92 | 9.57 | 6.93 | 0.49 | 0.1 | 0.29 | 0.1 | 0.68 | 4.88 | 3.32 | 4.88 | 22.27 | 4.79 | 19.63 | 9.67 | 22.95 | 9.67 | 22.46 | 9.67 | 21.97
126.336 | 0.1 | 14.45 | 0.1 | 14.16 | 0.1 | 13.67 | 4.79 | 11.82 | 4.59 | 4.88 | 5.66 | 4.88 | 8.2 | 9.77 | 21.49 | 9.77 | 20.31 | 9.57 | 7.72 | 0.49 | 1.07 | 0.29 | 0.1 | 0.59 | 4.88 | 4.1 | 4.88 | 21.88 | 4.79 | 20.31 | 9.67 | 23.44 | 9.67 | 22.95 | 9.67 | 22.46
126.368 | 0 | 14.45 | 0.1 | 14.65 | 0.1 | 14.16 | 4.88 | 12.4 | 4.59 | 4.88 | 5.27 | 4.88 | 7.42 | 9.77 | 21.49 | 9.77 | 20.7 | 9.57 | 8.5 | 0.39 | 2.25 | 0.2 | 0.78 | 0.49 | 4.88 | 5.47 | 4.88 | 21.68 | 4.79 | 20.9 | 9.67 | 24.02 | 9.67 | 23.63 | 9.67 | 23.05
126.4 | 0.1 | 14.45 | 0.1 | 15.23 | 0 | 14.75 | 4.88 | 12.89 | 4.49 | 4.79 | 4.88 | 4.79 | 10.64 | 9.77 | 21.39 | 9.77 | 21.19 | 9.57 | 9.28 | 0.29 | 4.3 | 0.2 | 1.76 | 0.49 | 4.88 | 6.25 | 4.79 | 21.58 | 4.79 | 21.39 | 9.67 | 24.51 | 9.67 | 24.02 | 9.67 | 23.54
126.432 | 0.1 | 14.36 | 0.1 | 15.63 | 0.1 | 15.14 | 4.88 | 13.48 | 4.49 | 4.79 | 4.88 | 4.79 | 11.43 | 9.77 | 21.49 | 9.77 | 21.49 | 9.57 | 10.16 | 0.29 | 5.66 | 0.2 | 3.32 | 0.39 | 4.79 | 5.86 | 4.88 | 21.49 | 4.79 | 22.07 | 9.67 | 24.9 | 9.67 | 24.61 | 9.67 | 24.02
126.464 | 0.1 | 14.36 | 0.1 | 16.11 | 0.1 | 15.72 | 4.88 | 13.97 | 4.49 | 4.88 | 4.88 | 4.88 | 10.55 | 9.77 | 21.39 | 9.77 | 21.88 | 9.67 | 10.74 | 0.29 | 6.64 | 0.2 | 5.27 | 0.39 | 4.88 | 5.37 | 4.88 | 21.39 | 4.79 | 22.66 | 9.77 | 25.49 | 9.67 | 25.1 | 9.67 | 24.61
126.496 | 0.1 | 14.36 | 0.1 | 16.6 | 0.1 | 16.21 | 4.88 | 14.45 | 4.49 | 5.57 | 4.88 | 4.88 | 10.06 | 9.77 | 21.39 | 9.86 | 22.27 | 9.67 | 11.72 | 0.2 | 7.52 | 0.1 | 6.45 | 0.29 | 4.88 | 5.18 | 4.88 | 21.39 | 4.79 | 23.15 | 9.77 | 26.08 | 9.67 | 25.59 | 9.77 | 25
126.528 | 0.1 | 14.36 | 0.1 | 17.09 | 0.1 | 16.6 | 4.88 | 14.94 | 4.49 | 6.54 | 4.88 | 4.88 | 9.96 | 9.77 | 21.39 | 9.77 | 22.66 | 9.67 | 12.4 | 0.2 | 8.5 | 0.1 | 7.23 | 0.29 | 4.88 | 5.08 | 4.88 | 21.29 | 4.79 | 23.73 | 9.77 | 26.56 | 9.67 | 26.17 | 9.77 | 25.59
126.56 | 0.1 | 14.36 | 0.1 | 13.77 | 0 | 17.09 | 4.88 | 15.53 | 4.49 | 8.01 | 4.69 | 4.79 | 10.25 | 9.77 | 16.7 | 9.77 | 23.05 | 9.67 | 13.28 | 0.1 | 9.28 | 0.1 | 8.11 | 0.2 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 24.22 | 9.77 | 27.25 | 9.67 | 26.66 | 9.77 | 26.17
126.592 | 0.1 | 7.62 | 0.1 | 13.97 | 0.1 | 13.77 | 4.88 | 16.02 | 4.59 | 8.89 | 4.69 | 4.88 | 10.25 | 9.77 | 8.4 | 9.86 | 23.44 | 9.67 | 13.97 | 0.1 | 9.96 | 0.1 | 8.89 | 0.2 | 4.79 | 4.98 | 4.88 | 21.19 | 4.88 | 24.61 | 9.77 | 27.54 | 9.67 | 27.05 | 9.77 | 26.66
126.624 | 0.1 | 2.25 | 0.1 | 14.16 | 0.1 | 14.16 | 4.88 | 16.5 | 4.59 | 9.67 | 4.69 | 4.88 | 9.86 | 9.77 | 7.32 | 9.86 | 23.73 | 9.67 | 14.65 | 0.1 | 10.74 | 0.1 | 9.67 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.88 | 25.2 | 9.77 | 28.03 | 9.77 | 27.54 | 9.77 | 27.15
126.656 | 0.1 | 1.95 | 0.1 | 14.45 | 0.1 | 14.26 | 4.88 | 16.99 | 4.59 | 10.45 | 4.69 | 4.88 | 9.86 | 9.77 | 10.45 | 9.86 | 24.22 | 9.67 | 15.33 | 0.1 | 11.52 | 0.1 | 10.45 | 0.1 | 4.88 | 4.88 | 4.88 | 21.19 | 4.79 | 25.68 | 9.77 | 28.52 | 9.77 | 28.03 | 9.77 | 27.54
126.688 | 0 | 1.66 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 17.38 | 4.69 | 11.04 | 4.59 | 4.88 | 9.96 | 9.77 | 11.43 | 9.86 | 22.27 | 9.67 | 16.02 | 0.1 | 12.31 | 0 | 11.23 | 0.1 | 4.88 | 4.88 | 4.88 | 11.62 | 4.79 | 26.27 | 9.77 | 28.91 | 9.77 | 28.52 | 9.77 | 28.03
126.72 | 0.1 | 1.46 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 17.87 | 4.69 | 11.52 | 4.59 | 4.79 | 9.86 | 9.77 | 10.55 | 9.86 | 20.9 | 9.67 | 16.7 | 0.1 | 13.09 | 0 | 12.01 | 0.1 | 4.79 | 4.79 | 4.88 | 2.83 | 4.79 | 23.15 | 9.77 | 29.49 | 9.77 | 29.01 | 9.77 | 28.42
126.752 | 0.1 | 1.37 | 0.1 | 14.45 | 0 | 14.36 | 4.88 | 18.26 | 4.69 | 12.11 | 4.59 | 4.79 | 9.77 | 9.77 | 10.16 | 9.86 | 21.29 | 9.67 | 17.48 | 0.1 | 13.87 | 0.1 | 12.7 | 0.1 | 4.79 | 4.69 | 4.88 | 3.22 | 4.79 | 22.07 | 9.77 | 25.29 | 9.77 | 29.4 | 9.77 | 29.01
126.784 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.75 | 4.69 | 12.79 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.86 | 21.49 | 9.67 | 18.07 | 0.1 | 14.55 | 0.1 | 13.48 | 0.1 | 4.88 | 4.69 | 4.88 | 3.61 | 4.79 | 21.68 | 9.77 | 24.71 | 9.77 | 26.17 | 9.86 | 29.49
126.816 | 0.1 | 1.17 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.24 | 4.69 | 13.18 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.39 | 9.77 | 18.75 | 0.1 | 15.23 | 0.1 | 14.26 | 0.1 | 5.76 | 4.69 | 4.88 | 4.59 | 4.79 | 21.58 | 9.77 | 24.51 | 9.77 | 25.2 | 9.77 | 26.17
126.848 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.73 | 4.79 | 13.77 | 4.59 | 4.88 | 9.77 | 9.77 | 10.45 | 9.86 | 21.49 | 9.77 | 19.24 | 0.1 | 15.92 | 0 | 14.94 | 0.1 | 6.45 | 4.69 | 4.88 | 6.05 | 4.79 | 21.49 | 9.86 | 24.41 | 9.77 | 24.81 | 9.86 | 25.2
126.88 | 0.1 | 0.78 | 0.1 | 14.36 | 0 | 14.45 | 4.88 | 20.12 | 4.79 | 14.36 | 4.49 | 5.57 | 9.67 | 9.77 | 9.86 | 9.86 | 21.39 | 9.77 | 19.82 | 0.1 | 16.7 | 0 | 15.72 | 0.1 | 8.2 | 4.49 | 4.79 | 6.05 | 4.79 | 21.49 | 9.77 | 24.51 | 9.77 | 24.61 | 9.86 | 24.61
126.912 | 0.1 | 0.78 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 20.51 | 4.79 | 14.84 | 4.59 | 6.54 | 9.67 | 9.77 | 9.86 | 9.86 | 21.39 | 9.67 | 20.51 | 0 | 17.38 | 0.1 | 16.41 | 0.1 | 10.16 | 4.49 | 4.88 | 5.57 | 4.79 | 21.29 | 9.77 | 24.41 | 9.77 | 24.51 | 9.77 | 24.51
126.944 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.45 | 4.88 | 20.9 | 4.79 | 15.33 | 4.59 | 7.91 | 9.67 | 9.96 | 9.96 | 9.86 | 21.49 | 9.77 | 21 | 0.1 | 17.97 | 0.1 | 17.09 | 0.1 | 11.52 | 4.59 | 4.88 | 5.27 | 4.79 | 21.39 | 9.86 | 24.32 | 9.77 | 24.51 | 9.86 | 24.51
126.976 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 17.58 | 4.79 | 15.82 | 4.59 | 8.89 | 9.67 | 10.94 | 9.96 | 9.86 | 21.49 | 9.77 | 21.68 | 0.1 | 18.56 | 0.1 | 17.77 | 0.1 | 12.5 | 4.59 | 4.88 | 5.18 | 4.79 | 21.29 | 9.86 | 24.32 | 9.77 | 24.32 | 9.86 | 24.41
127.008 | 0.1 | 0.59 | 0.1 | 4.79 | 0.1 | 14.16 | 4.88 | 17.87 | 4.79 | 16.31 | 4.59 | 9.67 | 9.67 | 11.82 | 9.86 | 9.86 | 21.39 | 9.77 | 22.27 | 0.1 | 19.34 | 0 | 18.46 | 0.1 | 13.38 | 4.59 | 4.79 | 5.27 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.41 | 9.86 | 24.22
127.04 | 0.1 | 0.49 | 0.1 | 1.86 | 0.1 | 3.81 | 4.88 | 18.07 | 4.79 | 16.8 | 4.59 | 10.35 | 9.67 | 13.28 | 9.77 | 9.86 | 21.49 | 9.77 | 17.68 | 0 | 20.02 | 0 | 19.14 | 0.1 | 14.26 | 4.49 | 4.79 | 5.08 | 4.79 | 21.19 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.22
127.072 | 0.1 | 0.39 | 0.1 | 1.95 | 0.1 | 1.86 | 4.88 | 18.07 | 4.79 | 17.29 | 4.69 | 10.94 | 9.57 | 14.06 | 9.77 | 9.86 | 21.39 | 9.67 | 16.7 | 0.1 | 20.51 | 0.1 | 19.73 | 0.1 | 14.94 | 4.49 | 4.88 | 4.98 | 4.79 | 21.19 | 9.77 | 24.22 | 9.77 | 24.22 | 9.77 | 24.22
127.104 | 0.1 | 0.29 | 0.1 | 1.76 | 0.1 | 1.76 | 4.88 | 18.07 | 4.88 | 17.77 | 4.69 | 11.52 | 9.67 | 14.65 | 9.77 | 9.86 | 21.39 | 9.77 | 17.29 | 0.1 | 21 | 0.1 | 20.22 | 0.1 | 15.63 | 4.59 | 4.88 | 4.98 | 4.79 | 21.19 | 9.86 | 24.22 | 9.86 | 24.22 | 9.86 | 24.22
127.136 | 0.78 | 0.29 | 0.1 | 1.37 | 0.1 | 1.86 | 4.88 | 18.16 | 4.88 | 18.16 | 4.69 | 12.11 | 9.67 | 15.23 | 9.77 | 9.86 | 11.52 | 9.77 | 17.58 | 0.1 | 21.68 | 0.1 | 20.9 | 0.1 | 16.31 | 4.59 | 4.88 | 4.98 | 4.79 | 14.16 | 9.86 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12
127.168 | 1.46 | 0.29 | 0.1 | 1.37 | 0.1 | 1.37 | 4.88 | 18.16 | 4.88 | 18.65 | 4.69 | 12.7 | 9.67 | 15.72 | 9.67 | 9.96 | 6.64 | 9.77 | 17.58 | 0.1 | 22.17 | 0 | 21.49 | 0.1 | 16.99 | 4.59 | 4.79 | 4.88 | 4.79 | 3.42 | 9.86 | 19.73 | 9.77 | 24.22 | 9.86 | 24.12
127.2 | 2.64 | 0.2 | 0.1 | 1.07 | 0 | 1.27 | 4.88 | 18.16 | 4.79 | 19.14 | 4.79 | 13.28 | 9.57 | 16.21 | 9.67 | 9.86 | 9.08 | 9.77 | 17.38 | 0.1 | 20.7 | 0.1 | 22.07 | 0.1 | 17.68 | 4.59 | 4.88 | 4.79 | 4.79 | 3.13 | 9.77 | 7.13 | 9.77 | 15.43 | 9.86 | 24.12
127.232 | 3.81 | 0.2 | 0.1 | 1.07 | 0 | 1.17 | 4.88 | 18.07 | 4.88 | 19.53 | 4.79 | 13.77 | 9.57 | 16.7 | 9.67 | 9.86 | 11.43 | 9.77 | 17.58 | 0 | 18.26 | 0.1 | 22.66 | 0.1 | 18.26 | 4.59 | 5.76 | 4.69 | 4.79 | 3.52 | 9.77 | 6.64 | 9.77 | 6.74 | 9.77 | 15.23
127.264 | 4.59 | 0.1 | 0.1 | 0.88 | 0.1 | 1.07 | 4.88 | 18.16 | 4.88 | 19.92 | 4.79 | 14.26 | 9.67 | 17.19 | 9.67 | 9.96 | 11.13 | 9.77 | 17.58 | 0.1 | 18.07 | 0.1 | 18.75 | 0.1 | 18.95 | 4.59 | 7.03 | 4.79 | 4.79 | 4.49 | 9.86 | 10.64 | 9.86 | 7.03 | 9.86 | 6.64
127.296 | 5.18 | 0.1 | 0.1 | 1.07 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.41 | 4.79 | 14.84 | 9.67 | 17.68 | 9.67 | 9.86 | 10.25 | 9.86 | 17.58 | 0.1 | 17.97 | 0.1 | 18.26 | 0.1 | 19.73 | 4.69 | 8.98 | 4.69 | 4.79 | 5.96 | 9.86 | 12.21 | 9.86 | 10.94 | 9.86 | 7.32
127.328 | 5.66 | 0.1 | 0.1 | 0.88 | 0.1 | 0.88 | 4.88 | 18.07 | 4.88 | 20.8 | 4.79 | 15.33 | 9.67 | 18.07 | 9.67 | 9.96 | 9.86 | 9.86 | 17.58 | 0.1 | 17.87 | 0 | 18.07 | 0.1 | 20.22 | 4.69 | 10.64 | 4.69 | 4.79 | 6.15 | 9.86 | 10.74 | 9.86 | 12.01 | 9.86 | 11.23
127.36 | 6.25 | 0 | 0.1 | 0.68 | 0 | 0.78 | 4.88 | 18.07 | 4.88 | 17.87 | 4.79 | 15.82 | 9.67 | 18.56 | 9.57 | 9.86 | 9.96 | 9.77 | 17.48 | 0.1 | 17.68 | 0 | 17.97 | 0.1 | 20.9 | 4.69 | 11.91 | 4.59 | 4.79 | 5.76 | 9.77 | 9.96 | 9.86 | 10.45 | 9.86 | 11.72
127.392 | 6.84 | 0 | 0.1 | 0.59 | 0.1 | 0.68 | 4.88 | 17.87 | 4.88 | 17.68 | 4.88 | 16.31 | 9.67 | 19.04 | 9.57 | 9.86 | 10.16 | 9.77 | 17.58 | 0 | 17.68 | 0.1 | 17.97 | 0.1 | 21.39 | 4.69 | 12.89 | 4.59 | 4.79 | 5.37 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.64
127.424 | 7.42 | 0.1 | 0.1 | 0.59 | 0.1 | 0.59 | 4.88 | 4.59 | 4.88 | 17.97 | 4.88 | 16.8 | 9.67 | 19.43 | 9.67 | 9.96 | 10.16 | 9.86 | 17.58 | 0.1 | 17.68 | 0.1 | 17.87 | 0.1 | 22.07 | 4.69 | 13.77 | 4.59 | 4.79 | 5.27 | 9.86 | 10.45 | 9.86 | 10.16 | 9.86 | 9.86
127.456 | 8.11 | 0.1 | 0.1 | 0.49 | 0.1 | 0.59 | 4.88 | 2.83 | 4.88 | 18.16 | 4.88 | 17.19 | 9.67 | 19.92 | 9.67 | 9.96 | 9.86 | 9.86 | 17.48 | 0.1 | 17.68 | 0.1 | 17.77 | 0.1 | 22.56 | 4.79 | 14.45 | 4.59 | 4.79 | 5.08 | 9.77 | 10.45 | 9.86 | 10.45 | 9.86 | 10.16
127.488 | 8.69 | 0.1 | 0.1 | 0.49 | 0.1 | 0.49 | 4.88 | 3.61 | 4.88 | 18.07 | 4.88 | 17.68 | 9.67 | 20.31 | 9.67 | 9.96 | 9.86 | 9.86 | 6.25 | 0.1 | 17.68 | 0 | 17.77 | 0.1 | 23.15 | 4.79 | 15.23 | 4.59 | 4.79 | 5.18 | 9.77 | 10.25 | 9.86 | 10.45 | 9.86 | 10.45
127.52 | 9.38 | 0 | 0.1 | 0.29 | 0 | 0.39 | 4.88 | 4.59 | 4.79 | 18.07 | 4.88 | 18.16 | 9.67 | 20.8 | 9.67 | 9.86 | 9.86 | 9.86 | 1.76 | 0 | 17.68 | 0 | 17.77 | 0.1 | 23.73 | 4.79 | 16.02 | 4.49 | 5.66 | 5.08 | 9.77 | 10.06 | 9.77 | 10.25 | 9.86 | 10.45
127.552 | 9.86 | 0 | 0.1 | 0.29 | 0.1 | 0.29 | 4.88 | 5.66 | 4.88 | 18.07 | 4.88 | 18.56 | 9.67 | 21.09 | 9.67 | 9.86 | 9.86 | 9.77 | 2.44 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.22 | 4.79 | 16.6 | 4.49 | 7.03 | 4.98 | 9.77 | 10.16 | 9.77 | 9.96 | 9.86 | 10.25
127.584 | 10.45 | 0.1 | 0.68 | 0.29 | 0.1 | 0.29 | 4.88 | 5.86 | 4.88 | 18.07 | 4.88 | 18.95 | 9.77 | 21.58 | 9.67 | 9.86 | 9.77 | 9.86 | 1.95 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 24.71 | 4.79 | 17.29 | 4.59 | 8.98 | 5.08 | 9.77 | 10.16 | 9.86 | 10.06 | 9.86 | 10.06
127.616 | 11.13 | 0 | 1.46 | 0.2 | 0.1 | 0.29 | 4.88 | 5.47 | 4.88 | 18.07 | 4.88 | 19.43 | 9.77 | 21.88 | 9.67 | 10.35 | 9.77 | 9.86 | 1.66 | 0.1 | 17.58 | 0.1 | 17.68 | 0.1 | 25.29 | 4.79 | 17.97 | 4.59 | 10.55 | 4.98 | 9.77 | 9.96 | 9.86 | 10.25 | 9.86 | 10.16
127.648 | 11.62 | 0.1 | 2.73 | 0.2 | 0.1 | 0.2 | 4.88 | 5.18 | 4.88 | 18.07 | 4.88 | 19.92 | 9.77 | 22.36 | 9.67 | 11.13 | 9.77 | 9.86 | 1.56 | 0.1 | 8.11 | 0 | 17.68 | 0.1 | 25.78 | 4.79 | 18.56 | 4.59 | 11.91 | 4.88 | 9.86 | 10.06 | 9.86 | 9.96 | 9.86 | 10.06
127.68 | 12.21 | 0 | 3.91 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.79 | 18.07 | 4.88 | 20.31 | 9.77 | 22.85 | 9.67 | 12.11 | 9.67 | 9.86 | 1.46 | 0.1 | 1.76 | 0.1 | 12.7 | 0.1 | 26.27 | 4.79 | 19.24 | 4.59 | 12.89 | 4.69 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86 | 9.96
127.712 | 12.79 | 0 | 4.59 | 0.1 | 0.1 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 20.7 | 9.77 | 23.15 | 9.67 | 13.38 | 9.67 | 9.86 | 1.46 | 0.1 | 2.34 | 0.1 | 3.52 | 0.1 | 22.46 | 4.79 | 19.92 | 4.59 | 13.67 | 4.69 | 9.77 | 9.86 | 9.77 | 9.86 | 9.86 | 9.96
127.744 | 13.28 | 0.1 | 5.18 | 0.1 | 0.78 | 0.1 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 19.53 | 9.77 | 23.54 | 9.67 | 14.26 | 9.67 | 9.86 | 1.37 | 0.1 | 1.95 | 0.1 | 2.34 | 0.1 | 21.39 | 4.79 | 20.51 | 4.59 | 14.45 | 4.79 | 9.86 | 9.96 | 9.86 | 9.86 | 9.86 | 9.86
127.776 | 13.87 | 0.1 | 5.76 | 0.1 | 1.56 | 0.1 | 4.88 | 4.88 | 4.88 | 17.97 | 4.88 | 17.77 | 9.77 | 23.93 | 9.77 | 14.84 | 9.67 | 9.86 | 1.17 | 0.1 | 1.76 | 0.1 | 2.05 | 0.1 | 21.49 | 4.88 | 21.09 | 4.69 | 15.14 | 4.69 | 9.77 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
127.808 | 14.36 | 0.1 | 6.35 | 0.1 | 2.83 | 0 | 5.57 | 4.79 | 4.88 | 7.42 | 4.88 | 18.07 | 9.77 | 24.22 | 9.67 | 15.33 | 9.67 | 9.86 | 1.17 | 0.1 | 1.66 | 0 | 1.76 | 0.1 | 21.39 | 4.79 | 21.68 | 4.69 | 15.92 | 4.59 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77
127.84 | 14.84 | 0 | 6.93 | 0 | 3.91 | 0.1 | 6.45 | 4.69 | 4.88 | 2.64 | 4.88 | 18.16 | 9.77 | 21.29 | 9.77 | 15.92 | 9.67 | 9.86 | 0.88 | 0.1 | 1.66 | 0 | 1.66 | 0.1 | 21.58 | 4.79 | 22.27 | 4.69 | 16.6 | 4.59 | 10.84 | 9.77 | 9.77 | 9.77 | 9.86 | 9.86
127.872 | 15.43 | 0 | 7.42 | 0 | 4.59 | 0 | 7.72 | 4.69 | 4.88 | 3.22 | 4.88 | 18.07 | 9.77 | 21 | 9.77 | 16.31 | 9.57 | 9.86 | 0.88 | 0 | 1.27 | 0.1 | 1.46 | 0.1 | 21.39 | 4.79 | 22.85 | 4.69 | 17.19 | 4.49 | 11.82 | 9.77 | 9.77 | 9.77 | 10.84 | 9.77
127.904 | 15.92 | 0.1 | 8.11 | 0.1 | 5.27 | 0.1 | 8.79 | 4.69 | 4.88 | 4.2 | 4.88 | 18.16 | 9.77 | 21.29 | 9.77 | 16.9 | 9.67 | 9.96 | 0.88 | 0.1 | 1.27 | 0.1 | 1.46 | 0.1 | 21.39 | 4.88 | 23.34 | 4.79 | 17.97 | 4.59 | 13.67 | 9.67 | 10.74 | 9.77 | 11.91 | 9.77
127.936 | 16.41 | 0.1 | 8.69 | 0.1 | 5.76 | 0.1 | 9.67 | 4.59 | 4.88 | 5.47 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 17.29 | 9.67 | 9.86 | 0.78 | 0.1 | 1.27 | 0.1 | 1.37 | 0.1 | 21.39 | 4.88 | 23.83 | 4.79 | 18.56 | 4.59 | 15.33 | 9.67 | 11.82 | 9.77 | 13.57 | 9.77
127.968 | 16.8 | 0.1 | 9.38 | 0.1 | 6.35 | 0.1 | 10.25 | 4.59 | 4.88 | 5.96 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 17.77 | 9.67 | 9.86 | 0.59 | 0.1 | 1.07 | 0 | 1.17 | 0.1 | 21.29 | 4.88 | 24.32 | 4.79 | 19.24 | 4.49 | 16.5 | 9.67 | 13.48 | 9.67 | 15.33 | 9.67
128 | 17.29 | 0 | 9.96 | 0 | 7.03 | 0 | 11.04 | 4.59 | 4.79 | 5.47 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 18.26 | 9.57 | 9.86 | 0.59 | 0 | 0.98 | 0.1 | 1.17 | 0.1 | 21.29 | 4.79 | 24.9 | 4.79 | 19.92 | 4.49 | 17.58 | 9.67 | 15.23 | 9.67 | 16.5 | 9.67
128.032 | 15.43 | 0 | 10.45 | 0.1 | 7.52 | 0 | 11.62 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 18.65 | 9.57 | 9.86 | 0.59 | 0 | 0.88 | 0.1 | 1.07 | 0.1 | 21.29 | 4.79 | 25.29 | 4.79 | 20.41 | 4.49 | 18.16 | 9.67 | 16.5 | 9.67 | 17.38 | 9.67
128.064 | 15.23 | 0.1 | 11.04 | 0.1 | 8.2 | 0 | 12.21 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 9.77 | 21.39 | 9.86 | 19.24 | 9.67 | 9.96 | 0.49 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.88 | 4.79 | 21 | 4.59 | 18.95 | 9.67 | 17.38 | 9.67 | 18.26 | 9.67
128.096 | 14.84 | 0.1 | 11.72 | 0.1 | 8.69 | 0.1 | 12.7 | 4.49 | 4.88 | 4.79 | 4.88 | 18.07 | 9.77 | 21.39 | 9.77 | 19.53 | 9.67 | 9.86 | 0.39 | 0.1 | 0.78 | 0.1 | 0.88 | 0.1 | 21.29 | 4.88 | 25.68 | 4.79 | 21.68 | 4.59 | 19.63 | 9.67 | 18.16 | 9.67 | 18.75 | 9.67
128.128 | 14.65 | 0.1 | 12.21 | 0.1 | 9.38 | 0.1 | 13.28 | 4.59 | 4.88 | 4.79 | 4.88 | 17.97 | 9.77 | 21.39 | 9.77 | 20.02 | 9.67 | 9.96 | 0.29 | 0.1 | 0.68 | 0 | 0.78 | 0.1 | 17.38 | 4.88 | 22.17 | 4.79 | 22.17 | 4.59 | 20.22 | 9.57 | 18.75 | 9.67 | 19.43 | 9.67
128.16 | 14.75 | 0 | 12.79 | 0 | 9.96 | 0 | 13.87 | 4.49 | 4.79 | 4.79 | 4.88 | 17.97 | 9.77 | 21.29 | 9.77 | 20.41 | 9.67 | 10.55 | 0.29 | 0 | 0.59 | 0 | 0.68 | 0.1 | 3.81 | 4.79 | 21.49 | 4.79 | 22.75 | 4.59 | 20.8 | 9.57 | 19.53 | 9.67 | 20.12 | 9.67
128.192 | 14.55 | 0 | 13.28 | 0.1 | 10.45 | 0 | 14.26 | 4.49 | 4.88 | 4.79 | 4.88 | 7.52 | 9.77 | 21.29 | 9.77 | 20.8 | 9.67 | 11.43 | 0.29 | 0.1 | 0.49 | 0.1 | 0.59 | 0.1 | 3.03 | 4.79 | 21.58 | 4.79 | 23.34 | 4.59 | 21.39 | 9.67 | 20.12 | 9.67 | 20.8 | 9.67
128.224 | 14.55 | 0 | 13.87 | 0.1 | 11.13 | 0.1 | 14.75 | 4.59 | 4.88 | 4.69 | 4.88 | 2.64 | 9.86 | 21.39 | 9.86 | 21.29 | 9.67 | 12.79 | 0.2 | 0.1 | 0.49 | 0.1 | 0.49 | 0.1 | 3.42 | 4.88 | 21.49 | 4.79 | 23.83 | 4.69 | 21.97 | 9.67 | 20.7 | 9.67 | 21.39 | 9.67
128.256 | 14.45 | 0.1 | 14.36 | 0.1 | 11.72 | 0.1 | 15.33 | 4.59 | 4.88 | 4.69 | 4.88 | 3.22 | 9.77 | 21.09 | 9.86 | 21.68 | 9.67 | 13.77 | 0.2 | 0.1 | 0.39 | 0.1 | 0.49 | 0.1 | 4.1 | 4.88 | 21.39 | 4.79 | 24.41 | 4.69 | 22.56 | 9.67 | 21.39 | 9.67 | 21.97 | 9.67
128.288 | 14.45 | 0.1 | 14.84 | 0.1 | 12.31 | 0.1 | 15.82 | 4.59 | 4.88 | 4.59 | 4.88 | 4.3 | 9.77 | 9.28 | 9.86 | 22.07 | 9.67 | 14.45 | 0.2 | 0.1 | 0.29 | 0.1 | 0.49 | 0.1 | 5.66 | 4.88 | 21.39 | 4.79 | 24.71 | 4.69 | 23.05 | 9.67 | 21.88 | 9.67 | 22.46 | 9.67
128.32 | 14.45 | 0 | 15.43 | 0.1 | 12.79 | 0 | 16.41 | 4.59 | 4.79 | 4.59 | 4.88 | 5.47 | 9.77 | 7.13 | 9.86 | 22.46 | 9.67 | 15.14 | 0.1 | 0 | 0.29 | 0 | 0.39 | 0.1 | 6.35 | 4.79 | 21.29 | 4.79 | 25.39 | 4.79 | 23.73 | 9.67 | 22.46 | 9.67 | 23.15 | 9.67
128.352 | 14.45 | 0 | 15.82 | 0.1 | 13.38 | 0 | 16.8 | 4.59 | 4.88 | 4.49 | 4.88 | 5.86 | 9.77 | 9.86 | 9.86 | 22.85 | 9.67 | 15.63 | 0.2 | 0.1 | 0.29 | 0.88 | 0.29 | 0.1 | 5.96 | 4.79 | 21.29 | 4.79 | 25.88 | 4.69 | 24.22 | 9.67 | 22.95 | 9.67 | 23.63 | 9.67
128.384 | 14.36 | 0.1 | 16.31 | 0.1 | 13.87 | 0.1 | 17.29 | 4.69 | 4.88 | 4.49 | 4.88 | 5.57 | 9.86 | 11.52 | 9.86 | 23.24 | 9.67 | 16.11 | 0.1 | 0.1 | 0.2 | 1.95 | 0.29 | 0.1 | 5.37 | 4.88 | 21.29 | 4.88 | 25.78 | 4.79 | 24.71 | 9.67 | 23.44 | 9.67 | 24.12 | 9.67
128.416 | 14.36 | 0.1 | 16.9 | 0.1 | 14.45 | 0.1 | 17.68 | 4.69 | 4.88 | 4.49 | 4.88 | 5.18 | 9.77 | 11.04 | 9.86 | 23.54 | 9.67 | 16.6 | 0.1 | 1.07 | 0.2 | 3.61 | 0.2 | 0.1 | 5.27 | 4.88 | 21.19 | 4.79 | 22.17 | 4.79 | 25.2 | 9.67 | 24.12 | 9.67 | 24.61 | 9.67
128.448 | 9.57 | 0.1 | 17.29 | 0.1 | 14.94 | 0.1 | 18.16 | 4.69 | 5.57 | 4.49 | 4.88 | 5.08 | 9.86 | 9.96 | 9.86 | 24.02 | 9.67 | 17.09 | 0.1 | 2.05 | 0.1 | 5.47 | 0.2 | 0.1 | 5.08 | 4.88 | 21.19 | 4.79 | 21.88 | 4.79 | 25.68 | 9.67 | 24.61 | 9.77 | 25.1 | 9.67
128.48 | 2.64 | 0 | 15.43 | 0.1 | 15.43 | 0 | 18.56 | 4.69 | 6.64 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 24.32 | 9.67 | 17.77 | 0.1 | 3.81 | 0.1 | 6.54 | 0.1 | 0.2 | 5.08 | 4.79 | 21.19 | 4.79 | 21.49 | 4.79 | 26.27 | 9.67 | 25.2 | 9.67 | 25.59 | 9.67
128.512 | 2.15 | 0 | 15.23 | 0.1 | 15.92 | 0 | 19.14 | 4.69 | 8.01 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 21.09 | 9.67 | 18.07 | 0.1 | 5.37 | 0.1 | 7.42 | 0.1 | 1.27 | 4.98 | 4.79 | 21.19 | 4.79 | 21.39 | 4.79 | 26.76 | 9.67 | 25.59 | 9.67 | 26.08 | 9.67
128.544 | 1.76 | 0.1 | 14.94 | 0.1 | 16.41 | 0.1 | 19.53 | 4.79 | 8.98 | 4.49 | 4.88 | 4.88 | 9.77 | 10.16 | 9.86 | 21.49 | 9.67 | 18.46 | 0.1 | 6.54 | 0.1 | 8.2 | 0.1 | 2.64 | 4.98 | 4.88 | 9.28 | 4.88 | 21.39 | 4.79 | 27.25 | 9.77 | 26.08 | 9.77 | 26.56 | 9.77
128.576 | 1.56 | 0.1 | 14.75 | 0.1 | 16.9 | 0.1 | 20.02 | 4.79 | 9.67 | 4.49 | 4.88 | 4.88 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 18.95 | 0.1 | 7.32 | 0.1 | 8.98 | 0.1 | 4.49 | 4.88 | 4.88 | 2.44 | 4.79 | 21.39 | 4.88 | 27.64 | 9.77 | 26.66 | 9.77 | 27.15 | 9.77
128.608 | 1.46 | 0.1 | 14.75 | 0.1 | 17.29 | 0 | 20.41 | 4.79 | 10.45 | 4.49 | 4.88 | 4.69 | 9.77 | 9.96 | 9.86 | 21.49 | 9.77 | 19.34 | 0.1 | 8.2 | 0.1 | 9.86 | 0.1 | 5.96 | 4.98 | 4.88 | 3.32 | 4.79 | 21.29 | 4.88 | 28.13 | 9.77 | 27.05 | 9.77 | 27.54 | 9.77
128.64 | 1.27 | 0 | 14.55 | 0 | 15.04 | 0.1 | 20.9 | 4.79 | 11.04 | 4.49 | 4.88 | 4.69 | 9.77 | 9.77 | 9.86 | 21.58 | 9.77 | 19.82 | 0.1 | 8.98 | 0.1 | 10.64 | 0.1 | 7.03 | 4.79 | 4.79 | 3.81 | 4.79 | 21.29 | 4.79 | 28.61 | 9.77 | 27.54 | 9.77 | 28.03 | 9.77
128.672 | 1.27 | 0 | 14.55 | 0 | 15.14 | 0 | 20.7 | 4.79 | 11.52 | 4.59 | 4.88 | 4.59 | 9.77 | 9.96 | 9.86 | 21.49 | 9.67 | 20.12 | 0.1 | 9.67 | 0 | 11.43 | 0.1 | 7.81 | 4.69 | 4.79 | 4.79 | 4.79 | 21.19 | 4.79 | 29.1 | 9.77 | 28.03 | 9.77 | 28.52 | 9.77
128.704 | 1.07 | 0.1 | 14.55 | 0.1 | 14.84 | 0.1 | 18.75 | 4.79 | 12.21 | 4.59 | 4.88 | 4.69 | 9.86 | 9.86 | 9.86 | 21.49 | 9.77 | 20.7 | 0.1 | 10.55 | 0.1 | 12.21 | 0.1 | 8.59 | 4.69 | 4.88 | 6.15 | 4.88 | 21.19 | 4.88 | 29.59 | 9.77 | 28.61 | 9.77 | 29.01 | 9.77
128.736 | 0.88 | 0.1 | 14.45 | 0.1 | 14.75 | 0.1 | 18.65 | 4.79 | 12.7 | 4.59 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.09 | 0.1 | 11.23 | 0.1 | 12.99 | 0.1 | 9.47 | 4.69 | 4.88 | 6.05 | 4.88 | 21.19 | 4.88 | 25.98 | 9.77 | 28.91 | 9.77 | 29.49 | 9.77
128.768 | 0.78 | 0.1 | 14.45 | 0.1 | 14.75 | 0.1 | 18.46 | 4.88 | 13.28 | 4.69 | 4.88 | 4.59 | 9.77 | 9.77 | 9.86 | 21.49 | 9.77 | 21.39 | 0.1 | 12.11 | 0.1 | 13.77 | 0.1 | 10.16 | 4.69 | 4.88 | 5.66 | 4.79 | 21.19 | 4.88 | 25.1 | 9.77 | 29.49 | 9.77 | 26.17 | 9.77
128.8 | 0.78 | 0 | 14.45 | 0 | 14.45 | 0 | 18.26 | 4.79 | 13.87 | 4.59 | 4.88 | 4.59 | 9.77 | 9.67 | 9.86 | 21.49 | 9.67 | 21.78 | 0.1 | 12.99 | 0.1 | 14.36 | 0.1 | 11.04 | 4.59 | 4.79 | 5.27 | 4.79 | 21.09 | 4.79 | 24.61 | 9.77 | 26.08 | 9.77 | 25.1 | 9.77
128.832 | 0.78 | 0 | 14.45 | 0 | 14.45 | 0 | 18.26 | 4.88 | 14.36 | 4.69 | 5.57 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.27 | 0.1 | 13.67 | 0 | 15.14 | 0 | 11.72 | 4.59 | 4.79 | 5.18 | 4.79 | 9.08 | 4.79 | 24.61 | 9.77 | 25.2 | 9.77 | 24.71 | 9.77
128.864 | 0.59 | 0.1 | 14.45 | 0.1 | 14.55 | 0.1 | 18.26 | 4.88 | 14.84 | 4.69 | 6.45 | 4.59 | 9.77 | 9.67 | 9.86 | 21.49 | 9.77 | 22.66 | 0.1 | 14.45 | 0.1 | 15.82 | 0.1 | 12.6 | 4.59 | 4.88 | 5.18 | 4.88 | 2.54 | 4.79 | 24.61 | 9.77 | 24.81 | 9.86 | 24.51 | 9.77
128.896 | 0.59 | 0.1 | 9.38 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.33 | 4.79 | 7.72 | 4.49 | 9.77 | 9.67 | 9.86 | 21.39 | 9.77 | 22.95 | 0.1 | 15.14 | 0 | 16.5 | 0.1 | 13.28 | 4.59 | 5.08 | 5.08 | 4.79 | 3.42 | 4.88 | 24.51 | 9.77 | 24.61 | 9.86 | 24.51 | 9.77
128.928 | 0.39 | 0.1 | 2.64 | 0.1 | 14.45 | 0.1 | 18.16 | 4.88 | 15.82 | 4.79 | 8.79 | 4.49 | 9.77 | 9.67 | 9.86 | 20.9 | 9.77 | 23.34 | 0.1 | 15.82 | 0.1 | 17.19 | 0.1 | 14.06 | 4.59 | 6.05 | 4.98 | 4.88 | 3.71 | 4.88 | 24.32 | 9.77 | 24.51 | 9.86 | 24.41 | 9.86
128.96 | 0.39 | 0 | 1.95 | 0 | 14.36 | 0 | 18.16 | 4.88 | 16.31 | 4.69 | 9.67 | 4.59 | 9.77 | 9.57 | 9.86 | 8.4 | 9.77 | 23.73 | 0.1 | 16.41 | 0.1 | 17.87 | 0.1 | 14.84 | 4.49 | 7.62 | 4.98 | 4.79 | 5.08 | 4.88 | 24.32 | 9.77 | 24.41 | 9.86 | 24.22 | 9.77
128.992 | 0.29 | 0 | 1.86 | 0.1 | 14.36 | 0 | 18.16 | 4.88 | 16.9 | 4.79 | 10.25 | 4.59 | 10.55 | 9.57 | 9.86 | 7.13 | 9.77 | 24.12 | 0.1 | 17.19 | 0 | 18.46 | 0 | 15.43 | 4.49 | 9.67 | 4.98 | 4.79 | 6.15 | 4.79 | 24.32 | 9.77 | 24.22 | 9.77 | 24.32 | 9.86
129.024 | 0.29 | 0.1 | 1.37 | 0.1 | 14.36 | 0.1 | 18.16 | 4.88 | 17.29 | 4.79 | 10.94 | 4.59 | 11.04 | 9.57 | 9.86 | 10.06 | 9.77 | 23.73 | 0.1 | 17.77 | 0.1 | 19.04 | 0.1 | 16.31 | 4.59 | 11.33 | 4.88 | 4.88 | 6.05 | 4.88 | 24.32 | 9.77 | 24.32 | 9.86 | 24.22 | 9.86
129.056 | 0.2 | 0.1 | 1.46 | 0.1 | 8.11 | 0.1 | 18.16 | 4.88 | 17.77 | 4.79 | 11.52 | 4.59 | 12.01 | 9.57 | 9.86 | 11.52 | 9.77 | 21.68 | 0.1 | 18.56 | 0.1 | 19.73 | 0.1 | 16.8 | 4.49 | 12.4 | 4.79 | 4.88 | 5.57 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.22 | 9.86
129.088 | 0.2 | 0.1 | 1.17 | 0.1 | 2.34 | 0.1 | 18.07 | 4.88 | 18.16 | 4.79 | 12.11 | 4.59 | 13.38 | 9.57 | 9.86 | 10.64 | 9.77 | 21.68 | 0.1 | 19.24 | 0.1 | 20.31 | 0.1 | 17.38 | 4.59 | 13.28 | 4.79 | 4.88 | 5.18 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
129.12 | 0.2 | 0 | 1.07 | 0 | 1.95 | 0 | 8.3 | 4.88 | 18.56 | 4.79 | 12.7 | 4.59 | 14.16 | 9.57 | 9.86 | 9.96 | 9.77 | 21.58 | 0.1 | 19.82 | 0 | 20.9 | 0.1 | 18.26 | 4.59 | 14.06 | 4.69 | 4.79 | 5.18 | 4.88 | 24.22 | 9.77 | 24.22 | 9.86 | 24.12 | 9.86
129.152 | 0.1 | 0 | 0.98 | 0.1 | 1.66 | 0 | 2.83 | 4.88 | 19.14 | 4.79 | 13.18 | 4.59 | 14.75 | 9.57 | 9.86 | 10.06 | 9.77 | 21.58 | 0.1 | 20.31 | 0.1 | 21.58 | 0 | 18.95 | 4.59 | 14.84 | 4.69 | 4.79 | 5.08 | 4.79 | 13.87 | 9.77 | 24.12 | 9.86 | 24.12 | 9.86
129.184 | 0.1 | 0.1 | 0.98 | 0.1 | 1.46 | 0.1 | 3.13 | 4.88 | 19.43 | 4.88 | 13.77 | 4.69 | 15.33 | 9.67 | 9.86 | 10.06 | 9.77 | 21.58 | 0.1 | 20.9 | 0.1 | 22.07 | 0.1 | 19.43 | 4.59 | 15.53 | 4.69 | 4.88 | 4.98 | 4.88 | 6.25 | 9.86 | 24.22 | 9.86 | 15.63 | 9.86
129.216 | 0.1 | 0.1 | 0.88 | 0.1 | 1.46 | 0.1 | 4.3 | 4.88 | 19.92 | 4.88 | 14.26 | 4.69 | 15.82 | 9.67 | 9.86 | 10.16 | 9.77 | 21.49 | 0.1 | 21.68 | 0.1 | 22.66 | 0.1 | 20.02 | 4.59 | 16.31 | 4.59 | 4.79 | 5.08 | 4.88 | 7.62 | 9.77 | 15.43 | 9.86 | 6.74 | 9.86
129.248 | 0 | 0.1 | 0.78 | 0.1 | 1.27 | 0.1 | 5.08 | 4.88 | 20.31 | 4.88 | 14.84 | 4.69 | 16.41 | 9.67 | 10.25 | 9.96 | 9.77 | 21.49 | 0.1 | 22.17 | 0.1 | 19.43 | 0.1 | 20.7 | 4.69 | 16.9 | 4.59 | 4.79 | 4.98 | 4.88 | 11.43 | 9.86 | 6.45 | 9.86 | 7.32 | 9.86
129.28 | 0.1 | 0.68 | 0.59 | 0 | 1.07 | 0 | 6.05 | 4.88 | 20.8 | 4.79 | 15.33 | 4.69 | 16.9 | 9.67 | 11.13 | 9.96 | 9.77 | 21.49 | 0.1 | 22.66 | 0 | 18.65 | 0.1 | 21.29 | 4.69 | 17.58 | 4.59 | 4.88 | 4.88 | 4.88 | 11.82 | 9.77 | 7.62 | 9.86 | 11.23 | 9.86
129.312 | 0.1 | 1.56 | 0.59 | 0 | 0.98 | 0 | 5.76 | 4.88 | 18.85 | 4.79 | 15.82 | 4.79 | 17.29 | 9.67 | 12.01 | 9.96 | 9.77 | 21.39 | 0.1 | 19.34 | 0 | 18.26 | 0 | 21.88 | 4.69 | 18.16 | 4.59 | 4.79 | 4.98 | 4.79 | 10.45 | 9.77 | 11.23 | 9.86 | 11.82 | 9.86
129.344 | 0.1 | 2.83 | 0.59 | 0.1 | 0.88 | 0.1 | 5.27 | 4.88 | 17.68 | 4.88 | 16.31 | 4.79 | 17.77 | 9.67 | 13.28 | 9.86 | 9.86 | 21.39 | 0.1 | 18.65 | 0.1 | 18.07 | 0.1 | 22.46 | 4.69 | 18.85 | 4.59 | 4.79 | 4.79 | 4.88 | 9.96 | 9.86 | 11.91 | 9.86 | 10.64 | 9.86
129.376 | 0.1 | 3.81 | 0.49 | 0.1 | 0.78 | 0.1 | 5.08 | 4.88 | 17.97 | 4.88 | 16.7 | 4.79 | 18.16 | 9.67 | 14.16 | 9.86 | 9.86 | 21.39 | 0.1 | 18.16 | 0.1 | 18.07 | 0.1 | 19.63 | 4.69 | 19.43 | 4.49 | 5.66 | 4.69 | 4.88 | 10.25 | 9.86 | 10.55 | 9.86 | 9.96 | 9.86
129.408 | 0.1 | 4.59 | 0.39 | 0.1 | 0.68 | 0.1 | 4.98 | 4.88 | 18.16 | 4.88 | 17.19 | 4.79 | 18.65 | 9.67 | 14.75 | 9.77 | 9.77 | 21.39 | 0.1 | 17.97 | 0.1 | 17.87 | 0.1 | 17.09 | 4.79 | 20.12 | 4.49 | 6.64 | 4.69 | 4.88 | 10.45 | 9.77 | 9.86 | 9.86 | 9.96 | 9.86
129.44 | 0.1 | 5.18 | 0.29 | 0 | 0.59 | 0.1 | 4.88 | 4.79 | 17.97 | 4.79 | 17.77 | 4.79 | 19.14 | 9.67 | 15.33 | 9.77 | 9.86 | 21.39 | 0 | 17.87 | 0 | 17.77 | 0.1 | 17.58 | 4.79 | 20.8 | 4.59 | 8.4 | 4.69 | 4.88 | 10.45 | 9.77 | 10.16 | 9.86 | 10.35 | 9.86
129.472 | 0.1 | 5.66 | 0.29 | 0 | 0.59 | 0 | 4.88 | 4.88 | 18.07 | 4.88 | 18.16 | 4.79 | 19.53 | 9.67 | 15.82 | 9.67 | 9.77 | 12.5 | 0.1 | 17.77 | 0 | 17.77 | 0 | 17.77 | 4.79 | 21.39 | 4.59 | 10.25 | 4.59 | 4.88 | 10.16 | 9.77 | 10.35 | 9.86 | 10.45 | 9.86
129.504 | 0.1 | 6.35 | 0.29 | 0.1 | 0.59 | 0.1 | 4.88 | 4.88 | 18.07 | 4.88 | 18.56 | 4.88 | 20.02 | 9.77 | 16.31 | 9.67 | 9.86 | 6.74 | 0.1 | 17.77 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 21.97 | 4.59 | 11.52 | 4.59 | 4.88 | 10.06 | 9.86 | 10.45 | 9.86 | 10.25 | 9.86
129.536 | 0.1 | 6.93 | 0.2 | 0.1 | 0.49 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 19.04 | 4.79 | 20.41 | 9.77 | 16.9 | 9.67 | 9.86 | 8.4 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 22.56 | 4.59 | 12.6 | 4.59 | 4.88 | 10.16 | 9.86 | 10.25 | 9.86 | 9.96 | 9.86
129.568 | 0 | 7.52 | 0.2 | 0.1 | 0.39 | 0.1 | 4.79 | 4.88 | 18.07 | 4.88 | 19.43 | 4.88 | 20.8 | 9.77 | 17.19 | 9.67 | 9.86 | 11.13 | 0.1 | 17.68 | 0.1 | 17.68 | 0.1 | 17.77 | 4.79 | 23.05 | 4.59 | 13.48 | 4.59 | 4.88 | 10.25 | 9.86 | 10.06 | 9.86 | 10.16 | 9.86
129.6 | 0.1 | 8.11 | 0.1 | 0.1 | 0.29 | 0 | 4.69 | 4.88 | 18.07 | 4.79 | 19.92 | 4.88 | 21.29 | 9.67 | 17.77 | 9.67 | 9.77 | 11.13 | 0.1 | 17.58 | 0.1 | 17.68 | 0 | 17.77 | 4.79 | 23.63 | 4.59 | 14.26 | 4.49 | 4.88 | 9.96 | 9.77 | 10.16 | 9.86 | 10.16 | 9.86
129.632 | 0.1 | 8.69 | 0.1 | 0 | 0.2 | 0 | 4.59 | 4.88 | 18.07 | 4.79 | 20.31 | 4.79 | 21.68 | 9.67 | 18.16 | 9.67 | 9.86 | 10.35 | 0.1 | 17.58 | 0 | 17.68 | 0 | 17.77 | 4.79 | 24.12 | 4.69 | 14.94 | 4.59 | 4.88 | 9.96 | 9.77 | 10.06 | 9.86 | 9.96 | 9.86
129.664 | 0.1 | 9.28 | 0.1 | 0.2 | 0.29 | 0.1 | 4.59 | 4.88 | 18.07 | 4.88 | 20.7 | 4.88 | 22.07 | 9.77 | 18.65 | 9.67 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 10.45 | 0.1 | 17.77 | 4.88 | 24.61 | 4.69 | 15.63 | 4.49 | 4.88 | 9.96 | 9.86 | 10.06 | 9.86 | 9.96 | 9.86
129.696 | 0.1 | 9.96 | 0.1 | 1.17 | 0.2 | 0.2 | 4.59 | 4.88 | 18.07 | 4.88 | 20.61 | 4.88 | 22.46 | 9.77 | 19.14 | 9.57 | 9.86 | 9.86 | 0.1 | 17.58 | 0.1 | 2.54 | 0.1 | 17.68 | 4.79 | 25.1 | 4.69 | 16.41 | 4.49 | 4.88 | 9.96 | 9.86 | 9.86 | 9.86 | 9.96 | 9.86
129.728 | 0.1 | 10.55 | 0.1 | 1.95 | 0.1 | 1.17 | 4.59 | 4.88 | 18.07 | 4.88 | 18.65 | 4.88 | 22.75 | 9.77 | 19.53 | 9.67 | 9.86 | 10.25 | 0.1 | 9.57 | 0.1 | 2.25 | 0.1 | 17.77 | 4.79 | 25.59 | 4.69 | 17.09 | 4.49 | 4.88 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86 | 9.86
129.76 | 0.1 | 11.13 | 0.1 | 3.42 | 0.1 | 2.15 | 4.59 | 4.79 | 6.74 | 4.79 | 18.26 | 4.88 | 23.15 | 9.77 | 20.02 | 9.67 | 9.77 | 10.06 | 0.1 | 2.15 | 0 | 2.05 | 0.1 | 17.68 | 4.79 | 26.17 | 4.69 | 17.68 | 4.49 | 4.88 | 9.67 | 9.77 | 9.86 | 9.86 | 9.86 | 9.86
129.792 | 0.1 | 11.62 | 0.1 | 4.2 | 0.1 | 3.42 | 4.59 | 4.88 | 2.64 | 4.79 | 18.46 | 4.88 | 23.54 | 9.77 | 20.41 | 9.57 | 9.86 | 9.86 | 0.1 | 2.34 | 0 | 1.66 | 0.1 | 17.68 | 4.79 | 21.39 | 4.79 | 18.26 | 4.59 | 4.79 | 9.77 | 9.77 | 9.86 | 9.86 | 9.77 | 9.86
129.824 | 0.1 | 12.31 | 0.1 | 4.88 | 0.1 | 4.3 | 4.59 | 4.88 | 3.42 | 4.88 | 18.26 | 4.88 | 23.93 | 9.77 | 20.8 | 9.67 | 10.45 | 9.86 | 0.1 | 1.95 | 0.1 | 1.86 | 0.1 | 6.74 | 4.88 | 20.31 | 4.79 | 19.14 | 4.59 | 5.86 | 9.77 | 9.86 | 9.77 | 9.86 | 9.77 | 9.86
129.856 | 0.1 | 12.79 | 0.1 | 5.47 | 0.1 | 4.88 | 4.49 | 4.88 | 4.3 | 4.88 | 18.26 | 4.88 | 24.32 | 9.77 | 21.19 | 9.67 | 11.52 | 9.86 | 0.1 | 1.66 | 0.1 | 1.46 | 0.1 | 1.66 | 4.79 | 21 | 4.79 | 19.63 | 4.59 | 6.45 | 9.67 | 10.74 | 9.77 | 9.86 | 9.77 | 9.86
129.888 | 0 | 13.38 | 0.1 | 6.05 | 0.1 | 5.47 | 4.59 | 4.88 | 5.47 | 4.88 | 18.16 | 4.88 | 21.29 | 9.77 | 21.58 | 9.67 | 12.7 | 9.86 | 0.1 | 1.66 | 0.1 | 1.46 | 0.1 | 2.25 | 4.88 | 21.19 | 4.79 | 20.22 | 4.59 | 8.01 | 9.67 | 11.43 | 9.77 | 10.74 | 9.67 | 9.86
129.92 | 0.1 | 13.87 | 0.1 | 6.54 | 0.1 | 6.05 | 4.59 | 4.88 | 5.86 | 4.79 | 18.07 | 4.88 | 22.07 | 9.77 | 22.07 | 9.67 | 13.77 | 9.67 | 0 | 1.56 | 0 | 1.17 | 0.1 | 1.95 | 4.88 | 21.19 | 4.79 | 20.9 | 4.59 | 10.25 | 9.67 | 13.18 | 9.67 | 11.82 | 9.67 | 10.94
129.952 | 0.1 | 14.45 | 0.1 | 7.13 | 0.1 | 6.54 | 4.59 | 4.88 | 5.57 | 4.79 | 18.16 | 4.79 | 21.49 | 9.77 | 22.46 | 9.67 | 14.36 | 9.67 | 0.1 | 1.27 | 0 | 1.17 | 0 | 1.86 | 4.79 | 21.19 | 4.79 | 21.49 | 4.69 | 11.52 | 9.67 | 15.04 | 9.67 | 13.38 | 9.67 | 11.72
129.984 | 0.1 | 14.94 | 0.1 | 7.81 | 0.1 | 7.13 | 4.59 | 4.88 | 5.27 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 22.75 | 9.67 | 15.04 | 9.77 | 0.1 | 1.37 | 0.1 | 1.17 | 0.1 | 1.66 | 4.88 | 21.29 | 4.79 | 22.07 | 4.69 | 12.6 | 9.67 | 16.31 | 9.67 | 15.23 | 9.67 | 13.38
130.016 | 0.1 | 15.43 | 0.1 | 8.3 | 0.1 | 7.81 | 4.59 | 4.88 | 4.98 | 4.88 | 18.07 | 4.88 | 21.58 | 9.77 | 23.15 | 9.67 | 15.63 | 9.67 | 0.1 | 1.27 | 0.1 | 0.98 | 0.1 | 1.46 | 4.88 | 21.39 | 4.79 | 22.56 | 4.69 | 13.38 | 9.67 | 17.19 | 9.67 | 16.41 | 9.67 | 15.14
130.048 | 0 | 15.92 | 0.1 | 8.98 | 0.1 | 8.4 | 4.69 | 4.88 | 4.88 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 23.54 | 9.67 | 16.11 | 9.67 | 0.1 | 1.17 | 0.1 | 0.88 | 0.1 | 1.46 | 4.88 | 21.29 | 4.79 | 23.15 | 4.69 | 14.16 | 9.67 | 17.97 | 9.67 | 17.29 | 9.67 | 16.41
130.08 | 0.1 | 16.41 | 0.1 | 9.57 | 0 | 8.98 | 4.69 | 4.98 | 4.79 | 4.88 | 18.07 | 4.88 | 21.49 | 9.77 | 24.02 | 9.67 | 16.6 | 9.67 | 0.1 | 0.98 | 0.1 | 0.88 | 0 | 1.27 | 4.88 | 21.29 | 4.79 | 23.73 | 4.69 | 14.94 | 9.67 | 18.65 | 9.67 | 18.07 | 9.67 | 17.29
130.112 | 0.1 | 16.9 | 0.1 | 10.16 | 0.1 | 9.57 | 4.69 | 5.66 | 4.79 | 4.79 | 18.07 | 4.88 | 21.49 | 9.77 | 24.32 | 9.67 | 17.09 | 9.67 | 0.1 | 0.88 | 0 | 0.78 | 0 | 1.07 | 4.88 | 21.19 | 4.79 | 24.22 | 4.79 | 15.72 | 9.67 | 19.43 | 9.57 | 18.75 | 9.57 | 18.07
130.144 | 0.1 | 17.29 | 0.1 | 10.74 | 0.1 | 10.16 | 4.79 | 6.74 | 4.88 | 4.88 | 8.69 | 4.88 | 21.39 | 9.77 | 21 | 9.77 | 17.58 | 9.67 | 0.1 | 0.88 | 0.1 | 0.68 | 0.1 | 1.17 | 4.88 | 21.29 | 4.88 | 24.71 | 4.79 | 16.31 | 9.67 | 20.02 | 9.67 | 19.43 | 9.67 | 18.75
130.176 | 0.1 | 15.33 | 0.1 | 11.33 | 0.1 | 10.74 | 4.69 | 8.2 | 4.69 | 4.88 | 2.83 | 4.88 | 21.49 | 9.77 | 21.29 | 9.77 | 17.97 | 9.57 | 1.17 | 0.68 | 0 | 0.68 | 0.1 | 1.07 | 4.88 | 21.29 | 4.79 | 25.29 | 4.79 | 17.09 | 9.67 | 20.61 | 9.67 | 20.12 | 9.67 | 19.43
130.208 | 0 | 15.04 | 0.1 | 11.91 | 0.1 | 11.33 | 4.79 | 9.08 | 4.69 | 4.88 | 3.13 | 4.88 | 21.39 | 9.86 | 21.39 | 9.77 | 18.46 | 9.57 | 2.44 | 0.68 | 0 | 0.49 | 0.1 | 0.88 | 4.88 | 21.19 | 4.79 | 25.68 | 4.79 | 17.77 | 9.67 | 21.29 | 9.67 | 20.7 | 9.67 | 20.02
130.24 | 0.1 | 14.75 | 0.1 | 12.5 | 0.1 | 12.01 | 4.79 | 9.86 | 4.59 | 4.79 | 4.2 | 4.79 | 21.39 | 9.77 | 21.49 | 9.77 | 18.95 | 9.57 | 4.3 | 0.59 | 0 | 0.49 | 0 | 0.78 | 4.88 | 4.98 | 4.79 | 26.17 | 4.79 | 18.46 | 9.67 | 21.97 | 9.57 | 21.39 | 9.67 | 20.7
130.272 | 0.1 | 14.65 | 0.1 | 12.99 | 0.1 | 12.4 | 4.79 | 10.45 | 4.59 | 4.88 | 5.27 | 4.79 | 21.39 | 9.77 | 21.49 | 9.77 | 19.24 | 9.57 | 5.76 | 0.49 | 0 | 0.49 | 0 | 0.78 | 4.79 | 2.64 | 4.79 | 22.75 | 4.79 | 18.95 | 9.67 | 22.36 | 9.57 | 21.88 | 9.57 | 21.19
130.304 | 0.1 | 14.65 | 0.1 | 13.57 | 0.1 | 12.99 | 4.79 | 11.13 | 4.59 | 4.88 | 5.96 | 4.88 | 18.95 | 9.86 | 21.49 | 9.77 | 19.73 | 9.57 | 6.74 | 0.49 | 0.1 | 0.29 | 0.1 | 0.59 | 4.88 | 3.52 | 4.88 | 22.07 | 4.79 | 19.63 | 9.67 | 22.85 | 9.67 | 22.56 | 9.67 | 21.88
130.336 | 0.1 | 14.55 | 0.1 | 14.06 | 0.1 | 13.57 | 4.79 | 11.72 | 4.59 | 4.88 | 5.66 | 4.88 | 8.4 | 9.77 | 21.49 | 9.77 | 20.12 | 9.57 | 7.62 | 0.49 | 1.07 | 0.29 | 0.1 | 0.59 | 4.88 | 4.1 | 4.79 | 21.68 | 4.79 | 20.22 | 9.67 | 23.44 | 9.67 | 23.05 | 9.67 | 22.27
130.368 | 0 | 14.45 | 0.1 | 14.65 | 0.1 | 14.06 | 4.79 | 12.21 | 4.59 | 4.88 | 5.27 | 4.88 | 7.23 | 9.77 | 21.49 | 9.86 | 20.61 | 9.57 | 8.3 | 0.29 | 2.25 | 0.29 | 0.68 | 0.49 | 4.88 | 5.37 | 4.88 | 21.49 | 4.79 | 20.9 | 9.67 | 24.02 | 9.67 | 23.63 | 9.67 | 22.95
130.4 | 0.1 | 14.45 | 0.1 | 15.04 | 0.1 | 14.55 | 4.79 | 12.89 | 4.49 | 4.88 | 5.08 | 4.88 | 10.55 | 9.77 | 21.49 | 9.77 | 21 | 9.57 | 9.08 | 0.29 | 4.2 | 0.2 | 1.66 | 0.49 | 4.88 | 6.15 | 4.79 | 21.49 | 4.79 | 21.39 | 9.67 | 24.51 | 9.67 | 24.12 | 9.67 | 23.44
130.432 | 0.1 | 14.36 | 0.1 | 15.53 | 0.1 | 15.04 | 4.88 | 13.38 | 4.49 | 4.79 | 4.88 | 4.79 | 11.52 | 9.77 | 21.49 | 9.86 | 21.39 | 9.57 | 9.96 | 0.29 | 5.57 | 0.2 | 3.13 | 0.49 | 4.88 | 5.96 | 4.88 | 21.39 | 4.79 | 21.97 | 9.77 | 24.9 | 9.67 | 24.51 | 9.67 | 23.93
130.464 | 0.1 | 14.45 | 0.1 | 16.11 | 0.1 | 15.53 | 4.88 | 13.87 | 4.49 | 4.88 | 4.88 | 4.88 | 10.74 | 9.77 | 21.39 | 9.86 | 21.78 | 9.67 | 10.74 | 0.29 | 6.64 | 0.2 | 5.18 | 0.39 | 4.88 | 5.47 | 4.88 | 21.39 | 4.79 | 22.56 | 9.77 | 25.59 | 9.67 | 25.1 | 9.67 | 24.51
130.496 | 0.1 | 14.45 | 0.1 | 16.6 | 0.1 | 16.11 | 4.88 | 14.36 | 4.49 | 5.57 | 4.88 | 4.88 | 9.96 | 9.86 | 21.39 | 9.86 | 22.27 | 9.67 | 11.52 | 0.2 | 7.52 | 0.2 | 6.35 | 0.29 | 4.88 | 5.18 | 4.88 | 21.39 | 4.79 | 23.15 | 9.77 | 26.08 | 9.67 | 25.68 | 9.67 | 25
130.528 | 0.1 | 14.36 | 0.1 | 17.09 | 0.1 | 16.6 | 4.88 | 14.94 | 4.59 | 6.54 | 4.79 | 4.88 | 9.96 | 9.77 | 21.39 | 9.86 | 22.56 | 9.67 | 12.31 | 0.2 | 8.3 | 0.1 | 7.23 | 0.29 | 4.88 | 5.08 | 4.88 | 21.29 | 4.79 | 23.63 | 9.77 | 26.56 | 9.67 | 26.17 | 9.77 | 25.49
130.56 | 0.1 | 14.36 | 0.1 | 13.77 | 0.1 | 17.09 | 4.88 | 15.43 | 4.49 | 8.01 | 4.69 | 4.79 | 10.16 | 9.77 | 19.34 | 9.86 | 22.95 | 9.67 | 13.09 | 0.1 | 9.08 | 0.1 | 8.11 | 0.2 | 4.88 | 5.18 | 4.79 | 21.29 | 4.79 | 24.22 | 9.77 | 27.05 | 9.67 | 26.66 | 9.77 | 26.08
130.592 | 0.1 | 8.01 | 0.1 | 13.97 | 0 | 13.67 | 4.88 | 15.92 | 4.59 | 8.89 | 4.69 | 4.88 | 10.25 | 9.77 | 8.69 | 9.86 | 23.34 | 9.67 | 13.77 | 0.1 | 9.96 | 0.1 | 8.98 | 0.2 | 4.79 | 5.08 | 4.88 | 21.19 | 4.79 | 24.61 | 9.77 | 27.54 | 9.67 | 26.95 | 9.77 | 26.47
130.624 | 0.1 | 2.34 | 0.1 | 14.26 | 0.1 | 13.97 | 4.88 | 16.41 | 4.59 | 9.67 | 4.69 | 4.88 | 9.96 | 9.86 | 7.13 | 9.86 | 23.73 | 9.67 | 14.45 | 0.1 | 10.64 | 0.1 | 9.67 | 0.2 | 4.88 | 5.08 | 4.88 | 21.19 | 4.79 | 25.1 | 9.77 | 28.03 | 9.77 | 27.54 | 9.77 | 27.05
130.656 | 0.1 | 2.05 | 0.1 | 14.45 | 0.1 | 14.06 | 4.88 | 16.9 | 4.59 | 10.35 | 4.69 | 4.88 | 9.86 | 9.77 | 10.35 | 9.86 | 24.12 | 9.67 | 15.23 | 0.1 | 11.52 | 0.1 | 10.45 | 0.2 | 4.88 | 4.98 | 4.88 | 21.19 | 4.79 | 25.68 | 9.77 | 28.52 | 9.77 | 28.03 | 9.77 | 27.54
130.688 | 0 | 1.66 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 17.38 | 4.59 | 10.94 | 4.59 | 4.88 | 9.96 | 9.77 | 11.52 | 9.86 | 23.63 | 9.67 | 15.92 | 0.1 | 12.31 | 0.1 | 11.33 | 0.2 | 4.88 | 4.88 | 4.88 | 14.26 | 4.79 | 26.17 | 9.77 | 28.91 | 9.77 | 28.42 | 9.77 | 28.03
130.72 | 0.1 | 1.37 | 0 | 14.36 | 0 | 14.36 | 4.88 | 17.77 | 4.59 | 11.52 | 4.59 | 4.79 | 9.86 | 9.77 | 10.55 | 9.86 | 21 | 9.67 | 16.7 | 0.1 | 13.18 | 0.1 | 12.11 | 0.1 | 4.79 | 4.88 | 4.88 | 3.13 | 4.79 | 22.85 | 9.77 | 29.4 | 9.77 | 29.1 | 9.77 | 28.52
130.752 | 0.1 | 1.37 | 0.1 | 14.45 | 0 | 14.36 | 4.88 | 18.26 | 4.69 | 12.11 | 4.59 | 4.79 | 9.77 | 9.77 | 9.96 | 9.86 | 21.58 | 9.67 | 17.29 | 0.1 | 13.87 | 0.1 | 12.7 | 0.1 | 4.79 | 4.79 | 4.88 | 3.03 | 4.79 | 21.39 | 9.77 | 25.2 | 9.67 | 29.49 | 9.77 | 28.91
130.784 | 0.1 | 1.27 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 18.75 | 4.69 | 12.7 | 4.59 | 4.88 | 9.77 | 9.77 | 9.86 | 9.96 | 21.68 | 9.67 | 17.87 | 0.1 | 14.65 | 0.1 | 13.57 | 0.1 | 4.88 | 4.69 | 4.88 | 3.42 | 4.79 | 21.49 | 9.86 | 24.41 | 9.77 | 25.88 | 9.77 | 29.4
130.816 | 0.1 | 1.07 | 0 | 14.45 | 0.1 | 14.45 | 4.88 | 19.14 | 4.69 | 13.18 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.86 | 21.58 | 9.77 | 18.46 | 0.1 | 15.23 | 0.1 | 14.26 | 0.1 | 5.18 | 4.69 | 4.88 | 4.39 | 4.79 | 21.39 | 9.86 | 24.51 | 9.77 | 25.1 | 9.77 | 25.88
130.848 | 0 | 1.07 | 0.1 | 14.45 | 0.1 | 14.45 | 4.88 | 19.53 | 4.79 | 13.77 | 4.59 | 4.88 | 9.77 | 9.77 | 10.16 | 9.96 | 21.49 | 9.77 | 19.24 | 0.1 | 15.92 | 0 | 14.94 | 0.1 | 6.25 | 4.59 | 4.88 | 5.86 | 4.79 | 21.39 | 9.86 | 24.41 | 9.77 | 24.71 | 9.77 | 25
130.88 | 0.1 | 0.88 | 0 | 14.36 | 0.1 | 14.45 | 4.88 | 20.02 | 4.69 | 14.26 | 4.49 | 5.47 | 9.67 | 9.77 | 9.96 | 9.86 | 21.39 | 9.77 | 19.82 | 0.1 | 16.7 | 0.1 | 15.72 | 0.1 | 7.91 | 4.59 | 4.79 | 6.15 | 4.79 | 21.39 | 9.77 | 24.51 | 9.77 | 24.61 | 9.77 | 24.71
130.912 | 0.1 | 0.88 | 0.1 | 14.45 | 0.1 | 14.36 | 4.88 | 20.41 | 4.79 | 14.84 | 4.49 | 6.35 | 9.67 | 9.77 | 9.86 | 9.86 | 21.39 | 9.67 | 20.41 | 0.1 | 17.29 | 0.1 | 16.31 | 0.1 | 9.86 | 4.59 | 4.88 | 5.57 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.51 | 9.77 | 24.51
130.944 | 0.1 | 0.78 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 20.9 | 4.79 | 15.33 | 4.59 | 7.72 | 9.67 | 9.86 | 9.96 | 9.96 | 21.39 | 9.77 | 21 | 0.1 | 17.97 | 0.1 | 16.99 | 0.1 | 11.43 | 4.59 | 4.88 | 5.47 | 4.79 | 21.39 | 9.86 | 24.32 | 9.77 | 24.41 | 9.86 | 24.51
130.976 | 0.1 | 0.68 | 0.1 | 14.36 | 0.1 | 14.36 | 4.88 | 17.48 | 4.79 | 15.82 | 4.59 | 8.69 | 9.67 | 10.64 | 9.86 | 9.96 | 21.39 | 9.77 | 21.58 | 0.1 | 18.56 | 0.1 | 17.68 | 0.1 | 12.4 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 9.86 | 24.41 | 9.77 | 24.32 | 9.77 | 24.41
131.008 | 0 | 0.59 | 0.1 | 5.66 | 0.1 | 14.36 | 4.88 | 17.77 | 4.79 | 16.31 | 4.59 | 9.67 | 9.67 | 11.72 | 9.86 | 9.96 | 21.39 | 9.77 | 22.27 | 0.1 | 19.24 | 0 | 18.26 | 0.1 | 13.28 | 4.59 | 4.88 | 5.08 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.41 | 9.86 | 24.32
131.04 | 0.1 | 0.49 | 0 | 1.76 | 0 | 4.39 | 4.88 | 17.87 | 4.79 | 16.8 | 4.59 | 10.35 | 9.67 | 13.18 | 9.77 | 9.86 | 21.39 | 9.77 | 19.63 | 0.1 | 20.02 | 0.1 | 19.14 | 0.1 | 14.16 | 4.49 | 4.79 | 5.08 | 4.79 | 21.29 | 9.77 | 24.32 | 9.77 | 24.22 | 9.86 | 24.32

## Comparison
Measured Static Method
Epoch | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36
VT |  | 302 | 301.97 | 301.74 | 301.94 | 301.21 |  | 300.31 | 300.48 | 301.93 | 300.74 | 301.9 |  | 301.88 | 301.18 | 301.04 | 302.18 | 303.87 |  | 401.07 | 401.24 | 402 | 402.1 | 399.8 |  | 400.72 | 395.66 | 400.12 | 399.72 | 400.6 |  | 399.32 | 401.39 | 400.28 | 401.43 | 400.91
Flow |  | 21.39 | 21.45 | 21.39 | 21.45 | 21.39 |  | 21.39 | 21.15 | 21.39 | 21.15 | 21.39 |  | 21.33 | 21.39 | 21.39 | 21.39 | 21.39 |  | 28.13 | 28.83 | 28.59 | 28.65 | 28.42 |  | 28.42 | 28.48 | 28.42 | 28.13 | 28.42 |  | 28.48 | 28.48 | 28.42 | 28.42 | 28.48
PEEP |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 |  | 4.88 | 4.88 | 4.88 | 4.88 | 4.88 |  | 9.86 | 9.86 | 9.86 | 9.86 | 9.86 |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 |  | 4.88 | 4.88 | 4.88 | 4.88 | 4.88 |  | 9.77 | 9.77 | 9.77 | 9.77 | 9.77
Peak |  | 17.19 | 17.19 | 17.29 | 17.29 | 17.29 |  | 21 | 21 | 21 | 21 | 21 |  | 24.41 | 24.41 | 24.12 | 24.32 | 24.41 |  | 22.46 | 22.66 | 22.56 | 22.46 | 22.66 |  | 26.37 | 26.37 | 26.17 | 26.27 | 26.37 |  | 29.3 | 29.49 | 29.3 | 29.49 | 29.3
C |  | 21.1781 | 21.0432 | 21.0272 | 21.0411 | 20.9902 |  | 22.768 | 22.7809 | 22.8908 | 22.8006 | 22.8886 |  | 26.1821 | 25.8968 | 26.1093 | 25.9828 | 26.1281 |  | 22.9445 | 22.9542 | 22.9977 | 22.8726 | 22.7418 |  | 24.4193 | 24.2587 | 24.5322 | 24.5077 | 24.4119 |  | 27.6346 | 27.7779 | 27.701 | 27.7806 | 27.7446
R |  | 7.9383 | 7.6643 | 7.9663 | 7.9441 | 7.9663 |  | 8.2188 | 8.3121 | 8.2188 | 8.3121 | 8.2188 |  | 8.4951 | 8.1907 | 7.6578 | 7.9383 | 8.1907 |  | 10.4088 | 10.5723 | 10.4512 | 10.0105 | 10.5137 |  | 10.7248 | 10.9129 | 10.5137 | 10.8354 | 10.7248 |  | 10.7022 | 11.1025 | 10.7248 | 11.126 | 10.7022
Plateau |  | 14.36 | 14.45 | 14.45 | 14.45 | 14.45 |  | 18.07 | 18.07 | 18.07 | 18.07 | 18.07 |  | 21.39 | 21.49 | 21.39 | 21.49 | 21.49 |  | 17.58 | 17.58 | 17.58 | 17.68 | 17.68 |  | 21.29 | 21.19 | 21.19 | 21.19 | 21.29 |  | 24.22 | 24.22 | 24.22 | 24.22 | 24.22
Computed Dynamic method
C |  | 21.555 | 21.473 | 21.4 | 21.503 | 21.413 |  | 23.114 | 23.134 | 23.232 | 23.275 | 23.285 |  | 26.429 | 26.385 | 26.4 | 26.555 | 26.294 |  | 23.082 | 22.998 | 23.082 | 22.979 | 22.846 |  | 24.621 | 24.774 | 24.599 | 24.77 | 24.708 |  | 27.706 | 28.098 | 27.95 | 28.206 | 28.106
R |  | 7.77 | 7.723 | 7.63 | 7.75 | 7.56 |  | 8.16 | 8.187 | 8.328 | 8.32 | 8.31 |  | 8.303 | 8.13 | 8.123 | 8.264 | 8.06 |  | 10.065 | 10.026 | 10.335 | 10.094 | 10.035 |  | 11.038 | 10.987 | 10.9 | 10.997 | 11.148 |  | 10.829 | 10.987 | 10.935 | 10.969 | 11.042
Plateau |  | 14.1107 | 14.1628 | 14.2 | 14.1418 | 14.1667 |  | 17.8926 | 17.8887 | 17.8963 | 17.8212 | 17.8654 |  | 21.2223 | 21.2148 | 21.303 | 21.3794 | 21.4566 |  | 17.4759 | 17.5467 | 17.5162 | 17.5986 | 17.5998 |  | 21.1755 | 20.8708 | 21.1657 | 20.9373 | 21.0134 |  | 24.3128 | 24.1854 | 24.1213 | 24.1321 | 24.1642
C dif |  | 0.37688 | 0.42979 | 0.37282 | 0.46189 | 0.42276 |  | 0.34599 | 0.35311 | 0.34117 | 0.47439 | 0.39645 |  | 0.24687 | 0.48818 | 0.29072 | 0.5722 | 0.16588 |  | 0.13749 | 0.043767 | 0.084288 | 0.10642 | 0.10425 |  | 0.20174 | 0.51526 | 0.066811 | 0.26234 | 0.29606 |  | 0.071398 | 0.32015 | 0.24896 | 0.42538 | 0.36136
R dif |  | -0.16829 | 0.058664 | -0.33634 | -0.19406 | -0.40634 |  | -0.058794 | -0.12506 | 0.10921 | 0.0079433 | 0.091206 |  | -0.19208 | -0.060743 | 0.46522 | 0.32571 | -0.13074 |  | -0.34382 | -0.54632 | -0.11621 | 0.083529 | -0.47872 |  | 0.31316 | 0.074079 | 0.38628 | 0.16159 | 0.42316 |  | 0.12675 | -0.11553 | 0.21016 | -0.15697 | 0.33975
Plateau dif |  | -0.24933 | -0.28722 | -0.25 | -0.30824 | -0.28331 |  | -0.17744 | -0.18133 | -0.1737 | -0.24884 | -0.20457 |  | -0.1677 | -0.27518 | -0.08697 | -0.1106 | -0.033371 |  | -0.10412 | -0.033266 | -0.063832 | -0.081414 | -0.080219 |  | -0.11446 | -0.31922 | -0.024298 | -0.25274 | -0.27663 |  | 0.092763 | -0.034642 | -0.098712 | -0.087922 | -0.055786
% C dif |  | 1.748% | 2.002% | 1.742% | 2.148% | 1.974% |  | 1.497% | 1.526% | 1.469% | 2.038% | 1.703% |  | 0.9341% | 1.85% | 1.101% | 2.155% | 0.6309% |  | 0.5957% | 0.1903% | 0.3652% | 0.4631% | 0.4563% |  | 0.8194% | 2.08% | 0.2716% | 1.059% | 1.198% |  | 0.2577% | 1.139% | 0.8907% | 1.508% | 1.286%
% R dif |  | -2.166% | 0.7596% | -4.408% | -2.504% | -5.375% |  | -0.7205% | -1.528% | 1.311% | 0.09547% | 1.098% |  | -2.313% | -0.7472% | 5.727% | 3.941% | -1.622% |  | -3.416% | -5.449% | -1.124% | 0.8275% | -4.771% |  | 2.837% | 0.6742% | 3.544% | 1.469% | 3.796% |  | 1.17% | -1.051% | 1.922% | -1.431% | 3.077%
% Plateau dif |  | -1.767% | -2.028% | -1.761% | -2.18% | -2% |  | -0.9917% | -1.014% | -0.9706% | -1.396% | -1.145% |  | -0.7902% | -1.297% | -0.4083% | -0.5173% | -0.1555% |  | -0.5958% | -0.1896% | -0.3644% | -0.4626% | -0.4558% |  | -0.5405% | -1.53% | -0.1148% | -1.207% | -1.316% |  | 0.3815% | -0.1432% | -0.4092% | -0.3643% | -0.2309%
C |  | 21.0411 | 21.0272 | 21.0432 | 20.9902 | 21.1781 |  | 22.8006 | 22.7809 | 22.8886 | 22.768 | 22.8908 |  | 26.1093 | 25.9828 | 26.1281 | 25.8968 | 26.1821 |  | 22.9445 | 22.8726 | 22.9542 | 22.7418 | 22.9977 |  | 24.4193 | 24.4119 | 24.5077 | 24.2587 | 24.5322 |  | 27.7446 | 27.701 | 27.7779 | 27.6346 | 27.7806
R |  | 7.9441 | 7.9383 | 7.9663 | 7.6643 | 7.9663 |  | 8.2188 | 8.2188 | 8.3121 | 8.2188 | 8.3121 |  | 8.1907 | 7.9383 | 8.1907 | 7.6578 | 8.4951 |  | 10.4512 | 10.4088 | 10.5137 | 10.0105 | 10.5723 |  | 10.7248 | 10.7248 | 10.8354 | 10.5137 | 10.9129 |  | 10.7248 | 10.7022 | 11.1025 | 10.7022 | 11.126
Plateau |  | 14.45 | 14.45 | 14.45 | 14.36 | 14.45 |  | 18.07 | 18.07 | 18.07 | 18.07 | 18.07 |  | 21.49 | 21.39 | 21.49 | 21.39 | 21.49 |  | 17.58 | 17.58 | 17.68 | 17.58 | 17.68 |  | 21.19 | 21.19 | 21.29 | 21.19 | 21.29 |  | 24.22 | 24.22 | 24.22 | 24.22 | 24.22
C |  | 21.473 | 21.413 | 21.503 | 21.4 | 21.555 |  | 23.232 | 23.134 | 23.275 | 23.114 | 23.285 |  | 26.4 | 26.385 | 26.429 | 26.294 | 26.555 |  | 22.998 | 22.979 | 23.082 | 22.846 | 23.082 |  | 24.708 | 24.621 | 24.77 | 24.599 | 24.774 |  | 28.098 | 27.95 | 28.106 | 27.706 | 28.206
R |  | 7.723 | 7.63 | 7.75 | 7.56 | 7.77 |  | 8.31 | 8.187 | 8.32 | 8.16 | 8.328 |  | 8.13 | 8.123 | 8.264 | 8.06 | 8.303 |  | 10.065 | 10.035 | 10.094 | 10.026 | 10.335 |  | 10.997 | 10.987 | 11.038 | 10.9 | 11.148 |  | 10.969 | 10.935 | 10.987 | 10.829 | 11.042
Plateau |  | 14.1628 | 14.1418 | 14.1667 | 14.1107 | 14.2 |  | 17.8887 | 17.8654 | 17.8926 | 17.8212 | 17.8963 |  | 21.303 | 21.2223 | 21.3794 | 21.2148 | 21.4566 |  | 17.5467 | 17.5162 | 17.5986 | 17.4759 | 17.5998 |  | 21.0134 | 20.9373 | 21.1657 | 20.8708 | 21.1755 |  | 24.1642 | 24.1321 | 24.1854 | 24.1213 | 24.3128
|  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max
C dif |  | 0.42276 | 0.37688 | 0.42979 | 0.37282 | 0.46189 |  | 0.35311 | 0.34599 | 0.39645 | 0.34117 | 0.47439 |  | 0.29072 | 0.24687 | 0.48818 | 0.16588 | 0.5722 |  | 0.10425 | 0.084288 | 0.10642 | 0.043767 | 0.13749 |  | 0.26234 | 0.20174 | 0.29606 | 0.066811 | 0.51526 |  | 0.32015 | 0.24896 | 0.36136 | 0.071398 | 0.42538
R dif |  | -0.19406 | -0.33634 | -0.16829 | -0.40634 | 0.058664 |  | 0.0079433 | -0.058794 | 0.091206 | -0.12506 | 0.10921 |  | -0.060743 | -0.13074 | 0.32571 | -0.19208 | 0.46522 |  | -0.34382 | -0.47872 | -0.11621 | -0.54632 | 0.083529 |  | 0.31316 | 0.16159 | 0.38628 | 0.074079 | 0.42316 |  | 0.12675 | -0.11553 | 0.21016 | -0.15697 | 0.33975
Plateau dif |  | -0.28331 | -0.28722 | -0.25 | -0.30824 | -0.24933 |  | -0.18133 | -0.20457 | -0.17744 | -0.24884 | -0.1737 |  | -0.1106 | -0.1677 | -0.08697 | -0.27518 | -0.033371 |  | -0.080219 | -0.081414 | -0.063832 | -0.10412 | -0.033266 |  | -0.25274 | -0.27663 | -0.11446 | -0.31922 | -0.024298 |  | -0.055786 | -0.087922 | -0.034642 | -0.098712 | 0.092763
|  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max |  | Median | 1st Q | 3rd Q | Min | Max
% C dif |  | 1.974% | 1.748% | 2.002% | 1.742% | 2.148% |  | 1.526% | 1.497% | 1.703% | 1.469% | 2.038% |  | 1.101% | 0.9341% | 1.85% | 0.6309% | 2.155% |  | 0.4563% | 0.3652% | 0.4631% | 0.1903% | 0.5957% |  | 1.059% | 0.8194% | 1.198% | 0.2716% | 2.08% |  | 1.139% | 0.8907% | 1.286% | 0.2577% | 1.508%
% R dif |  | -2.504% | -4.408% | -2.166% | -5.375% | 0.7596% |  | 0.09547% | -0.7205% | 1.098% | -1.528% | 1.311% |  | -0.7472% | -1.622% | 3.941% | -2.313% | 5.727% |  | -3.416% | -4.771% | -1.124% | -5.449% | 0.8275% |  | 2.837% | 1.469% | 3.544% | 0.6742% | 3.796% |  | 1.17% | -1.051% | 1.922% | -1.431% | 3.077%
% Plateau dif |  | -2% | -2.028% | -1.767% | -2.18% | -1.761% |  | -1.014% | -1.145% | -0.9917% | -1.396% | -0.9706% |  | -0.5173% | -0.7902% | -0.4083% | -1.297% | -0.1555% |  | -0.4558% | -0.4626% | -0.3644% | -0.5958% | -0.1896% |  | -1.207% | -1.316% | -0.5405% | -1.53% | -0.1148% |  | -0.2309% | -0.3643% | -0.1432% | -0.4092% | 0.3815%
[14 empty] | 300 |  |  | 400
| Statistics | [10 empty] | PEEP | 0 | 5 | 10 | 0 | 5 | 10 |  | Overall
| Measured | 24.1662 | 2.2531 | C | 23.6282 | 22.7858 | 26.0777 | 20.9902 | 27.7806 |  | Crs | Measured | 21.0[21.0,21.0] | 22.8[22.8,22.9] | 26.1[26.0,26.1] | 22.9[22.9,23.0] | 24.4[24.4,24.5] | 27.7[27.7,27.8] |  | 23.6[22.8,26.1]
|  | 9.3753 | 1.3399 | R | 9.2528 | 8.1907 | 10.7022 | 7.6578 | 11.126 |  |  | Calculated | 21.5[21.4,21.5] | 23.2[23.1,23.3] | 26.4[26.4,26.3] | 23.0[23.0,23.1] | 24.7[24.6,24.8] | 28.1[28.0,28.1] |  | 23.9[23.0,26.4]
|  | 19.5037 | 3.2256 | Plateau | 19.63 | 17.605 | 21.465 | 14.36 | 24.22
[11 empty] | Rrs | Measured | 7.9[7.9,8.0] | 8.2[5.2,8.3] | 8.2[7.9,8.2] | 10.5[10.4,10.5] | 10.7[10.7,10.8] | 10.7[10.7,11.1] |  | 9.3[8.2,10.7]
[12 empty] | Calculated | 7.7[7.6,7.8] | 8.3[8.2,8.3] | 8.1[8.1,8.3] | 10.1[10.0,10.1] | 11.0[11.0,11.0] | 11.0[10.9,11.0] |  | 9.2[8.1,10.9]
| Calculated | 24.4657 | 2.244 | C | 23.942 | 23.019 | 26.3962 | 21.4 | 28.206
|  | 9.3668 | 1.3952 | R | 9.177 | 8.1375 | 10.9262 | 7.56 | 11.148 |  | Plateau | Measured | 14.5[14.5,14.5] | 18.1[18.1,18.1] | 21.5[21.4,21.5] | 17.6[17.6,17.7] | 21.2[21.2,21.3] | 24.2[24.2,24.2] |  | 19.6[17.6,21.5]
|  | 19.3513 | 3.2824 | Plateau | 19.3835 | 17.5597 | 21.2828 | 14.1107 | 24.3128 |  |  | Calculated | 14.2[14.1,14.2] | 17.9[17.9,17.9] | 21.3[21.2,21.4] | 17.5[17.5,17.6] | 21.0[20.9,21.2] | 24.2[24.1,24.2] |  | 19.4[17.6,21.3]
|  | Mean | SD |  | Median | 1st Q | 3rd Q | Min | Max
| C dif | 0.29949 | 0.14937 | C dif | 0.33066 | 0.17485 | 0.41618 | 0.043767 | 0.5722
| R dif | -0.0084532 | 0.26929 | R dif | -0.025425 | -0.16546 | 0.15288 | -0.54632 | 0.46522
| Plateau dif | -0.15241 | 0.10572 | Plateau dif | -0.14108 | -0.24983 | -0.080518 | -0.31922 | 0.092763 |  |  |  | 0: Crs; 300 0: Rrs | Plateau
[12 empty] | PEEP | Measured | Calculated |  | Measured | Calculated |  | Measured | Calculated
|  | Mean | SD |  | Median | 1st Q | 3rd Q | Min | Max |  |  | 0 | 21.0[21.0,21.0] | 21.5[21.4,21.5] |  | 7.9[7.9,8.0] | 7.7[7.6,7.8] |  | 14.5[14.5,14.5] | 14.2[14.1,14.2]
| % C dif | 1.237% | 0.6358% | % C dif | 1.242% | 0.678% | 1.747% | 0.1903% | 2.155% |  | 300 | 5 | 22.8[22.8,22.9] | 23.2[23.1,23.3] |  | 8.2[5.2,8.3] | 8.3[8.2,8.3] |  | 18.1[18.1,18.1] | 17.9[17.9,17.9]
| % R dif | -0.2125% | 2.892% | % R dif | -0.3125% | -2.03% | 1.43% | -5.449% | 5.727% |  |  | 10 | 26.1[26.0,26.1] | 26.4[26.4,26.3] |  | 8.2[7.9,8.2] | 8.1[8.1,8.3] |  | 21.5[21.4,21.5] | 21.3[21.2,21.4]
| % Plateau dif | -0.8655% | 0.6668% | % Plateau dif | -0.693% | -1.312% | -0.3754% | -2.18% | 0.3815%
[12 empty] | 0 | 22.9[22.9,23.0] | 23.0[23.0,23.1] |  | 10.5[10.4,10.5] | 10.1[10.0,10.1] |  | 17.6[17.6,17.7] | 17.5[17.5,17.6]
[11 empty] | 400 | 5 | 24.4[24.4,24.5] | 24.7[24.6,24.8] |  | 10.7[10.7,10.8] | 11.0[11.0,11.0] |  | 21.2[21.2,21.3] | 21.0[20.9,21.2]
[12 empty] | 10 | 27.7[27.7,27.8] | 28.1[28.0,28.1] |  | 10.7[10.7,11.1] | 11.0[10.9,11.0] |  | 24.2[24.2,24.2] | 24.2[24.1,24.2]
[12 empty] | Overall | 23.6[22.8,26.1] | 23.9[23.0,26.4] |  | 9.3[8.2,10.7] | 9.2[8.1,10.9] |  | 19.6[17.6,21.5] | 19.4[17.6,21.3]