Apheresis Procedure: QIMR Study cohort 4

Version 5: 18/10/2018

Bone Marrow Transplant Program

Metro North Health Hospital and Health Service

Royal Brisbane & Women’s Hospital

Children's Health Queensland Hospital and Health Service

Lady Cilento Children’s Hospital

# Purpose

## This SOP describes the apheresis procedures used for this study. It is intended to be used in conjunction with:

### Spectra Optia Essentials

# Scope

## The target audience for this resource are the registered nurses and medical officers who care for patients in the Apheresis Unit, Royal Brisbane & Women’s Hospital (RBWH), who are involved and trained in apheresis.

## Patient safety takes priority during all apheresis procedures, therefore only competent nurses can perform this procedure.

## All apheresis procedures will comply with 05009/WUG: Apheresis

## All deviations and associated corrective actions shall be documented in the patient medical record.

# Principle / Background

## Apheresis refers to a number of procedures involving the removal of whole blood, separating it into the various blood components through centrifugation, removing a desired component (white/red cells, platelets and plasma with or without replacement fluids) and returning the rest of the components. Access to the intravascular compartment is required and efficacy of the treatment will depend on the volume of blood processed.

## This apheresis procedure will consist of 2 stages:

### Stage 1: RBC Depletion of QIMR volunteer

### Stage 2: PMN procedure of concentrated red blood cell product

## During a red cell depletion, red blood cells are removed from the patient until the target haematocrit has been attained, and replaced with desired replacement fluid

## All patients requiring an apheresis procedure are under the direct care of a Haematologist.

## The RBWH Apheresis Unit is accredited by:

### Australian Bone Marrow Donor Registry (ABMDR)

### National Association of Testing Authorities (NATA)

### Foundation for the Accreditation of Cellular Therapies (FACT)

## The Royal Brisbane & Women’s Hospital (RBWH) is accredited by the Australian Council on Healthcare Standards (ACHS)

## The Disaster Plan (Section 11) addresses those elements most considered to be at risk and provides guidelines to cover some of the potential serious emergencies that may arise. All staff involved in apheresis procedures are responsible for following this plan as directed by senior personnel.

# Objectives / Endpoints

## All apheresis procedures shall be performed safely and appropriately.

# Documentation

## Policy and Standard/s

### ACHS EQUiP National

### Blood and Blood Products, Management of PROC103 Doc 110/15 (MNHHS)

### Clinical Incident and Disclosure Management PROC007 Document 13/14 (MNHHS)

## Procedures, Guidelines, Protocols

### Australian Guidelines for the Prevention and Control of Infection in Healthcare (2010)

### COLGDE001: RBWH Guidelines in the Management of Potential Adverse Events in Apheresis

### 02002/Proc: Medications Management

### 21605/Proc: Blood collection (Adult)

### 05009/Proc: Apheresis

### 05450/Proc: Peripheral Intravenous Cannulation, Venepuncture and Infusions – Adult and Paediatrics

### 74100/Proc: Documentation in the Patient Record

### 81000/Proc: Aseptic Non Touch Technique

### 000342/Proc: Standard Precautions

### 003139/Proc: Cleaning and Decontamination – patient environment and clinical equipment

## Forms and Templates

### CLIFRM013: RBWH Donor Consent for Bone Marrow Harvest or Apheresis Procedure

### COLFRM009: RBWH Apheresis Patient Checklist label

### COLFRM028: RBWH Apheresis Service Request

### COLFRM030: RBWH Apheresis Product Collection Form

### COLFRM053: RBWH Apheresis Procedure Run Sheet: Spectra Optia

# Definitions and Abbreviations

## AIM		Automated interface management

## ABMDR		Australian Bone Marrow Donor Registry

## ACD-A		Acid citrate dextrose - solution A

## BMP		Bone Marrow Processing

## BTS		Biomedical Technical Services

## ELFT		Electrolyte & liver function tests

## FBC		Full blood count

## Hb		Haemoglobin

## HBV		Hepatitis B virus

## HCV		Hepatitis C virus

## HIV		Human immunodeficiency virus

## IDM		Infectious disease marker

## IV			Intravenous

## MNHHS		Metro North Hospital and Health Service

## NUM		Nurse Unit Manager

## ODTU		Oncology Day Therapy Unit

## PMN		Polymorphonuclear (abbreviation used for granulocytes)

## QIMR		Queensland Institute of Medical Research

## RBC		Red blood cell

## RBCX		Red blood cell exchange

## RBWH		Royal Brisbane & Women’s Hospital

## RN		Registered nurse

## SOP		Standard operating procedure

## TBV		Total blood volume

# Materials

## Equipment

### Spectra Optia Apheresis System (version 11.3)

RBCX (Depletion)

PMN procedure

### IDL filler

### Standard filler

### Electronic blood pressure monitor

### Volumetric infusion pump

### Heat sealer

## Disposables and Reagents

### Spectra Optia IDL Set (Ref: 10310).

### Spectra Optia Exchange Set (Ref: 12220)

### BMP Accessory Set (Ref: 11300)

### ACD-A anticoagulant (750ml)

### 0.9% normal saline (1000ml)

### Alcohol swab

### 3 way taps

### IV cannulation or central line equipment. Refer to procedures: 05450/Proc: Peripheral Intravenous Cannulation, Venepuncture and Infusions – Adult and Paediatrics; 05501/Proc: Cannulation of Haemodialysis Access; 05600/Proc: Central Venous Access Devices (CVADs), Management – Adult. Ensure use of sufficient gauge cannulas (backeye needles are preferable for inlet venous access) to allow for adequate flow required for the procedure (preferably at least 16G for inlet and 18G for return)

### Blood collection tubes

### Intravenous giving set

### Burette

# Special Considerations

## This procedure describes how to perform:

### Stage 1 - Red cell depletion procedure.

### Stage 2 - PMN procedure

## Ensure haematocrit is accurate, as this is used to calculate:

### The limit for collect volumes

### The plasma pump flow rate before the AIM system manages the concentration of cells in the collect port

### Changing the haematocrit after the interface has been established will not change the interface position

## The packing factor during a red cell depletion defaults to 20

## The packing factor during a PMN procedure defaults to 4.5

## The Inlet:AC ratio defaults to 13:1 during a RBC depletion procedure

## There MUST be FBC and ELFT blood test results available (taken within the previous 24hrs) PRIOR to commencement of procedure.

## The minimum FBC criteria required in order to proceed with the procedure is described below.

| QIMR volunteer | Platelet > 100 x 109/L |
| --- | --- |
|  | Hb > 100G/L |

## There is a risk of deranged electrolytes due to the effects of the procedure. This effect is usually transient and resolves quickly upon cessation of CMNC collection. Electrolyte replacement may be required. Calcium gluconate is the most frequently required electrolyte replacement and the amount and rate is dependent on symptoms of hypocalcaemia during the procedure. Magnesium and potassium replacement can also be required, below is a guide for when supplementation may be required:

| Pre-Procedure Electrolyte | Replacement guide |
| --- | --- |
| Magnesium < 60 mmol/L | 10 – 20 mmols IV magnesium sulphate may be required during procedure |
| Potassium between 3.0 – 3.5 mmol/L | Oral potassium may be required e.g 1-2 chlorvescent tablets |
| Potassium < 3.0 mmol | IV potassium chloride 20 mmols may be required during the procedure and / or 3 oral chlorvescent |

# Stage 1 - Red Cell Depletion Procedure

## Confirm the following forms are current and valid:

### CLIFRM013: RBWH Donor Consent for Bone Marrow harvest or Apheresis Procedure

### COLFRM028: RBWH Apheresis Service Request

## Prior to commencement of procedure, ensure patient has had appropriate IDMs taken.

### HIV serology

### Hepatitis B serology

### Hepatitis C serology

### Syphilis

## All stock used to collect the product must be visually examined prior to use for signs of damage and contamination, as per COLSOP004: Apheresis Inventory Control Procedure. Record expiry date and lot number of all stock used on COLFRM030: RBWH Apheresis Product Collection form. Check expiry dates and lot numbers of procedure kit and intravenous fluid bags with 2nd RN and record on COLFRM053: RBWH Apheresis Procedure Run Sheet: Spectra Optia

### Confirm that all consumables and reagents used during the procedure are satisfactory for use; inspect for damage or evidence of contamination and mark the appropriate check box on COLFRM030: RBWH Apheresis Product Collection form

## On initial set up screen, choose RBCX procedure

## Set up, load and prime Spectra Optia Exchange Set (Ref: 12220) as per on screen instructions

## Follow machine prompts and enter patient data.

## Exchange type - Depletion

## Depletion fluid type – Saline / Albumin

## Fluid balance – 100%

## The default end point is target hct

## Enter target hct required to collect 450ml of concentrated red cell product

### Collect volume should not exceed 500ml

### Target hct shall not exceed 10% drop from pre procedure hct

## Educate patient of signs and symptoms of potential adverse reactions, emphasising importance of informing nursing staff if he/she experiences any such reactions

## Perform baseline observations (temperature, pulse, respirations and blood pressure) and record on run sheet

## Ensure COLFRM009: RBWH Apheresis Patient Checklist has been completed

## Perform Cannulation according to 05450/Proc: Peripheral Intravenous Cannulation, Venepuncture and Infusions – Adult and Paediatrics

## Collect blood samples if required.

## Follow machine prompts to commence procedure

## Review run screen to ensure pump flow rates are appropriate for the procedure. Monitor the AC infusion rate, ensuring that it does not exceed 0.9 ml/min

## The Spectra Optia system will attempt to establish the interface at commencement of the procedure, during this time, the collect valve will remain closed.

## The AIM system monitors for RBC spillover.

## The replace pump pumps replacement fluid into the reservoir where it combines with the plasma for return to the patient

## Monitor and record patient’s vital signs and run parameters q15 minutes or as clinically indicated. Monitor for signs of adverse reactions and record on procedure run sheet, reporting adverse events on Riskman.

## Ensure patient comfort and safety throughout the procedure.

## During a RBC depletion procedure, RBC are removed until the target Hct has been attained.

## At the end of the procedure, follow the system prompts to rinseback

## Seal the collection bag prior to disconnection of the patient, in preparation for transfer to BMP Accessory Set

## The time of collection completion shall be recorded on COLFRM030: RBWH Apheresis Product Collection form.

## On completion of rinseback, collect routine blood samples (usually FBC and ELFT).

## Perform final set of observations and record on procedure run sheet

## Remove venous access as appropriate

## Record final values on COLFRM053: RBWH Apheresis Procedure Run Sheet and complete COLFRM030: RBWH Apheresis Product Collection form.

## Unload set and dispose of waste appropriately

## Document all care given in patient medical notes. Include all completed paperwork with patient’s medical notes.

## Record relevant statistical information in: G:\Oncology\Apheresis\Log and Statistics

# Stage 2 - PMN Procedure

## Transfer collected red cell concentrate into BMP Accessory Set (Ref: 12220)

## All stock used to collect the product must be visually examined prior to use for signs of damage and contamination, as per COLSOP004: Apheresis Inventory Control Procedure. Check expiry dates and lot numbers of consumables with 2nd RN and record on COLFRM053: RBWH Apheresis Procedure Run Sheet: Spectra Optia

## Set up, load and prime Spectra Optia IDL Set (Ref: 10310) as per on screen instructions for PMN procedure

### If custom prime is required, use Albumin 4% to complete

## This procedure is intended for patients, not a bag of concentrated red cells, therefore requires patient data to be entered into program in order to proceed with procedure.

### Follow machine prompts and enter patient data.

### When prompted, enter 60% for patient hct

## Follow machine prompts to commence procedure

## The Spectra Optia system will attempt to establish the interface at commencement of procedure, during this time, the collect valve will remain closed

## The AIM system monitors the collect port and changes the plasma pump flow rate to manage the concentration of cells that flow through he collect port. When the AIM system detects cells in the collect port, the collect valve opens. Ensure the collect flow rate is set at 2ml/min.

## Ensure the main run screen accurately displays that no HES is being used

## Review AC infusion rate regularly, ensuring that it does not exceed 0.9mls/min

## The CP defaults to 60 at commencement of procedure. The apheresis nurse can set the CP to collect lighter or darker from the buffy coat. Use the collection preference tool for the appropriate CP

### See Appendix  1 for the appropriate CP, using the collection preference tool, for PMN collections

### A higher CP results in collecting lighter in the buffy coat

### A lower CP results in collecting darker in the buffy coat

## If the interface is lost during the procedure, the collect flow rate can be reduced to 1ml/min until interface has been regained. Ensure the collect flow rate is then returned to 2ml/min

## QIMR staff are responsible for labelling of collection bags.

## Seal the collection bag using heat sealer - QIMR staff will indicate the most appropriate point to heat seal line.

## On completion of procedure select DISCONNECT when prompted

## Record final procedural values on run sheet

## QIMR staff are responsible for transporting collected product to QIMR for further testing.

## Record relevant statistical information located in: G:\Oncology\Apheresis\Log and Statistics

## Unload set and dispose of waste appropriately

## Decontaminate Spectra Optia as per protocol (see COLSOP005: RBWH Apheresis Equipment Cleaning Procedure). Decontaminate all other equipment involved in performing the procedure, in accordance with 003139/Proc: Cleaning and Decontamination – patient environment and clinical equipment. Once all equipment used during the procedure has been cleaned, tick appropriate area at the end of the run sheet, confirming completion of cleaning tasks.

## Once paperwork is complete, scan / photocopy all procedural paperwork and send to QIMR.

# End Points

## The total collect volume of red cell depletion will be approximately 450ml

### Total collect volume shall not exceed 500ml

### Discuss with QIMR staff if total collect volume is expected to exceed 500ml. QIMR staff will advise apheresis nursing staff of most appropriate action

## Target hct shall not exceed 10% drop from pre procedure hct

## RBC depletion procedure time should not exceed 180 minutes

### Discuss with QIMR staff if total procedure time is expected to exceed 180 minutes. QIMR staff will advise apheresis nursing staff of most appropriate action

# Disaster Plan

## This section describes the actions to be taken in the event of serious problems which may arise in the Apheresis Unit which may impact on either:

### The ability to perform the CMNC collection to the required specification

### The quality of the collected product

## Air-Conditioning System Failure – in the event that the air-conditioning system malfunctions, contact hospital maintenance and arrange for urgent repairs (ext. 67963).

## Loss of Power – All critical equipment used to perform apheresis procedures should be connected to red emergency power points. These power points are supplied with the electricity from the emergency backup power generators, and in the event of a power failure, are automatically activated.

### Critical equipment used within the apheresis unit includes:

Cell separators

Blood warmers

Electronic blood pressure monitors

Intravenous volumetric pumps

Heat sealers

### In the event of an interruption to normal power supply, there may be a temporary loss of power before the emergency generator is activated. If there is an interruption of the power supply:

Close clamps on collect line immediately

If the emergency power supply is activated, power will resume, and the Spectra Optia system will reset. After system resets, a screen will appear with instructions for how to proceed.

To restart the procedure, follow onscreen instructions

To abort procedure, press RINSEBACK (if appropriate) or DISCONNECT and follow on screen instructions

### In the event of a complete loss of power with no emergency power supply:

The Spectra Optia cell separator will shut down completely. Close clamps on collect line immediately. Seal the collection bags using a heat sealer (these should still be functional if battery charged)

Contact the patient physician and inform him of events and that the procedure may be aborted

Consult with the Director of the Haematology & BMT

The extracorporeal volume in the kit will be 297ml

The patient may be disconnected without performing RINSEBACK or perform MANUAL RINSEBACK as per instructions in Spectra Optia Essentials Guide, pp.73-76 (in the event that a blood prime was performed, RINSEBACK is not necessary)

## Immediate evacuation – in the event that an immediate evacuation is required:

### Close clamps on collect line if possible. NURSE AND PATIENT SAFETY IS PARAMOUNT, therefore do not complete this step if completion may endanger the health and safety of the nurse or the patient

### Disconnect patient immediately and evacuate as per RBWH Code Orange Evacuation procedure. N.B. since RINSEBACK has not been performed, the patient will have a 297ml fluid deficit and may become hypotensive, monitor patient accordingly

## Structural damage (e.g. fire, water / flood damage) – in the event of structural damage, relocate product, material and patient files as required, to another area of ODTU if possible. Depending on the scale of the problem, a final decision on a temporary location may need to be determined on a case-by-case assessment by a structural engineer and the Director of the BMT program

### All critical equipment and/or consumables which may have incurred damage should be discarded and replaced as soon as possible

### Contact hospital maintenance (ext. 67963) and organise for them to inspect critical equipment as soon as possible

### Suspend all operational activities until repairs are complete and the facility is signed off as safe to reoccupy

## Critical equipment failure – Perform troubleshooting as per relevant equipment manuals. If unable to resolve problem, contact BTS or manufacturer representative.

## Critical consumable shortage – if a critical consumable shortage occurs, contact the following Apheresis Units to source consumables:

### Greenslopes Private Hospital (Speed dial *2521).

### Lady Cilento Childrens’ Hospital (Jo Ritchie – 3068 5686, or Jill Shergold – 3068 5767)

### Once alternative supply has been sourced, organise transport of the required consumables. Liaise with the NUM, ODTU to organise payment of delivery e.g. courier service

## Critical staff shortage – The Apheresis Nurse Practitioner shall liaise with the NUM, ODTU regularly to ensure staffing requirements of the apheresis unit are met. In the event that key personnel are unavailable to perform critical procedures:

### Contact Apheresis Nurse Practitioner (pager# 26299) and NUM, ODTU (ext. 68754 / 73991) informing them of the situation

### The NUM / Apheresis Nurse Practitioner will attempt to adequately staff the Apheresis Unit

### If appropriate, the situation can be escalated to the Director of the BMT program and the Nursing Director, CCS

### If adequate staffing cannot be accomplished, the Apheresis Nurse Practitioner shall liaise with the Director of the BMT Program to discuss requesting the assistance of other apheresis units in the area. Alternative units may be affiliated with the RBWH e.g. Greenslopes Private Hospital, or other similarly accredited Apheresis Units e.g. Princess Alexandra Hospital or the Mater Public Hospital

# Responsibility

|  | Final Approval | Implement | Quality Assurance | Review | Perform Procedure |
| --- | --- | --- | --- | --- | --- |
| Apheresis Nurse Practitioner |  | X | X | X | X |
| Medical Director | X | X |  |  |  |
| QIMR staff |  |  |  | X |  |
| Nursing Staff |  |  |  |  | X |

# Occupational Health and Safety

## Standard precautions shall be maintained throughout the procedure. Refer to 000342/Proc: Standard Precautions.

# Risk

## Low to Medium (case dependant).

# Consultation

## Director Haematology & Bone Marrow Transplant		Glen Kennedy

## Apheresis Nurse Practitioner					Kari Mudie

## QIMR Registrar							Anand Odedra

## Appendix 1 – Collection preference guide

1%

7.5%

3%

2%

5%

END OF PROCEDURE