PREPARATION OF INVESTIGATED MATERIALS AND μ-CTE™250 SYSTEM - Samples weighting – average mass of a sample: 1.73 ± 0.20 g; - Preparation of µ-CTETM250 system – baking-out chambers at temperature 100oC for 0.5 h under the inert gas constant flow rate (nitrogen – 50 mL·min-1)

ANALYTES SAMPLING - μ-CTE™250 WORKING PARAMETERS - Sampling device – stainless steel tube filled with sorption medium Tenax TA (60/80 mesh); - Samples seasoning/conditioning time inside a single stainless steel chamber – 30 min; - Inert dry gas flow rate – nitrogen, 12 mL·min-1, RH – 0%; - Seasoning/conditioning temperature – 40oC

ANALYTES EXTRACTION FROM SORPTION MEDIUM – TWO STAGE THERMAL DESORPTION TECHNIQUE - Conditions of the 1st stage of thermal desorption: (i) Tenax TA tube extraction temp. - 295°C; (ii) microtrap (37 mg of Tenax TA and 27 mg of Carbotrap) temp. - 0°C; (iii) desorption time – 10 min; (iv) inert gas flow rate (helium) – 50 mL·min-1 - Conditions of the 2nd stage of thermal desorption: (i) microtrap desorption temperature - 300°C; (ii) microtrap ballistic heating time – 5 min; (iii) transferring liberated analytes directly to the GC capillary column

ANALYTES SEPARATION, IDENTIFICATION AND FINAL DETERMINATION CONDITIONS - GC-FID system main working parameters: (i) GC capillary column type – DB-1: 30 m × 320 μm × 5 μm; (ii) oven working program: initial temperature – 45°C maintained for 1 min, next increased 15°C⋅min-1 up to 120°C and hold for 2 min, next increased with the rate 10°C∙min-1 up to 250°C and maintained for 5 min; (iii) FID temperature – 280oC; (iv) TD-GC transfer line temperature – 180oC; (v) gas flow rate – helium, 2 mL·min-1; - GC-MSD system main working parameters: (i) GC capillary column type – HP-1ms: 30 m × 250 μm × 1 μm; (ii) oven working program: initial temperature – 50°C maintained for 1 min, next increased 15°C⋅min-1 up to 120°C, and maintained for 2 min, after this increased with the rate 10°C∙min-1 up to 250°C and held for 5 min; (iii) GC-MSD transfer line temperature – 280oC; (iv) TD-GC transfer line temperature – 180oC; (v) gas flow rate – helium, 1 mL·min-1

DATA ANALYSIS - Identification and quantitative determination of emitted organic compounds based on reference solutions and obtained calibration curves; - Additional identification of emitted organic compounds based on GC-MS chemical compounds mass spectral library (NIST Mass Spectral Library)