Supplementary Table S1. NGS run summary

| Parameter | Results |
| --- | --- |
| Uniquely mapped reads number | 2,839,398 |
| Average depth | 18,858× |
| Uniquely mapped reads% | 88.30% |

Supplementary Table S2. Assay precision

| NGSIS (Log10 scale) | Repeatability (n=3) |  |  |  | Total imprecision (n=9) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Mean | SD | CV (%) |  | Mean | SD | CV (%) |
| 1 | 0.5 | 0.3 | 47.2 |  | 0.8 | 0.3 | 44.7 |
| 2 | 1.6 | 0.2 | 9.4 |  | 1.6 | 0.2 | 14.0 |
| 3 | 2.7 | 0.1 | 3.7 |  | 2.7 | 0.2 | 6.0 |
| 4 | 3.4 | 0.2 | 6.1 |  | 3.6 | 0.2 | 5.8 |
| 5 | 3.9 | 0.3 | 7.8 |  | 4.2 | 0.3 | 7.2 |
| MR: molecular response; SD: standard deviation; CV: coefficient of variation. |  |  |  |  |  |  |  |
| MR=log10(100%IS)-log10(%IS) =2-log10(%IS) |  |  |  |  |  |  |  |

Supplementary Table S3. Assay linearity evaluation Relative ratio; Replicate; NGSIS (Log10 scale); SD; CV (%). 0.1; 1; 0.8; 0.17; 19.25. 2; 0.8. 3; 1.1. 0.01; 1; 1.6; 0.06; 3.77. 2; 1.5. 3; 1.5. 0.001; 1; 2.9; 0.06; 2.04. 2; 2.8. 3; 2.8. 0.0001; 1; 3.6; 0.06; 1.62. 2; 3.6. 3; 3.5. 0.00001; 1; 4.2; 0.15; 3.67. 2; 4.0. 3; 4.3. MR: molecular response; SD: standard deviation; CV: coefficient of variation.. MR=log10(100%IS)-log10(%IS) =2-log10(%IS).

| Supplementary Table S4. Assay |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Sample | IS Level (%) | Day 1 |  | Day2 |  | Day3 |  |  |  |  | Day4 |  |  |  |  | Observed proportion (%) |  |
|  |  |  | #1 | #2 | #1 | #2 | #1 | #2 | #3 | #4 | #1 |  | #2 | #3 | #4 |  |  |
| LOB | N1 | Negative | ND | ND | ND | ND | ND | ND | ND | ND | 0.00056 |  | ND | ND | ND | 91.7 |  |
|  | N2 |  | ND | ND | ND | ND | 0.0002 | ND | ND | ND | ND |  | ND | ND | ND |  |  |
| LOB: Limit of Blank; ND: Not detected |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

Supplementary Figure S1. Detection of BCR-ABL1 transcript type visualized by Arriba. (A) The representative case of e13a2 transcript type: Patient 14 and (B) The representative case of e14a2 transcript type: Patient 9.