Supplementary Material

The Association between Urine Elements and Fasting Glucose Levels in a Community-Based Elderly People in Beijing

Archives of Environmental Contamination and Toxicology

Authors:

Liu Liu1,2, Ang Li3,4, Qun Xu3,4, Qin Wang1, Feng Han1,5, Chunyu Xu1, Zhe Liu1, Dongqun Xu1, *, Donggang Xu6, †

1. China CDC Key Laboratory of Environment and Population Health, National Institute of Environmental Health, Chinese Center for Disease Control and Prevention, Beijing 100021, China

2. Chaoyang District Center for Disease Control and Prevention, Beijing 100021, China

3. Department of Epidemiology and Biostatistics, Institute of Basic Medical Sciences Chinese Academy of Medical Sciences, School of Basic Medicine Peking Union Medical College, Beijing, 100005, China

4. Center of Environmental and Health Sciences, Chinese Academy of Medical Sciences, Peking Union Medical College, Beijing, 100005, China

5. The National Institute for Occupational Health and Poison Control, Chinese Center for Disease Control and Prevention, Beijing, 102206, China

6. Department of Genome Engineering, Beijing Institute of Basic Medical Sciences, Taiping Road 27, Beijing

100850, China.

* Corresponding author, E-mail: xudq@chinacdc.cn

† Co-corresponding author

Table S1 Instrument parameters for ICP-MS

| Instrument parameter | Value/setting |
| --- | --- |
| RF power (kW) | 1.5 |
| Nebulizer gas flow (L/min) | 0.92-0.96 (optimized daily) |
| Detector mode | Dual |
| Scan mode | Peak hopping |
| Sweeps/reading | 20 |
| Readings/replicate | 1 |
| Replicates | 3 |
| Calibration Regression Type | Simple liner |
| Interface | Nickel cones |

Table S2 The condition of Dynamic Reaction Cell (DRC)

| Element | Mode | Gas | Flow rate (mL/min) | Rejection parameter q (Rpq) |
| --- | --- | --- | --- | --- |
| 27Al | DRC，Channel A | NH3 | 0.7 | 0.35 |
| 52Cr | DRC，Channel A | NH3 |  | 0.45 |
| 55Mn | DRC，Channel A | NH3 |  | 0.45 |
| 56Fe | DRC，Channel A | NH3 |  | 0.55 |
| 59Co | DRC，Channel A | NH3 |  | 0.25 |
| 60Ni | DRC，Channel A | NH3 |  | 0.45 |
| 65Cu | DRC，Channel A | NH3 |  | 0.45 |
| 66Zn | DRC，Channel A | NH3 |  | 0.45 |
| 91As | DRC，Channel B | O2 | 0.7 | 0.65 |
| 78Se | DRC，Channel B | O2 |  | 0.45 |
| 88Sr | Vented | NA |  | 0.45 |
| 111Cd | Vented | NA |  | 0.45 |
| 133Cs | Vented | NA | - | 0.45 |
| 138Ba | Vented | NA |  | 0.45 |
| 208Pb | Vented | NA |  | 0.45 |

Table S3 Limits of detection, intra-assay and inter-assay coefficients of variation for routine analysis in urine

| Element | LOD (µg/L) | Intra-assay CV% | Inter-assay CV% |
| --- | --- | --- | --- |
| Al | 0.05 | 4.08% | 3.94% |
| Cr | 0.025 | 2.96% | 6.54% |
| Mn | 0.014 | 2.55% | 3.08% |
| Fe | 0.045 | 2.16% | 6.65% |
| Co | 0.013 | 2.78% | 4.30% |
| Ni | 0.028 | 2.12% | 5.68% |
| Cu | 0.165 | 2.01% | 4.87% |
| Zn | 0.173 | 3.44% | 4.35% |
| As | 0.018 | 1.35% | 4.41% |
| Se | 0.137 | 1.51% | 4.74% |
| Sr | 0.064 | 1.64% | 4.00% |
| Cd | 0.003 | 1.63% | 2.17% |
| Cs | 0.002 | 3.54% | 3.65% |
| Ba | 0.015 | 2.63% | 4.70% |
| Pb | 0.020 | 2.88% | 4.36% |