Reprocessing manual for a FFP2 mask with V-PRO® maX Low Temperature Sterilization

- Collecting used FFP 2 Masks with a specialized mailbox-type container on wards

- Delivery to the central sterilization (CS) department in plastic containers meeting requirement for transport of infectious waste (eg.g UN approval for infectious waste of Class 6.2, UN 3291)

- Store for at least 24 hours before opening the container

- Visual inspection by staff of Central Sterilization with appropriate personal protecting equipment using a magnifying glass for visual contamination and defects (e.g. holes). Rejection of defective and dirty masks, or contaminated masks (most commonly by residual make-up and lipstick)

- Film packaging: FFP2 respirators are either individually or sets a 10 to 20 packed in sterile bags suitable for H2O2 sterilization. 
The packaging is carried out according to the recommendation of the manufacturer .

- Methodology of preparation:

  - Low temperature process Hydrogen peroxide (H2O2)

  - V-PRO® maX Low Temperature Sterilization
 (STERIS 5960 Heisley Road Mentor, OH 44060, USA)

  - Short program (Materials without lumen – "28 minute Non Lumen Cycle can process loads up to 50 lbs including cameras, light cords, non-lumened endoscopes and batteries", https://www.steris.com/healthcare/products/v-pro-sterilizers/v-pro-max-low-temperature-sterilization-system)

- Keep masks in a non-airtight container for aeration, or if individually packed, have the healthcare worker open the enveloped 1 hour  before use

- Shipment to central distribution center for stockpiling as individually packed items in boxes to 10-20 masks

- The material is delivered as a non-sterile product, no double packaging

- Estimated amount of hands-on time for reprocessing 2 min/mask, for visual inspection, packaging and shipment.