## Findings summary syst review
Supplementary Table 1. Summary of the findings from the systematic literature research
Title | Authors, Year | Short Method Description | Annotation Time Perspective | Data Sources | Investigation of Clinical Actionability (incl. definition when available)? | Personal Resources Involved in Annotation | Size of the Annotated Output | Annotation Quality
Real-time evaluation of patient monitoring algorithms for critical care at the bedside15 | Zhang et al, 2007 | Real-time annotation at bedside using a system storing both physiological data and the clinical annotation. Annotation as “True Positive, Clinically Relevant”, “True Positive, Clinically Irrelevant” and “False Positive” (based on the definitions by Tsien 2000) by asking the healthcare staff at bedside. Recording of patient movements, occurrence of a procedure and response of the personnel to the alarm during and post-alarm. The annotation box is either triggered by the system when an alarm occurs or can be opened by the annotator in case a patient requires the attention of the personnel or in case of medication management interventions. | Prospective | Physiological data (focus on numeric parameters; recorded at 1 Hz) from 16 patients from a multidisciplinary Intensive Care Unit of Children’s Hospital in Boston. Total of 300 hours for testing sessions. Focus on 196 hours in the analysis of the system. | Yes – “True Positive, Clinically Relevant” as defined by Tsien (2000): “the alarm was appropriate given the actual data value as compared to the set threshold value, and the patient's condition required prompt attention” | One annotator entering the data, and asking the healthcare personnel at bedside to sort alarm-sounding events | 325 clinical alarms (within 196 h; 290 true positive and clinically relevant, 20 true positive but clinically irrelevant, 15 false positive) (1,768 technical alarms: recorded but not annotated) | “2-tier gold standard": 1) initial classification by a nurse or physician; 2) possibility to revise the initial classification within 30 minutes based on following events
Reducing false alarm rates for critical arrhythmias using the arterial blood pressure waveform17 | Aboukhalil et al, 2008 | Independent, manual annotation of the dataset by two persons, from different locations, by evaluating visually ECG and ABP waveforms surrounding the alarm timestamp (with the possibility to change time and amplitude scale), using an open-source tool. 1) Annotation as true, false or ambiguous 2) Comparison of the produced datasets and creation of a set for matched alarms and one for mismatched alarms 3) Reevaluation and adjudication of the mismatched alarms by one experienced physician or research engineer 4) Reevaluation of entire set of matched alarms | Retrospective | Selected records from MIMIC II that showed a critical arrhythmia alarm and had an ECG and ABP waveform at this moment: 496 records selected from adult patients with 8,636 alarms (total of 45,370 hours of simultaneous ECG & ABP waveforms) | No | Eleven volunteers, with different amounts of experience, divided into two teams: 1) one physician and four signal processing experts, 2) six graduate students with training in cardiac electrophysiology (graduate level) | 5,386 annotated alarms (3,084 true, 2,302 false) from 447 patients during a total of 41,301 hours (after exclusion of records from 49 patients with active intra-aortic balloon pumps, and of alarms that were not linked to waveforms) | In the discussion: errors or “points of contention” might be possible; events might be missed
Intensive care unit alarms—how many do we need?3 | Siebig et al, 2010 | Annotation, based on the review of video recordings, as clinically relevant/irrelevant or alerting, technically valid/not valid. Simple rules to determine clinical relevance and technical validity. The presence/absence of healthcare staff and the occurrence of manipulation was also noted. | Retrospective | Data from a standard cardiovascular monitoring system (incl. waveforms, measurements, alarms and alarm settings, messages; data is "saved to files at 1-sec intervals") in combination with video recordings from a total of 68 patients of a 12-bed medical ICU of a university hospital (period: January 2006 to May 2007; total of 982 hours of monitoring) | Yes – an alarm is “clinically relevant” if it is technically true and if “a diagnostic or therapeutic decision was necessary or when a persisting technical problem had to be solved by the staff”. Technically true alarms not immediately followed by an intervention are “alerting”. | Physician with at least 3 years of experience in intensive care medicine | 5,820 annotated alarms (of 5,934 recorded ones; 3,422 technically true, 2,398 technically false with 885 alarms relevant, 2,712 irrelevant and 2,223 alerting alarms) | n/a
Collection of annotated data in a clinical validation study for alarm algorithms in intensive care—a methodologic framework23 | Siebig et al, 2010 | Annotation, based on the review of video recordings, 1) as technically valid/not valid, 2) as clinically relevant/irrelevant or alerting. Simple rules/criteria (“annotation scheme”) to determine clinical relevance and technical validity. Information on the presence/absence of healthcare staff and the occurrence of manipulation was also noted. Data entry in a specially developed program based on JavaScript. | Retrospective | Data from a standard cardiovascular monitoring system (incl. waveforms, measurements, alarms and alarm settings, messages; data is “saved to files at 1-sec intervals”) in combination with video recordings (patient, surrounding monitoring devices, healthcare staff actions at bedside) from a 12-bed medical ICU of a university hospital. Currently included: 38 patient cases and 515 monitoring hours | Yes – same definition as in Siebig et al3 | Two intensivists | 3,682 alarms recorded and annotated (54.5% of the alarms are technically true, 43.6% technically false, 1.9% not assessable; 16.6% of the alarms are clinically relevant, 46.5% not relevant, 35.0% alerting) | Inter- and intraobserver variability calculated between two intensivists; 1) Interobserver variability: re-annotation of a sample of 200 alarm situations by a second person. Agreement on the technical validity for 95% of the alarms, and on the clinical relevance for 85% of the alarms. 2) Intraobserver variability: re-annotation of a sample of 100 alarms by the same annotator three months later. 99% identical annotations related to technical validity and 97% related to clinical relevance.
The wolf is crying in the operating room: patient monitor and anesthesia workstation alarming patterns during cardiac surgery24 | Schmid et al, 2011 | Alarm annotation, based on the evaluation of waveforms and measurements in combination with the video recordings, as technically true/false, and clinically relevant or not. Also recorded: information regarding whether a medical reaction followed after the alarm. | Retrospective | Data from anesthesia workplace and patient monitors (including alarm message, alarm grade, parameter triggering the alarm) in combination with video recordings (two angles) of the workplace for 25 patients undergoing elective cardiac surgery; 124 hours recorded with a total of 8,975 alarms | Yes – “[C]linically relevant if there was a need for medical intervention.” “Medical reaction: yes” if the alarm was followed by a medical reaction to it. | Mainly one anesthesiologist; if unclear, reevaluation and discussion by a group of three physicians | 8,975 alarms (incl. 6,386 serious and life-threatening alarms and 2,589 advisory or static blood pressure alarms; of 6,386 alarms, 4,438 valid and 1,948 not valid/technically false alarms; of the 4,438 valid alarms, 1,735 clinically relevant and 2,703 not relevant). | n/a
Reducing false alarms of intensive care online-monitoring systems: an evaluation of two signal extraction algorithms25 | Borowski et al, 2011 | Manual annotation by a physician by evaluating “graphical representations of the collected data from the monitoring system” in combination with video recordings (show the patients and the monitoring system screens) with the aid of a JavaScript program. Assessment if the alarm is technically true or false; assessment if the alarm is relevant/true, advisory or not relevant/false based on simple rules. | Retrospective | Data from one ICU of the University Hospital Regensburg of adult patients with continuous monitoring (“at least invasive arterial blood pressure, heart rate, and oxygen saturation”). Total of 85 cases (period: January 2006 to May 2008). Overall monitoring time of 1,245:52:28 hours with 9,290 alarms (with 4,825 threshold alarms; annotation focusing on threshold alarms from four sensors triggering most of these) | Yes – “alarm relevant or true” is assigned if the alarm is “technically true” and requires “a diagnostic or therapeutic decision or the correction of a technical problem.” | One “experienced physician” | Total of 4,359 alarms annotated as true or false after reassessment of advisory alarms (1,650 true, 2,709 false; with 2,032 advisory alarms being divided into 1,201 true and 831 false alarms); later on, 418 false alarms were excluded from the analysis | n/a
Reducing false intracranial pressure alarms using morphological waveform features26 | Scalzo et al, 2013 | Rule-based, manual annotation (with the aid of an annotation software) of intracranial pressure (ICP) alarms as true, false, and noise by an expert. Intervention of interest: cerebrospinal fluid drainage. The annotation software shows randomly selected ICP waveform segments surrounding an ICP alarm to the expert. | Retrospective | 63,954 ICP alarms of 154 included patients, recorded by bedside monitors (period: August 2010 to October 2011) of a 24-bed NeuroICU (University of California, Los Angeles); Extraction of hour-long ICP/ECG waveform segments around the ICP alarm timestamps | Yes – true alarm, if “[cerebrospinal fluid] drainage” “to treat ICP elevation within 15 min following the alarm“. | One “expert researcher” | Of 63,954 ICP alarms, after exclusion of alarms associated with defined situations: annotation of 4,791 alarms (1,739 true, 3,052 false) of 108 patients, associated with good signal quality | n/a
Semi-supervised detection of intracranial pressure alarms using waveform dynamics27 | Scalzo et al, 2013 | Rule-based, manual annotation (with the aid of an annotation software) of intracranial pressure (ICP) alarms as true, false, and noise by an expert. Intervention of interest: Cerebrospinal fluid drainage. The annotation software shows randomly selected ICP waveform segments surrounding an ICP alarm to the expert. Development of semi-supervised learning models to detect alarms also from unlabeled dataset | Retrospective | 63,954 ICP alarms of 154 included patients, recorded by bedside monitors (period: August 2010 to October 2011) of a 24-bed NeuroICU (University of California, Los Angeles); Extraction of hour-long ICP/ECG waveform segments around the ICP alarm timestamps; Convenience sample of 8,000 alarms | Yes – same as in Scalzo et al.26 | One “expert researcher” | Of 8,000 alarms, after exclusion of alarms associated with defined situations: annotation of 4,791 alarms (1,739 true, 3,052 false) of 108 patients, associated with good signal quality | n/a
The proportion of clinically relevant alarms decreases as patient clinical severity decreases in intensive care units: a pilot study28 | Inokuchi et al, 2013 | 1) Rule-based technical annotation (by nurses) as technically true, false or indeterminable - use of waveforms recordings and pulse rate 2) Rule-based clinical annotation (by physicians) as relevant, helpful but not relevant, and irrelevant - use of medical and video records 3) Evaluation of the intra- and interobserver variability; re-annotation of 300 alarms six months later by same annotator to investigate the intraobserver variability. | Retrospective | Alarm data (including settings and messages), video recordings and waveforms from 18 ICU patients aged 18 years or older (period: January-February 2012) in a 6-bed mixed ICU at University of Tokyo hospital; 2,697 patient-monitored hours with a total of 11,591 alarms | Yes – “[R]elevant when an immediate clinical examination plus diagnostic or therapeutic decision (e.g., ECG, echocardiography or drug administration) were necessary” | Two nurses and two intensivists with at least 6 years of experience in intensive care medicine annotated the data | Annotation of 11,591 alarms (8,224 technically true, 2,479 false, 888 indeterminable; 740 clinically relevant, 3,800 helpful, 7,049 irrelevant, 2 indeterminable) | Interobserver variability of the annotations estimated by the κ coefficient: 0.98 (technical) and 0.68 (clinical) Intraobserver validity of the annotations: 0.95 (technical) and 0.73 (clinical) “These values are within the range of substantial (0.61–0.80) or almost perfect (0.81–1.00) agreement.”
Insights into the problem of alarm fatigue with physiologic monitor devices: a comprehensive observational study of consecutive intensive care unit patients29 | Drew et al, 2014 | Annotation as arrhythmia, parameter or technical alarms based on a defined annotation protocol (documenting concrete criteria and rules as proof of true or false positive). For arrhythmia alarms, annotation also of their signal quality and accuracy. | Retrospective | All monitor data, including user settings, alarms, ECG leads, pressures, SpO2 and respiration waveforms, of 461 intensive care unit patients (period: 31 days, March 2013) from five ICUs with a total of 2,558,760 alarms (48,173 hours of patient monitoring time; includes 1,051,054 arrhythmia alarms, 612,927 parameter alarms and 791,632 technical alarms). For annotation: display of two pages of waveforms for each alarm (with the possibility to extend the rhythm strips) | Yes (for a subset of alarms) – a ventricular tachycardia alarm is “actionable” if it “lasted 30 seconds or longer” and therefore meets the practice guideline for treatment; an “ST-segment monitoring alarm” is “actionable” if it “persisted for more than one minute”. | Mainly four nurses with clinical experience with monitor devices and PhD training and further courses. Weekly discussion rounds and almost daily email communication with the principal investigator to reevaluate cases and for adjudication. | 12,671 arrhythmia alarms (1,405 true positive, 11,251 false positive) | Interrater reliability calculated as Cohen’s Kappa for a random sample of 300 alarms rated twice: 95%
The PhysioNet/Computing in Cardiology Challenge 2015: Reducing False Arrhythmia Alarms in the ICU30 | Clifford et al, 2015 | Independent expert annotation by means of visual evaluation of waveform recordings surrounding the alarm timestamp (15 s before, 5 s after the alarm timestamp; with the possibility for the annotators to see more record parts). Patient list to review is randomized. Annotation as true, false, reject or uncertain. Inclusion of alarms in the dataset, if common decision taken by at least two-thirds of annotators. | Retrospective | PhysioNet Challenge 2015: ICU patients’ arrhythmia waveform data from 2-lead ECG, BP, and the PPG from four hospitals in the United States and Europe; includes at least 5 minutes of data before and after each alarm and alarm messages | No | 2 or more experts per alarm | 1,564 annotated alarms (later: random selection of 750 alarms for the training set and 500 for the test set) | n/a
False alarm reduction in critical care31 | Clifford et al, 2016 | Independent expert annotation by means of visual evaluation of waveform recordings surrounding the alarm timestamp (15 s before, 5 s after the alarm timestamp; with the possibility for the annotators to see more record parts). Patient list to review is randomized. Annotation as true, false, reject or uncertain. Inclusion of alarms in the dataset, if common decision taken by at least two-thirds of annotators. | Retrospective | PhysioNet Challenge 2015: ICU patients’ arrhythmia waveform data from 2-lead ECG, BP, and the PPG from four hospitals in the United States and Europe; includes at least 5 minutes of data before and after each alarm and alarm messages | No | 2 or more experts per alarm | 1,564 annotated alarms (later: random selection of 750 alarms for the training set and 500 for the test set) | n/a
A practical algorithm to reduce false critical ECG alarms using arterial blood pressure and/or photoplethysmogram waveforms32 | Zong et al, 2016 | Manual review and annotation of each ECG alarm as true positive (TP), false positive (FP) and uncertain. | Retrospective | Selected records of MIMIC II with a minimum length of 2 hours and including continuous waveforms of ECG, PPG, and/or ABP, and one to multiple critical ECG alarms (number of records: 283 in the development set, 290 in the test set; the records in both sets are from different ICU patient groups) | No | Six human experts - two annotators (knowledgeable and experienced biomedical scientists) and a third expert to adjudicate annotation conflicts per sets. The expert groups differ for each set. | 2,408 critical ECG alarms (1,414 in the development set with 916 TP, 30 uncertain and 468 FP; 994 in the test set with 387 TP, 607 FP) | n/a
Real alerts and artifact classification in archived multi-signal vital sign monitoring data: implications for mining big data33 | Hravnak et al, 2016 | Independent annotation as real alert or artifact, using a rule set with examples developed (based on visual evaluation and clinical experience) during a first phase of data analysis by clinical investigators. Using the rules was not mandatory. Iterative development and refinement of the rules. | Retrospective | Continuous non-invasive vital sign monitoring data (heart rate, respiratory rate, oxygen saturation) and oscillometric blood pressure (1/20 Hz) from a total of 634 admissions in a 24-bed adult surgical trauma unit with a total recording of 41,634 hours (two periods of eight weeks). Data is split into two sets (training and test) | No | Annotation by two expert clinicians; development of annotation rule set by three clinical investigators | 1) Training set: 812 events with 493 real alerts and 319 artifacts 2) Test set: 1521 events with 906 real alerts and 615 artifacts | “Ground truth” ascertained in case of disagreement by “annotation in concert”
Using Supervised Machine Learning to Classify Real Alerts and Artifact in Online Multisignal Vital Sign Monitoring Data16 | Chen et al, 2016 | Independent annotation by rating alarms as real or artifacts using a scale (+3 to -3) via a web application. Agreement on an annotation defined as cumulative score of two annotators of ≥4 or ≤-4 equals. Alarm batches to annotate were chosen with aid of “Active Regression for Informative Projection Retrieval” (with initial samples being randomly selected) to decrease human effort needed. | Retrospective | Continuous non-invasive vital sign monitoring data (heart rate, respiratory rate, oxygen saturation) and oscillometric blood pressure (1/20 Hz) from a total of 2351 admissions in a 24-bed adult surgical trauma unit with a total recording of 179,157 hours (period: November 2006 to September 2008; 21 months). Data is split into two sets (training and test) | No | Four persons in total | Annotation of a subset of 973 alerts (576 for training, 397 for test set) | Kappa coefficient calculated for “paired review”, with increase of the coefficient after full committee re-adjudication. For BP and SpO2 respectively, in case of agreement (0.28±0.05 and 0.29±0.03), of disagreement (-0.23±0.05 and -0.15±0.01), and after adjudication (0.34±0.03 and 0.23±0.008). A “formal alert adjudication protocol” including cumulative scores by several reviewers was used to ensure and improve annotation accuracy.
Reduction of clinically irrelevant alarms in patient monitoring by adaptive time delays34 | Schmid et al, 2017 | Independent, manual alarm event annotation as technically true/false, clinically relevant/irrelevant, warning, serious or life threatening by evaluation and correlation of the video recordings, the physiological signals and alarm data | Retrospective | 1) 124 h of intraoperative patient physiological signals and alarm data recorded during elective cardiac surgery by the monitor and anesthesia devices. 2) Video recordings from the anesthesia setting (two angles) | Yes – with “clinically relevant alarms” being “alarms with direct or indirect clinical consequence” | Three experienced anesthesiologists (annotation by two, adjudication after consultation of the third one) | 645 annotated events with (1) 4,893 alarms (3515 clinically irrelevant, 1378 relevant) triggered by the conventional alarm algorithm, (2) 1,729 alarms (with 931 clinically irrelevant) triggered by the new alarm algorithm | Interrater reliability calculated as the Cohen’s Kappa coefficient: κ = 0.87 of events; third experienced anesthesiologist required in 13% of events
Patient characteristics associated with false arrhythmia alarms in intensive care35 | Harris et al, 2017 | Annotation (true, false, indeterminate) of arrhythmia alarms (accelerated ventricular, asystole, pause, ventricular bradycardia, ventricular fibrillation, ventricular tachycardia) using a standardized protocol, and re-annotation of a randomly selected sample by a further annotator to evaluate the interrater reliability. (Additionally: visual evaluation and rating of the signal quality.) | Retrospective | Continuous waveform and alarm data from 461 unique intensive care unit patients (period: 31 days in 2013) from five ICUs (77 beds) from an urban medical center with a total of 2,558,760 alarms (48,173 hours of patient monitoring time; includes 1,051,054 arrhythmia alarms, 612,927 parameter alarms and 791,632 technical alarms) | Yes – as defined by Drew et al.29, with a focus on non-actionable (or false) alarms | Nurse scientists with expertise in cardiac monitoring | Annotation of 12,671 arrhythmia alarms | Interrater reliability calculated as Cohen’s Kappa coefficient for a randomly selected sample to be annotated by several annotators: 0.86 at 95%
CEA: Clinical Event Annotator mHealth Application for Real-time Patient Monitoring36 | Nizami et al, 2018 | Real-time annotation at the bedside using an app on a tablet of clinical events, including monitor alarms and clinical interventions but also routine care, movements, physiological or instantaneous events | Prospective | Six patient trials in a neonatal intensive care unit | No | Research assistant or nurse | 122 alarms (and 1,188 events) | n/a
Individualized parameterization of multiparametric monitors alarms in infarcted patients37 | Assis et al, 2019 | Manual annotation at the bedside of the alarms as true or false judging the “clinical consistency” (“Does the alarm reflect any relevant clinical condition?”). Information of the nurse in charge in case of a relevant, health-threatening clinical condition. | Prospective | All the alarms triggered by the variables heart rate, ST segment, blood pressure, respiratory rate and peripheral oxygen saturation within one hour per patient (total number of patients = 32). | Yes – the alarm is “consistent” if it “reflect[s] any relevant clinical condition” through a physical examination at bedside | One nurse "with expertise and experience in the area" | Total of 460 alarms (295 pre- and 165 post-intervention). | n/a
ECG Monitoring during End of Life Care: Implications on Alarm Fatigue38 | Suba et al, 2019 | Annotation based on the alarm type, the audio level, and the hospital's own alarm priority settings | Retrospective | Discharge summaries, progress and significant event notes, consultation notes, documented vital signs, flowsheet documentation, orders, medication reports, and other clinically relevant documentation from ICU patients of the University of San Francisco during a 31-day study period (March 2013) of selected patients (who died and had comfort care order; total = 11). Data include 151 arrhythmia alarms. | No | n/a | 151 annotated arrythmia alarms | n/a
Detecting False Alarms by Analyzing Alarm-Context Information: Algorithm Development and Validation39 | Fernandes et al, 2020 | (Simulation of the) Alarm labelling depending on a calculated “false alarm probability” based on four different false alarm indicators integrated in a rule-based algorithm (as shown in a state machine diagram). Assignment of a label, if false alarm probability (FAP) >= 75.0%, else: UNDEFINED | n/a | Simulated data from Philips Efficia CM100 Monitor and eHealth Sensor Platform Kit with a focus on two sensors: heart rate and temperature | No | n/a | Total of 20: 10 labels each for temperature and heart rate alarms | n/a
Real-time machine learning-based intensive care unit alarm classification without prior knowledge of the underlying rhythm40 | Au-Yeung et al, 2021 | Re-annotation of each of the alarms [“events”] and sorting in the categories: (i) asystole, (ii) extreme bradycardia, (iii) extreme tachycardia, (iv) ventricular fibrillation, (v) ventricular tachycardia, (vi) normal sinus rhythm, (vii) sinus tachycardia, and (viii) noise/artefact, based among others on definitions from the PhysioNet 2015 challenge. | Retrospective | Recordings from the training set of the PhysioNet Challenge 2015: 300s of records prior to an alarm; signals included were 2-lead ECG, BP, and the PPG | No | n/a | Number not precised: Probably 750 recordings (training set of the PhysioNet Challenge 2015) | n/a
Utilizing Intensive Care Alarms for Machine Learning41 | Flint et al, 2022 | Alarm annotation (alarm start stamp as a reference point) based on a rule set including different mappings, and implementation through Python scripts and SQL statements. | Retrospective | 1) Alarm logs between March 2019 and October 2021 from a German university hospital containing data from 37,325 patients (43,603 ICU stays) and including 40,942,722 alarm log entries with among others 15,551,660 alarm starts 2) Patient health data | Yes – alarm is “actionable” if it requires a medical intervention | Focus groups with medical experts; technical experts | n/a | n/a