Table 1: Description of the criteria and forms of analysis used to validate the analytical method for LYC quantification. Source: (International conference on harmonization, 2005; Mayeaux et al., 2006; Perez, 2010; Shi and Maguer, 2000).

| Criteria | Form of analysis | Acceptance Criteria |
| --- | --- | --- |
| Accuracy | Solutions of three different concentration levels (25, 75 and 125 µg/mL) were prepared in triplicate | Agreement degree on LYC recovery percentage (from 80% to 120%) |
| Precision | Solutions were prepared, in triplicate, at three different concentration levels (25, 75 and 125 µg/mL), which were analyzed at different days. In order to assess the intermediate accuracy, two different analysts carried out the above process | Determination of the proximity between the results obtained by calculating the relative standard deviation (acceptable: less than or equal to 5) |
| Range | Linearity, precision and accuracy results interpretation | Compatibility with the intended purpose |
| Selectivity (specificity) | Assessment of non-interference of substances present or possibly present in the sample, they are: -mobile phase blank; -MSN-R and MSN-C blank dispersed in solvent equal to the mobile phase composition and centrifuged at 4000 RPM for 15 min; -Different pH: neutral and basic; -Degradation in 24 h by exposure in: visible light; UV light; oxygen; heat (100 ºC for 60 min); Additionally, the reference chemical LYC was compared with the one used in the applicability assay | Demonstration of the ability to identify and quantify LYC even in the presence of the mentioned factors and components by comparing the chromatograms |
| DL | Calculated through the residual’s standard deviation, obtained in the linearity statistical analysis and through the intersection of the straight line of a linear regression (constructed by three linearity curves): DL:3.3 x residuals standard deviationcalibration curve intersection | Analysis of the result obtained in the calculation |
| QL | Calculated through the residual’s standard deviation, obtained in the linearity statistical analysis and through the intersection of the straight line of a linear regression (constructed by three linearity curves): QL:10 x residuals standard deviation calibration curve intersection | Analysis of the result obtained in the calculation |
| Linearity | Solutions at 5 concentration levels (25, 50, 75, 100 and 125 µg/mL) were prepared in triplicate | The square of the correlation coefficient (R2), obtained through linear regression analysis, must be above 0.990. Additionally, data homoscedasticity proof was performed |
| Robustness | Small variations were analyzed in: - proportion of the mobile phase; - mobile phase flow; -injection volume; -column oven temperature These analyzes were performed at the average concentration (75 µg/mL) | Verification of the method's ability to resist small but deliberate variations using the accuracy criteria |